## Documentation
Date | Author | Change to spreadsheet
2011-12-05 | J. Barbe | First version to come into J. Wright's possession from Jean Barbe of LNE.
2016-01-21 | J. Wright | Added degree of equivalence plots and tables showing results from Criteria A, B, and D that are explained in Wright, J., Toman, B., Mickan, B., Wubbeler, G., Bodnar, O., and Elster, C., Transfer Standard Uncertainty Can Cause Inconclusive Inter-Laboratory Comparisons, Metrologia 53, pp. 1243 to 1258, October, 2016.
2020-03-11 | J. Wright | User can remove discrepant results by entering negative values of Ubase in Table 2.
2020-06-12 | J. Wright | Added conditional formatting so that discrepant results turn red in column AB. The user should then change the participant's uncertainty for the discrepant measurement to a negative value in the appropriate cell in Table 2. This removes that measurement from the CRV calculation and uses the uncorrelated equation for calculating u(di). This spreadsheet was validated by comparison of results from an independently programmed spreadsheet made by Bodo Mickan and/or Rainer Kramer of PTB called CCM-K6-2017_RV-calculations.xlsx when applied to the K6-2017 gas flow comparison.
2020-09-01 | J. Wright | Added traceability dependence (Table 4) where +1 or -1 is entered depending upon whether the lab's flow traceability is independent (+1) or dependent (-1) of other participants. These values are used during calculation of the uncertainty of the degree of equivalence and also in uncertainty weighting (columns X and Y), Chi-squared calculations (AA and AB). Added columns for entry of suggested transfer standard uncertainty components (root-sum-squared in column G).
2020-10-06 | J. Wright | Fixed conditional formatting for column W, Lab J. Fixed criterion D formulas.
2021-04-15 | J. Wright | Improved graph at column CF.
2022-04-14 | J. Wright | Changed the numerator in Criterion B to use Ucomp i/Ubase i instead of UTS/Ubase,i where Ucomp i is the RSS of the transfer standard uncertainty (UTS) and the repeatability measured in each participating lab (UR,i). This change was made based on suggestions from Frahm and Buker.

## Comparison calcs
This spreadsheet will calculate a comparison reference value (CRV), degrees of equivalence, uncertainties, etc. following the methods of Cox 2002.
| It will handle up to 15 labs and 18 set points: use [-] in unused cells. Cells shaded yellow can be used as inputs, the rest are outputs. | [62 empty] | 1 | 1.9 | 2.9 | 3.9 | 4.9 | 5.9 | 6.9 | 7.9 | 8.9 | 9.9
| The user can remove discrepant results by entering negative values of Ubase in Table 2. Chi-squared test is applied at cell U321:U339. | [32 empty] | increment | 0.05 | [28 empty] | -1 | 1.9 | 2.9 | 3.9 | 4.9 | 5.9 | 6.9 | 7.9 | 8.9 | 9.9
Look for notes on individual cells and the document called "Laboratory Comparison Calculations Spreadsheet Template Following Cox 2002 Method" for instructions on using this spreadsheet.
[32 empty] | Results by set point | [16 empty] | This section handles only 10 labs so far (not 15)
| Table 0. Uncertainty due to the transfer standard | [61 empty] | [-]
| Set Point | UT | UP | Uother | Udrift | UTS | [11 empty] | Comparison Reference Value (CRV) calculations | [15 empty] | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | [14 empty] | [-]
| (mL/min) | (k=2, %) | (k=2, %) | (k=2, %) | (k=2, %) | (k=2, %) | [26 empty] | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O | [14 empty] | [-]
| 10000 | 0 | 0 | 0 | 0.03 | 0.03 | [25 empty] | xi (%) | 10000 | -0.21067 | -0.1511 | -0.12 | -0.12779 | -0.18 | -0.11061 | -0.25 | -0.13057 | -0.16284 | -0.061563 | [-] | [-] | [-] | [-] | [-]
| 7500 | 0 | 0 | 0 | 0.03 | 0.03 | [25 empty] | Ubase i , k=2 (%) | 10000 | 0.12 | 0.062333 | 0.15 | 0.1 | 0.09 | 0.13 | 0.12 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-]
| 5000 | 0 | 0 | 0 | 0.03 | 0.03 | [25 empty] | UR i , k=2 (%) | 10000 | 0.06157 | 0.011816 | 0.0017948 | 0.00085392 | 0.00066531 | 0.0042614 | 0.0029942 | 0.0032141 | 0.0025319 | 0 | [-] | [-] | [-] | [-] | [-]
| 2500 | 0 | 0 | 0 | 0.03 | 0.03 | [25 empty] | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 10000 | 0.13817 | 0.070179 | 0.15298 | 0.10441 | 0.094871 | 0.13348 | 0.12373 | 0.042548 | 0.039133 | 0.1044 | [-] | [-] | [-] | [-] | [-]
| 1000 | 0 | 0 | 0 | 0.03 | 0.03 | [25 empty] | Ucomp i / Ubase, i | 10000 | 0.57075 | 0.51727 | 0.20036 | 0.30012 | 0.33342 | 0.23309 | 0.25124 | 1.0057 | 1.2043 | 0.3 | [-] | [-] | [-] | [-] | [-]
| 1000 | 0 | 0 | 0 | 0.04 | 0.04 | [25 empty] | di (%) | 10000 | -0.062091 | -0.0025235 | 0.028579 | 0.020793 | -0.031421 | 0.03797 | -0.10142 | 0.018012 | -0.01426 | 0.087016 | [-] | [-] | [-] | [-] | [-]
| 750 | 0 | 0 | 0 | 0.04 | 0.04 | [25 empty] | U(di) (k=2, %) | 10000 | 0.13627 | 0.066364 | 0.15127 | 0.10188 | 0.092084 | 0.13152 | 0.12161 | 0.035908 | 0.031788 | 0.10188 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| 500 | 0 | 0 | 0 | 0.04 | 0.04 | [25 empty] | En = di / U(di) | 10000 | -0.45565 | -0.038026 | 0.18893 | 0.20409 | -0.34122 | 0.2887 | -0.83401 | 0.50161 | -0.44859 | 0.85412 | [-] | [-] | [-] | [-] | [-] | 10000 |  |  | 10000 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| 250 | 0 | 0 | 0 | 0.04 | 0.04 | [25 empty] | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 100 | 0 | 0 | 0 | 0.04 | 0.04 | [25 empty] | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 100 | 0 | 0 | 0 | 0.06 | 0.06 | [25 empty] | di / Ubase i |  | -0.51743 | -0.040484 | 0.19053 | 0.20793 | -0.34912 | 0.29208 | -0.84517 | 0.60039 | -0.57039 | 0.87016 | [-] | [-] | [-] | [-] | [-] | -0.14858 | 0.011412 | 0.022824 | Criterion D | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 75 | 0 | 0 | 0 | 0.06 | 0.06 | [25 empty] | Pi |  | 1 | 1 | 1 | 1 | 1 | 1 | 0.92184 | 0.84081 | 0.81487 | 0.83205 | [-] | [-] | [-] | [-] | [-]
| 50 | 0 | 0 | 0 | 0.06 | 0.06 | [25 empty] | Criterion D |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-]
| 25 | 0 | 0 | 0 | 0.06 | 0.06 | [27 empty] | 1.05 | 2.05 | 3.05 | 4.05 | 5.05 | 6.05 | 7.05 | 8.05 | 9.05 | 10.05 | 11.05 | 12.05 | 13.05 | 14.05 | 15.05
| 10 | 0 | 0 | 0 | 0.06 | 0.06 | [26 empty] | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
| 10 | 0 | 0 | 0 | 0.06 | 0.06 | [25 empty] | xi (%) | 7500 | -0.15367 | -0.11631 | -0.07 | -0.10253 | -0.14 | -0.092829 | -0.21 | -0.096055 | -0.12081 | -0.018991 | [-] | [-] | [-] | [-] | [-]
| 5 | 0 | 0 | 0 | 0.06 | 0.06 | [25 empty] | Ubase i , k=2 (%) | 7500 | 0.12 | 0.062333 | 0.15 | 0.1 | 0.09 | 0.13 | 0.12 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-]
| 2 | 0 | 0 | 0 | 0.06 | 0.06 | [25 empty] | UR i , k=2 (%) | 7500 | 0.030847 | 0.0072601 | 0.0092036 | 0.0024187 | 0.0012506 | 0.0011724 | 0.0013006 | 0.0097008 | 0.0018473 | 0 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 7500 | 0.12748 | 0.069557 | 0.15325 | 0.10443 | 0.094877 | 0.13342 | 0.1237 | 0.043521 | 0.039095 | 0.1044 | [-] | [-] | [-] | [-] | [-]
| Table 1. Reported relative error of the TS, xi (%) | [15 empty] | Lab |  | CALCULATED | [12 empty] | Ucomp i / Ubase, i | 7500 | 0.35858 | 0.49518 | 0.2092 | 0.30097 | 0.33362 | 0.23095 | 0.25023 | 1.051 | 1.2023 | 0.3 | [-] | [-] | [-] | [-] | [-]
| Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O | Count | Set Point | xCRV | uCRV | UCRV | [10 empty] | di (%) | 7500 | -0.042697 | -0.0053315 | 0.040977 | 0.0084484 | -0.029023 | 0.018148 | -0.099023 | 0.014922 | -0.0098294 | 0.091986 | [-] | [-] | [-] | [-] | [-]
[19 empty] | (%) | (k=1, %) | (k=2, %) | [10 empty] | U(di) (k=2, %) | 7500 | 0.12541 | 0.065684 | 0.15153 | 0.10189 | 0.092075 | 0.13144 | 0.12156 | 0.037017 | 0.031695 | 0.10186 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| 10000 | -0.21067 | -0.1511 | -0.12 | -0.12779 | -0.18 | -0.11061 | -0.25 | -0.13057 | -0.16284 | -0.061563 | [-] | [-] | [-] | [-] | [-] | 10 | 10000 | -0.14858 | 0.011412 | 0.022824 | [10 empty] | En = di / U(di) | 7500 | -0.34046 | -0.081169 | 0.27042 | 0.082915 | -0.31521 | 0.13806 | -0.81457 | 0.4031 | -0.31012 | 0.90303 | [-] | [-] | [-] | [-] | [-] | 7500 |  |  | 7500 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| 7500 | -0.15367 | -0.11631 | -0.07 | -0.10253 | -0.14 | -0.092829 | -0.21 | -0.096055 | -0.12081 | -0.018991 | [-] | [-] | [-] | [-] | [-] | 10 | 7500 | -0.11098 | 0.011443 | 0.022887 | [10 empty] | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 5000 | -0.11858 | -0.095908 | -0.053333 | -0.094759 | -0.13 | -0.061248 | -0.18 | -0.076252 | -0.10597 | -0.0080488 | [-] | [-] | [-] | [-] | [-] | 10 | 5000 | -0.09274 | 0.011441 | 0.022882 | [10 empty] | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 2500 | -0.16175 | -0.10655 | -0.053333 | -0.10661 | -0.14 | -0.049939 | -0.13 | -0.085114 | -0.11948 | -0.017299 | [-] | [-] | [-] | [-] | [-] | 10 | 2500 | -0.10294 | 0.011363 | 0.022726 | [10 empty] | di / Ubase i |  | -0.35581 | -0.085531 | 0.27318 | 0.084484 | -0.32248 | 0.1396 | -0.82519 | 0.49739 | -0.39318 | 0.91986 | [-] | [-] | [-] | [-] | [-] | -0.11098 | 0.011443 | 0.022887 | Criterion D | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 1000 | -0.13227 | -0.055968 | -0.013333 | -0.071106 | -0.1 | -0.038821 | -0.05 | -0.054796 | -0.070417 | 0.0099693 | [-] | [-] | [-] | [-] | [-] | 10 | 1000 | -0.062332 | 0.011721 | 0.023442 | [10 empty] | Pi |  | 1 | 1 | 1 | 1 | 1 | 1 | 0.94772 | 0.89704 | 0.89872 | 0.70035 | [-] | [-] | [-] | [-] | [-]
| 1000 | -0.16272 | -0.12052 | -0.086337 | -0.12018 | -0.16 | -0.036642 | -0.11 | -0.099178 | -0.16891 | -0.058587 | [-] | [-] | [-] | [-] | [-] | 10 | 1000 | -0.12915 | 0.012259 | 0.024518 | [10 empty] | Criterion D |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-]
| 750 | -0.1515 | -0.12429 | -0.062173 | -0.10452 | -0.15 | -0.032024 | -0.09 | -0.081305 | -0.15857 | -0.051762 | [-] | [-] | [-] | [-] | [-] | 10 | 750 | -0.11818 | 0.01213 | 0.02426 | [12 empty] | 1.1 | 2.1 | 3.1 | 4.1 | 5.1 | 6.1 | 7.1 | 8.1 | 9.1 | 10.1 | 11.1 | 12.1 | 13.1 | 14.1 | 15.1
| 500 | -0.16122 | -0.14246 | -0.087777 | -0.1215 | -0.15 | -0.044826 | -0.12 | -0.095601 | -0.17128 | -0.064837 | [-] | [-] | [-] | [-] | [-] | 10 | 500 | -0.13137 | 0.012339 | 0.024677 | [11 empty] | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
| 250 | -0.15969 | -0.15709 | -0.052024 | -0.12771 | -0.17 | -0.058241 | -0.14 | -0.10021 | -0.16797 | -0.060436 | [-] | [-] | [-] | [-] | [-] | 10 | 250 | -0.13558 | 0.01225 | 0.0245 | [10 empty] | xi (%) | 5000 | -0.11858 | -0.095908 | -0.053333 | -0.094759 | -0.13 | -0.061248 | -0.18 | -0.076252 | -0.10597 | -0.0080488 | [-] | [-] | [-] | [-] | [-]
| 100 | -0.16608 | -0.16971 | -0.023316 | -0.13621 | -0.19 | -0.059378 | -0.09 | -0.10736 | -0.16181 | -0.075526 | [-] | [-] | [-] | [-] | [-] | 10 | 100 | -0.13583 | 0.012554 | 0.025109 | [10 empty] | Ubase i , k=2 (%) | 5000 | 0.12 | 0.062333 | 0.15 | 0.1 | 0.1 | 0.13 | 0.12 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-]
| 100 | 0.064789 | 0.093756 | 0.069371 | 0.19237 | 0.01 | 0.040187 | 0.182 | 0.099323 | 0.037763 | 0.10243 | [-] | [-] | [-] | [-] | [-] | 10 | 100 | 0.080595 | 0.015304 | 0.030608 | [10 empty] | UR i , k=2 (%) | 5000 | 0.030558 | 0.0060985 | 0.0065843 | 0.0069315 | 0.00035187 | 0.0027853 | 0.0023839 | 0.0055016 | 0.00056092 | 0 | [-] | [-] | [-] | [-] | [-]
| 75 | 0.077454 | 0.096093 | 0.088777 | 0.19951 | 0.03 | 0.08393 | 0.21 | 0.11915 | 0.059982 | 0.12389 | [-] | [-] | [-] | [-] | [-] | 10 | 75 | 0.099297 | 0.015366 | 0.030731 | [10 empty] | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 5000 | 0.12741 | 0.069445 | 0.15311 | 0.10463 | 0.1044 | 0.13345 | 0.12372 | 0.042782 | 0.039055 | 0.1044 | [-] | [-] | [-] | [-] | [-]
| 50 | 0.066144 | 0.082737 | 0.11203 | 0.17497 | 0.03 | 0.086481 | 0.22 | 0.15819 | 0.064609 | 0.11242 | [-] | [-] | [-] | [-] | [-] | 10 | 50 | 0.10529 | 0.015618 | 0.031236 | [10 empty] | Ucomp i / Ubase, i | 5000 | 0.35686 | 0.49113 | 0.20476 | 0.3079 | 0.30002 | 0.23176 | 0.25079 | 1.0167 | 1.2002 | 0.3 | [-] | [-] | [-] | [-] | [-]
| 25 | 0.035121 | 0.06138 | 0.10804 | 0.12014 | -0.02 | 0.092929 | 0.234 | 0.14755 | 0.03636 | 0.083042 | [-] | [-] | [-] | [-] | [-] | 10 | 25 | 0.076865 | 0.015928 | 0.031855 | [10 empty] | di (%) | 5000 | -0.025838 | -0.0031678 | 0.039407 | -0.0020187 | -0.03726 | 0.031492 | -0.08726 | 0.016488 | -0.013232 | 0.084691 | [-] | [-] | [-] | [-] | [-]
| 10 | 0.00031182 | 0.10012 | 0.092156 | 0.03001 | -0.02 | 0.083396 | 0.256 | 0.046977 | 0.045974 | 0.056086 | [-] | [-] | [-] | [-] | [-] | 10 | 10 | 0.050836 | 0.017292 | 0.034585 | [10 empty] | U(di) (k=2, %) | 5000 | 0.12534 | 0.065567 | 0.15139 | 0.1021 | 0.10187 | 0.13147 | 0.12158 | 0.036148 | 0.03165 | 0.10186 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| 10 | 0.090148 | 0.17016 | -0.024617 | 0.14548 | -0.1 | [-] | 0.352 | -0.034976 | 0.14755 | 0.094871 | [-] | [-] | [-] | [-] | [-] | 7 | 10 | 0.089814 | 0.021402 | 0.042804 | [10 empty] | En = di / U(di) | 5000 | -0.20614 | -0.048313 | 0.26029 | -0.019772 | -0.36578 | 0.23954 | -0.71771 | 0.45611 | -0.41808 | 0.83141 | [-] | [-] | [-] | [-] | [-] | 5000 |  |  | 5000 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| 5 | 0.20072 | 0.31433 | 0.0033333 | 0.20943 | -0.19561 | [-] | 0.347 | 0.12689 | 0.25487 | 0.15711 | [-] | [-] | [-] | [-] | [-] | 7 | 5 | 0.21154 | 0.022087 | 0.044173 | [10 empty] | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 2 | 0.22053 | 0.54472 | 0.036667 | 0.15055 | -0.43342 | [-] | 0.204 | 0.071712 | 0.39279 | 0.24382 | [-] | [-] | [-] | [-] | [-] | 6 | 2 | 0.10851 | 0.027082 | 0.054163 | [10 empty] | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
[32 empty] | di / Ubase i |  | -0.21532 | -0.05082 | 0.26271 | -0.020187 | -0.3726 | 0.24224 | -0.72717 | 0.54959 | -0.52929 | 0.84691 | [-] | [-] | [-] | [-] | [-] | -0.09274 | 0.011441 | 0.022882 | Criterion D | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| Table 2. Reported participating lab uncertainty, Ubase i , k=2 (%) | [30 empty] | Pi |  | 1 | 1 | 1 | 1 | 1 | 1 | 0.996 | 0.87043 | 0.83715 | 0.8776 | [-] | [-] | [-] | [-] | [-]
| Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O |  | ID | Reported relative difference | Standard uncertainty of reported value |  |  |  |  | Chi2 | Degrees of Equivalence | [5 empty] | Criterion D |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-]
[18 empty] | Laboratory | xi | u(xi) | xi/u2(xi) | 1/u2(xi) | wi | wi xi | (xi-xCRV)2/u2(xi) | di | u2(di) | u2cor(di) | U(di) | Ucor(di) | En |  |  | 1.15 | 2.15 | 3.15 | 4.15 | 5.15 | 6.15 | 7.15 | 8.15 | 9.15 | 10.15 | 11.15 | 12.15 | 13.15 | 14.15 | 15.15
| 10000 | 0.12 | 0.062333 | 0.15 | 0.1 | 0.09 | 0.13 | 0.12 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-] |  | Lab A | -0.21067 | 0.069085 | -44.1408 | 209.5247 | 0.027288 | -0.0057488 | 0.80779 | -0.062091 | 0.0046425 | 0.0048851 | 0.13627 | 0.13979 | -0.45565 |  | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
| 7500 | 0.12 | 0.062333 | 0.15 | 0.1 | 0.09 | 0.13 | 0.12 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.15367 | 0.063741 | -37.8239 | 246.1306 | 0.032231 | -0.0049531 | 0.44871 | -0.042697 | 0.0039319 | 0.0041338 | 0.12541 | 0.12859 | -0.34046 | xi (%) | 2500 | -0.16175 | -0.10655 | -0.053333 | -0.10661 | -0.14 | -0.049939 | -0.13 | -0.085114 | -0.11948 | -0.017299 | [-] | [-] | [-] | [-] | [-]
| 5000 | 0.12 | 0.062333 | 0.15 | 0.1 | 0.1 | 0.13 | 0.12 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.11858 | 0.063706 | -29.2175 | 246.3992 | 0.032253 | -0.0038245 | 0.1645 | -0.025838 | 0.0039276 | 0.0040954 | 0.12534 | 0.12799 | -0.20614 | Ubase i , k=2 (%) | 2500 | 0.12 | 0.062333 | 0.15 | 0.1 | 0.08 | 0.13 | 0.12 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-]
| 2500 | 0.12 | 0.062333 | 0.15 | 0.1 | 0.08 | 0.13 | 0.12 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.16175 | 0.064053 | -39.4237 | 243.7349 | 0.031471 | -0.0050904 | 0.84282 | -0.058804 | 0.0039737 | 0.0041426 | 0.12607 | 0.12873 | -0.46643 | UR i , k=2 (%) | 2500 | 0.033336 | 0.0070404 | 0.0025021 | 0.0037977 | 0.0032322 | 0.0021735 | 0.00094919 | 0.0095612 | 0.00035371 | 0.0066667 | [-] | [-] | [-] | [-] | [-]
| 1000 | 0.12 | 0.063667 | 0.15 | 0.1 | 0.09 | 0.13 | 0.12 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-] |  |  | -0.13227 | 0.063769 | -32.5265 | 245.9144 | 0.033784 | -0.0044685 | 1.2028 | -0.069936 | 0.0039291 | 0.0039908 | 0.12536 | 0.12635 | -0.55786 | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 2500 | 0.12811 | 0.069534 | 0.15299 | 0.10447 | 0.085501 | 0.13343 | 0.1237 | 0.04349 | 0.039053 | 0.10462 | [-] | [-] | [-] | [-] | [-]
| 1000 | 0.05 | 0.063667 | 0.15 | 0.1 | 0.09 | 0.13 | 0.12 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.16272 | 0.034225 | -138.9139 | 853.7079 | 0.12829 | -0.020876 | 0.96203 | -0.033569 | 0.0010211 | 0.0013538 | 0.063909 | 0.073588 | -0.52527 | Ucomp i / Ubase, i | 2500 | 0.37373 | 0.49436 | 0.20069 | 0.30239 | 0.37717 | 0.23137 | 0.25013 | 1.0496 | 1.2001 | 0.30732 | [-] | [-] | [-] | [-] | [-]
| 750 | 0.05 | 0.063433 | 0.15 | 0.1 | 0.09 | 0.13 | 0.11 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.1515 | 0.032461 | -143.7776 | 949.0163 | 0.13963 | -0.021155 | 1.0536 | -0.03332 | 0.00090659 | 0.0012744 | 0.060219 | 0.071397 | -0.55332 | di (%) | 2500 | -0.058804 | -0.003609 | 0.04961 | -0.0036705 | -0.037056 | 0.053004 | -0.027056 | 0.01783 | -0.016541 | 0.085645 | [-] | [-] | [-] | [-] | [-]
| 500 | 0.05 | 0.064 | 0.15 | 0.1 | 0.1 | 0.13 | 0.12 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.16122 | 0.034321 | -136.8671 | 848.9331 | 0.12924 | -0.020837 | 0.75676 | -0.029857 | 0.0010257 | 0.001348 | 0.064053 | 0.073431 | -0.46612 | U(di) (k=2, %) | 2500 | 0.12607 | 0.065716 | 0.15129 | 0.10197 | 0.082426 | 0.13148 | 0.12159 | 0.03708 | 0.031759 | 0.10212 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| 250 | 0.05 | 0.069333 | 0.15 | 0.1 | 0.08 | 0.13 | 0.12 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-] |  |  | -0.15969 | 0.033375 | -143.3627 | 897.7597 | 0.13473 | -0.021514 | 0.52193 | -0.024112 | 0.00096382 | 0.0012837 | 0.062091 | 0.071658 | -0.38833 | En = di / U(di) | 2500 | -0.46643 | -0.054918 | 0.32791 | -0.035996 | -0.44957 | 0.40312 | -0.22252 | 0.48085 | -0.52083 | 0.83869 | [-] | [-] | [-] | [-] | [-] | 2500 |  |  | 2500 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| 100 | 0.05 | 0.082667 | 0.15 | 0.1 | 0.09 | 0.13 | 0.12 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-] |  |  | -0.16608 | 0.035033 | -135.3143 | 814.7722 | 0.12842 | -0.021327 | 0.74548 | -0.030248 | 0.0010697 | 0.0014333 | 0.065413 | 0.075718 | -0.46242 | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 100 | 0.05 | 0.081667 | 0.15 | 0.1 | 0.09 | 0.13 | 0.11 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-] |  |  | 0.064789 | 0.039364 | 41.8127 | 645.3661 | 0.15115 | 0.009793 | 0.16122 | -0.015805 | 0.0013153 | 0.0019157 | 0.072534 | 0.087538 | -0.2179 | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 75 | 0.05 | 0.085667 | 0.15 | 0.1 | 0.09 | 0.13 | 0.11 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-] |  |  | 0.077454 | 0.039967 | 48.4888 | 626.0363 | 0.14781 | 0.011448 | 0.2987 | -0.021843 | 0.0013612 | 0.0019121 | 0.07379 | 0.087455 | -0.29602 | di / Ubase i |  | -0.49004 | -0.057898 | 0.33074 | -0.036705 | -0.4632 | 0.40773 | -0.22547 | 0.59433 | -0.66164 | 0.85645 | [-] | [-] | [-] | [-] | [-] | -0.10294 | 0.011363 | 0.022726 | Criterion D | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 50 | 0.05 | 0.096667 | 0.15 | 0.1 | 0.1 | 0.13 | 0.11 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-] |  |  | 0.066144 | 0.040401 | 40.5245 | 612.6692 | 0.14944 | 0.0098845 | 0.93883 | -0.039145 | 0.0013883 | 0.0020729 | 0.074519 | 0.091059 | -0.52531 | Pi |  | 1 | 1 | 1 | 1 | 0.99986 | 1 | 1 | 0.84568 | 0.758 | 0.86151 | [-] | [-] | [-] | [-] | [-]
| 25 | 0.05 | 0.11921 | 0.15 | 0.1 | 0.08 | 0.13 | 0.12 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-] |  |  | 0.035121 | 0.039446 | 22.571 | 642.6676 | 0.16304 | 0.0057261 | 1.1199 | -0.041744 | 0.0013023 | 0.0023789 | 0.072175 | 0.097548 | -0.57837 | Criterion D |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-]
| 10 | 0.05 | 0.21333 | 0.15 | 0.1 | 0.16 | 0.13 | 0.14 | 0.03 | 0.05 | 0.11 | [-] | [-] | [-] | [-] | [-] |  |  | 0.00031182 | 0.040362 | 0.19141 | 613.8526 | 0.18356 | 5.7237e-05 | 1.567 | -0.050524 | 0.00133 | 0.0020554 | 0.072939 | 0.090673 | -0.69269 |  |  | 1.2 | 2.2 | 3.2 | 4.2 | 5.2 | 6.2 | 7.2 | 8.2 | 9.2 | 10.2 | 11.2 | 12.2 | 13.2 | 14.2 | 15.2
| 10 | 0.05 | 0.22667 | 0.1 | 0.1 | 0.16 | [-] | -0.14 | -0.03 | 0.05 | 0.11 | [-] | [-] | [-] | [-] | [-] |  |  | 0.090148 | 0.043512 | 47.6153 | 528.1901 | 0.24194 | 0.02181 | 5.9108e-05 | 0.00033452 | 0.0014352 | 0.002825 | 0.075768 | 0.1063 | 0.0044151 |  | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
| 5 | 0.05 | 0.44333 | -0.1 | 0.12 | -0.16 | [-] | 0.19 | 0.03 | 0.05 | 0.11 | [-] | [-] | [-] | [-] | [-] |  |  | 0.20072 | 0.043546 | 105.8513 | 527.3659 | 0.25726 | 0.051636 | 0.061723 | -0.010819 | 0.0014084 | 0.0019142 | 0.075057 | 0.087503 | -0.14414 | xi (%) | 1000 | -0.13227 | -0.055968 | -0.013333 | -0.071106 | -0.1 | -0.038821 | -0.05 | -0.054796 | -0.070417 | 0.0099693 | [-] | [-] | [-] | [-] | [-]
| 2 | -0.05 | 0.48 | 0.1 | 0.25 | -0.2 | [-] | 0.32 | 0.03 | -0.08 | 0.11 | [-] | [-] | [-] | [-] | [-] |  |  | 0.22053 | -0.045808 | [5 empty] | 0.11201 | 0.0028318 |  | 0.10643 |  | 1.0525 | Ubase i , k=2 (%) | 1000 | 0.12 | 0.063667 | 0.15 | 0.1 | 0.09 | 0.13 | 0.12 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-]
[18 empty] | Lab B | -0.1511 | 0.035089 | -122.7215 | 812.1715 | 0.10578 | -0.015983 | 0.0051721 | -0.0025235 | 0.001101 | 0.0013436 | 0.066364 | 0.073311 | -0.038026 | UR i , k=2 (%) | 1000 | 0.031078 | 0.0094408 | 0.0092677 | 0.015787 | 0.0053141 | 0.0011431 | 0.0011919 | 0.01871 | 0.0023501 | 0 | [-] | [-] | [-] | [-] | [-]
| Table 3. Reproducibility, UR i , k=2 (%) | [17 empty] | -0.11631 | 0.034778 | -96.1592 | 826.7617 | 0.10827 | -0.012592 | 0.0235 | -0.0053315 | 0.0010786 | 0.0012804 | 0.065684 | 0.071566 | -0.081169 | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 1000 | 0.12754 | 0.071011 | 0.15325 | 0.10559 | 0.095017 | 0.13342 | 0.1237 | 0.046369 | 0.039122 | 0.11402 | [-] | [-] | [-] | [-] | [-]
| Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O |  |  | -0.095908 | 0.034723 | -79.548 | 829.4218 | 0.10857 | -0.010413 | 0.0083231 | -0.0031678 | 0.0010748 | 0.0012426 | 0.065567 | 0.070501 | -0.048313 | Ucomp i / Ubase, i | 1000 | 0.35996 | 0.49399 | 0.20933 | 0.339 | 0.33852 | 0.23094 | 0.2502 | 1.1785 | 1.2037 | 0.27273 | [-] | [-] | [-] | [-] | [-]
[19 empty] | -0.10655 | 0.034767 | -88.151 | 827.2988 | 0.10682 | -0.011382 | 0.010775 | -0.003609 | 0.0010796 | 0.0012485 | 0.065716 | 0.070669 | -0.054918 | di (%) | 1000 | -0.069936 | 0.0063635 | 0.048998 | -0.0087743 | -0.037668 | 0.023511 | 0.012332 | 0.007536 | -0.0080853 | 0.072301 | [-] | [-] | [-] | [-] | [-]
| 10000 | 0.06157 | 0.011816 | 0.0017948 | 0.00085392 | 0.00066531 | 0.0042614 | 0.0029942 | 0.0032141 | 0.0025319 | 0 | [-] | [-] | [-] | [-] | [-] |  |  | -0.055968 | 0.035506 | -44.3966 | 793.2458 | 0.10898 | -0.0060993 | 0.032122 | 0.0063635 | 0.0011233 | 0.001185 | 0.06703 | 0.068849 | 0.094935 | U(di) (k=2, %) | 1000 | 0.12536 | 0.06703 | 0.15145 | 0.10295 | 0.09208 | 0.13135 | 0.12146 | 0.040007 | 0.031321 | 0.11158 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| 7500 | 0.030847 | 0.0072601 | 0.0092036 | 0.0024187 | 0.0012506 | 0.0011724 | 0.0013006 | 0.0097008 | 0.0018473 | 0 | [-] | [-] | [-] | [-] | [-] |  |  | -0.12052 | 0.038035 | -83.3094 | 691.2474 | 0.10388 | -0.01252 | 0.051469 | 0.0086289 | 0.0012964 | 0.0016291 | 0.072011 | 0.080724 | 0.11983 | En = di / U(di) | 1000 | -0.55786 | 0.094935 | 0.32353 | -0.085225 | -0.40908 | 0.179 | 0.10153 | 0.18837 | -0.25815 | 0.64796 | [-] | [-] | [-] | [-] | [-] | 1000 |  |  | 1000 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| 5000 | 0.030558 | 0.0060985 | 0.0065843 | 0.0069315 | 0.00035187 | 0.0027853 | 0.0023839 | 0.0055016 | 0.00056092 | 0 | [-] | [-] | [-] | [-] | [-] |  |  | -0.12429 | 0.037577 | -88.0204 | 708.194 | 0.1042 | -0.012951 | 0.026414 | -0.0061072 | 0.0012649 | 0.0016327 | 0.071131 | 0.080813 | -0.085858 | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 2500 | 0.033336 | 0.0070404 | 0.0025021 | 0.0037977 | 0.0032322 | 0.0021735 | 0.00094919 | 0.0095612 | 0.00035371 | 0.0066667 | [-] | [-] | [-] | [-] | [-] |  |  | -0.14246 | 0.038091 | -98.1855 | 689.2006 | 0.10493 | -0.014948 | 0.084874 | -0.011097 | 0.0012987 | 0.001621 | 0.072075 | 0.080524 | -0.15397 | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 1000 | 0.031078 | 0.0094408 | 0.0092677 | 0.015787 | 0.0053141 | 0.0011431 | 0.0011919 | 0.01871 | 0.0023501 | 0 | [-] | [-] | [-] | [-] | [-] |  |  | -0.15709 | 0.040374 | -96.3691 | 613.4815 | 0.092064 | -0.014462 | 0.28379 | -0.021508 | 0.00148 | 0.0017999 | 0.076941 | 0.08485 | -0.27954 | di / Ubase i |  | -0.5828 | 0.09995 | 0.32666 | -0.087743 | -0.41854 | 0.18085 | 0.10276 | 0.2512 | -0.32341 | 0.65728 | [-] | [-] | [-] | [-] | [-] | -0.062332 | 0.011721 | 0.023442 | Criterion D | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 1000 | 0.024196 | 0.011541 | 0.010062 | 0.00093884 | 0.0013301 | 0.0081497 | 0.015315 | 0.0041494 | 0.0025319 | 0.0066667 | [-] | [-] | [-] | [-] | [-] |  |  | -0.16971 | 0.046277 | -79.2466 | 466.9511 | 0.073598 | -0.01249 | 0.53608 | -0.033883 | 0.0019839 | 0.0023475 | 0.089083 | 0.096902 | -0.38035 | Pi |  | 0.99998 | 1 | 1 | 1 | 0.99999 | 1 | 1 | 0.96812 | 0.91659 | 0.99877 | [-] | [-] | [-] | [-] | [-]
| 750 | 0.010719 | 0.0049378 | 0.012731 | 0.0020421 | 8.8789e-05 | 0.0036544 | 0.0024246 | 0.006953 | 0.0018473 | 9.8131e-18 | [-] | [-] | [-] | [-] | [-] |  |  | 0.093756 | 0.051052 | 35.9731 | 383.6872 | 0.089864 | 0.0084253 | 0.066466 | 0.013162 | 0.0023721 | 0.0029725 | 0.097408 | 0.10904 | 0.13512 | Criterion D |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-]
| 500 | 0.024735 | 0.010384 | 0.0092317 | 0.0036078 | 0.0057501 | 0.0024173 | 0.0026804 | 0.0069529 | 0.00056092 | 0.0066667 | [-] | [-] | [-] | [-] | [-] |  |  | 0.096093 | 0.052617 | 34.7083 | 361.1938 | 0.085279 | 0.0081948 | 0.0037073 | -0.0032037 | 0.0025325 | 0.0030834 | 0.10065 | 0.11106 | -0.031831 |  |  | 1.25 | 2.25 | 3.25 | 4.25 | 5.25 | 6.25 | 7.25 | 8.25 | 9.25 | 10.25 | 11.25 | 12.25 | 13.25 | 14.25 | 15.25
| 250 | 0.018856 | 0.010633 | 0.0043627 | 0.0091973 | 0.0019118 | 0.0053189 | 0.0020495 | 0.0041589 | 0.00035371 | 0 | [-] | [-] | [-] | [-] | [-] |  |  | 0.082737 | 0.057823 | 24.7453 | 299.0828 | 0.072951 | 0.0060357 | 0.15212 | -0.022552 | 0.0030996 | 0.0037843 | 0.11135 | 0.12303 | -0.20254 |  | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
| 100 | 0.028449 | 0.011508 | 0.011508 | 0.0068012 | 0.0067104 | 0.0083878 | 0.00098988 | 0.0030124 | 0.0023501 | 0.0066667 | [-] | [-] | [-] | [-] | [-] |  |  | 0.06138 | 0.069952 | 12.5437 | 204.3627 | 0.051845 | 0.0031822 | 0.049006 | -0.015485 | 0.0046396 | 0.0057162 | 0.13623 | 0.15121 | -0.11367 | xi (%) | 1000 | -0.16272 | -0.12052 | -0.086337 | -0.12018 | -0.16 | -0.036642 | -0.11 | -0.099178 | -0.16891 | -0.058587 | [-] | [-] | [-] | [-] | [-]
| 100 | 0.0099012 | 0.012479 | 0.0047755 | 0.0023183 | 0.0087268 | 0.00071574 | 0.00029136 | 0.02161 | 0.00050632 | 0.013333 | [-] | [-] | [-] | [-] | [-] |  |  | 0.10012 | 0.11625 | 7.4083 | 73.9907 | 0.022125 | 0.0022153 | 0.17975 | 0.049289 | 0.013216 | 0.013942 | 0.22992 | 0.23615 | 0.21437 | Ubase i , k=2 (%) | 1000 | 0.05 | 0.063667 | 0.15 | 0.1 | 0.09 | 0.13 | 0.12 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-]
| 75 | 0.017012 | 0.011645 | 0.006175 | 0.00089244 | 0.011666 | 0.0099942 | 0.0032298 | 0.020505 | 0.00073995 | 0 | [-] | [-] | [-] | [-] | [-] |  |  | 0.17016 | 0.11782 | 12.2584 | 72.0393 | 0.032997 | 0.0056149 | 0.46508 | 0.080349 | 0.013423 | 0.014813 | 0.23172 | 0.24342 | 0.34675 | UR i , k=2 (%) | 1000 | 0.024196 | 0.011541 | 0.010062 | 0.00093884 | 0.0013301 | 0.0081497 | 0.015315 | 0.0041494 | 0.0025319 | 0.0066667 | [-] | [-] | [-] | [-] | [-]
| 50 | 0.020708 | 0.020731 | 0.01174 | 0.001338 | 0.0098033 | 0.005993 | 0.0030386 | 0.022251 | 0.00072837 | 0.011547 | [-] | [-] | [-] | [-] | [-] |  |  | 0.31433 | 0.22499 | 6.2094 | 19.7545 | 0.0096366 | 0.0030291 | 0.20874 | 0.10279 | 0.050134 | 0.050639 | 0.44781 | 0.45006 | 0.22955 | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 1000 | 0.06845 | 0.07607 | 0.15557 | 0.10771 | 0.098498 | 0.13626 | 0.12741 | 0.050172 | 0.047238 | 0.10791 | [-] | [-] | [-] | [-] | [-]
| 25 | 0.011138 | 0.041965 | 0.0072638 | 0.0046299 | 0.029529 | 0.0072966 | 0.0018811 | 0.039137 | 0.0015364 | 0.0066667 | [-] | [-] | [-] | [-] | [-] |  |  | 0.54472 | 0.25266 | 8.5333 | 15.6653 | 0.011489 | 0.0062584 | 2.9808 | 0.43621 | 0.063102 | 0.06492 | 0.5024 | 0.50959 | 0.86825 | Ucomp i / Ubase, i | 1000 | 0.93497 | 0.6539 | 0.27497 | 0.40011 | 0.44469 | 0.31401 | 0.35693 | 1.3405 | 1.6032 | 0.40552 | [-] | [-] | [-] | [-] | [-]
| 10 | 0.020402 | 0.070354 | 0.030887 | 0.011554 | 0.033584 | 0.0042166 | 0.0052316 | 0.020458 | 0.002399 | 0.017638 | [-] | [-] | [-] | [-] | [-] |  | Lab C | -0.12 | 0.076491 | -20.51 | 170.9166 | 0.02226 | -0.0026712 | 0.1396 | 0.028579 | 0.0057206 | 0.0059631 | 0.15127 | 0.15444 | 0.18893 | di (%) | 1000 | -0.033569 | 0.0086289 | 0.042813 | 0.0089674 | -0.030851 | 0.092507 | 0.019149 | 0.029971 | -0.039757 | 0.070563 | [-] | [-] | [-] | [-] | [-]
| 10 | 0.03838 | 0.023398 | 0.01395 | 0.048085 | 0.01811 | [-] | 0.0011262 | 0.044032 | 0.0058484 | 0.0066667 | [-] | [-] | [-] | [-] | [-] |  |  | -0.07 | 0.076624 | -11.9227 | 170.3236 | 0.022304 | -0.0015613 | 0.28599 | 0.040977 | 0.0057402 | 0.0059421 | 0.15153 | 0.15417 | 0.27042 | U(di) (k=2, %) | 1000 | 0.063909 | 0.072011 | 0.15362 | 0.10488 | 0.095397 | 0.13403 | 0.12503 | 0.043773 | 0.040377 | 0.10509 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| 5 | 0.038534 | 0.048385 | 0.015327 | 0.013603 | 0.062884 | [-] | 0.0026014 | 0.097972 | 0.0082079 | 0.017638 | [-] | [-] | [-] | [-] | [-] |  |  | -0.053333 | 0.076556 | -9.0999 | 170.6241 | 0.022334 | -0.0011912 | 0.26496 | 0.039407 | 0.0057299 | 0.0058978 | 0.15139 | 0.15359 | 0.26029 | En = di / U(di) | 1000 | -0.52527 | 0.11983 | 0.27869 | 0.085501 | -0.32339 | 0.69017 | 0.15315 | 0.68469 | -0.98464 | 0.67147 | [-] | [-] | [-] | [-] | [-] | 1000 |  |  | 1000 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| 2 | 0.04789 | 0.14608 | 0.022711 | 0.028587 | 0.095278 | [-] | 0.0065771 | 0.03283 | 0.01638 | 0.017638 | [-] | [-] | [-] | [-] | [-] |  |  | -0.053333 | 0.076496 | -9.1144 | 170.8945 | 0.022066 | -0.0011769 | 0.4206 | 0.04961 | 0.0057224 | 0.0058913 | 0.15129 | 0.15351 | 0.32791 | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
[19 empty] | -0.013333 | 0.076626 | -2.2709 | 170.315 | 0.023398 | -0.00031198 | 0.4089 | 0.048998 | 0.0057341 | 0.0057959 | 0.15145 | 0.15226 | 0.32353 | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| Table 4. Participating lab's traceability dependence: 1=independent of other participants, -1=dependent on another participant so that their measurements are not included in the CRV calculation. | [17 empty] | -0.086337 | 0.077784 | -14.2698 | 165.2807 | 0.024838 | -0.0021444 | 0.30295 | 0.042813 | 0.0059 | 0.0062327 | 0.15362 | 0.1579 | 0.27869 | di / Ubase i |  | -0.67138 | 0.13553 | 0.28542 | 0.089674 | -0.34279 | 0.71159 | 0.15958 | 0.99903 | -1.5903 | 0.70563 | [-] | [-] | [-] | [-] | [-] | -0.12915 | 0.012259 | 0.024518 | Criterion D | pass | pass | pass | pass | pass | pass | pass | pass | ? | pass
| Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O |  |  | -0.062173 | 0.077881 | -10.2503 | 164.8664 | 0.024257 | -0.0015082 | 0.51717 | 0.056008 | 0.0059184 | 0.0062862 | 0.15386 | 0.15857 | 0.36402 | Pi |  | 0.89593 | 0.99999 | 1 | 1 | 1 | 0.99779 | 1 | 0.48141 | 0.10664 | 0.98739 | [-] | [-] | [-] | [-] | [-]
[19 empty] | -0.087777 | 0.077758 | -14.5174 | 165.3902 | 0.025179 | -0.0022102 | 0.31424 | 0.043589 | 0.0058941 | 0.0062164 | 0.15355 | 0.15769 | 0.28388 | Criterion D |  | pass | pass | pass | pass | pass | pass | pass | pass | ? | pass | [-] | [-] | [-] | [-] | [-]
| 10000 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.052024 | 0.077652 | -8.6279 | 165.8441 | 0.024888 | -0.0012948 | 1.1578 | 0.083554 | 0.0058797 | 0.0061996 | 0.15336 | 0.15747 | 0.54483
| 7500 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.023316 | 0.077834 | -3.8486 | 165.068 | 0.026017 | -0.0006066 | 2.0896 | 0.11251 | 0.0059005 | 0.0062641 | 0.15363 | 0.15829 | 0.73236 |  |  | 1.3 | 2.3 | 3.3 | 4.3 | 5.3 | 6.3 | 7.3 | 8.3 | 9.3 | 10.3 | 11.3 | 12.3 | 13.3 | 14.3 | 15.3
| 5000 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | 0.069371 | 0.080813 | 10.6223 | 153.1229 | 0.035863 | 0.0024879 | 0.019288 | -0.011223 | 0.0062965 | 0.0068969 | 0.1587 | 0.1661 | -0.07072 |  | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
| 2500 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | 0.088777 | 0.080836 | 13.5858 | 153.0331 | 0.036132 | 0.0032076 | 0.016938 | -0.010521 | 0.0062984 | 0.0068493 | 0.15873 | 0.16552 | -0.066281 | xi (%) | 750 | -0.1515 | -0.12429 | -0.062173 | -0.10452 | -0.15 | -0.032024 | -0.09 | -0.081305 | -0.15857 | -0.051762 | [-] | [-] | [-] | [-] | [-]
| 1000 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | 0.11203 | 0.08099 | 17.0786 | 152.4517 | 0.037185 | 0.0041657 | 0.0069186 | 0.0067366 | 0.0063155 | 0.0070002 | 0.15894 | 0.16733 | 0.042384 | Ubase i , k=2 (%) | 750 | 0.05 | 0.063433 | 0.15 | 0.1 | 0.09 | 0.13 | 0.11 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-]
| 1000 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | 0.10804 | 0.080859 | 16.5248 | 152.9475 | 0.038802 | 0.0041922 | 0.14867 | 0.031177 | 0.0062845 | 0.0073611 | 0.15855 | 0.17159 | 0.19664 | UR i , k=2 (%) | 750 | 0.010719 | 0.0049378 | 0.012731 | 0.0020421 | 8.8789e-05 | 0.0036544 | 0.0024246 | 0.006953 | 0.0018473 | 9.8131e-18 | [-] | [-] | [-] | [-] | [-]
| 750 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | 0.092156 | 0.082241 | 13.6255 | 147.8524 | 0.044211 | 0.0040744 | 0.25244 | 0.04132 | 0.0064645 | 0.0071898 | 0.1608 | 0.16959 | 0.25696 | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 750 | 0.064922 | 0.075154 | 0.15576 | 0.10772 | 0.098489 | 0.13606 | 0.11707 | 0.050481 | 0.047206 | 0.1077 | [-] | [-] | [-] | [-] | [-]
| 500 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.024617 | 0.058725 | -7.1381 | 289.9686 | 0.13282 | -0.0032696 | 3.797 | -0.11443 | 0.0029906 | 0.0043803 | 0.10937 | 0.13237 | -1.0462 | Ucomp i / Ubase, i | 750 | 0.82822 | 0.63537 | 0.27985 | 0.40052 | 0.44445 | 0.30897 | 0.3643 | 1.3533 | 1.6017 | 0.4 | [-] | [-] | [-] | [-] | [-]
| 250 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | 0.0033333 | -0.058811 | [5 empty] | -0.2082 | 0.0039466 |  | 0.12564 |  | -1.6571 | di (%) | 750 | -0.03332 | -0.0061072 | 0.056008 | 0.013663 | -0.031819 | 0.086158 | 0.028181 | 0.036877 | -0.040391 | 0.06642 | [-] | [-] | [-] | [-] | [-]
| 100 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | 0.036667 | 0.059405 | 10.3902 | 283.3703 | 0.20783 | 0.0076203 | 1.4627 | -0.071847 | 0.0027955 | 0.0046139 | 0.10575 | 0.13585 | -0.67943 | U(di) (k=2, %) | 750 | 0.060219 | 0.071131 | 0.15386 | 0.10496 | 0.095454 | 0.13388 | 0.11453 | 0.04427 | 0.040495 | 0.10494 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| 100 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  | Lab D | -0.12779 | 0.052203 | -46.891 | 366.9479 | 0.047791 | -0.006107 | 0.15865 | 0.020793 | 0.0025949 | 0.0028375 | 0.10188 | 0.10654 | 0.20409 | En = di / U(di) | 750 | -0.55332 | -0.085858 | 0.36402 | 0.13018 | -0.33334 | 0.64353 | 0.24606 | 0.833 | -0.99743 | 0.63296 | [-] | [-] | [-] | [-] | [-] | 750 |  |  | 750 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| 75 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.10253 | 0.052216 | -37.6049 | 366.7756 | 0.04803 | -0.0049245 | 0.026179 | 0.0084484 | 0.0025955 | 0.0027974 | 0.10189 | 0.10578 | 0.082915 | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 50 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.094759 | 0.052316 | -34.6212 | 365.362 | 0.047825 | -0.0045318 | 0.0014889 | -0.0020187 | 0.0026061 | 0.002774 | 0.1021 | 0.10534 | -0.019772 | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 25 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.10661 | 0.052236 | -39.0728 | 366.4876 | 0.047321 | -0.0050451 | 0.0049375 | -0.0036705 | 0.0025995 | 0.0027684 | 0.10197 | 0.10523 | -0.035996 | di / Ubase i |  | -0.66641 | -0.096277 | 0.37339 | 0.13663 | -0.35354 | 0.66275 | 0.25619 | 1.2292 | -1.6157 | 0.6642 | [-] | [-] | [-] | [-] | [-] | -0.11818 | 0.01213 | 0.02426 | Criterion D | pass | pass | pass | pass | pass | pass | pass | ? | ? | pass
| 10 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.071106 | 0.052795 | -25.5106 | 358.769 | 0.049288 | -0.0035047 | 0.027621 | -0.0087743 | 0.0026499 | 0.0027117 | 0.10295 | 0.10415 | -0.085225 | Pi |  | 0.90192 | 1 | 1 | 1 | 1 | 0.99966 | 1 | 0.2688 | 0.095074 | 0.99538 | [-] | [-] | [-] | [-] | [-]
| 10 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.12018 | 0.053854 | -41.4389 | 344.8014 | 0.051816 | -0.0062274 | 0.027727 | 0.0089674 | 0.0027499 | 0.0030826 | 0.10488 | 0.11104 | 0.085501 | Criterion D |  | pass | pass | pass | pass | pass | pass | pass | ? | ? | pass | [-] | [-] | [-] | [-] | [-]
| 5 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.10452 | 0.053861 | -36.028 | 344.7037 | 0.050718 | -0.0053009 | 0.064345 | 0.013663 | 0.0027539 | 0.0031217 | 0.10496 | 0.11174 | 0.13018 |  |  | 1.35 | 2.35 | 3.35 | 4.35 | 5.35 | 6.35 | 7.35 | 8.35 | 9.35 | 10.35 | 11.35 | 12.35 | 13.35 | 14.35 | 15.35
| 2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | [-] | [-] | [-] | [-] | [-] |  |  | -0.1215 | 0.053882 | -41.8496 | 344.4411 | 0.052438 | -0.0063713 | 0.033525 | 0.0098656 | 0.002751 | 0.0030733 | 0.1049 | 0.11088 | 0.094048 |  | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
[19 empty] | -0.12771 | 0.054048 | -43.7185 | 342.3312 | 0.051373 | -0.0065608 | 0.021202 | 0.0078697 | 0.0027711 | 0.003091 | 0.10528 | 0.11119 | 0.074749 | xi (%) | 500 | -0.16122 | -0.14246 | -0.087777 | -0.1215 | -0.15 | -0.044826 | -0.12 | -0.095601 | -0.17128 | -0.064837 | [-] | [-] | [-] | [-] | [-]
| Table 5. Uncertainty of the reported value Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | [17 empty] | -0.13621 | 0.053959 | -46.7816 | 343.458 | 0.054134 | -0.0073734 | 4.9493e-05 | -0.00037961 | 0.002754 | 0.0031175 | 0.10496 | 0.11167 | -0.0036168 | Ubase i , k=2 (%) | 500 | 0.05 | 0.064 | 0.15 | 0.1 | 0.1 | 0.13 | 0.12 | 0.03 | 0.025 | 0.1 | [-] | [-] | [-] | [-] | [-]
| Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O |  |  | 0.19237 | 0.058321 | 56.5569 | 294.0015 | 0.068858 | 0.013246 | 3.6731 | 0.11177 | 0.0031671 | 0.0037675 | 0.11255 | 0.12276 | 0.99307 | UR i , k=2 (%) | 500 | 0.024735 | 0.010384 | 0.0092317 | 0.0036078 | 0.0057501 | 0.0024173 | 0.0026804 | 0.0069529 | 0.00056092 | 0.0066667 | [-] | [-] | [-] | [-] | [-]
[19 empty] | 0.19951 | 0.058311 | 58.6755 | 294.1004 | 0.069438 | 0.013853 | 2.9534 | 0.10021 | 0.0031641 | 0.003715 | 0.1125 | 0.1219 | 0.89076 | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 500 | 0.068643 | 0.076183 | 0.15552 | 0.10776 | 0.10786 | 0.13604 | 0.12652 | 0.050481 | 0.047173 | 0.10791 | [-] | [-] | [-] | [-] | [-]
| 10000 | 0.13817 | 0.070179 | 0.15298 | 0.10441 | 0.094871 | 0.13348 | 0.12373 | 0.042548 | 0.039133 | 0.1044 | [-] | [-] | [-] | [-] | [-] |  |  | 0.17497 | 0.058313 | 51.4557 | 294.0789 | 0.07173 | 0.012551 | 1.428 | 0.069683 | 0.0031565 | 0.0038412 | 0.11237 | 0.12395 | 0.62014 | Ucomp i / Ubase, i | 500 | 0.94059 | 0.64572 | 0.27368 | 0.40162 | 0.40411 | 0.30825 | 0.33408 | 1.3533 | 1.6002 | 0.40552 | [-] | [-] | [-] | [-] | [-]
| 7500 | 0.12748 | 0.069557 | 0.15325 | 0.10443 | 0.094877 | 0.13342 | 0.1237 | 0.043521 | 0.039095 | 0.1044 | [-] | [-] | [-] | [-] | [-] |  |  | 0.12014 | 0.058355 | 35.2809 | 293.6548 | 0.074498 | 0.0089505 | 0.55004 | 0.043279 | 0.0031517 | 0.0042283 | 0.11228 | 0.13005 | 0.38546 | di (%) | 500 | -0.029857 | -0.011097 | 0.043589 | 0.0098656 | -0.018634 | 0.086539 | 0.011366 | 0.035765 | -0.039912 | 0.066528 | [-] | [-] | [-] | [-] | [-]
| 5000 | 0.12741 | 0.069445 | 0.15311 | 0.10463 | 0.1044 | 0.13345 | 0.12372 | 0.042782 | 0.039055 | 0.1044 | [-] | [-] | [-] | [-] | [-] |  |  | 0.03001 | 0.058595 | 8.7406 | 291.2587 | 0.087093 | 0.0026136 | 0.12633 | -0.020826 | 0.0031343 | 0.0038597 | 0.11197 | 0.12425 | -0.186 | U(di) (k=2, %) | 500 | 0.064053 | 0.072075 | 0.15355 | 0.1049 | 0.105 | 0.13378 | 0.12409 | 0.044038 | 0.040204 | 0.10505 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| 2500 | 0.12811 | 0.069534 | 0.15299 | 0.10447 | 0.085501 | 0.13343 | 0.1237 | 0.04349 | 0.039053 | 0.10462 | [-] | [-] | [-] | [-] | [-] |  |  | 0.14548 | 0.063072 | 36.5715 | 251.3794 | 0.11514 | 0.016751 | 0.77906 | 0.05567 | 0.00352 | 0.0049097 | 0.11866 | 0.14014 | 0.46916 | En = di / U(di) | 500 | -0.46612 | -0.15397 | 0.28388 | 0.094048 | -0.17748 | 0.64688 | 0.091592 | 0.81213 | -0.99275 | 0.6333 | [-] | [-] | [-] | [-] | [-] | 500 |  |  | 500 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| 1000 | 0.12754 | 0.071011 | 0.15325 | 0.10559 | 0.095017 | 0.13342 | 0.1237 | 0.046369 | 0.039122 | 0.11402 | [-] | [-] | [-] | [-] | [-] |  |  | 0.20943 | 0.067426 | 46.067 | 219.961 | 0.1073 | 0.022472 | 0.00097295 | -0.0021032 | 0.0040584 | 0.0045642 | 0.12741 | 0.13512 | -0.016507 | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 1000 | 0.06845 | 0.07607 | 0.15557 | 0.10771 | 0.098498 | 0.13626 | 0.12741 | 0.050172 | 0.047238 | 0.10791 | [-] | [-] | [-] | [-] | [-] |  |  | 0.15055 | 0.12934 | 8.9991 | 59.7754 | 0.04384 | 0.0066 | 0.10562 | 0.042035 | 0.015996 | 0.017814 | 0.25295 | 0.26694 | 0.16618 | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 750 | 0.064922 | 0.075154 | 0.15576 | 0.10772 | 0.098489 | 0.13606 | 0.11707 | 0.050481 | 0.047206 | 0.1077 | [-] | [-] | [-] | [-] | [-] |  | Lab E | -0.18 | 0.047435 | -79.9961 | 444.4226 | 0.057881 | -0.010419 | 0.43876 | -0.031421 | 0.0021199 | 0.0023625 | 0.092084 | 0.09721 | -0.34122 | di / Ubase i |  | -0.59714 | -0.17339 | 0.29059 | 0.098656 | -0.18634 | 0.66569 | 0.094714 | 1.1922 | -1.5965 | 0.66528 | [-] | [-] | [-] | [-] | [-] | -0.13137 | 0.012339 | 0.024677 | Criterion D | pass | pass | pass | pass | pass | pass | pass | ? | ? | pass
| 500 | 0.068643 | 0.076183 | 0.15552 | 0.10776 | 0.10786 | 0.13604 | 0.12652 | 0.050481 | 0.047173 | 0.10791 | [-] | [-] | [-] | [-] | [-] |  |  | -0.14 | 0.047438 | -62.2114 | 444.3672 | 0.058191 | -0.0081467 | 0.37431 | -0.029023 | 0.0021194 | 0.0023213 | 0.092075 | 0.096359 | -0.31521 | Pi |  | 0.9396 | 0.99999 | 1 | 1 | 1 | 0.99954 | 1 | 0.30296 | 0.1058 | 0.99462 | [-] | [-] | [-] | [-] | [-]
| 250 | 0.06675 | 0.080748 | 0.1553 | 0.1081 | 0.089463 | 0.13612 | 0.12651 | 0.050173 | 0.047171 | 0.11705 | [-] | [-] | [-] | [-] | [-] |  |  | -0.13 | 0.052202 | -47.7059 | 366.9683 | 0.048035 | -0.0062446 | 0.50947 | -0.03726 | 0.0025941 | 0.002762 | 0.10187 | 0.10511 | -0.36578 | Criterion D |  | pass | pass | pass | pass | pass | pass | pass | ? | ? | pass | [-] | [-] | [-] | [-] | [-]
| 100 | 0.070067 | 0.092554 | 0.15567 | 0.10792 | 0.098717 | 0.13627 | 0.12649 | 0.050091 | 0.047228 | 0.11724 | [-] | [-] | [-] | [-] | [-] |  |  | -0.14 | 0.042751 | -76.6027 | 547.1621 | 0.07065 | -0.009891 | 0.75134 | -0.037056 | 0.0016985 | 0.0018674 | 0.082426 | 0.086426 | -0.44957 |  |  | 1.4 | 2.4 | 3.4 | 4.4 | 5.4 | 6.4 | 7.4 | 8.4 | 9.4 | 10.4 | 11.4 | 12.4 | 13.4 | 14.4 | 15.4
| 100 | 0.078728 | 0.1021 | 0.16163 | 0.11664 | 0.10852 | 0.14318 | 0.1253 | 0.070477 | 0.065002 | 0.12601 | [-] | [-] | [-] | [-] | [-] |  |  | -0.1 | 0.047509 | -44.3054 | 443.0542 | 0.060868 | -0.0060868 | 0.62865 | -0.037668 | 0.0021197 | 0.0021815 | 0.09208 | 0.093412 | -0.40908 |  | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
| 75 | 0.079934 | 0.10523 | 0.16167 | 0.11662 | 0.10879 | 0.14353 | 0.12534 | 0.070146 | 0.065004 | 0.1253 | [-] | [-] | [-] | [-] | [-] |  |  | -0.16 | 0.049249 | -65.9674 | 412.2959 | 0.061959 | -0.0099135 | 0.39241 | -0.030851 | 0.0022752 | 0.0026079 | 0.095397 | 0.10213 | -0.32339 | xi (%) | 250 | -0.15969 | -0.15709 | -0.052024 | -0.12771 | -0.17 | -0.058241 | -0.14 | -0.10021 | -0.16797 | -0.060436 | [-] | [-] | [-] | [-] | [-]
| 50 | 0.080801 | 0.11565 | 0.16198 | 0.11663 | 0.11703 | 0.1433 | 0.12534 | 0.070676 | 0.065004 | 0.12583 | [-] | [-] | [-] | [-] | [-] |  |  | -0.15 | 0.049244 | -61.8556 | 412.3708 | 0.060674 | -0.0091011 | 0.41749 | -0.031819 | 0.0022779 | 0.0026457 | 0.095454 | 0.10287 | -0.33334 | Ubase i , k=2 (%) | 250 | 0.05 | 0.069333 | 0.15 | 0.1 | 0.08 | 0.13 | 0.12 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-]
| 25 | 0.078893 | 0.1399 | 0.16172 | 0.11671 | 0.10427 | 0.14336 | 0.13418 | 0.077664 | 0.065018 | 0.12548 | [-] | [-] | [-] | [-] | [-] |  |  | -0.15 | 0.053928 | -51.5771 | 343.8475 | 0.052348 | -0.0078522 | 0.1194 | -0.018634 | 0.002756 | 0.0030783 | 0.105 | 0.11097 | -0.17748 | UR i , k=2 (%) | 250 | 0.018856 | 0.010633 | 0.0043627 | 0.0091973 | 0.0019118 | 0.0053189 | 0.0020495 | 0.0041589 | 0.00035371 | 0 | [-] | [-] | [-] | [-] | [-]
| 10 | 0.080723 | 0.23251 | 0.16448 | 0.11719 | 0.17415 | 0.14324 | 0.15241 | 0.070132 | 0.078139 | 0.12654 | [-] | [-] | [-] | [-] | [-] |  |  | -0.17 | 0.044732 | -84.9612 | 499.7717 | 0.075 | -0.01275 | 0.59217 | -0.034422 | 0.0018508 | 0.0021707 | 0.086043 | 0.093182 | -0.40006 | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 250 | 0.06675 | 0.080748 | 0.1553 | 0.1081 | 0.089463 | 0.13612 | 0.12651 | 0.050173 | 0.047171 | 0.11705 | [-] | [-] | [-] | [-] | [-]
| 10 | 0.087023 | 0.23564 | 0.11745 | 0.12614 | 0.17184 | [-] | -0.15232 | -0.080242 | 0.078321 | 0.12548 | [-] | [-] | [-] | [-] | [-] |  |  | -0.19 | 0.049358 | -77.9885 | 410.4657 | 0.064695 | -0.012292 | 1.2046 | -0.054172 | 0.0022786 | 0.0026422 | 0.09547 | 0.10281 | -0.56742 | Ucomp i / Ubase, i | 250 | 0.88443 | 0.59696 | 0.26825 | 0.41044 | 0.50057 | 0.3104 | 0.33377 | 1.3405 | 1.6001 | 0.36364 | [-] | [-] | [-] | [-] | [-]
| 5 | 0.087091 | 0.44998 | -0.11762 | 0.13485 | -0.18208 | [-] | 0.19927 | 0.11874 | 0.078533 | 0.12654 | [-] | [-] | [-] | [-] | [-] |  |  | 0.01 | 0.054259 | 3.3967 | 339.6694 | 0.079554 | 0.00079554 | 1.6928 | -0.070595 | 0.0027098 | 0.0033102 | 0.10411 | 0.11507 | -0.67806 | di (%) | 250 | -0.024112 | -0.021508 | 0.083554 | 0.0078697 | -0.034422 | 0.077337 | -0.0044222 | 0.035363 | -0.032392 | 0.075142 | [-] | [-] | [-] | [-] | [-]
| 2 | -0.091616 | 0.50531 | 0.11881 | 0.25868 | -0.22952 | [-] | 0.32564 | 0.074685 | -0.10133 | 0.12654 | [-] | [-] | [-] | [-] | [-] |  |  | 0.03 | 0.054397 | 10.1385 | 337.949 | 0.079791 | 0.0023937 | 1.6229 | -0.069297 | 0.0027229 | 0.0032738 | 0.10436 | 0.11443 | -0.664 | U(di) (k=2, %) | 250 | 0.062091 | 0.076941 | 0.15336 | 0.10528 | 0.086043 | 0.1339 | 0.12411 | 0.043784 | 0.040309 | 0.11445 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
[19 empty] | 0.03 | 0.058515 | 8.7616 | 292.0539 | 0.071236 | 0.0021371 | 1.6555 | -0.07529 | 0.0031801 | 0.0038648 | 0.11278 | 0.12433 | -0.66755 | En = di / U(di) | 250 | -0.38833 | -0.27954 | 0.54483 | 0.074749 | -0.40006 | 0.57759 | -0.03563 | 0.80768 | -0.80358 | 0.65653 | [-] | [-] | [-] | [-] | [-] | 250 |  |  | 250 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| Table 6. Ucomp i / Ubase i | [17 empty] | -0.02 | 0.052134 | -7.3584 | 367.9187 | 0.093338 | -0.0018668 | 3.4521 | -0.096865 | 0.0024643 | 0.0035409 | 0.099283 | 0.11901 | -0.97564 | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O |  |  | -0.02 | 0.087075 | -2.6378 | 131.8918 | 0.039439 | -0.00078878 | 0.6618 | -0.070836 | 0.007283 | 0.0080083 | 0.17068 | 0.17898 | -0.41502 | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
[19 empty] | -0.1 | 0.085919 | -13.5465 | 135.4647 | 0.062049 | -0.0062049 | 4.8807 | -0.18981 | 0.0069239 | 0.0083137 | 0.16642 | 0.18236 | -1.1406 | di / Ubase i |  | -0.48223 | -0.31021 | 0.55703 | 0.078697 | -0.43028 | 0.5949 | -0.036852 | 1.1788 | -1.2957 | 0.68311 | [-] | [-] | [-] | [-] | [-] | -0.13558 | 0.01225 | 0.0245 | Criterion D | pass | pass | pass | pass | pass | pass | pass | ? | ? | pass
| 10000 | 0.57075 | 0.51727 | 0.20036 | 0.30012 | 0.33342 | 0.23309 | 0.25124 | 1.0057 | 1.2043 | 0.3 | [-] | [-] | [-] | [-] | [-] |  |  | -0.19561 | -0.091042 | [5 empty] | -0.40715 | 0.0087764 |  | 0.18737 |  | -2.173 | Pi |  | 0.9789 | 0.99992 | 1 | 1 | 0.99983 | 0.99998 | 1 | 0.31319 | 0.25971 | 0.99616 | [-] | [-] | [-] | [-] | [-]
| 7500 | 0.35858 | 0.49518 | 0.2092 | 0.30097 | 0.33362 | 0.23095 | 0.25023 | 1.051 | 1.2023 | 0.3 | [-] | [-] | [-] | [-] | [-] |  |  | -0.43342 | -0.11476 | [5 empty] | -0.54194 | 0.013903 |  | 0.23582 |  | -2.2981 | Criterion D |  | pass | pass | pass | pass | pass | pass | pass | ? | ? | pass | [-] | [-] | [-] | [-] | [-]
| 5000 | 0.35686 | 0.49113 | 0.20476 | 0.3079 | 0.30002 | 0.23176 | 0.25079 | 1.0167 | 1.2002 | 0.3 | [-] | [-] | [-] | [-] | [-] |  | Lab F | -0.11061 | 0.066742 | -24.8308 | 224.4901 | 0.029237 | -0.0032339 | 0.32365 | 0.03797 | 0.0043243 | 0.0045669 | 0.13152 | 0.13516 | 0.2887 |  |  | 1.45 | 2.45 | 3.45 | 4.45 | 5.45 | 6.45 | 7.45 | 8.45 | 9.45 | 10.45 | 11.45 | 12.45 | 13.45 | 14.45 | 15.45
| 2500 | 0.37373 | 0.49436 | 0.20069 | 0.30239 | 0.37717 | 0.23137 | 0.25013 | 1.0496 | 1.2001 | 0.30732 | [-] | [-] | [-] | [-] | [-] |  |  | -0.092829 | 0.066711 | -20.8589 | 224.7017 | 0.029425 | -0.0027315 | 0.074001 | 0.018148 | 0.0043194 | 0.0045212 | 0.13144 | 0.13448 | 0.13806 |  | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
| 1000 | 0.35996 | 0.49399 | 0.20933 | 0.339 | 0.33852 | 0.23094 | 0.2502 | 1.1785 | 1.2037 | 0.27273 | [-] | [-] | [-] | [-] | [-] |  |  | -0.061248 | 0.066723 | -13.7577 | 224.6212 | 0.029402 | -0.0018008 | 0.22276 | 0.031492 | 0.004321 | 0.0044889 | 0.13147 | 0.134 | 0.23954 | xi (%) | 100 | -0.16608 | -0.16971 | -0.023316 | -0.13621 | -0.19 | -0.059378 | -0.09 | -0.10736 | -0.16181 | -0.075526 | [-] | [-] | [-] | [-] | [-]
| 1000 | 0.93497 | 0.6539 | 0.27497 | 0.40011 | 0.44469 | 0.31401 | 0.35693 | 1.3405 | 1.6032 | 0.40552 | [-] | [-] | [-] | [-] | [-] |  |  | -0.049939 | 0.066717 | -11.2193 | 224.6595 | 0.029008 | -0.0014486 | 0.63118 | 0.053004 | 0.0043221 | 0.0044909 | 0.13148 | 0.13403 | 0.40312 | Ubase i , k=2 (%) | 100 | 0.05 | 0.082667 | 0.15 | 0.1 | 0.09 | 0.13 | 0.12 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-]
| 750 | 0.82822 | 0.63537 | 0.27985 | 0.40052 | 0.44445 | 0.30897 | 0.3643 | 1.3533 | 1.6017 | 0.4 | [-] | [-] | [-] | [-] | [-] |  |  | -0.038821 | 0.066711 | -8.7231 | 224.7026 | 0.03087 | -0.0011984 | 0.12421 | 0.023511 | 0.0043129 | 0.0043747 | 0.13135 | 0.13228 | 0.179 | UR i , k=2 (%) | 100 | 0.028449 | 0.011508 | 0.011508 | 0.0068012 | 0.0067104 | 0.0083878 | 0.00098988 | 0.0030124 | 0.0023501 | 0.0066667 | [-] | [-] | [-] | [-] | [-]
| 500 | 0.94059 | 0.64572 | 0.27368 | 0.40162 | 0.40411 | 0.30825 | 0.33408 | 1.3533 | 1.6002 | 0.40552 | [-] | [-] | [-] | [-] | [-] |  |  | -0.036642 | 0.068129 | -7.8943 | 215.4427 | 0.032376 | -0.0011863 | 1.8437 | 0.092507 | 0.0044913 | 0.004824 | 0.13403 | 0.13891 | 0.69017 | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 100 | 0.070067 | 0.092554 | 0.15567 | 0.10792 | 0.098717 | 0.13627 | 0.12649 | 0.050091 | 0.047228 | 0.11724 | [-] | [-] | [-] | [-] | [-]
| 250 | 0.88443 | 0.59696 | 0.26825 | 0.41044 | 0.50057 | 0.3104 | 0.33377 | 1.3405 | 1.6001 | 0.36364 | [-] | [-] | [-] | [-] | [-] |  |  | -0.032024 | 0.068032 | -6.9191 | 216.0602 | 0.03179 | -0.001018 | 1.6038 | 0.086158 | 0.0044812 | 0.004849 | 0.13388 | 0.13927 | 0.64353 | Ucomp i / Ubase, i | 100 | 0.9817 | 0.5035 | 0.27748 | 0.40574 | 0.45066 | 0.31438 | 0.33344 | 1.3371 | 1.6028 | 0.36865 | [-] | [-] | [-] | [-] | [-]
| 100 | 0.9817 | 0.5035 | 0.27748 | 0.40574 | 0.45066 | 0.31438 | 0.33344 | 1.3371 | 1.6028 | 0.36865 | [-] | [-] | [-] | [-] | [-] |  |  | -0.044826 | 0.068018 | -9.6891 | 216.1479 | 0.032907 | -0.0014751 | 1.6187 | 0.086539 | 0.0044742 | 0.0047965 | 0.13378 | 0.13851 | 0.64688 | di (%) | 100 | -0.030248 | -0.033883 | 0.11251 | -0.00037961 | -0.054172 | 0.07645 | 0.045828 | 0.028466 | -0.025982 | 0.060302 | [-] | [-] | [-] | [-] | [-]
| 100 | 1.2162 | 0.75041 | 0.40126 | 0.60045 | 0.67368 | 0.46157 | 0.54546 | 2.1258 | 2.4001 | 0.55876 | [-] | [-] | [-] | [-] | [-] |  |  | -0.058241 | 0.068059 | -12.5734 | 215.8861 | 0.032398 | -0.0018869 | 1.2912 | 0.077337 | 0.004482 | 0.0048019 | 0.1339 | 0.13859 | 0.57759 | U(di) (k=2, %) | 100 | 0.065413 | 0.089083 | 0.15363 | 0.10496 | 0.09547 | 0.13394 | 0.12398 | 0.043343 | 0.040001 | 0.11452 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| 75 | 1.2473 | 0.71346 | 0.40211 | 0.60007 | 0.67915 | 0.4679 | 0.54624 | 2.1136 | 2.4002 | 0.54545 | [-] | [-] | [-] | [-] | [-] |  |  | -0.059378 | 0.068137 | -12.7898 | 215.3971 | 0.033949 | -0.0020159 | 1.2589 | 0.07645 | 0.004485 | 0.0048486 | 0.13394 | 0.13926 | 0.57078 | En = di / U(di) | 100 | -0.46242 | -0.38035 | 0.73236 | -0.0036168 | -0.56742 | 0.57078 | 0.36965 | 0.65677 | -0.64953 | 0.52658 | [-] | [-] | [-] | [-] | [-] | 100 |  |  | 100 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| 50 | 1.2695 | 0.6567 | 0.40758 | 0.60015 | 0.60796 | 0.46384 | 0.54615 | 2.1331 | 2.4002 | 0.55546 | [-] | [-] | [-] | [-] | [-] |  |  | 0.040187 | 0.07159 | 7.8412 | 195.1171 | 0.045698 | 0.0018365 | 0.31858 | -0.040407 | 0.0048909 | 0.0054913 | 0.13987 | 0.14821 | -0.28889 | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 25 | 1.2205 | 0.61418 | 0.40292 | 0.60178 | 0.83591 | 0.46494 | 0.50025 | 2.3879 | 2.4008 | 0.54881 | [-] | [-] | [-] | [-] | [-] |  |  | 0.08393 | 0.071763 | 16.2972 | 194.1759 | 0.045846 | 0.0038478 | 0.045853 | -0.015367 | 0.0049139 | 0.0054647 | 0.1402 | 0.14785 | -0.10961 | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 10 | 1.2675 | 0.43343 | 0.44989 | 0.61102 | 0.42975 | 0.46268 | 0.4302 | 2.1131 | 1.201 | 0.56854 | [-] | [-] | [-] | [-] | [-] |  |  | 0.086481 | 0.071652 | 16.8449 | 194.7807 | 0.04751 | 0.0041087 | 0.068903 | -0.018808 | 0.0048901 | 0.0055747 | 0.13986 | 0.14933 | -0.13448 | di / Ubase i |  | -0.60496 | -0.40987 | 0.75008 | -0.0037961 | -0.60191 | 0.58808 | 0.3819 | 0.94888 | -1.0393 | 0.5482 | [-] | [-] | [-] | [-] | [-] | -0.13583 | 0.012554 | 0.025109 | Criterion D | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 10 | 1.4245 | 0.28412 | 0.616 | 0.76891 | 0.39171 | [-] | 0.42865 | 2.4808 | 1.2057 | 0.54881 | [-] | [-] | [-] | [-] | [-] |  |  | 0.092929 | 0.071682 | 18.0855 | 194.6165 | 0.049373 | 0.0045882 | 0.05022 | 0.016064 | 0.0048846 | 0.0059612 | 0.13978 | 0.15442 | 0.11492 | Pi |  | 0.93236 | 0.99991 | 0.99699 | 1 | 0.99664 | 0.99998 | 1 | 0.52962 | 0.45298 | 0.99992 | [-] | [-] | [-] | [-] | [-]
| 5 | 1.4262 | 0.17386 | 0.61927 | 0.51269 | 0.54323 | [-] | 0.31609 | 3.8295 | 1.2112 | 0.56854 | [-] | [-] | [-] | [-] | [-] |  |  | 0.083396 | 0.07162 | 16.2584 | 194.9529 | 0.058296 | 0.0048616 | 0.20668 | 0.03256 | 0.0048304 | 0.0055558 | 0.139 | 0.14907 | 0.23424 | Criterion D |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-]
| 2 | 1.5354 | 0.32901 | 0.64155 | 0.26585 | 0.56298 | [-] | 0.18862 | 2.2798 | 0.77745 | 0.56854 | [-] | [-] | [-] | [-] | [-] | [17 empty] | 1.5 | 2.5 | 3.5 | 4.5 | 5.5 | 6.5 | 7.5 | 8.5 | 9.5 | 10.5 | 11.5 | 12.5 | 13.5 | 14.5 | 15.5
max (Ucomp/Ubase, i) |  | 1.5354 | 0.75041 | 0.64155 | 0.76891 | 0.83591 | 0.4679 | 0.54624 | 3.8295 | 2.4008 | 0.56854 | 0 | 0 | 0 | 0 | 0 | [16 empty] | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
[32 empty] | xi (%) | 100 | 0.064789 | 0.093756 | 0.069371 | 0.19237 | 0.01 | 0.040187 | 0.182 | 0.099323 | 0.037763 | 0.10243 | [-] | [-] | [-] | [-] | [-]
| Table 7. Degree of equivalence, di (%) | [16 empty] | Lab G | -0.25 | 0.061865 | -65.3212 | 261.2848 | 0.034029 | -0.0085073 | 2.6876 | -0.10142 | 0.003697 | 0.0039396 | 0.12161 | 0.12553 | -0.83401 | Ubase i , k=2 (%) | 100 | 0.05 | 0.081667 | 0.15 | 0.1 | 0.09 | 0.13 | 0.11 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-]
| Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O |  |  | -0.21 | 0.06185 | -54.8959 | 261.409 | 0.034232 | -0.0071888 | 2.5633 | -0.099023 | 0.0036945 | 0.0038963 | 0.12156 | 0.12484 | -0.81457 | UR i , k=2 (%) | 100 | 0.0099012 | 0.012479 | 0.0047755 | 0.0023183 | 0.0087268 | 0.00071574 | 0.00029136 | 0.02161 | 0.00050632 | 0.013333 | [-] | [-] | [-] | [-] | [-]
[19 empty] | -0.18 | 0.061858 | -47.0414 | 261.3408 | 0.034209 | -0.0061576 | 1.9899 | -0.08726 | 0.0036955 | 0.0038634 | 0.12158 | 0.12431 | -0.71771 | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 100 | 0.078728 | 0.1021 | 0.16163 | 0.11664 | 0.10852 | 0.14318 | 0.1253 | 0.070477 | 0.065002 | 0.12601 | [-] | [-] | [-] | [-] | [-]
| 10000 | -0.062091 | -0.0025235 | 0.028579 | 0.020793 | -0.031421 | 0.03797 | -0.10142 | 0.018012 | -0.01426 | 0.087016 | [-] | [-] | [-] | [-] | [-] |  |  | -0.13 | 0.061848 | -33.9849 | 261.4225 | 0.033755 | -0.0043882 | 0.19137 | -0.027056 | 0.0036961 | 0.003865 | 0.12159 | 0.12434 | -0.22252 | Ucomp i / Ubase, i | 100 | 1.2162 | 0.75041 | 0.40126 | 0.60045 | 0.67368 | 0.46157 | 0.54546 | 2.1258 | 2.4001 | 0.55876 | [-] | [-] | [-] | [-] | [-]
| 7500 | -0.042697 | -0.0053315 | 0.040977 | 0.0084484 | -0.029023 | 0.018148 | -0.099023 | 0.014922 | -0.0098294 | 0.091986 | [-] | [-] | [-] | [-] | [-] |  |  | -0.05 | 0.061849 | -13.0707 | 261.4136 | 0.035913 | -0.0017957 | 0.039754 | 0.012332 | 0.003688 | 0.0037497 | 0.12146 | 0.12247 | 0.10153 | di (%) | 100 | -0.015805 | 0.013162 | -0.011223 | 0.11177 | -0.070595 | -0.040407 | 0.10141 | 0.018729 | -0.042832 | 0.021834 | [-] | [-] | [-] | [-] | [-]
| 5000 | -0.025838 | -0.0031678 | 0.039407 | -0.0020187 | -0.03726 | 0.031492 | -0.08726 | 0.016488 | -0.013232 | 0.084691 | [-] | [-] | [-] | [-] | [-] |  |  | -0.11 | 0.063707 | -27.1027 | 246.3882 | 0.037027 | -0.0040729 | 0.090349 | 0.019149 | 0.0039084 | 0.0042411 | 0.12503 | 0.13025 | 0.15315 | U(di) (k=2, %) | 100 | 0.072534 | 0.097408 | 0.1587 | 0.11255 | 0.10411 | 0.13987 | 0.1215 | 0.063484 | 0.057345 | 0.12223 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| 2500 | -0.058804 | -0.003609 | 0.04961 | -0.0036705 | -0.037056 | 0.053004 | -0.027056 | 0.01783 | -0.016541 | 0.085645 | [-] | [-] | [-] | [-] | [-] |  |  | -0.09 | 0.058536 | -26.2661 | 291.8456 | 0.04294 | -0.0038646 | 0.23178 | 0.028181 | 0.0032793 | 0.0036471 | 0.11453 | 0.12078 | 0.24606 | En = di / U(di) | 100 | -0.2179 | 0.13512 | -0.07072 | 0.99307 | -0.67806 | -0.28889 | 0.83458 | 0.29502 | -0.74692 | 0.17862 | [-] | [-] | [-] | [-] | [-] | 100 |  |  | 100 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| 1000 | -0.069936 | 0.0063635 | 0.048998 | -0.0087743 | -0.037668 | 0.023511 | 0.012332 | 0.007536 | -0.0080853 | 0.072301 | [-] | [-] | [-] | [-] | [-] |  |  | -0.12 | 0.06326 | -29.9865 | 249.8878 | 0.038043 | -0.0045652 | 0.03228 | 0.011366 | 0.0038496 | 0.0041719 | 0.12409 | 0.12918 | 0.091592 | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 1000 | -0.033569 | 0.0086289 | 0.042813 | 0.0089674 | -0.030851 | 0.092507 | 0.019149 | 0.029971 | -0.039757 | 0.070563 | [-] | [-] | [-] | [-] | [-] |  |  | -0.14 | 0.063254 | -34.9908 | 249.9344 | 0.037507 | -0.005251 | 0.0048877 | -0.0044222 | 0.003851 | 0.0041709 | 0.12411 | 0.12916 | -0.03563 | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | ? | ? | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | ? | ? | pass
| 750 | -0.03332 | -0.0061072 | 0.056008 | 0.013663 | -0.031819 | 0.086158 | 0.028181 | 0.036877 | -0.040391 | 0.06642 | [-] | [-] | [-] | [-] | [-] |  |  | -0.09 | 0.063247 | -22.4986 | 249.9847 | 0.039401 | -0.0035461 | 0.52502 | 0.045828 | 0.0038426 | 0.0042062 | 0.12398 | 0.12971 | 0.36965 | di / Ubase i |  | -0.31611 | 0.16116 | -0.074822 | 1.1177 | -0.78438 | -0.31083 | 0.92187 | 0.62429 | -1.7133 | 0.19849 | [-] | [-] | [-] | [-] | [-] | 0.080595 | 0.015304 | 0.030608 | Criterion D | pass | pass | pass | ? | pass | pass | pass | pass | ? | pass
| 500 | -0.029857 | -0.011097 | 0.043589 | 0.0098656 | -0.018634 | 0.086539 | 0.011366 | 0.035765 | -0.039912 | 0.066528 | [-] | [-] | [-] | [-] | [-] |  |  | 0.182 | 0.06265 | 46.3692 | 254.7757 | 0.059671 | 0.01086 | 2.6199 | 0.10141 | 0.0036908 | 0.0042912 | 0.1215 | 0.13101 | 0.83458 | Pi |  | 0.98495 | 0.99999 | 1 | 0.18401 | 0.87499 | 1 | 0.66192 | 0.75635 | 0.11547 | 1 | [-] | [-] | [-] | [-] | [-]
| 250 | -0.024112 | -0.021508 | 0.083554 | 0.0078697 | -0.034422 | 0.077337 | -0.0044222 | 0.035363 | -0.032392 | 0.075142 | [-] | [-] | [-] | [-] | [-] |  |  | 0.21 | 0.062671 | 53.4677 | 254.6079 | 0.060114 | 0.012624 | 3.1203 | 0.1107 | 0.0036915 | 0.0042424 | 0.12152 | 0.13027 | 0.91102 | Criterion D |  | pass | pass | pass | ? | pass | pass | pass | pass | ? | pass | [-] | [-] | [-] | [-] | [-]
| 100 | -0.030248 | -0.033883 | 0.11251 | -0.00037961 | -0.054172 | 0.07645 | 0.045828 | 0.028466 | -0.025982 | 0.060302 | [-] | [-] | [-] | [-] | [-] |  |  | 0.22 | 0.062668 | 56.018 | 254.6273 | 0.062107 | 0.013664 | 3.3505 | 0.11471 | 0.0036834 | 0.0043681 | 0.12138 | 0.13218 | 0.94504 |  |  | 1.55 | 2.55 | 3.55 | 4.55 | 5.55 | 6.55 | 7.55 | 8.55 | 9.55 | 10.55 | 11.55 | 12.55 | 13.55 | 14.55 | 15.55
| 100 | -0.015805 | 0.013162 | -0.011223 | 0.11177 | -0.070595 | -0.040407 | 0.10141 | 0.018729 | -0.042832 | 0.021834 | [-] | [-] | [-] | [-] | [-] |  |  | 0.234 | 0.067089 | 51.9898 | 222.1785 | 0.056365 | 0.013189 | 5.4859 | 0.15713 | 0.0042472 | 0.0053238 | 0.13034 | 0.14593 | 1.2056 |  | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
| 75 | -0.021843 | -0.0032037 | -0.010521 | 0.10021 | -0.069297 | -0.015367 | 0.1107 | 0.019852 | -0.039315 | 0.024589 | [-] | [-] | [-] | [-] | [-] |  |  | 0.256 | 0.076203 | 44.0859 | 172.2106 | 0.051495 | 0.013183 | 7.2487 | 0.20516 | 0.0055078 | 0.0062332 | 0.14843 | 0.1579 | 1.3822 | xi (%) | 75 | 0.077454 | 0.096093 | 0.088777 | 0.19951 | 0.03 | 0.08393 | 0.21 | 0.11915 | 0.059982 | 0.12389 | [-] | [-] | [-] | [-] | [-]
| 50 | -0.039145 | -0.022552 | 0.0067366 | 0.069683 | -0.07529 | -0.018808 | 0.11471 | 0.052897 | -0.04068 | 0.0071347 | [-] | [-] | [-] | [-] | [-] |  |  | 0.352 | -0.07616 | [5 empty] | 0.26219 | 0.0062584 |  | 0.15822 |  | 1.6571 | Ubase i , k=2 (%) | 75 | 0.05 | 0.085667 | 0.15 | 0.1 | 0.09 | 0.13 | 0.11 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-]
| 25 | -0.041744 | -0.015485 | 0.031177 | 0.043279 | -0.096865 | 0.016064 | 0.15713 | 0.070681 | -0.040505 | 0.0061768 | [-] | [-] | [-] | [-] | [-] |  |  | 0.347 | 0.099633 | 34.9563 | 100.7385 | 0.049142 | 0.017052 | 1.8486 | 0.13546 | 0.0094389 | 0.0099447 | 0.19431 | 0.19945 | 0.69716 | UR i , k=2 (%) | 75 | 0.017012 | 0.011645 | 0.006175 | 0.00089244 | 0.011666 | 0.0099942 | 0.0032298 | 0.020505 | 0.00073995 | 0 | [-] | [-] | [-] | [-] | [-]
| 10 | -0.050524 | 0.049289 | 0.04132 | -0.020826 | -0.070836 | 0.03256 | 0.20516 | -0.0038594 | -0.0048622 | 0.0052498 | [-] | [-] | [-] | [-] | [-] |  |  | 0.204 | 0.16282 | 7.695 | 37.7205 | 0.027665 | 0.0056436 | 0.34392 | 0.095487 | 0.025777 | 0.027596 | 0.32111 | 0.33224 | 0.29737 | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 75 | 0.079934 | 0.10523 | 0.16167 | 0.11662 | 0.10879 | 0.14353 | 0.12534 | 0.070146 | 0.065004 | 0.1253 | [-] | [-] | [-] | [-] | [-]
| 10 | 0.00033452 | 0.080349 | -0.11443 | 0.05567 | -0.18981 | [-] | 0.26219 | -0.12479 | 0.057738 | 0.0050578 | [-] | [-] | [-] | [-] | [-] |  | Lab H | -0.13057 | 0.021274 | -288.4947 | 2209.541 | 0.28777 | -0.037573 | 0.71683 | 0.018012 | 0.00032234 | 0.00056493 | 0.035908 | 0.047536 | 0.50161 | Ucomp i / Ubase, i | 75 | 1.2473 | 0.71346 | 0.40211 | 0.60007 | 0.67915 | 0.4679 | 0.54624 | 2.1136 | 2.4002 | 0.54545 | [-] | [-] | [-] | [-] | [-]
| 5 | -0.010819 | 0.10279 | -0.2082 | -0.0021032 | -0.40715 | [-] | 0.13546 | -0.084643 | 0.04333 | -0.054427 | [-] | [-] | [-] | [-] | [-] |  |  | -0.096055 | 0.021761 | -202.8506 | 2111.815 | 0.27655 | -0.026564 | 0.47021 | 0.014922 | 0.00034257 | 0.00054442 | 0.037017 | 0.046666 | 0.4031 | di (%) | 75 | -0.021843 | -0.0032037 | -0.010521 | 0.10021 | -0.069297 | -0.015367 | 0.1107 | 0.019852 | -0.039315 | 0.024589 | [-] | [-] | [-] | [-] | [-]
| 2 | 0.11201 | 0.43621 | -0.071847 | 0.042035 | -0.54194 | [-] | 0.095487 | -0.036802 | 0.28428 | 0.13531 | [-] | [-] | [-] | [-] | [-] |  |  | -0.076252 | 0.021391 | -166.6475 | 2185.4723 | 0.28607 | -0.021814 | 0.5941 | 0.016488 | 0.00032667 | 0.00049452 | 0.036148 | 0.044475 | 0.45611 | U(di) (k=2, %) | 75 | 0.07379 | 0.10065 | 0.15873 | 0.1125 | 0.10436 | 0.1402 | 0.12152 | 0.063056 | 0.057281 | 0.12147 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
[19 empty] | -0.085114 | 0.021745 | -180.0001 | 2114.8178 | 0.27307 | -0.023242 | 0.67232 | 0.01783 | 0.00034373 | 0.00051261 | 0.03708 | 0.045282 | 0.48085 | En = di / U(di) | 75 | -0.29602 | -0.031831 | -0.066281 | 0.89076 | -0.664 | -0.10961 | 0.91102 | 0.31483 | -0.68635 | 0.20243 | [-] | [-] | [-] | [-] | [-] | 75 |  |  | 75 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| Table 8. Uncertainty of degree of equivalence, U(di) (k=2, %) | [17 empty] | -0.054796 | 0.023184 | -101.9419 | 1860.3998 | 0.25558 | -0.014005 | 0.10566 | 0.007536 | 0.00040014 | 0.00046191 | 0.040007 | 0.042984 | 0.18837 | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O |  |  | -0.099178 | 0.025086 | -157.5998 | 1589.0563 | 0.2388 | -0.023684 | 1.4274 | 0.029971 | 0.00047903 | 0.00081173 | 0.043773 | 0.056982 | 0.68469 | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | ? | ? | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | ? | ? | pass
[19 empty] | -0.081305 | 0.025241 | -127.6194 | 1569.6468 | 0.23095 | -0.018777 | 2.1346 | 0.036877 | 0.00048995 | 0.00085774 | 0.04427 | 0.058574 | 0.833 | di / Ubase i |  | -0.43687 | -0.037398 | -0.070137 | 1.0021 | -0.76997 | -0.11821 | 1.0064 | 0.66174 | -1.5726 | 0.22354 | [-] | [-] | [-] | [-] | [-] | 0.099297 | 0.015366 | 0.030731 | Criterion D | pass | pass | pass | pass | pass | pass | pass | pass | ? | pass
| 10000 | 0.13627 | 0.066364 | 0.15127 | 0.10188 | 0.092084 | 0.13152 | 0.12161 | 0.035908 | 0.031788 | 0.10188 | [-] | [-] | [-] | [-] | [-] |  |  | -0.095601 | 0.025241 | -150.0592 | 1569.6475 | 0.23897 | -0.022845 | 2.0078 | 0.035765 | 0.00048484 | 0.00080716 | 0.044038 | 0.056821 | 0.81213 | Pi |  | 0.96141 | 1 | 1 | 0.44274 | 0.89067 | 1 | 0.42503 | 0.73214 | 0.16746 | 1 | [-] | [-] | [-] | [-] | [-]
| 7500 | 0.12541 | 0.065684 | 0.15153 | 0.10189 | 0.092075 | 0.13144 | 0.12156 | 0.037017 | 0.031695 | 0.10186 | [-] | [-] | [-] | [-] | [-] |  |  | -0.10021 | 0.025086 | -159.2416 | 1589.0061 | 0.23846 | -0.023897 | 1.9871 | 0.035363 | 0.00047926 | 0.00079914 | 0.043784 | 0.056538 | 0.80768 | Criterion D |  | pass | pass | pass | pass | pass | pass | pass | pass | ? | pass | [-] | [-] | [-] | [-] | [-]
| 5000 | 0.12534 | 0.065567 | 0.15139 | 0.1021 | 0.10187 | 0.13147 | 0.12158 | 0.036148 | 0.03165 | 0.10186 | [-] | [-] | [-] | [-] | [-] |  |  | -0.10736 | 0.025045 | -171.1575 | 1594.2134 | 0.25127 | -0.026977 | 1.2918 | 0.028466 | 0.00046966 | 0.00083324 | 0.043343 | 0.057732 | 0.65677 |  |  | 1.6 | 2.6 | 3.6 | 4.6 | 5.6 | 6.6 | 7.6 | 8.6 | 9.6 | 10.6 | 11.6 | 12.6 | 13.6 | 14.6 | 15.6
| 2500 | 0.12607 | 0.065716 | 0.15129 | 0.10197 | 0.082426 | 0.13148 | 0.12159 | 0.03708 | 0.031759 | 0.10212 | [-] | [-] | [-] | [-] | [-] |  |  | 0.099323 | 0.035239 | 79.9863 | 805.3131 | 0.18861 | 0.018734 | 0.28247 | 0.018729 | 0.0010075 | 0.001608 | 0.063484 | 0.080199 | 0.29502 |  | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
| 1000 | 0.12536 | 0.06703 | 0.15145 | 0.10295 | 0.09208 | 0.13135 | 0.12146 | 0.040007 | 0.031321 | 0.11158 | [-] | [-] | [-] | [-] | [-] |  |  | 0.11915 | 0.035073 | 96.8605 | 812.9347 | 0.19194 | 0.022869 | 0.32038 | 0.019852 | 0.00099401 | 0.0015449 | 0.063056 | 0.07861 | 0.31483 | xi (%) | 50 | 0.066144 | 0.082737 | 0.11203 | 0.17497 | 0.03 | 0.086481 | 0.22 | 0.15819 | 0.064609 | 0.11242 | [-] | [-] | [-] | [-] | [-]
| 1000 | 0.063909 | 0.072011 | 0.15362 | 0.10488 | 0.095397 | 0.13403 | 0.12503 | 0.043773 | 0.040377 | 0.10509 | [-] | [-] | [-] | [-] | [-] |  |  | 0.15819 | 0.035338 | 126.6737 | 800.7862 | 0.19532 | 0.030898 | 2.2407 | 0.052897 | 0.0010049 | 0.0016895 | 0.063399 | 0.082207 | 0.83435 | Ubase i , k=2 (%) | 50 | 0.05 | 0.096667 | 0.15 | 0.1 | 0.1 | 0.13 | 0.11 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-]
| 750 | 0.060219 | 0.071131 | 0.15386 | 0.10496 | 0.095454 | 0.13388 | 0.11453 | 0.04427 | 0.040495 | 0.10494 | [-] | [-] | [-] | [-] | [-] |  |  | 0.14755 | 0.038832 | 97.8468 | 663.1614 | 0.16824 | 0.024823 | 3.313 | 0.070681 | 0.0012542 | 0.0023308 | 0.07083 | 0.096557 | 0.99789 | UR i , k=2 (%) | 50 | 0.020708 | 0.020731 | 0.01174 | 0.001338 | 0.0098033 | 0.005993 | 0.0030386 | 0.022251 | 0.00072837 | 0.011547 | [-] | [-] | [-] | [-] | [-]
| 500 | 0.064053 | 0.072075 | 0.15355 | 0.1049 | 0.105 | 0.13378 | 0.12409 | 0.044038 | 0.040204 | 0.10505 | [-] | [-] | [-] | [-] | [-] |  |  | 0.046977 | 0.035066 | 38.2039 | 813.2524 | 0.24318 | 0.011424 | 0.012113 | -0.0038594 | 0.00093061 | 0.001656 | 0.061012 | 0.081387 | -0.063256 | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 50 | 0.080801 | 0.11565 | 0.16198 | 0.11663 | 0.11703 | 0.1433 | 0.12534 | 0.070676 | 0.065004 | 0.12583 | [-] | [-] | [-] | [-] | [-]
| 250 | 0.062091 | 0.076941 | 0.15336 | 0.10528 | 0.086043 | 0.1339 | 0.12411 | 0.043784 | 0.040309 | 0.11445 | [-] | [-] | [-] | [-] | [-] |  |  | -0.034976 | -0.040121 | [5 empty] | -0.12479 | 0.0020678 |  | 0.090945 |  | -1.3721 | Ucomp i / Ubase, i | 50 | 1.2695 | 0.6567 | 0.40758 | 0.60015 | 0.60796 | 0.46384 | 0.54615 | 2.1331 | 2.4002 | 0.55546 | [-] | [-] | [-] | [-] | [-]
| 100 | 0.065413 | 0.089083 | 0.15363 | 0.10496 | 0.09547 | 0.13394 | 0.12398 | 0.043343 | 0.040001 | 0.11452 | [-] | [-] | [-] | [-] | [-] |  |  | 0.12689 | 0.059369 | 36.0017 | 283.718 | 0.1384 | 0.017562 | 2.0327 | -0.084643 | 0.0030368 | 0.0035426 | 0.11021 | 0.11904 | -0.76799 | di (%) | 50 | -0.039145 | -0.022552 | 0.0067366 | 0.069683 | -0.07529 | -0.018808 | 0.11471 | 0.052897 | -0.04068 | 0.0071347 | [-] | [-] | [-] | [-] | [-]
| 100 | 0.072534 | 0.097408 | 0.1587 | 0.11255 | 0.10411 | 0.13987 | 0.1215 | 0.063484 | 0.057345 | 0.12223 | [-] | [-] | [-] | [-] | [-] |  |  | 0.071712 | 0.037342 | 51.4265 | 717.1307 | 0.52595 | 0.037717 | 0.97127 | -0.036802 | 0.00066103 | 0.0024794 | 0.051421 | 0.099588 | -0.7157 | U(di) (k=2, %) | 50 | 0.074519 | 0.11135 | 0.15894 | 0.11237 | 0.11278 | 0.13986 | 0.12138 | 0.063399 | 0.057008 | 0.12189 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| 75 | 0.07379 | 0.10065 | 0.15873 | 0.1125 | 0.10436 | 0.1402 | 0.12152 | 0.063056 | 0.057281 | 0.12147 | [-] | [-] | [-] | [-] | [-] |  | Lab I | -0.16284 | 0.019567 | -425.3311 | 2611.9714 | 0.34018 | -0.055394 | 0.53111 | -0.01426 | 0.00025261 | 0.0004952 | 0.031788 | 0.044506 | -0.44859 | En = di / U(di) | 50 | -0.52531 | -0.20254 | 0.042384 | 0.62014 | -0.66755 | -0.13448 | 0.94504 | 0.83435 | -0.71359 | 0.058533 | [-] | [-] | [-] | [-] | [-] | 50 |  |  | 50 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| 50 | 0.074519 | 0.11135 | 0.15894 | 0.11237 | 0.11278 | 0.13986 | 0.12138 | 0.063399 | 0.057008 | 0.12189 | [-] | [-] | [-] | [-] | [-] |  |  | -0.12081 | 0.019547 | -316.1612 | 2617.0942 | 0.34272 | -0.041402 | 0.25285 | -0.0098294 | 0.00025115 | 0.000453 | 0.031695 | 0.042567 | -0.31012 | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | pass | pass | pass | pass
| 25 | 0.072175 | 0.13623 | 0.15855 | 0.11228 | 0.099283 | 0.13978 | 0.13034 | 0.07083 | 0.05668 | 0.12137 | [-] | [-] | [-] | [-] | [-] |  |  | -0.10597 | 0.019528 | -277.9028 | 2622.4098 | 0.34327 | -0.036377 | 0.45917 | -0.013232 | 0.00025043 | 0.00041828 | 0.03165 | 0.040904 | -0.41808 | Criterion B |  | pass | pass | pass | pass | pass | pass | pass | ? | ? | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | pass | ? | ? | pass
| 10 | 0.072939 | 0.22992 | 0.1608 | 0.11197 | 0.17068 | 0.139 | 0.14843 | 0.061012 | 0.070069 | 0.12172 | [-] | [-] | [-] | [-] | [-] |  |  | -0.11948 | 0.019526 | -313.377 | 2622.7356 | 0.33865 | -0.040463 | 0.71759 | -0.016541 | 0.00025216 | 0.00042104 | 0.031759 | 0.041039 | -0.52083 | di / Ubase i |  | -0.78291 | -0.2333 | 0.044911 | 0.69683 | -0.7529 | -0.14468 | 1.0428 | 1.7632 | -1.6272 | 0.064861 | [-] | [-] | [-] | [-] | [-] | 0.10529 | 0.015618 | 0.031236 | Criterion D | pass | pass | pass | pass | pass | pass | ? | ? | ? | pass
| 10 | 0.075768 | 0.23172 | 0.10937 | 0.11866 | 0.16642 | [-] | 0.15822 | 0.090945 | 0.06559 | 0.11795 | [-] | [-] | [-] | [-] | [-] |  |  | -0.070417 | 0.019561 | -184.034 | 2613.4856 | 0.35904 | -0.025283 | 0.17085 | -0.0080853 | 0.00024525 | 0.00030702 | 0.031321 | 0.035044 | -0.25815 | Pi |  | 0.73596 | 1 | 1 | 0.96509 | 0.92703 | 1 | 0.32903 | 0.06622 | 0.15007 | 1 | [-] | [-] | [-] | [-] | [-]
| 5 | 0.075057 | 0.44781 | 0.12564 | 0.12741 | 0.18737 | [-] | 0.19431 | 0.11021 | 0.064931 | 0.11857 | [-] | [-] | [-] | [-] | [-] |  |  | -0.16891 | 0.023619 | -302.7791 | 1792.5883 | 0.26939 | -0.045501 | 2.8334 | -0.039757 | 0.00040757 | 0.00074028 | 0.040377 | 0.054416 | -0.98464 | Criterion D |  | pass | pass | pass | pass | pass | pass | ? | ? | ? | pass | [-] | [-] | [-] | [-] | [-]
| 2 | 0.10643 | 0.5024 | 0.10575 | 0.25295 | 0.23582 | [-] | 0.32111 | 0.051421 | 0.1149 | 0.11436 | [-] | [-] | [-] | [-] | [-] |  |  | -0.15857 | 0.023603 | -284.6381 | 1794.9996 | 0.26411 | -0.04188 | 2.9285 | -0.040391 | 0.00040997 | 0.00077775 | 0.040495 | 0.055776 | -0.99743 |  |  | 1.65 | 2.65 | 3.65 | 4.65 | 5.65 | 6.65 | 7.65 | 8.65 | 9.65 | 10.65 | 11.65 | 12.65 | 13.65 | 14.65 | 15.65
[19 empty] | -0.17128 | 0.023587 | -307.8722 | 1797.4986 | 0.27365 | -0.046871 | 2.8634 | -0.039912 | 0.00040409 | 0.0007264 | 0.040204 | 0.053904 | -0.99275 |  | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
| Table 9. En values | [17 empty] | -0.16797 | 0.023586 | -301.9507 | 1797.6517 | 0.26977 | -0.045313 | 1.8861 | -0.032392 | 0.00040621 | 0.0007261 | 0.040309 | 0.053892 | -0.80358 | xi (%) | 25 | 0.035121 | 0.06138 | 0.10804 | 0.12014 | -0.02 | 0.092929 | 0.234 | 0.14755 | 0.03636 | 0.083042 | [-] | [-] | [-] | [-] | [-]
| Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O |  |  | -0.16181 | 0.023614 | -290.1739 | 1793.3014 | 0.28265 | -0.045735 | 1.2106 | -0.025982 | 0.00040002 | 0.0007636 | 0.040001 | 0.055267 | -0.64953 | Ubase i , k=2 (%) | 25 | 0.05 | 0.11921 | 0.15 | 0.1 | 0.08 | 0.13 | 0.12 | 0.03 | 0.025 | 0.11 | [-] | [-] | [-] | [-] | [-]
[19 empty] | 0.037763 | 0.032501 | 35.7493 | 946.6881 | 0.22172 | 0.0083729 | 1.7368 | -0.042832 | 0.0008221 | 0.0014225 | 0.057345 | 0.075433 | -0.74692 | UR i , k=2 (%) | 25 | 0.011138 | 0.041965 | 0.0072638 | 0.0046299 | 0.029529 | 0.0072966 | 0.0018811 | 0.039137 | 0.0015364 | 0.0066667 | [-] | [-] | [-] | [-] | [-]
| 10000 | -0.45565 | -0.038026 | 0.18893 | 0.20409 | -0.34122 | 0.2887 | -0.83401 | 0.50161 | -0.44859 | 0.85412 | [-] | [-] | [-] | [-] | [-] |  |  | 0.059982 | 0.032502 | 56.7802 | 946.6229 | 0.2235 | 0.013406 | 1.4632 | -0.039315 | 0.00082028 | 0.0013711 | 0.057281 | 0.074058 | -0.68635 | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 25 | 0.078893 | 0.1399 | 0.16172 | 0.11671 | 0.10427 | 0.14336 | 0.13418 | 0.077664 | 0.065018 | 0.12548 | [-] | [-] | [-] | [-] | [-]
| 7500 | -0.34046 | -0.081169 | 0.27042 | 0.082915 | -0.31521 | 0.13806 | -0.81457 | 0.4031 | -0.31012 | 0.90303 | [-] | [-] | [-] | [-] | [-] |  |  | 0.064609 | 0.032502 | 61.1608 | 946.6267 | 0.2309 | 0.014918 | 1.5666 | -0.04068 | 0.00081247 | 0.0014971 | 0.057008 | 0.077385 | -0.71359 | Ucomp i / Ubase, i | 25 | 1.2205 | 0.61418 | 0.40292 | 0.60178 | 0.83591 | 0.46494 | 0.50025 | 2.3879 | 2.4008 | 0.54881 | [-] | [-] | [-] | [-] | [-]
| 5000 | -0.20614 | -0.048313 | 0.26029 | -0.019772 | -0.36578 | 0.23954 | -0.71771 | 0.45611 | -0.41808 | 0.83141 | [-] | [-] | [-] | [-] | [-] |  |  | 0.03636 | 0.032509 | 34.4044 | 946.2169 | 0.24005 | 0.0087281 | 1.5524 | -0.040505 | 0.00080315 | 0.0018797 | 0.05668 | 0.086712 | -0.71463 | di (%) | 25 | -0.041744 | -0.015485 | 0.031177 | 0.043279 | -0.096865 | 0.016064 | 0.15713 | 0.070681 | -0.040505 | 0.0061768 | [-] | [-] | [-] | [-] | [-]
| 2500 | -0.46643 | -0.054918 | 0.32791 | -0.035996 | -0.44957 | 0.40312 | -0.22252 | 0.48085 | -0.52083 | 0.83869 | [-] | [-] | [-] | [-] | [-] |  |  | 0.045974 | 0.03907 | 30.1183 | 655.1196 | 0.1959 | 0.0090061 | 0.015488 | -0.0048622 | 0.0012274 | 0.0019528 | 0.070069 | 0.08838 | -0.069392 | U(di) (k=2, %) | 25 | 0.072175 | 0.13623 | 0.15855 | 0.11228 | 0.099283 | 0.13978 | 0.13034 | 0.07083 | 0.05668 | 0.12137 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| 1000 | -0.55786 | 0.094935 | 0.32353 | -0.085225 | -0.40908 | 0.179 | 0.10153 | 0.18837 | -0.25815 | 0.64796 | [-] | [-] | [-] | [-] | [-] |  |  | 0.14755 | 0.039161 | 96.2159 | 652.0814 | 0.29868 | 0.044071 | 2.1739 | 0.057738 | 0.0010755 | 0.0024652 | 0.06559 | 0.099303 | 0.8803 | En = di / U(di) | 25 | -0.57837 | -0.11367 | 0.19664 | 0.38546 | -0.97564 | 0.11492 | 1.2056 | 0.99789 | -0.71463 | 0.050894 | [-] | [-] | [-] | [-] | [-] | 25 |  |  | 25 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
| 1000 | -0.52527 | 0.11983 | 0.27869 | 0.085501 | -0.32339 | 0.69017 | 0.15315 | 0.68469 | -0.98464 | 0.67147 | [-] | [-] | [-] | [-] | [-] |  |  | 0.25487 | 0.039266 | 165.2995 | 648.5747 | 0.31639 | 0.080636 | 1.2177 | 0.04333 | 0.001054 | 0.0015598 | 0.064931 | 0.078989 | 0.66732 | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | X | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | X | pass | pass | pass
| 750 | -0.55332 | -0.085858 | 0.36402 | 0.13018 | -0.33334 | 0.64353 | 0.24606 | 0.833 | -0.99743 | 0.63296 | [-] | [-] | [-] | [-] | [-] |  |  | 0.39279 | -0.050666 | [5 empty] | 0.28428 | 0.0033005 |  | 0.1149 |  | 2.4741 | Criterion B |  | pass | pass | pass | pass | pass | pass | X | ? | ? | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | X | ? | ? | pass
| 500 | -0.46612 | -0.15397 | 0.28388 | 0.094048 | -0.17748 | 0.64688 | 0.091592 | 0.81213 | -0.99275 | 0.6333 | [-] | [-] | [-] | [-] | [-] |  | Lab J | -0.061563 | 0.052202 | -22.592 | 366.9725 | 0.047794 | -0.0029423 | 2.7786 | 0.087016 | 0.0025948 | 0.0028373 | 0.10188 | 0.10653 | 0.85412 | di / Ubase i |  | -0.83489 | -0.1299 | 0.20785 | 0.43279 | -1.2108 | 0.12357 | 1.3095 | 2.356 | -1.6202 | 0.056152 | [-] | [-] | [-] | [-] | [-] | 0.076865 | 0.015928 | 0.031855 | Criterion D | pass | pass | pass | pass | ? | pass | X | ? | ? | pass
| 250 | -0.38833 | -0.27954 | 0.54483 | 0.074749 | -0.40006 | 0.57759 | -0.03563 | 0.80768 | -0.80358 | 0.65653 | [-] | [-] | [-] | [-] | [-] |  |  | -0.018991 | 0.052202 | -6.9692 | 366.9725 | 0.048056 | -0.00091264 | 3.1051 | 0.091986 | 0.002594 | 0.0027959 | 0.10186 | 0.10575 | 0.90303 | Pi |  | 0.67562 | 1 | 1 | 0.9997 | 0.12315 | 1 | 0.0065274 | 0.004774 | 0.15745 | 1 | [-] | [-] | [-] | [-] | [-]
| 100 | -0.46242 | -0.38035 | 0.73236 | -0.0036168 | -0.56742 | 0.57078 | 0.36965 | 0.65677 | -0.64953 | 0.52658 | [-] | [-] | [-] | [-] | [-] |  |  | -0.0080488 | 0.052202 | -2.9537 | 366.9725 | 0.048036 | -0.00038663 | 2.6321 | 0.084691 | 0.0025941 | 0.0027619 | 0.10186 | 0.10511 | 0.83141 | Criterion D |  | pass | pass | pass | pass | ? | pass | X | ? | ? | pass | [-] | [-] | [-] | [-] | [-]
| 100 | -0.2179 | 0.13512 | -0.07072 | 0.99307 | -0.67806 | -0.28889 | 0.83458 | 0.29502 | -0.74692 | 0.17862 | [-] | [-] | [-] | [-] | [-] |  |  | -0.017299 | 0.052308 | -6.3224 | 365.4822 | 0.047191 | -0.00081635 | 2.6808 | 0.085645 | 0.002607 | 0.0027759 | 0.10212 | 0.10537 | 0.83869 |  |  | 1.7 | 2.7 | 3.7 | 4.7 | 5.7 | 6.7 | 7.7 | 8.7 | 9.7 | 10.7 | 11.7 | 12.7 | 13.7 | 14.7 | 15.7
| 75 | -0.29602 | -0.031831 | -0.066281 | 0.89076 | -0.664 | -0.10961 | 0.91102 | 0.31483 | -0.68635 | 0.20243 | [-] | [-] | [-] | [-] | [-] |  |  | 0.0099693 | 0.057009 | 3.0675 | 307.6923 | 0.042271 | 0.00042141 | 1.6084 | 0.072301 | 0.0031126 | 0.0031744 | 0.11158 | 0.11268 | 0.64796 |  | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
| 50 | -0.52531 | -0.20254 | 0.042384 | 0.62014 | -0.66755 | -0.13448 | 0.94504 | 0.83435 | -0.71359 | 0.058533 | [-] | [-] | [-] | [-] | [-] |  |  | -0.058587 | 0.053955 | -20.1252 | 343.5115 | 0.051622 | -0.0030244 | 1.7104 | 0.070563 | 0.0027608 | 0.0030935 | 0.10509 | 0.11124 | 0.67147 | xi (%) | 10 | 0.00031182 | 0.10012 | 0.092156 | 0.03001 | -0.02 | 0.083396 | 0.256 | 0.046977 | 0.045974 | 0.056086 | [-] | [-] | [-] | [-] | [-]
| 25 | -0.57837 | -0.11367 | 0.19664 | 0.38546 | -0.97564 | 0.11492 | 1.2056 | 0.99789 | -0.71463 | 0.050894 | [-] | [-] | [-] | [-] | [-] |  |  | -0.051762 | 0.053852 | -17.8488 | 344.8276 | 0.050736 | -0.0026262 | 1.5212 | 0.06642 | 0.0027529 | 0.0031207 | 0.10494 | 0.11173 | 0.63296 | Ubase i , k=2 (%) | 10 | 0.05 | 0.21333 | 0.15 | 0.1 | 0.16 | 0.13 | 0.14 | 0.03 | 0.05 | 0.11 | [-] | [-] | [-] | [-] | [-]
| 10 | -0.69269 | 0.21437 | 0.25696 | -0.186 | -0.41502 | 0.23424 | 1.3822 | -0.063256 | -0.069392 | 0.043131 | [-] | [-] | [-] | [-] | [-] |  |  | -0.064837 | 0.053955 | -22.2724 | 343.5115 | 0.052297 | -0.0033908 | 1.5204 | 0.066528 | 0.0027589 | 0.0030812 | 0.10505 | 0.11102 | 0.6333 | UR i , k=2 (%) | 10 | 0.020402 | 0.070354 | 0.030887 | 0.011554 | 0.033584 | 0.0042166 | 0.0052316 | 0.020458 | 0.002399 | 0.017638 | [-] | [-] | [-] | [-] | [-]
| 10 | 0.0044151 | 0.34675 | -1.0462 | 0.46916 | -1.1406 | [-] | 1.6571 | -1.3721 | 0.8803 | 0.042881 | [-] | [-] | [-] | [-] | [-] |  |  | -0.060436 | 0.058523 | -17.6455 | 291.9708 | 0.043816 | -0.002648 | 1.6486 | 0.075142 | 0.0032749 | 0.0035948 | 0.11445 | 0.11991 | 0.65653 | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 10 | 0.080723 | 0.23251 | 0.16448 | 0.11719 | 0.17415 | 0.14324 | 0.15241 | 0.070132 | 0.078139 | 0.12654 | [-] | [-] | [-] | [-] | [-]
| 5 | -0.14414 | 0.22955 | -1.6571 | -0.016507 | -2.173 | [-] | 0.69716 | -0.76799 | 0.66732 | -0.45901 | [-] | [-] | [-] | [-] | [-] |  |  | -0.075526 | 0.058618 | -21.9801 | 291.0267 | 0.04587 | -0.0034644 | 1.0583 | 0.060302 | 0.0032785 | 0.0036421 | 0.11452 | 0.1207 | 0.52658 | Ucomp i / Ubase, i | 10 | 1.2675 | 0.43343 | 0.44989 | 0.61102 | 0.42975 | 0.46268 | 0.4302 | 2.1131 | 1.201 | 0.56854 | [-] | [-] | [-] | [-] | [-]
| 2 | 1.0525 | 0.86825 | -0.67943 | 0.16618 | -2.2981 | [-] | 0.29737 | -0.7157 | 2.4741 | 1.1832 | [-] | [-] | [-] | [-] | [-] |  |  | 0.10243 | 0.063004 | 25.8042 | 251.9244 | 0.059003 | 0.0060436 | 0.12009 | 0.021834 | 0.0037352 | 0.0043356 | 0.12223 | 0.13169 | 0.17862 | di (%) | 10 | -0.050524 | 0.049289 | 0.04132 | -0.020826 | -0.070836 | 0.03256 | 0.20516 | -0.0038594 | -0.0048622 | 0.0052498 | [-] | [-] | [-] | [-] | [-]
[19 empty] | 0.12389 | 0.06265 | 31.5634 | 254.7771 | 0.060154 | 0.0074522 | 0.15405 | 0.024589 | 0.0036889 | 0.0042398 | 0.12147 | 0.13023 | 0.20243 | U(di) (k=2, %) | 10 | 0.072939 | 0.22992 | 0.1608 | 0.11197 | 0.17068 | 0.139 | 0.14843 | 0.061012 | 0.070069 | 0.12172 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
[19 empty] | 0.11242 | 0.062915 | 28.4019 | 252.6316 | 0.061621 | 0.0069277 | 0.01286 | 0.0071347 | 0.0037144 | 0.0043991 | 0.12189 | 0.13265 | 0.058533 | En = di / U(di) | 10 | -0.69269 | 0.21437 | 0.25696 | -0.186 | -0.41502 | 0.23424 | 1.3822 | -0.063256 | -0.069392 | 0.043131 | [-] | [-] | [-] | [-] | [-] | 10 |  |  | 10 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
[19 empty] | 0.083042 | 0.062738 | 21.0975 | 254.0579 | 0.064453 | 0.0053523 | 0.0096929 | 0.0061768 | 0.0036824 | 0.004759 | 0.12137 | 0.13797 | 0.050894 | Criterion A, En<=1? |  | pass | pass | pass | pass | pass | pass | X | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | pass | pass | pass | pass | X | pass | pass | pass
[19 empty] | 0.056086 | 0.063268 | 14.0117 | 249.8265 | 0.074704 | 0.0041898 | 0.0068854 | 0.0052498 | 0.0037038 | 0.0044291 | 0.12172 | 0.1331 | 0.043131 | Criterion B |  | pass | pass | pass | pass | pass | pass | X | ? | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | pass | pass | pass | pass | X | ? | pass | pass
[19 empty] | 0.094871 | 0.062738 | 24.1028 | 254.0579 | 0.11637 | 0.01104 | 0.0064993 | 0.0050578 | 0.0034781 | 0.0048678 | 0.11795 | 0.13954 | 0.042881 | di / Ubase i |  | -1.0105 | 0.23104 | 0.27547 | -0.20826 | -0.44273 | 0.25046 | 1.4655 | -0.12865 | -0.097245 | 0.047726 | [-] | [-] | [-] | [-] | [-] | 0.050836 | 0.017292 | 0.034585 | Criterion D | pass | pass | pass | pass | pass | pass | X | pass | pass | pass
[19 empty] | 0.15711 | 0.063268 | 39.2499 | 249.8265 | 0.12187 | 0.019147 | 0.74006 | -0.054427 | 0.003515 | 0.0040207 | 0.11857 | 0.12682 | -0.45901 | Pi |  | 0.46486 | 1 | 1 | 1 | 1 | 1 | 4.2395e-05 | 0.90294 | 0.99373 | 1 | [-] | [-] | [-] | [-] | [-]
[19 empty] | 0.24382 | 0.063268 | 60.9127 | 249.8265 | 0.18323 | 0.044674 | 4.5738 | 0.13531 | 0.0032694 | 0.0050878 | 0.11436 | 0.14266 | 1.1832 | Criterion D |  | pass | pass | pass | pass | pass | pass | X | pass | pass | pass | [-] | [-] | [-] | [-] | [-]
[18 empty] | Lab K | [15 empty] | 1.75 | 2.75 | 3.75 | 4.75 | 5.75 | 6.75 | 7.75 | 8.75 | 9.75 | 10.75 | 11.75 | 12.75 | 13.75 | 14.75 | 15.75
[33 empty] | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
[32 empty] | xi (%) | 10 | 0.090148 | 0.17016 | -0.024617 | 0.14548 | -0.1 | [-] | 0.352 | -0.034976 | 0.14755 | 0.094871 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Ubase i , k=2 (%) | 10 | 0.05 | 0.22667 | 0.1 | 0.1 | 0.16 | [-] | -0.14 | -0.03 | 0.05 | 0.11 | [-] | [-] | [-] | [-] | [-]
[32 empty] | UR i , k=2 (%) | 10 | 0.03838 | 0.023398 | 0.01395 | 0.048085 | 0.01811 | [-] | 0.0011262 | 0.044032 | 0.0058484 | 0.0066667 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 10 | 0.087023 | 0.23564 | 0.11745 | 0.12614 | 0.17184 | [-] | -0.15232 | -0.080242 | 0.078321 | 0.12548 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Ucomp i / Ubase, i | 10 | 1.4245 | 0.28412 | 0.616 | 0.76891 | 0.39171 | [-] | 0.42865 | 2.4808 | 1.2057 | 0.54881 | [-] | [-] | [-] | [-] | [-]
[32 empty] | di (%) | 10 | 0.00033452 | 0.080349 | -0.11443 | 0.05567 | -0.18981 | [-] | 0.26219 | -0.12479 | 0.057738 | 0.0050578 | [-] | [-] | [-] | [-] | [-]
[32 empty] | U(di) (k=2, %) | 10 | 0.075768 | 0.23172 | 0.10937 | 0.11866 | 0.16642 | [-] | 0.15822 | 0.090945 | 0.06559 | 0.11795 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
[32 empty] | En = di / U(di) | 10 | 0.0044151 | 0.34675 | -1.0462 | 0.46916 | -1.1406 | [-] | 1.6571 | -1.3721 | 0.8803 | 0.042881 | [-] | [-] | [-] | [-] | [-] | 10 |  |  | 10 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
[32 empty] | Criterion A, En<=1? |  | pass | pass | X | pass | X | [-] | X | X | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | X | pass | X |  | X | X | pass | pass
[32 empty] | Criterion B |  | pass | pass | X | pass | X | [-] | X | ? | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | X | pass | X |  | X | ? | pass | pass
[32 empty] | di / Ubase i |  | 0.0066905 | 0.35448 | -1.1443 | 0.5567 | -1.1863 | [-] | 1.8728 | -4.1596 | 1.1548 | 0.04598 | [-] | [-] | [-] | [-] | [-] | 0.089814 | 0.021402 | 0.042804 | Criterion D | pass | pass | X | pass | X |  | X | X | ? | pass
[32 empty] | Pi |  | 0.97793 | 1 | 0.2213 | 0.97602 | 0.061454 | [-] | 2.6093e-09 | 4.1547e-06 | 0.34151 | 1 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Criterion D |  | pass | pass | X | pass | X | [-] | X | X | ? | pass | [-] | [-] | [-] | [-] | [-]
[34 empty] | 1.8 | 2.8 | 3.8 | 4.8 | 5.8 | 6.8 | 7.8 | 8.8 | 9.8 | 10.8 | 11.8 | 12.8 | 13.8 | 14.8 | 15.8
[33 empty] | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
[18 empty] | Lab L | [13 empty] | xi (%) | 5 | 0.20072 | 0.31433 | 0.0033333 | 0.20943 | -0.19561 | [-] | 0.347 | 0.12689 | 0.25487 | 0.15711 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Ubase i , k=2 (%) | 5 | 0.05 | 0.44333 | -0.1 | 0.12 | -0.16 | [-] | 0.19 | 0.03 | 0.05 | 0.11 | [-] | [-] | [-] | [-] | [-]
[32 empty] | UR i , k=2 (%) | 5 | 0.038534 | 0.048385 | 0.015327 | 0.013603 | 0.062884 | [-] | 0.0026014 | 0.097972 | 0.0082079 | 0.017638 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 5 | 0.087091 | 0.44998 | -0.11762 | 0.13485 | -0.18208 | [-] | 0.19927 | 0.11874 | 0.078533 | 0.12654 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Ucomp i / Ubase, i | 5 | 1.4262 | 0.17386 | 0.61927 | 0.51269 | 0.54323 | [-] | 0.31609 | 3.8295 | 1.2112 | 0.56854 | [-] | [-] | [-] | [-] | [-]
[32 empty] | di (%) | 5 | -0.010819 | 0.10279 | -0.2082 | -0.0021032 | -0.40715 | [-] | 0.13546 | -0.084643 | 0.04333 | -0.054427 | [-] | [-] | [-] | [-] | [-]
[32 empty] | U(di) (k=2, %) | 5 | 0.075057 | 0.44781 | 0.12564 | 0.12741 | 0.18737 | [-] | 0.19431 | 0.11021 | 0.064931 | 0.11857 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
[32 empty] | En = di / U(di) | 5 | -0.14414 | 0.22955 | -1.6571 | -0.016507 | -2.173 | [-] | 0.69716 | -0.76799 | 0.66732 | -0.45901 | [-] | [-] | [-] | [-] | [-] | 5 |  |  | 5 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
[32 empty] | Criterion A, En<=1? |  | pass | pass | X | pass | X | [-] | pass | pass | pass | pass | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | pass | pass | X | pass | X |  | pass | pass | pass | pass
[32 empty] | Criterion B |  | pass | pass | X | pass | X | [-] | pass | ? | pass | pass | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | pass | pass | X | pass | X |  | pass | ? | pass | pass
[32 empty] | di / Ubase i |  | -0.21637 | 0.23187 | -2.082 | -0.017526 | -2.5447 | [-] | 0.71297 | -2.8214 | 0.8666 | -0.49479 | [-] | [-] | [-] | [-] | [-] | 0.21154 | 0.022087 | 0.044173 | Criterion D | pass | pass | X | pass | X |  | pass | ? | pass | pass
[32 empty] | Pi |  | 0.95468 | 1 | 3.0248e-07 | 1 | 4.4108e-30 | [-] | 0.98919 | 0.0061881 | 0.60127 | 0.99216 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Criterion D |  | pass | pass | X | pass | X | [-] | pass | ? | pass | pass | [-] | [-] | [-] | [-] | [-]
[34 empty] | 1.85 | 2.85 | 3.85 | 4.85 | 5.85 | 6.85 | 7.85 | 8.85 | 9.85 | 10.85 | 11.85 | 12.85 | 13.85 | 14.85 | 15.85
[33 empty] | Set Point | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J | Lab K | Lab L | Lab M | Lab N | Lab O
[32 empty] | xi (%) | 2 | 0.22053 | 0.54472 | 0.036667 | 0.15055 | -0.43342 | [-] | 0.204 | 0.071712 | 0.39279 | 0.24382 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Ubase i , k=2 (%) | 2 | -0.05 | 0.48 | 0.1 | 0.25 | -0.2 | [-] | 0.32 | 0.03 | -0.08 | 0.11 | [-] | [-] | [-] | [-] | [-]
[18 empty] | Lab M | [13 empty] | UR i , k=2 (%) | 2 | 0.04789 | 0.14608 | 0.022711 | 0.028587 | 0.095278 | [-] | 0.0065771 | 0.03283 | 0.01638 | 0.017638 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Ux i = RSS [Ubase i, UTS, UR i], k=2 (%). | 2 | -0.091616 | 0.50531 | 0.11881 | 0.25868 | -0.22952 | [-] | 0.32564 | 0.074685 | -0.10133 | 0.12654 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Ucomp i / Ubase, i | 2 | 1.5354 | 0.32901 | 0.64155 | 0.26585 | 0.56298 | [-] | 0.18862 | 2.2798 | 0.77745 | 0.56854 | [-] | [-] | [-] | [-] | [-]
[32 empty] | di (%) | 2 | 0.11201 | 0.43621 | -0.071847 | 0.042035 | -0.54194 | [-] | 0.095487 | -0.036802 | 0.28428 | 0.13531 | [-] | [-] | [-] | [-] | [-]
[32 empty] | U(di) (k=2, %) | 2 | 0.10643 | 0.5024 | 0.10575 | 0.25295 | 0.23582 | [-] | 0.32111 | 0.051421 | 0.1149 | 0.11436 | [-] | [-] | [-] | [-] | [-] |  |  |  | Lab | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
[32 empty] | En = di / U(di) | 2 | 1.0525 | 0.86825 | -0.67943 | 0.16618 | -2.2981 | [-] | 0.29737 | -0.7157 | 2.4741 | 1.1832 | [-] | [-] | [-] | [-] | [-] | 2 |  |  | 2 | Lab A | Lab B | Lab C | Lab D | Lab E | Lab F | Lab G | Lab H | Lab I | Lab J
[32 empty] | Criterion A, En<=1? |  | X | pass | pass | pass | X | [-] | pass | pass | X | X | [-] | [-] | [-] | [-] | [-] | xCRV | uCRV | UCRV | Criterion A, ¦Eni¦≤ 1 | X | pass | pass | pass | X |  | pass | pass | X | X
[32 empty] | Criterion B |  | X | pass | pass | pass | X | [-] | pass | ? | X | X | [-] | [-] | [-] | [-] | [-] | (%) | (k=1,%) | (k=2,%) | Criterion B | X | pass | pass | pass | X |  | pass | ? | X | X
[32 empty] | di / Ubase i |  | 2.2403 | 0.90877 | -0.71847 | 0.16814 | -2.7097 | [-] | 0.2984 | -1.2267 | 3.5535 | 1.2301 | [-] | [-] | [-] | [-] | [-] | 0.10851 | 0.027082 | 0.054163 | Criterion D | X | pass | pass | pass | X |  | pass | pass | X | X
[32 empty] | Pi |  | 0.0099879 | 0.89656 | 0.83289 | 1 | 1.1449e-37 | [-] | 1 | 0.38504 | 1.4544e-14 | 0.15487 | [-] | [-] | [-] | [-] | [-]
[32 empty] | Criterion D |  | X | pass | pass | pass | X | [-] | pass | pass | X | X | [-] | [-] | [-] | [-] | [-]
Lab N
Lab O
[20 empty] | Nominal Flow Rate |  | Reference Value |  |  | (Standard uncertainty)2 |  | Expanded uncertainty (k=2)
[18 empty] | CHIINV(0.05,n-1) |  | mL/min | Sum(1/u2(xi)) | xref | c2obs | maxChi | u2(xref) | u2cor(xref) | U(xref) | Ucor(xref)
[18 empty] | 16.919 | Okay! | 10000 | 7678.2431 | -0.14858 | 8.5878 | 2.7786 | 0.00013024 | 0.00037282 | 0.022824 | 0.038617
[18 empty] | 16.919 | Okay! | 7500 | 7636.3512 | -0.11098 | 7.6241 | 3.1051 | 0.00013095 | 0.0003328 | 0.022887 | 0.036486
[18 empty] | 16.919 | Okay! | 5000 | 7639.5919 | -0.09274 | 6.8468 | 2.6321 | 0.0001309 | 0.00029874 | 0.022882 | 0.034568
[18 empty] | 16.919 | Okay! | 2500 | 7744.6955 | -0.10294 | 6.9238 | 2.6808 | 0.00012912 | 0.000298 | 0.022726 | 0.034525
[18 empty] | 16.919 | Okay! | 1000 | 7278.9925 | -0.062332 | 4.349 | 1.6084 | 0.00013738 | 0.00019916 | 0.023442 | 0.028225
[18 empty] | 16.919 | Okay! | 1000 | 6654.3203 | -0.12915 | 9.6417 | 2.8334 | 0.00015028 | 0.00048298 | 0.024518 | 0.043954
[18 empty] | 16.919 | Okay! | 750 | 6796.531 | -0.11818 | 10.499 | 2.9285 | 0.00014713 | 0.00051492 | 0.02426 | 0.045384
[18 empty] | 16.919 | Okay! | 500 | 6568.5059 | -0.13137 | 9.3514 | 2.8634 | 0.00015224 | 0.00047456 | 0.024677 | 0.043569
[18 empty] | 16.919 | Okay! | 250 | 6663.6374 | -0.13558 | 9.3948 | 1.9871 | 0.00015007 | 0.00046996 | 0.0245 | 0.043357
[18 empty] | 16.919 | Okay! | 100 | 6344.6381 | -0.13583 | 9.9204 | 2.0896 | 0.00015761 | 0.0005212 | 0.025109 | 0.045659
[18 empty] | 16.919 | Okay! | 100 | 4269.6654 | 0.080595 | 10.6907 | 3.6731 | 0.00023421 | 0.00083462 | 0.030608 | 0.05778
[18 empty] | 16.919 | Okay! | 75 | 4235.4311 | 0.099297 | 9.9994 | 3.1203 | 0.0002361 | 0.00078696 | 0.030731 | 0.056106
[18 empty] | 16.919 | Okay! | 50 | 4099.789 | 0.10529 | 11.4209 | 3.3505 | 0.00024391 | 0.00092857 | 0.031236 | 0.060945
[18 empty] | 16.919 | Okay! | 25 | 3941.7825 | 0.076865 | 15.731 | 5.4859 | 0.00025369 | 0.0013303 | 0.031855 | 0.072946
[18 empty] | 16.919 | Okay! | 10 | 3344.2082 | 0.050836 | 10.2772 | 7.2487 | 0.00029902 | 0.0010244 | 0.034585 | 0.064012
[18 empty] | 12.5916 | Okay! | 10 | 2183.1815 | 0.089814 | 12.1022 | 4.8807 | 0.00045805 | 0.0018478 | 0.042804 | 0.085972
[18 empty] | 12.5916 | Okay! | 5 | 2049.9391 | 0.21154 | 6.1105 | 2.0327 | 0.00048782 | 0.00099361 | 0.044173 | 0.063043
[18 empty] | 11.0705 | Okay! | 2 | 1363.4887 | 0.10851 | 10.4381 | 4.5738 | 0.00073341 | 0.0025518 | 0.054163 | 0.10103
Note on AN3: jdwright: This values sets the separation of points in the plot of all results.
Note on B6: JDW: This table allows the user to enter transfer standard uncertainty components and sensitivities related to the transfer standard, e.g. calibration stability, property effects, sensitivity to temperature, etc. Column G is the RSS of those components.
Note on C7: JDW: Transfer standard uncertainty due to temperature sensitivity and the temperature differences between participants' labs.
Note on D7: JDW: Transfer standard uncertainty due to pressure sensitivity.
Note on E7: JDW: Transfer standard uncertainty due to other components, i.e. differences in fluid properties between participants' labs, leaks, etc.
Note on F7: JDW: Long term calibration stability of the transfer standard usually quantified by repeated calibrations performed in the pilot lab.
Note on G7: JDW: This is the root-sum-of-squares of transfer standard uncertainty components.
Note on AL17: Wright, John D. (Fed): Criterion A: Participant i passes if ¦〖En〗_i ¦≤1 and fails if ¦〖En〗_i ¦> 1.
Note on AL18: Wright, John D. (Fed): Criterion B: Participant i passes if u_comp⁄(u_(base i) )≤ R_th and ¦〖En〗_i ¦≤ 1, fails if ¦〖En〗_i ¦> 1, and the comparison results are inconclusive for participant i otherwise.
Note on AL20: Wright, John D. (Fed): = NORMDIST(NORM.INV(0.975, x_i, u_(base i) ), x_CRV,〖 u〗_(x_CRV ), TRUE) – NORMDIST(NORM.INV(0.025, x_i, u_(base i) ), x_CRV,u_(x_CRV ), TRUE).
Note on AL21: Wright, John D. (Fed): Criterion D: Participant i passes if content probability P_i≥ P_th and ¦〖En〗_i ¦≤ 1, fails if ¦〖En〗_i ¦> 1, and the comparison results are inconclusive for participant i otherwise.
Note on B28: JDW: Reported relative error of the TS, x (%), Error of the transfer standard relative to the participating laboratory's reference standard, in percent.
Note on T28: jdwright: The lab count is the number of participants at each set point that do not have discrepant results or dependent traceability
Note on C29: jdwright: Enter participating lab names in this yellow highlighted row and they will be automatically updated elsewhere.
Note on B50: JDW: k = 2 uncertainty, in %, of the participating lab's reference standard without uncertainty due to the best existing device (BED) or transfer standard reproducibility. Use a negative uncertainty value if 1) the participant's value is found discrepant by the CHi-squared test or 2) if this is a Key Comparison and the participant is not an MRA signatory or Designated Institute.
Note on U51: Don't drag cells in this part of the spreadsheet indiscriminately!
Note on AB51: jdwright: Conditional formatting highlights cells in this column red if they are a discrepant result. Such results should be removed from the CRV calculation by entering a negative uncertainty value in the appropriate cell in Table 2.
Note on W52: jdwright: Conditional formatting turns these cells red if the value has been removed by entering a negative uncertainty in Table 2 or -1 in Table4.
Note on Z52: jdwright: Beware of dragging cells!
Note on AB52: JDW: Don't drag this column's cells indiscriminately!
Note on AC52: JDW: Don't drag this column's cells indiscriminately! Error in Barbe's sheet corrected here: u2xref should not not squared again.
Note on AD52: JDW: Don't drag this column's cells indiscriminately!
Note on AE52: JDW: Don't drag this column's cells indiscriminately!
Note on B72: JDW: k = 2 reproducibility, in %, = (2 s) / (xbar sqrt (n)) 100 for the n repeated measurements of the transfer standard in the participant's lab.
Note on B94: jdwright: Enter 1 for labs that have independent traceability. Enter -1 for labs with dependent traceability.
Note on B116: JDW: Optional Input!!, k = 2 uncertainty of the reported value, in %, should be the RSS of base unc., transfer standard unc., and reproducibility in participant's lab… but it is better to input the values in tables 0, 2, and 3 and use the RSS formula here so you can see how big the values are in Table 5.
Note on B138: jdwright: Ucomp is the RSS of UTS and UR i measured in the participating lab.
Note on V325: jdwright: If this column says OOPs, the chi squared consistency check did not pass for one or more labs.