## CRM
JCTLM-DBWG/P-02-I-02
Checklist of frequently observed non-conformances by review teams with suggestions for a successful submission to JCTLM.
15194:2009 Paragraph / Clause | Clause subject | Guidance for the nomination submission form
4.1.4 | Quantity value | Clause 4.1.4 refers to the systematic format of quantity values reported in the supporting documentation, i.e. the certificate and certification report (or equivalent information). The submitter should consistently use a number of significant figures depending on the uncertainty (sub-clause 4.1.4.2). Examples are given in Clause 4.2 of 15194 or in 7.2.2 of "Guide to the expression of Uncertainty in Measurement" GUM JCGM 100:2008.
4.2 | Construction of systematic designations | Clause 4.2 gives examples of systematic format of name and value of the submitted CRM to be adopted in the supporting documentation, i.e. the certificate and certification report (or equivalent information). The submitter must include all the elements specified in clause 4.1, i.e. material, analyte(s) and kind-of-quantity.
5.1 | Hierarchical position | Clause 5.1 refers to the hierarchical position of the submitted certified reference material (CRM) in the reference measurement system. The submitter must state in the certificate if the CRM is a primary reference material in position m.1 in ISO 17511 (clause 5.2.4), a primary calibrator in position m.2 in ISO 17511 (clause 5.2.5), a secondary calibrator in position m.3 in ISO 17511 (clause 5.2.6) or an international conventional calibrator position m.3 in ISO 17511 (clause 5.5). When a CRM is intended for use as a trueness control for a reference measurement procedure (position p.3 in ISO 17511), the hierarchial position sholuld be stated as "not applicable because intended use is as a trueness control."
5.2 DBWG-P-04A | Extent-of-equivalence | Clause 5.2 requires a CRM to have metrological and commutability properties, allowing it to act as a higher metrological order measurement standard within a calibration hierarchy, or as a trueness control material of higher metrological order. JCTLM requires an extent-of-equivalence to existing CRMs listed in its database when such CRMs are listed. The JCTLM document DBWG-P-04A "Process for comparing certified values of the same measurand in multiple reference materials (CRMs)" refers to the JCTLM process for demonstrating the extent-of-equivalence of the value assigned to CRMs that are fit for the identical purpose. When a CRM under nomination is suitable for the same use as a CRM already listed in the JCTLM database, the submitter must supply a document showing the extent-of-equivalence data. The information in JCTLM procedure DBWG-P-04A must be included. Note that DBWG-P-04A clause 6.2 requires the extent-of-equivalence of CRMs to include considerations about the impact of a new CRM when used for its intended purpose in the calibration hierarchies of existing end-user measuring systems. Examples of possible ways to display comparison data demonstrating the extent-of-equivalence are given in attachments of DBWG-P-04A. When applicable, fill out the section titled “ Hyperlink to Final Report of extent-of-equivalence studies among listed RMs" on the RM Nomination Template.
5.3 | Production and characterization | The submitter should state the CRM is produced and characterized in accordance with ISO 17034 (replaced Guide 34) and ISO 33405 (replaced Guide 35).
6.1 | Supporting documentation | Clause 6.1 refers to the documentation to be provided together with a CRM. The submitter must supply a certificate consistent with a certification report (or equivalent information) also supplied. The essential items to be included in the certificate are listed in clause 6.3 and the mandatory elements to be included in the certification report (or equivalent information) are listed in Table 1 of clause 6.4.1. It is advisable to submit a picture of the CRM packaging with a clear display of the label.
6.2 | Label | Clause 6.2 refers to the label attached to the packaging of a CRM unit. The submitter should limit the information on the label to (i) name and identification code of the producer, (ii) name of the material, (iii) batch number (when available) which may include a unit number (unique for each item in a batch) which represents the filling order, and (iv) relevant health and safety warnings. It is advisable to submit a picture of the CRM packaging with a clear display of the label.
6.3 | Certificate | Clause 6.3 refers to the essential items to be included in the certificate of the submitted CRM. Particularly, the submitter must (i) declare the intended use of the CRM, i.e. if it is a calibrator and/or a control material, (ii) give information on its commutability (item d of clause 6.3) considering that a CRM shall have commutability properties supporting its function (Paragraph 5.2) and noting that commutability is not required for primary pure substance CRMs in position m.1 in ISO 17511; (iii) refer to the certification report or equivalent information as a source of the data reported in the certificate (item l of clause 6.3) in particular with respect to the intended use of the CRM (item d of clause 6.3) and (iv) state the measurement procedure(s) used to obtain property values (item j of clause 6.3), including full details where values are dependent on the measurement procedure(s), noting that the full details include the metrological traceability of the meaurement procedure used to value assign the CRM.
6.4.2 | Warning and safety precautions | Clause 6.4.2 refers to any danger associated with the CRM and its use. The submitter should describe all necessary precautions to follow considering typical regional, national and local legislation and regulations.
6.4.4 | Scope of application for the CRM | Clause 6.4.4 refers to limits of applicability of the CRM. The submitter must include (i) the measurement procedures for which the CRM is produced and intended for use, and those for which the CRM is known to be unsuitable, (ii) the influence quantities and (iii) any major pre-treatment which is not performed on biological samples, e.g. reconstitution of lyophilized material.
6.4.6 | General properties | Clause 6.4.6 refers to the general properties of the submitted CRM reported in the certification report (or equivalent information). Particularly, the submitter must (i) describe details of preparation of the starting material as necessary for its application as a CRM (sub-clause 6.4.6.3), (ii) specify compounds and contents, e.g. concentrations, of any additives to the starting material (sub-clause 6.4.6.4), (iii) clearly state the physical state of the CRM (sub-clause 6.4.6.5), examples of preferred format to adopt are given in clause 4.1.1, (iv) state the within-sample and between-sample homogeneity, including the sample size on which the homogeneity has been established and the minimum sample size for use (sub-clause 6.4.6.6) and (v) state the stability of the CRM in storage conditions, e.g. temperature, humidity and light, based on collected experimental data (sub-clause 6.4.6.10).
6.4.7 | Specific properties | Clause 6.4.7 refers to the specific properties of the submitted CRM reported in the certification report (or equivalent information). Particularly, the submitter must (i) state, in case of a "pure" CRM, the purity of the main component as mass fraction, volume fraction, amount-of-substance fraction or number fraction (sub-clause 6.4.7.4), (ii) describe the matrix of the material (sub-clause 6.4.7.5), especially those components that influence the certified property value, (iii) supply data for the extent of commutability (sub-clause 6.4.7.7), assessed according to the intended use of the CRM, (iv) use SI units wherever possible and appropriate (sub-clause 6.4.7.9), (v) express the measurement uncertainty of the certified property value (item h of the certificate) either as a combined standard uncertainty or as an interval derived from an expanded uncertainty with a stated level of confidence (sub-clause 6.4.7.10) noting that the uncertainty components caused by (a) inhomogeneity, (b) instability and (c) value assignment should be given according to the data supplied in sub-clauses 6.4.8.3, 6.4.8.5 and 6.4.8.6, respectively, of the certification report (or equivalent information) and (vi) describe in detail the metrological traceability of values assigned to the analytes (sub-clause 6.4.7.11).
6.4.7.7 | Commutability | Clause 6.4.7.7 requires the commutablity of the submitted CRMs on routine systems should be comprehensively evaluated. The evaluation of commutablity must (i) be performed according to the CLSI EP30, 2nd edition, or IFCC Working Group Recommendations for Assessing Commutability, (ii) include major manufacturer’s measurement procedures as much as possible and (iii) include the data for the extent of commutablity. When a CRM is value assigned using a RMP, that RMP should be included in the commutability assessment. Note that clause 5.3 states the suitability of the CRM for calibration or trueness control purposes shall be assessed by evaluating its commutability, such that it is consistent with its intended use as stated in 6.4.4 and 6.4.9. Relevant information shall be given on the certificate or in supporting documentation. Note that primary pure substance CRMs (position m.1 in ISO 17511) do not require commutability assessment. Calibrators prepared from pure substance CRMs do not require commutability assessment when they are used as calibrators (position m.2 in ISO 17511) for a RMP (position p.3 in ISO 17511) assuming the RMP is highly selective for the measurand and not influenced by the matrix of the solution. A justification for the measurement procedures included in a commutability assessment should be provided. Note that the IFCC guidelines for commutability assessment recommend that measurement procedures representative of commonly used end-user procedures or technologies be included in commutability assessment for a reference material.
6.4.8 | Characterization | Clause 6.4.8 refers to the characterization of the submitted CRM reported in the certification report (or equivalent information). Note that characterization is used to mean value assignment. Particularly, the submitter must (i) use only technically valid measurement procedures suitable to characterize the CRM (sub-clause 6.4.8.1) and describe the measurement procedures used, (ii) describe the homogeneity study that was undertaken, evaluated and reported in accordance with the ISO 33405 (replaced Guide 35) (sub-clause 6.4.8.3); the collected experimental data and adopted measurement procedures must be included, possibly in annexes of the certification report (item a of clause 6.4.13), (iii) make a statistical evaluation of the homogeneity data (sub-clause 6.4.8.4); adopted methods should be described and preferably applied in agreement with the ISO 33405 (replaced Guide 35), (iv) describe the stability study that was undertaken, evaluated and reported in accordance with the ISO 17034 and ISO 33405 (replaced Guide 35) (sub-clause 6.4.8.5); collected experimental data and adopted measurement procedures must be included, possibly in annexes of the certification report (item b of clause 6.4.13), (v) describe the experimental plan and all measurement procedures used to assign the value to the analyte (sub-clause 6.4.8.6); a list of the laboratories participating in value assignment should be provided; when available, a RMP listed in the JCTLM database must be used to value-assign the nominated measurand; if an unlisted measurement procedure was used to value-assign the nominated measurand, the submitter is advised to nominate the measurement procedure as a RMP before or at the same time of the CRM nomination, (vi) report the quantity value and its measurement uncertainty based on an uncertainty budget for each measurement procedure participating in value assignment (sub-clause 6.4.8.7) noting that the uncertainty of the calibrators used for measurement procedures must be included in the combined uncertainty of the value assigned. Additional information to be supplied is listed in sub-clause 6.4.8.6, e.g. number of replicates within-run and number of runs; collected experimental data must be included, possibly in annexes of the certification report (item c of clause 6.4.13).
6.4.8.3 | Assessment of homogeneity | Clause 6.4.8.3 refers to how the homogeneity study is described. The submitter must state (i)sampling schemes including how many units were used for homogeneity testing and how these units were picked (ii)measurement scheme including the details of measurement methods used for homogeneity and the repeatability conditions, (iii) statistical analysis including between-bottle homogeneity study and within-bottle homogeneity study (when the units of a CRM can be subsampled) and (iv) provide the raw data (Both graphs and tables are accpetable). The raw data on homogeneity can be included in Annexes (note a) of sub-clause 6.4.13.
6.4.8.5 | Assessment of Stability | Clause 6.4.8.5 refers to how the stability study is described. The submitter must state (i) the short term stability at stated transfer conditions, (ii) the long term stability at given storage conditions and (ii) the effect of freeze-thaw or repeated use if CRMs are not for single use. In addition, the submitter must provide the details of measurement methods, the raw data (Both graphs and tables are accpetable) and statistics results used for the stability study. The raw data on stability can be included in Annexes (note b) of sub-clause 6.4.13.
6.4.8.6 | Value assignment | Clause 6.4.8.6 requires comprehensive description of value assignment. The submitter must provide sufficient evidence to demonstrate the accuracy of measurement procedures used for certified value assignment. If available, the results from interlaboratory comparison studies (CCQM Key comparisons, RELA, national study, etc.) should be provided. When the certified values are the consensus values from multiple laboratories, the measurement procedures and calibration traceability information for each of the laboratories participating in the value assignment shall be stated.
6.4.8.7 | Quantity value and measurement uncertainty | The overall combined uncertainty of the value assigned to the CRM must include, at a minimum, the variability from long term stability, short term stability, homogeneity, and value assignment. Note that the uncertainty of the calibrator(s) must be included in the combined uncertainty of the assigned value. A detailed uncertainty budget and an explanation of each individual component of the combined uncertainty must be provided in the certification report. Particularly, when the certified values are the consensus values from multiple laboratories, the uncertainty budget of value assignment from each laboratory must be reported.
6.4.9 | Intended use | Clause 6.4.9 refers to the intended use of the CRM. The submitter must state the intended use of the CRM, i.e. if it is suitable as a calibration material and/or a control material. The description in the "Scope" and "Intended Use" sections of the certification report (or equivalent information) must be in accordance with the Certificate of the CRM. Particularly, the differences between the natural/endogeneous materials and the artificial/recombinant materials must be taken into consideration. Known limitations based on the commutability assessment (sub-clause 6.4.7.7) shall be documented. Information on CRM function and applicability limits must be included in item d of the certificate.
6.4.14 | Dates of authorization and revision | Clause 6.4.14 refers to the dates for the submitted version of the certification report (or equivalent information). The submitter must provide this information.
General requirement |  | When citing a publication for any technical descriptions or information, the submitter shall carefully check that the publication is correct and complete, and any differences between a publication and the actual CRM submited are fully described. When a measurement procedure used for value assignment is modified from what is in a supporting publication, the submitters is encouraged to submit a SOP or similar document to describe any methodological information or performance metrics which are not fully described in the publication. The layout of this document should be aligned to the clauses of ISO 15193.

## RMP
JCTLM-DBWG/P-02-I-02
Checklist of frequently observed non-conformances by review teams with suggestions for a successful submission to JCTLM.
15193:2009 Paragraph / Clause | Clause subject | Guidance for the nomination submission form
4.3 | Introduction | Clause 4.3 refers to the introduction which is a general overview of the RMP. Ensure that sufficient detail is provided for clarity regarding what is being measured by the RMP. For example, the mutation identified is required for genetic tests.
4.4 | Scope | Clause 4.4 requires the scope of the RMP to be clearly stated and any limitations for using the RMP to also be stated. The expected position of the RMP in a calibration hierarchy should be stated, e.g. primary RMP (p.2 in ISO 17511), secondary RMP (p.3 in ISO 17511). The scope should clearly identify the measurand and the intended use of the RMP. The scope must state the sample material/matrix for which the RMP is suitable for use and any required sample preservation or pretreatment. The scope must state if the calibration of the RMP is SI traceable or not. The scope must identify known interfering substances, including isoforms or other derivative molecules, drugs, metabolites, etc. See clause 4.14.6 where details of analytical influence quantities that must be provided are described. Any modifications that are allowed to eliminate or minimize the influence of interferng substances must be described in a clause “Special cases” (see 4.15). The measurement interval must be stated.
4.6 | Measurement principle and measurement method | Clause 4.6 refers to the description of the measurement principle and method choosen which must be described in sufficient detail that there is no ambiguity in the procedure. When signals from potentially interfering substances are possible, e.g. co-eluting substances, absence of these substances must be confirmed or a process for correction applied. When referring to a publication for these technical descriptions, the submitter shall carefully check that the cited publication is correct and complete, and any differences between a publication and the actual measurement procedure submitted as the RMP are fully described; e.g. the measurement interval, number or concentrations of calibrators. Submitters are encouraged to submit an SOP or similar document to describe any methodological information or performance metrics which are not fully described in the publication. The layout of this document should be aligned to the clauses of ISO 15193. The submitter should include SOP details that describe differences to a publication that will be provided to end-users in the "Comments to be published on the JCTLM DB listing” field in the nomination form.
4.11 | Preparation of measuring system and analytical portion | Clause 4.11 refers to the steps needed before the actual measurement. When referring to a publication for these technical descriptions, the submitter shall carefully check that the publication is correct and complete, and any differences between a publication and the actual measurement procedure submited as the RMP are fully described.
4.12 | Operation of measuring system | Clause 4.12 refers to the measurement process and validation of its results. The information must be in sufficient detail that the procedure is fully described. For example, the purity and suitability of calibrators must be stated; the sample types must be stated (e.g. serum, plasma, etc.); stability and storage conditions of samples, calibrators and reagents must be stated; and other technical details provided as appropriate. Calibrators must be suitable for the matrix of the samples and their preparation clearly described. When referring to a publication for these technical descriptions, the submitter shall carefully check that the cited publication is correct and complete, and any differences between a publication and the actual measurement procedure submitted as the RMP are fully described.
4.13 | Data processing | Clause 4.13 refers to processing of the measurement data to produce a reportable result. A clear description of the measured signal, calculations, including used formula(s), rounding, conversion equations used, if appropriate, number of significant digits, and other relevant details must be provided. Calculation of measurement uncertainty must be provided (see 4.14.9).
4.13.3 | Comparison with measurement results obtained by other measurement procedures | Clause 4.13.3 refers to validation that the RMP is suitable for all sample types claimed, e.g. serum and plasma, or serum and CSF, etc. Clause 4.13.3 also requires results to be compared to other measurment procedures which may have alternative measurement principles and may include measurement procedures used by medical laboratories. Results from such alternative measurement procedures are intended to be informative and to demonstrate that at least similar results are obtained. Also clarify if the RMP is intended for use to develop a standardization scheme to reduce the disagreement of results from medical laboratory procedures. Also see 4.16.
4.14.3 | Analytical sensitivity | Clause 4.14.3 is the analytical sensitivity typically expressed as the change in signal for a stated change in concentration. It is not the detection limit, nor the lower and higher measurement limits which are reported in separate clauses. The submisison must provide the analytical sensitivity and explain how the analytical sensitivity was determined.
4.14.6 | Analytical influence quantities | Clause 4.14.6 requires that quantitative information on the influence of all anticipated interferences or other influence quantities be provided. The substances examined must be appropriate for what is likely to be encountered in a clinical sample for which the measurand is measured. The concentrations examined and the quantitative influence must be provided. An assessment must be provided that the RMP is fit for purpose when there are relevant influence quantities present; such influence quantities must be identified. Note that the scope, item c of clause 4.4, must identify interfering substances and if modifications are permitted to eliminate or minimize the influence of interfering substances. In Nucleic Acid Testing/ genetic testing, analytical influence quantities are often described as the specificity or cross-reactivity of the method (when describing molecular approaches such as PCR). The cross-reactivity of the RMP against homologous sequences to the genetic target of interest should be considered which are relevant to the genome and/or clinical matrix. For example, gene isoforms/pseudogenes within the same genome or viruses/bacteria present in the clinical specimen type. Evidence of a lack of cross-reactivity should be presented based on both in silico (e.g. PrimerBlast) and wet-lab testing (e.g. purified DNA of alternative bacteria species or clinical samples which are negative for the species of interest).
4.14.8 | Recovery measurement | Clause 4.14.8 is the recovery of measurand from samples and must identify and include all pre-measurement and measurement steps. The submission must explain how the recovery was determined and explain if any of the measurand was lost in preparatory steps.
4.14.9 | Measurement uncertainty | Clause 4.14.9 requires the combined uncertainty of the measured value from the RMP to be determined. The uncertainty must include (where possible) the influence of all sources of variability that contribute in a meaningful way to the combined uncertainty. The uncertainty of the value assigned to the calibrator and the uncertainty of the measurement process are the two major contributors. The experimental design to estimate uncertainty must be described in sufficient detail to demonstrate that all meaningful contributors to the uncertainty of the measured value were captured. The combined uncertainty of the measured value must be given as standard uncertainty and expanded uncertainty with a stated coverage factor. The combined uncertainty of the measured value by the RMP must include the uncertainty of calibrators used for the RMP. If the uncertainty varies over the measuring interval, it must be determined at representative concentrations of the measurand.
4.14.11 | Measurement precision | Clause 4.14.11 requires providing an estimate of the 3 types of precision described in 4.14.12, repeatability SD; 4.14.13, intermediate precision SD; and 4.14.14, reproducibility SD; and if possible the uncertainty of each SD. An explanation how the SDs were determined must be provided. If the uncertainties of the SDs are determined, an explanation of the process must be included. A qualitative assessment if the precision is suitable for the intended use of the RMP in a calibration hierarchy of an end-user MP must be provided (see Clin Chem Lab Med 2015;53:905-12 that describes a proposed model for allocating uncertainty to higher-order references based on the allowable uncertainty of the clinical sample result).
4.14.15 | Detection limit | Clause 4.14.15 is frequently called the limit of the blank (LoB), see definition 3.13. The submitter must provide the detection limit and explain how the detection limit was determined. For example, LoB may be established using samples which are known to be negative for the targeted molecule e.g. matrix blanks or (for genetic tests) wild-type DNA controls.
4.14.16 | Lower and higher measurement limits | Clause 4.14.16 refers to what are frequently called the lower limit of detection (LLoD) and the upper limit of detection (ULoD). ISO 15193 does not define these terms. The submitter must provide the limits, explain how the limits were evaluated, and an assessment that the limits are suitable for the intended use of the RMP. Note that for many measurement procedures, the limits of quantitation are more important than the limits of detection. An explanation of the relationship between these types of limits should be provided when relevant to the intended use.
4.16 DBWG-P-04B | Validation of a RMP | Clause 4.16 requires that evidence be presented that the RMP is fit for its intended use. When another RMP for the same measurand exists and is listed by JCTLM, an extent-of-equivalence of results to the listed RMP is required. When another RMP is not listed by JCTLM, a comparison study of its nominated RMP with at least one existing well validated candidate RMP or globally available end-user measurement procedure used by a routine measurement laboratory is required. See JCTLM procedure DBWG-P-04B for requirements to demonstrate extent-of-equivalence of a nominated RMP to existing RMPs or other comparator MPs. Note that DBWG-P-04B clause 5.5.2 requires to determine if the extent-of-equivalence data demonstrates that the nominated RMP will produce fit-for-purpose results when used in the calibration hierarchies of end-user measurement procedures. DBWG-P-04B clause 7.1 states if the results of an extent-of-equivalence study indicate discordance among the evaluated RMM/Ps relative to fit-for-purpose criteria, the nominated RMM/P under review will not be accepted for listing in the JCTLM database until the discrepancy is explained. When applicable, fill out the section titled “ Hyperlink to Final Report of extent-of-equivalence studies among listed RMM/Ps" on the RMP Nomination Template.
general requirement |  | When citing a publication for any technical descriptions or information, the submitter shall carefully check that the publication (citation and its technical content) is correct and complete, and any differences between a publication and the actual information submited are fully described. The submitter should include details that describe differences to a publication that will be provided to end-users in the "Comments to be published on the JCTLM DB listing” field in the nomination form.

## RMS
JCTLM-DBWG/P-02-I-02
Checklist of frequently observed non-conformances by review teams with suggestions for a successful submission to JCTLM.
Clause in JCTLM DBWG P-03B1 (CMC-NMI) or DBWG P-03B2 (RMS lab) | Clause subject | Guidance for the nomination submission form
JCTLM DB WG P-03B - 6.1.1. | Is the laboratory’s accreditation status appropriate for listing of the reference measurement service? | The JCTLM database identification number for the RMP (Column AB) should be input. The uncertainty in the accreditation document must be the same as the uncertainty claimed in the Excel nomination file. If the claimed uncertainty is different (smaller or larger) from what is stated in the accreditation certificate, a validation report supporting the claimed uncertainty is needed for review. If the claimed measurement range (lower limit or upper limit) is different from what is stated in the accreditation certificate, a validation report supporting the claimed measurement range is needed for review. Note that modifications to RMPs that define the measurand, e.g. enzymes, are usually not acceptable.
JCTLM DB WG P-03B - 6.1.2. | Is the laboratory operating a procedure based on a method listed in the JCTLM database? | If the RMP used is not listed in the JCTLM database, then a simultaneous nomination of the RMP must be made. An RMP Service (RMS) will not be listed by JCTLM unless a JCTLM listed RMP is used.
JCTLM DB WG P-03B - 6.1.2. | Is the laboratory operating a procedure based on a method listed in the JCTLM database? | It is acceptable that a RMS provider improves/modifies a JCTLM listed reference measurement procedure provided the basic principle of the measurement procedure (e.g. IDMS) is maintained. A peer-reviewed publication of a modified method is the preferred route for fulfilling JCTLM requirements. An alternative, if the former is not possible, is providing a method validation report confirming that there is no bias observed between the measurement results obtained from the original and modified method. Note that modifications to RMPs that define the measurand, e.g. enzymes, are usually not acceptable. In addition, any differences in uncertainty or measurement range must be fully described in a peer-reviewed publication or in a method validation report. Note that modifications to RMPs that define the measurand, e.g. enzymes, are usually not acceptable. Failure to provide a complete method validation report for changes in a JCTLM listed RMP is a commonly observed deficiency.
JCTLM DB WG P-03B - 6.1.2. | Is the laboratory operating a procedure based on a method listed in the JCTLM database? | Examples of changes to a JCTLM listed RMP that require a peer-reviewed publication or a method validation report include (non-inclusive list, other changes may also require validation): (a) change in sample preparation, extraction procedure or other isolation technique for the measurand (b) change in the derivatization used in the procedure (c) change in the internal standard used (d) change in the stationary phase or gradient in liquid chromatography (e) change from gas chromatography to liquid chromatography, or the reverse (f) change in the gas in gas chromatography (g) change in a transition used in mass spectrometry (h) change in matrix of the measurand
JCTLM DB WG P-03B - 6.1.2. | Is the laboratory operating a procedure based on a method listed in the JCTLM database? | The standard operating procedure (SOP) used in the reference measurement service laboratory must be provided for the review team to verify that the RMP listed by JCTLM is being followed without modifications. The SOP must identify the JCTLM listed RMP by its database identification number.
JCTLM DB WG P-03B - 6.1.2. | Is the laboratory operating a procedure based on a method listed in the JCTLM database? | A validation report is required when the uncertainty of an implemented or modified a JCTLM listed RMP is lower or higher than what is included in the JCTLM listed procedure. Ensure that the uncertainty evaluation is complete and accounts for all sources of uncertainty in the measured result.
JCTLM DB WG P-03B - 6.1.2. | Is the laboratory operating a procedure based on a method listed in the JCTLM database? | A validation report is required when the measurement range is modified to include lower or higher values than what is included in the JCTLM listed procedure. Include documentation that the uncertainty is validated at the lower or higher values of the measurand. Note that when the RMS is based on an RMP defined measurand, the measurement range is defined by the RMP and changes are typically not expected.
JCTLM DB WG P-03B - 6.1.2. | Is the laboratory operating a procedure based on a method listed in the JCTLM database? | If the CRM used for calibration is changed from what is used in the JCTLM listed RMP, a validation report must be provided to demonstrate the new CRM as suitable for use. An extent of equivalence report is appropriate, see JCTLM-DBWG/P-04A. Examples include a CRM is out of stock or does not exist. The purity, assigned value with uncertainty, homogeniety, stability and other requirements of ISO 15194 must be documented.
JCTLM DB WG P-03B – 6.1.3. | Is the laboratory participating in regular inter-laboratory comparisons? | Results from recent years participation in RELA or an equivalent EQA scheme must be provided. If there was a modification to the RMP, indicate which version of the RMP was used for which years of EQA results.
JCTLM DB WG P-03B – 6.1.3. | Is the laboratory participating in regular inter-laboratory comparisons? | The EQA scheme must be intended for use with RMPs. The program information, accreditation, and results must be publically available internationally.
JCTLM DB WG P-03B – 6.1.3. | Is the laboratory participating in regular inter-laboratory comparisons? | Regular inter-laboratory comparisons in a network of RMP laboratories are acceptable when the results are publically available.

## JCTLM DB procedures
Listing of JCTLM procedures that must be followed when submitting nominations
Documents are available at https://www.bipm.org/en/committees/jc/jctlm/wg/jctlm-dbwg/publications, and by a link from jctlm.org.
Red font documents are directly related to the nomination of CRMs, RMPs, and RMS laboratories
General information
JCTLM-DBWG/P-01A | Outline of JCTLM procedures for material and method nomination review
JCTLM-DBWG/P-02A | Reference material and procedure nomination requests
JCTLM-DBWG/P-02A-I-01 | Instructions for completing nomination templates (Excel spreadsheets)
JCTLM-DBWG/P-03A | Review and approval of nominations for materials and methods - used by Review Teams
JCTLM-DBWG/P-03A-F-03 | Review report form for material and method nomination - used by Review Teams
CRM submissions
JCTLM-DBWG/P-02A-F-01 | Reference material nomination template
JCTLM-DBWG/P-04A | Multiple CRM comparison process (extent-of-equivalence)
RMP submissions
JCTLM-DBWG/P-02A-F-02 | Reference measurement procedure nomination template
JCTLM-DBWG/P-04B | Multiple reference method/procedure comparison process (extent-of-equivalence)
RMS submissions
JCTLM-DBWG/P-02B | Nomination process for reference measurement laboratory services
JCTLM-DBWG/P-02B-F-02 | Template for nomination of reference laboratory measurement services
JCTLM-DBWG/P-03B1 | Review of NMI measurement services
JCTLM-DBWG/P-03B2 | Review of measurement services from accredited laboratories