## Slide 1: Length Activities at NIS
Presented by:
Dr. Osama Terra
Head of Primary Length Standard lab.
Vice Chair of AFRIMETS TCL
National Institute of Standards (NIS)
Egypt

## Slide 2
Overview
Overview about NIS
Length division at NIS
Laboratories capabilities
  - Primary Length standard and laser technology
  - End and line standard
  - Engineering and Surface Metrology
Research and publications
Comparisons and CMCs

## Slide 3
Establish, maintain and disseminate the SI units
Provide calibration, training and consulting services.
Research and develop new and improved measurement procedures
NIS Mission

## Slide 4
2700BC
Metrology begins
in Egypt.
The Royal
Egyptian Cubit
1963
NIS is
Founded
1992
NIS moved to recent premises in Haram, Giza
2000
NIS Timeline
1962
Egypt signed
The meter convention
NIS Signed
CIPM MRA

## Slide 5
NIS area ~ 150,000 m2
NIS Location
3 km

## Slide 6
NIS Structure
NIS has 21 laboratories (in 6 divisions ). NIS Head: Prof. Dr. Mohamed A. Amer
Notes: Each division related to quantity
Electrical ampere and second

## Slide 7
NIS published CMCs
NIS has published 39 CMCs in different metrology areas such as:
  1- Length
  2- Mass
  3- Pressure
  4- Time and Frequency
  5- DC Current, resistance, voltage and capacitance
  6- Viscosity
  7- Force
  8- Torque
  9- Chemistry
  10- Ionizing radiation
I expect a significant increase in the number of published CMCs in the
near future, since the current NIS top management is currently promoting the
active participation in the CIPM MRA activities.

## Slide 8
Length Division

## Slide 9
Length Division
Primary length Standard
and Laser technology
End and Line Standards
Engineering and Surface Metrology
Optical frequency measurement.
Laser sources characterization.
Fiber optics metrology.
Laser Doppler velocimeter calibration
EDM calibration
GPS receivers calibration.
Head: Dr. Osama Terra
GB Calibration by Interferometry.
Line Standards Calibration.
Polarimetric measurement.
Head: Dr. Niveen Fareed
GB Calibration by mech. comparison.
Roundness Measurements.
Optical and Mechanical surface Measurements.
Angle measurement.
Dimensional measurement (CMM, universal microscope, Profile projector).
Hand measurement tools calibration (caliper, Micrometer, dial gauges).
Head: Prof. Dr. Salah Ali
(1)
(2)
(3)
Head: Prof. Dr. Hatem Hussein
Division staff
10 researchers (PhD holder)
10 assistant researchers (MSc, BSc)
4 technicians

## Slide 10
(First Laboratory)
Primary Length Standard
& laser technology Lab.

## Slide 11
Primary Length Standard & Laser Technology (Laboratory 1)
1.1 - Optical Frequency Measurement
Iodine stabilized He-Ne Lasers
Stability 2.5 x 10-11
Femtosecond frequency comb
Wavelengths: 500 nm – 2100 nm
Stability 2 x 10-12

## Slide 12
Primary Length Standard &
Laser Technology (Laboratory 1)
1.2 - Laser Sources Characterization
3- Spatial modes characterization (beam width)
4- Longitudinal modes characterization
2- Spectrum Analysis
Wavelength range: 600 to 1700nm
Wavelength accuracy: ±0.02 nm
1- Wavelength Measurement
Wavelength range: 400 to 1700nm
Uncertainty: ±0.3 pm

## Slide 13
Primary Length Standard &
Laser Technology (Laboratory 1)
1.3 – Optical fiber Metrology
2- Optical spectrum analyzer calibration
Using H13C14N absorption spectrum
range 1540 -1560 nm
Uncertainty: ± 0.6 pm
1- OTDR calibration
Range: 1- 50 km
Uncertainty: ± 30 cm
Using Recirculating loop artifact
3- DWDM laser characterization
range 800 -1700 nm
Accuracy: ± 0.02 nm

## Slide 14
Primary Length Standard &
Laser Technology (Laboratory 1)
1.4 – Calibration of Laser Doppler Velocimeters (LDV)

## Slide 15
Primary Length Standard &
Laser Technology (Laboratory 1)
1.5 – Calibration of large scale measuring equipment
According to (ISO 17123-4:2012)
According to ISO 17123-8:2015
EDMs instruments
GNSS receivers
300 m
300 m
Notes: Master and PhD thesis

## Slide 16
(Second Laboratory)
End and line standard

## Slide 17
End and line standard (Laboratory 2)
2.1 –Gauge block calibration by Interferometry
1-Köster Interferometer
  Range: 0.5 – 100 mm
  Uncertainty: ±20-70 nm
2- Double-Sided calibration
(Project with NMISA, South Africa )
  Range: 0.5 - 100 mm
  Uncertainty: ±60 nm

## Slide 18
End and line standard (Laboratory 2)
2.2 – Long-gauge block calibration
Lab Master
  Range: up to 1 m
  Uncertainty: ± 0.5 µm
2.3 – Line scale measurement
Displacement interferometer
  Range: 1.2 m
  Uncertainty: ± 2 µm
Notes: Uncertainty of lab master, displacement interferometer

## Slide 19
End and line standard (Laboratory 2)
2.3 – Refractometry and Polarimetry
-Calibration of auto-refractometers
Range: 1.3 – 1.7
Calibration of polarimeters and Quartz plates

## Slide 20
(Third Laboratory)
Engineering and Surface Metrology

## Slide 21
Engineering and Surface Metrology (Laboratory 3)
3.1 – Gauge- Block calibration by comparison
    Electromechanical Comparator
    Range : 250 mm
    Uncertainty: ±80nm
    Gauge block set
    (122 piece , Ceramic, steel , Tungsten carbide)
    inch, meter).
    Range: 0-100 mm (0-4 inch) Accuracy: grade k
3.2 – Diameter standard calibration
-Universal Length measuring machine
- Reference setting rings
Range : 1 m
Uncertainty < ±1 µm
Notes: GB uncertainty
How many accredited labs in hand tools

## Slide 22
Engineering and Surface Metrology (Laboratory 3)
3.3 – Dimensional Metrology
1- CMM (3D)
Regular and irregular form
Range : 700x900x500 mm.
Uncertainty: ±2 um.
2- Profile projector (2D)
Range : 100 x 200 mm.
uncertainty: ±5 um.
Notes: You don’t need to mention accurac
Model enough

## Slide 23
Engineering and Surface Metrology (Laboratory 3)
3.4 –Machine tools metrology
- Laser displacement interferometer Calibration of CMM, CNC and universal microscope
Range 40 m: .
uncertainty: ±1 ppm.
3.5 – Hand Instruments calibration (caliper, Micrometer, dial gauges).
Steel Gauge blocks set ( Range: 0-1000 mm, accuracy: Grade:1)

## Slide 24
Engineering and Surface Metrology (Laboratory 3)
3.6 –Mechanical form measurements
1- Stylus (Talysurf i120)
Range: H: 60 mm V: 1mm
uncertainty: ± 60 nm.
2- Roundnes measurement
(Taylor and Hobson Talyrond 73)

## Slide 25
Engineering and Surface Metrology (Laboratory 3)
3.7 – Surface measurement
Large object Interferometers (Zygo GPIXP)
-used in: Flatness, parallelism , and radius of
Curvature & aberration (for lenses)
Range: (H: 100 mm, V: 100 um)
Uncertainty: V: ±2 nm
Interference Microscope (Zygo Maxim GP200)
- Using phase shifting and white light scanning
- Used in: Surface texture and Thin film analysis
Range: (H: 3×2 mm , V: 100 um) .
Uncertainty: (H: ±2 µm, V: ± 1 nm )
Atomic Force Microscope (Thermomicroscope Ap100)
- Used in: Nanometrology, research
(scanner1) Range: (H: 100 µm, V: 8 um) .
Resolution: (H: 0.25 A, V: 0.025 A )
(scanner2) Range: (H: 5 µm, V:2.5 µm)
Resolution: (H: 0.013 A, V: 0.009A)

## Slide 26
Engineering and Surface Metrology (Laboratory 3)
3.8 – Angle measurements
1- Autocollimator (2)
  Measurement range: ±1000 arcsec.
  Uncertainty: ± 0.05 arcsec.
3- Reference angle gauge blocks (12 steel pieces)
  Range: 2 - 90O
  Uncertainty: ± 2 sec.
2- Indexing table
  Range: 360O
  Uncertainty: ± 0.5 arcsecond.

## Slide 27
Highlights on the Research Activities
Length division at NIS has published more than
60 publications in journals indexed in SCOPUS
during the period from 2010 to 2018.

## Slide 28
Title: An ultra-stable optical frequency standard for telecommunication purposes based upon the 5S1/2  5D5/2 two-photon transition in Rubidium
Authors: Osama Terra, Hatem Hussein
Journal: Applied Physics B (2016)
Femtosecond frequency
comb

## Slide 29
Title: LASER RANGING IN AIR UTILIZING MODE-LOCKING IN LASER CAVITIES
Authors: Osama Terra
Journal: IEEE photonics technology letters (2017)

## Slide 30
Title: Distance measurement using frequency scanning interferometry with mode-hoped laser
Authors: M. Medhat, M. Sobee, H. M. Hussein, and Osama Terra
Journal: Optics and Laser Technology (2016)
Notes: Master thesis

## Slide 31
Title: Absolute gauge block calibration using ultra-precise optical frequency synthesizer locked to a femtosecond comb.
Authors: Hatem Hussein, Niveern Farid, Osama Terra
Journal: Applied Optics (2015)
Multiwavelength GB length measurement equation

## Slide 32
CMCs and Comparisons

## Slide 33
Published CMCs in Length

## Slide 34
International Comparisons (1)
CCL

## Slide 35
International Comparisons (2)
CCL

## Slide 36
International Comparisons (3)
CCPR

## Slide 37
BIPM.L-K11 (Laser radiation)

## Slide 38
COOMET.PR-S8 (Wavelength for fiber optics)

## Slide 39
APMP-K8 (Surface roughness)

## Slide 40
AFRIMETS.L-S3 (gauge block by mechanical comparison)

## Slide 41
Thank you for attention