Guide to the Analysis of CCT Key Comparison Results

Prepared by the CCT Working Group for Key Comparisons

February 2026

1 Introduction

This document provides guidance on the analysis of results of CCT Key Comparisons (KCs). Since different approaches are followed by other CIPM Consultative Committees, this document is applicable only to CCT KCs.

The analysis of the results of a KC includes many steps and, in this document, these steps are discussed in the same chronological order in which the KC pilot and the participants have to take them (see Figure 1).

Given the terseness and ambiguity of the information of CIPM MRA documents on the key comparison reference value (KCRV) and the degrees of equivalence (DoEs), we deemed necessary a clarification of the definitions and interpretations of the KCRV and DoEs in CCT KCs as the first and preliminary step (Section 2).

Moreover, because some decisions about the methods of analysis need to be taken already when designing the KC Technical Protocol, we discuss the general principles to be followed in the analysis that should be set in the Technical Protocol (Section 3).

Section 4 describes the inspection of the initial data set using simple graphical methods. Such inspection provides a qualitative assessment of the dispersion of the results and allows to identify anomalous results.

Section 5 provides guidance on how to deal with anomalous results, amendment and withdrawal of results, and selection of the final data set to be used in the analysis.

Section 6 describes the statistical tests to be performed on the final data set to obtain useful information for the following analysis.

Section 7 describes how bilateral degrees of equivalence can be directly computed from the final data set, without needing to select a procedure for computing the KCRV (and without computing the KCRV), provided correlations and dark uncertainty (if detected) are properly taken into account.

Section 8 recalls the four commonly used statistical models in key comparisons and describes the random effects model, which is considered the most appropriate statistical model for CCT KCs.

Section 9 summarizes the most commonly used statistical procedures for computing the KCRV.

Finally, Section 10 provides guidance on the selection of the statistical procedure to be used for computing the KCRV.

This document draws inspiration from two major sources: 1) the CCQM Guidance note “Estimation of a consensus KCRV and associated Degrees of Equivalence” [1], which is currently under revision by a CCQM Task Group, and 2) the NIST Decision Tree User’s Manual (3rd Edition) and its background article, see Appendix I and [2].

This document is not intended to be prescriptive or mechanically applied to all CCT KCs. Rather, it provides a useful framework to allow the pilots and the participants to take the appropriate decision on the selection of the method to be used in the determination of the KCRV, basing it on their expertise, measurement science and statistical science.

Figure 1: Flow chart of a CCT Key Comparison. The sections of this document mirror this flow chart.

2 Degrees of Equivalence and Key Comparison Reference Value

In the framework of the CIPM MRA [3], the primary objective of a Key Comparison (KC) is to obtain a quantitative measure of the degrees of equivalence (DoEs) of national measurement standards maintained by the National Measurement Institutes (NMIs). The MRA defines (though merely in words, without using any formula) two forms of DoEs, often referred to as unilateral DoEs and bilateral DoEs. The definition of DoEs involves the Key Comparison Reference Value (KCRV) and the meaning to be attributed to it. However, the information provided by the MRA documents (essentially, [3] and [4]) on DoEs and KCRV is terse, ambiguous and under-determined, so in this section we clarify how the CCT interprets the DoEs and the KCRV.

2.1 Degrees of equivalence in CCT Key Comparisons

2.1.1 Unilateral Degrees of Equivalence

The unilateral DoE, Di, is the degree to which the standard of participant i is consistent with the KCRV. Indicating with xi the value measured by participant i in the KC and with xKCRV the value of the KCRV, Di is quantitatively expressed by the two following terms:

Di ≡ xi – xKCRV

U(Di) ≡ U(xi – xKCRV)

where U(Di) is the uncertainty of Di at a 95% level of confidence.

2.1.2 Bilateral Degrees of Equivalence

The bilateral (or pairwise) DoE, Dij, between any pair of participants i and j is the degree to which the standards of participants i and j are consistent with each other, and is defined quantitatively by the following two terms:

Dij ≡ xi – xj

U(Dij) = U(xi – xj)

where U(Dij) is the uncertainty in Dij at a 95% level of confidence.

Note how, while the definition of the unilateral DoEs involves the KCRV, our definition of bilateral DoEs does not depend on the KCRV. The only difficulty is the evaluation of the uncertainty Uxi-xj, because it is not simply given by the combination of the individual participant expanded uncertainties, U(xi) and U(xj), in quadrature. Depending on the statistical model adopted for the measured values xi and on the topology of the KC (for example when xi and xj are performed in different loops), UDij can involve other uncertainty components besides the reported uncertainties for xi and xj, which prevents presenting here a formula that is universally valid.

2.2 Key Comparison Reference Value in CCT Key Comparisons

The CIPM MRA documents do not give a definition of KCRV nor do they specify how the KCRV should be computed, they only deal with how the KCRV should be interpreted. In this section, we first discuss the meaning to be attributed to the KCRV in CCT KCs (section 2.2.1), and then we give a definition of KCRV in CCT KCs. The subject of how the KCRV should be computed will be addressed in the remaining sections of this Guide.

2.2.1 The meaning to be attributed to the KCRV in CCT KCs

2.2.1.1 The meaning attributed to the KCRV in the CIPM MRA and in the literature

The CIPM MRA [3] states that, in most cases, the KCRV can be considered to be a close approximation to the SI value of the measurand, while CIPM MRA-G11 document [4] states that each KCRV is considered to be a close approximation of the true value of the measurand. However, as acknowledged by the Technical Supplement to the Arrangement, some Consultative Committees (CCs) have difficulties in relating the KCRV to the SI value of the measurand. Some authors (see, for example, [5]) even question the possibility for any KC to quantify the consistency of a standard to the true value of the measurand, arguing that, in order to achieve such a goal, additional information and assumptions would be needed. Other authors state that obtaining both consistency with the true value of the measurand and mutual consistency among the participants in a KC is attainable in rather few cases [6].

2.2.1.2 The meaning attributed to the KCRV in the past CCT Key Comparisons

It appears that the large majority of the pilots and participants in the past CCT KCs shared the concerns expressed in [5] since, in nearly all past CCT KCs, the participants refrained from attributing the KCRV the meaning of best (or close) approximation to the SI value (or true value). In Table 1 we report the statements made by the CCT KC participants on the meaning of the calculated KCRV.

| KC | Ref. | Statements on the meaning attributed to the KCRV |
| --- | --- | --- |
| CCT-K1 | [7] | No additional significance is attached to the KCRV with respect to ‘best estimate’ of T90 |
| CCT-K2 | [8] | The KCRV represents a common baseline against which all laboratory values can be compared |
| CCT-K3 | [9] | The participants decided not to compute the KCRV |
| CCT-K4 | [10] | The value of the so-defined KCRV must not be related to the ITS-90 value of the corresponding fixed point. The KCRV has no physical meaning and is used only as a notational shorthand for presenting a common baseline against which all laboratory values can be compared |
| CCT-K5 | [11] | No statement on the meaning of the KCRV |
| CCT-K6 | [12] | The KCRV has no absolute significance outside the comparison |
| CCT-K7 | [13] | It is clear that this KCRV is not the best possible approximation of the true SI value |
| CCT-K8 | [14] | The KCRV given in this comparison has no absolute significance, that means it does not represent a reference value in the SI, only a comparison parameter |
| CCT-K9 | [15] | The KCRV is not meant to have any intrinsic meaning – it is a computational tool allowing comparisons across labs linked through measurements at the pilot lab |
| CCT-K10 | [16] | No statement on the meaning of the KCRV |
| CCT-K11 | [17] | No statement on the meaning of the KCRV |

Table 1: Statements on the meaning attributed to the KCRV in CCT-KCs (from the corresponding final reports).

2.2.1.3 The meaning to be attributed to the KCRV in CCT Key Comparisons

In order to establish the meaning that can be reasonably attributed to the KCRV in CCT KCs, we analyze in detail the measurand in CCT thermometry and hygrometry KCs.

First of all, in all completed and running CCT KCs, the measured quantity (the measurand) is not an SI quantity, because the CCT KCs are based on the International Temperature Scale of 1990 (ITS-90) [18]. The ITS-90 is a practical temperature scale, that has been constructed in such a way that the quantity T90 that it identifies is a close approximation to another quantity T (thermodynamic temperature, which is the SI quantity) according to the best estimates at the time the scale was adopted (1990). However, since the adoption of the ITS-90, the substantial advancements in the measurement of T (primary thermometry) have evidenced significant differences between T and T90 [19], to the extent that nowadays T90 cannot be considered anymore a close approximation to T. Moreover, the same technological advancements in primary thermometry are allowing direct and practical measurements of T in some temperature ranges, and the CCT is considering to start running KCs in terms of T in parallel to the traditional KCs (in terms of T90).

In conclusion, in the past and current CCT KCs, the KCRV cannot be considered to be a close approximation to the SI value of the measurand (as done in [3]). The statement in [4], which is more general than that in [3], and refers to the true value of the measurand (independently from it being an SI quantity or not), is regarded as more appropriate for CCT KCs. In the future, when the CCT will run KCs in terms of thermodynamic temperature (SI), the statement in [3] will be applicable to their measurand.

The meaning of the KCRV in CCT Thermometry Key Comparisons

Let’s consider a CCT KC of SPRTs calibrated at the ITS-90 fixed points (this is the case of CCT-K2, CCT-K3, CCT-K4, CCT-K7, CCT-K9 and CCT-K7.2021, so the majority of thermometry KCs). In such a KC, each participant realizes its local ITS-90 temperature scale, based on its local set of fixed-point cells, then calibrates its transfer artifact (SPRT) against its realized fixed-point temperatures, and finally transports the SPRT to the pilot’s laboratory, where the pilot compares the travelling SPRT to its own fixed-point cells.

The measurand in such a KC is the resistance ratio W between the SPRT resistance at a given fixed-point (for example, the indium freezing point) and the SPRT resistance at the triple point of water. The ITS-90 definition assigned to the freezing point of indium a reference temperature T90 = 429.7485 K and a reference resistance ratio Wr,In = 1.60980185. This does not mean that the true value of the measurand in the KC is 1.60980185. The resistance ratio of the specific SPRT used as transfer artifact has a true value which is different from the reference value.

In fact, for the purposes of the KC, the true value of the SPRT ratio has a very marginal relevance. In the KC the SPRT is only a means to store the information about the local scale realizations of the participants. The participants are measuring different values Wi,In of the resistance ratio at the indium freezing point not because of their different abilities in measuring electrical resistances, but (primarily) because they realize different freezing point temperatures of indium (because of different batches of high-purity materials, different thermal environments produced by different furnaces, …). The goal of the KC is to find out the differences (in temperature!) between the individual ITS-90 scale realizations of the participants. In fact, the KCRV and the DoEs are expressed in terms of temperature and not in terms of resistance ratio. From this point of view, it can be said that the measurand is, in fact, the temperature of the freezing point of indium, which is an unequivocally defined thermodynamic equilibrium state (freezing point of an ideally pure material), although a procedure-defined measurand (all participants are using the same method: fixed-point cells with thermometer wells accommodated in furnaces/cryostats/thermal baths).

Note that the value of the measurand is assigned by the ITS-90 definition (for indium: T90 = 429.7485 K). All participants assign to the indium triple point the same value of T90 = 429.7485 K. So, obviously, the participants are not measuring T90 values at the fixed points. What they are doing is, in fact, measuring the resistance ratio Wi,In of the transfer SPRT, corresponding to their realized local fixed-point temperature. In doing so, they determine the temperature differences between their different ITS-90 realizations.

What is the meaning of the KCRV in this case?

Obviously, the KCRV is not an estimate of the SI value, because the ITS-90 temperature T90 is not an SI quantity. On the other hand, it is not correct to state that the KCRV has no physical meaning. Because the KCRV is obtained by pooling the individual independent results of the participants using an appropriate statistical procedure, the KCRV gives the statistically most likely position of the ideal fixed-point realization. The deviation from the KCRV of each participant (Di ≡ xi – xKCRV) also has a physical meaning: it tells the participant its deviation from the statistically most likely position of the ideal fixed-point transition. In conclusion, if we define the measurand as the ideal freezing point temperature of indium when measured with the procedure defined by the ITS-90 (fixed point cells, furnaces, SPRT, …), we can say that the KCRV can be considered as an estimate of the true value of the measurand (though not SI), as in [3].

The meaning of the KCRV in CCT Hygrometry Key Comparisons

A similar reasoning to that applied above to thermometry KCs can be also applied to CCT hygrometry KCs (CCT-K6 and CCT-K8). In hygrometry, similarly to thermometry, each participant realizes its local dew/frost-point ITS-90 temperature scale of humid gas (air), striving to realize the same thermodynamic equilibrium state: complete water saturation of air at a given pressure (atmospheric) and temperature.

The difference from a thermometry KC is that, in a thermometry KC, the measurand (temperature) is undergoing a first-order phase transition, and the phase transition acts as a temperature controller at the temperature of the phase transition itself. In a dew/frost-point KC, the measurand (dew/frost point temperature) is not undergoing a first-order phase transition. For this reason, while the measurand values in thermometry KCs are “forced by nature” at the phase transitions of the materials, in dew/frost point KCs nominal measurand values are to be agreed in the protocol (for example, 30 °C, 50 °C, 65 °C, 80 °C, 85 °C, 90 °C and 95 °C in CCT-K8), and the participants are asked to realize temperatures within ±0.5 °C of the agreed nominal values.

The transfer artifact in dew/frost point KCs is a state-of-the-art chilled-mirror type of dew point hygrometer. The measurand, the resistance of a PRT, embedded under the surface of the chilled-mirror, at the agreed nominal values of dew/frost point temperature, is transformed into a temperature value using the IEC 60751 standard. To cope with the fact that different participants, when measuring at the same nominal dew/frost point temperature, are in fact measuring at different temperatures (within ±0.5 °C), the difference ∆t between the measured dew/frost point temperature (PRT resistance transformed into IEC 60751 temperature) tdhyg and the applied dew/frost point temperature (generated by the local primary dew-point generator) tdref: ∆t=tdhyg-tdref is formed. In this way, the fact that each participant measures the transfer artifact at slightly different dew/frost point temperatures becomes irrelevant, under the assumption that the behavior of the transfer artifact is the same within ±0.5 °C. Again, in the same way as in a thermometry KC, the transfer artifact is used to store the information on the individual local dew/frost-point scales, and what is actually compared in the KC are the local dew/frost-point scales of the participants. Although the participants are not realizing the same measurand, the use of a stable transfer artifact allows to obtain an estimate of the differences between the realized primary dew-point standards (the generators) for the same measurand (though nominal).

Again, although the numerical value of the KCRV has no meaning of SI value, but the position of the KCRV represents the most likely position of the ideal dew/frost point realization, and the deviations from the KCRV (Di ≡ xi – xKCRV) are the deviations of the participants’ realizations from the statistically most likely position of the ideal realization. In this sense, we can say that the KCRV is an estimate of the true value of the measurand, as in [3].

2.2.3 Definition of KCRV in CCT Key Comparisons

We define the KCRV in CCT KCs as:

the best estimate that can be made of the true value of the measurand from the KC measurements.

2.3 Consistency of the results with the KCRV, with one another and mutual consistency

Before discussing further, we clarify the exact meaning of different concepts of consistency used by the MRA in [3] and [4]: consistency of the results with the KCRV, consistency of a pair of results with one another, and mutual consistency of a set of results:

(A) Consistency of a measurement result with the KCRV:

A measurement result xi is said to be consistent with the KCRV xKCRV when the difference xi – xKCRV does not differ (statistically and substantively) significantly from zero. The statistical significance depends on the probability distribution of the difference. The substantive significance depends on a determination, made by the relevant Consultative Committee (CCT in our case) about how large the absolute value of that difference can be without it impacting the comparability of measurement results in practice.

(B) Consistency of two measurement results with one another:

Two measurement results xi and xj are said to be consistent with one another when the difference xi – xj does not differ (statistically and substantively) significantly from zero. If there are N participants, then N(N-1)/2 such differences can be contemplated. Deciding whether they are significantly different from zero therefore involves N(N-1)/2 tests of the hypothesis of no difference. The conclusions of these tests should take into account the fact that the larger the number of tests, the larger the probability that at least one of the tests will suggest an incorrect conclusion.

(C) Mutual consistency of a set of measurement results:

A set of measurement results is said to be mutually consistent (with one another, regardless of what the KCRV or the true value of the measurand may be), when the dispersion of the measured values is not (statistically and substantively) larger than what the reported uncertainties suggest that it should be. Cochran’s Q test is often used to determine the statistical significance of any apparent mutual inconsistency.

While [3] gives the impression that, if goal A is satisfied (consistency of all results with the KCRV), then it follows that goal B (consistency between each pair of results) is also satisfied, this is not true. In general, given a set of results, there is no definite logical implication (for example, (A) ⇒ (B), or (A) ⇒ (C)). In fact, it is possible for all the individual results to be consistent with the KCRV, and at least one pair of participants to have inconsistent results (in this case (A) ⇒ (B) is not true). It is possible also to have cases in which (A) and (C) are true and (B) is not true for at least one pair.

Obviously, an ideal set of KC results would satisfy simultaneously all three types of consistency: consistency of all results with the KCRV, consistency of all pair of results with one another, and consistency of the whole set of results. This is hardly ever achieved.

3 Principles to be adopted and Technical Protocol

The design of the Technical Protocol is crucial for the success of a KC, because no analysis method can save poor data collection.

Moreover, some decisions taken in the early phase of the life of a KC, namely when designing the Technical Protocol, have an impact on the later phase of the analysis of the measurement results. In this section we provide guidance on such decisions.

CIPM-MRA-G11 document [4] states that the Technical Protocol, for CIPM KCs, should include a description of the method to be used to determine the KCRV. However, as clarified by Koepke et al. [20], the appropriate method for data reduction cannot be selected before examining the KC measurement results because the data reduction should be consistent with a statistical model that is demonstrably adequate for the measurement results (see principle 4 below). Therefore, the method to be used to determine the KCRV cannot be specified in the Technical Protocol in advance of obtaining and examining the measurement results from all the participants.

However, the Technical Protocol should at least set the general principles to be followed in the analysis of the results. The CCT endorses the 6 principles enunciated in [20], reformulated by Possolo et al. [21], and reproduced here, as follows:

1) Degrees of equivalence (DoEs) will be computed for all participants regardless of whether or not their results are included in the calculation of the KCRV and associated uncertainty, except for any participants that will have chosen to withdraw from the KC. The selection of some results for the computation of the KCRV, and the exclusion of others, should be based on substantive reasons that are documented in the KC Final Report. The mere fact that a measured value lies far from the bulk of the others alone is insufficient reason to set it aside, even if a statistical test suggests that it is an “outlier” [22-23].

2) No measured value should dominate the KCRV value “automatically,” simply because the associated measurement uncertainty is much smaller than the uncertainties associated with the other measured values. To this end, procedures for data reduction should include some damping mechanism to limit the influence that measured values with unusually small associated uncertainties will have upon the KCRV. This provision is particularly influential when the measurement results are mutually inconsistent, and the uncertainties reported for some of them are much smaller than the uncertainties reported for the bulk of the others.

3) Measurement methods should be sufficiently well characterized to warrant confidence in the belief that the measured values, taken as a group, are roughly centered at the true value of the of the measurand. Participating laboratories should have previously demonstrated sufficient competence based on satisfactory performance in previous KCs, pilot studies, or proficiency tests. If all the measured values would tend to be too low or too high relative to the true value of the measurand, no statistical procedure that relies on the data alone will be able to detect this and correct the KCRV accordingly, but the DoEs may, even in such cases, still be comparable and informative.

4) A model should be formulated that explicitly relates the measured values to the true value of the measurand. Furthermore, the estimation of the true value, and the evaluation of the associated uncertainty, should be consistent with the statistical model and with some principle of estimation whose general reliability is widely recognized. The calculation of the KCRV is contingent on such a model and on the choice of the optimality criterion that the KCRV is intended to satisfy. Suppose that the model specifies that the measured values are equal to the measurand plus random, mutually independent, Gaussian measurement errors. This alone does not suffice to justify combining the measured values in a particular way (say, as weighted average of the measured values). An additional criterion is needed: for example, one that seeks to minimize the mean squared error of the KCRV, or one that requires a KCRV with minimal absolute error. Therefore, the choices that one needs to make to determine the KCRV should take into account the purpose that the KC is intended to serve.

5) The statistical model underlying data reductions should be able to detect, evaluate and propagate dark uncertainty, which manifests itself as dispersion of the measured values in excess of what the reported uncertainties suggest that it should be, and thereby accounts for mutual inconsistency of the measurement results. Failure to recognize and propagate dark uncertainty generally yields DoEs with uncertainties that are too small. Both the average and the median of the measured values, with or without the conventional weights (which are proportional to the reciprocal, squared reported uncertainties) ignore dark uncertainty, and thus tacitly assume that the results are mutually consistent. Furthermore, the reliability of the reported uncertainties, often expressed in terms of the effective numbers of degrees of freedom that support them, should be taken into account, hence should be reported routinely, as required by CIPM guidance for KCs.

6) Degrees of equivalence should be computed consistently with their primary goal of identifying participants with “unusual” results, in the sense that their measured values lie beyond the range allowed by the model. In particular, each DoE should include: i) the uncertainty reported for the corresponding measured value; ii) the uncertainty associated with the KCRV; iii) the correlation between the measured value and the KCRV; and (iv) any dark uncertainty that will have been detected during data reduction. In most cases, addressing these four requirements involves resorting to either classical or Bayesian Monte Carlo methods for uncertainty propagation.

4 Screening of the initial data set

Upon completion of the KC measurements, and after receiving the measurement reports from all the participants, the pilot assembles the initial data set. For a KC with N participants, the initial data set is a set of N triplets Ti, ui,νi, where Ti is the result of participant i, ui is the combined standard uncertainty of participant i, and νi is the number of degrees of freedom associated to ui.

It is recommended that the degrees of freedom of each uncertainty component contributing to ui also be reported, regardless of how they may be used in the analysis of the results (see section 6). Each participant should also classify the uncertainty components into Type A and Type B estimates, because this information is needed to properly take correlations into account.

The initial data set is inspected by the pilot using appropriate graphical methods (see CCQM Guidance Note [1]). At a minimum, the measured values Ti and their corresponding standard or expanded uncertainty bars should be depicted in a plot whose caption must state whether the uncertainty bars represent standard or expanded uncertainty, to identify:

- Measured values Ti whose location appears inconsistent with the majority of the measured values;

- Uncertainties that are unusually large or small;

- Measured values Ti that deviate substantially from some preliminary estimate of the KCRV relative to their reported uncertainties.

Even without resorting to the use of outlier tests, such a simple visualization of the measurement results allows the identification of anomalous results and enables a visual assessment of the dispersion of the results. Figure 2, taken from CCQM Guidance Note [1], shows some of the main features of a metrology comparison data set. One of the 4 regions depicted in the figure will best represent the data set:

A = mutually consistent results (no excess variance, no outliers);

B = general overdispersion or understatement of uncertainty affecting most of the participants;

C = generally consistent results with a small number of apparent outliers;

D = over-dispersion combined with some particularly extreme apparent outliers.

However, Figure 2 does not depict a general configuration that can be sometimes encountered in KCs: when the measured values have an asymmetric distribution, trailing either to the left, or to the right of the center.

Visually-identified outliers can be confirmed with statistical outlier tests, or tests based on non-parametric or robust statistics. A common non-parametric indicator of an outlier is a result outside the inter-quartile range [Q1,Q3] that lies more than 1.5Q3-Q1 away from the nearest quartile. This implies that a value sampled from a Gaussian (or normal) distribution will be (falsely) declared as an outlier with probability 0.7 %.

Figure 2: Features of importance in metrology comparison data. Vertical bars denote 95 % coverage intervals or expanded uncertainty intervals. A: mutually consistent data; B: evidence of general over-dispersion (or underestimation of uncertainty affecting most of the participants); C: generally consistent results with a small number of outlying values; D: general over-dispersion combined with some particularly extreme values [1].

5 Dealing with anomalous results, amendments and withdrawal of results

In case the initial screening of the measurement results using graphical methods and statistical tests of outliers identify anomalous results, the pilot should follow up with the participant as detailed in [4]. When the participant shows evidence of arithmetic, typographical, or transcription error, the result can be corrected. When the participant cannot identify any such error, and also there is no substantive reason to set the results aside (for example, incorrect calibration, traceability through another participant, notable deviation from the KC’s technical protocol, etc.), then the participant’s result stands and is not removed from the data set.

Statistical tests alone are not sufficient to justify removal of any measurement result from the data set used to calculate the KCRV. There must be a substantive reason for excluding any measurement result from the calculation of the KCRV. Laboratories excluded from KCRV calculation will, however, be included in data summaries and plots, and DoEs will be calculated for them after calculation of the KCRV.

The participants are allowed to withdraw their results only if a reason not attributable to the performance of the laboratory can be identified (for example, if an excessive drift or a malfunction is detected in the transfer standard).

The initial data set, eventually modified by correcting arithmetic, typographical or transcription errors, and after excluding results based on substantive reasons, becomes the final data set. The final data set is used in the following analysis (and in the determination of the KCRV).

6 Statistical tests

As will become evident in the following sections, statistical tests performed on the final data set are a fundamental step in the selection of the procedure for calculating the KCRV. The goal of the statistical tests is to listen to the data set, and to coax it to tell you how it would like to be analyzed.

In this section we describe 3 statistical tests, the same that are used in the NIST Decision Tree [21] as it develops a recommendation for how the measurement results should be modeled and analyzed, and we explain the meaning to be attributed to the results of such tests. Obviously, the KC pilots and participants are free to apply other statistical tests to help them navigate the NIST Decision Tree, or to select a model and data reduction procedure that is not among those implemented in the NIST Decision Tree.

The NIST Decision Tree recommends that the following statistical tests be performed: a homogeneity test (that is, tests of mutual consistency), which tests the mutual consistency of the measurement results using Cochran’s Q test [24]; a normality (Gaussian shape) test, which tests if the results are normally distributed using the Shapiro-Wilk test [25]; and a symmetry test, which tests if the results are symmetrical using the Miao-Gel-Gastwirth test [26].

For each of these tests, the p-value is the probability that a set of measurement results that complies with the hypothesis under test will, owing to the vagaries of sampling, be at least as deviant from this hypothesis as the measurement results under test are. A small p-value casts doubt on the data actually satisfying the hypothesis under test.

For example, suppose the outcome of the homogeneity test is p = 0.03. Then, if the data set is really homogeneous, there is a probability of 3 % to obtain a data set that is farther away from satisfying homogeneity than the observed data set. For this reason, small p-values are interpreted as speaking against the validity of the hypothesis under test.

The question then becomes: how small does the p-value need to be to warrant rejecting the hypothesis under test? This is up to the user to decide. In the example above (p = 0.03), if the user rejects the hypothesis of homogeneity, then the user does so knowing that there is a probability of 3 % of this being the wrong decision.

The user needs to consider another feature of statistical tests of hypothesis: the probability of failing to reject a hypothesis that is false. The complementary of this probability, which is the probability of rejecting a false hypothesis, is called the power of the test. The smaller the data set a statistical test is based on, the lower its power. The power of a test also depends on the manner in which the data deviates from what the hypothesis under test describes.

The NIST Decision Tree was designed to offer substantial protection against nefarious consequences of wrong decisions that the user can make when interpreting the p-values issuing from the three tests it uses. The NIST Decision Tree suggests the following thresholds:

- Homogeneity test: It is safe to reject the hypothesis of homogeneity even for fairly high p-values. In this conformity, it is recommended that the hypothesis be rejected rather freely, when the p-value is 10 % or smaller.

- Normality test: The generally small numbers of participants in KCs call for a moderately conservative stance regarding distributional shape. Accordingly, it is recommended that the hypothesis of Gaussian shape be rejected only when the p-value is smaller than 5 %.

- Symmetry test: Since asymmetry can easily be confused with heaviness of the tails of the probability distribution of the laboratory or method effects, especially in small data sets as commonly arise in KCs, it is recommended that the hypothesis of symmetry be rejected only when the p-value is smaller than 1 %.

7 Bilateral degrees of equivalence

The bilateral degree of equivalence (also known as pair equivalence), Dij, between two participants is simply the difference between the values measured by any pair of participants, together with its associated expanded uncertainty for 95 % coverage, UDij:

Dij=Ti-Tj

UDij=U(Ti-Tj)

Once the final data set is fixed, the bilateral degrees of equivalence can be computed from the final data set. The definition and computation of a KCRV is not necessary for calculating the bilateral degrees of equivalence.

The only difficulty is the evaluation of the standard uncertainty uDij, because it is not simply given by the combination of the individual participant standard uncertainties, ui and uj, in quadrature.

Depending on the statistical model adopted for the measured values Ti and on the topology of the KC, uDij can involve other uncertainty components besides the reported uncertainties for Ti and Tj, which prevents presenting here a formula that is universally valid. Also, it may not be possible to evaluate either the uDi or the uDij in closed form, and only Monte Carlo based evaluation is possible.

A comprehensive treatment of this problem is beyond the scope of this document, so here we provide only an example of a situation that is often encountered in CCT KCs. For this specific example, we provide the formal approach (basically, the approach provided in the “Guide to the expression of uncertainty in measurement” [27], and the approach usually adopted by the CCT KC pilots. The NIST Decision Tree computes unilateral, but not bilateral degrees of equivalence. The NIST Consensus builder (available at https://consensus.nist.gov) computes both, including a correct evaluation of the expanded uncertainties that are an integral part of the degrees of equivalence.

We consider a KC in which each participant calibrates a transfer standard SPRT against its reference temperature and delivers the transfer SPRT to the pilot. The result of participant i is the resistance ratio Wil of the transfer standard l with standard uncertainty uil and degrees of freedom νil. Note that uil includes not only the uncertainty in generating the temperature Ti, but also the uncertainty in measuring the resistance ratio Wil. In the same way, another participant j measures the resistance ratio Wjm of the transfer standard m when it is in thermal equilibrium with its reference temperature Tj, with standard uncertainty ujm and degrees of freedom νjm.

The pilot P measures the resistance ratios WPl and WPm of transfer standards l and m measured by participants i and j, respectively, when they are in thermal equilibrium with the pilot’s reference temperature TP (to distinguish the two separate measurements, we denote TP with TPl and TPm), with standard uncertainties uPl and uPm and degrees of freedom νPl and νPm, respectively.

Although this is not the most complex case encountered in CCT KCs, it shows that, when computing uDij, correlations need to be considered. The history of CCT KCs reveals that the CCT community has demonstrated good competence in dealing with correlations.

7.1 Usual CCT approach

The bilateral degrees of freedom are obtained from the following equations (dWrdT is the derivative of the SPRT reference function with respect to temperature, calculated at the reference temperature):

Til-TPl=Wil-WPldWrdT

uTil-TPl=uil2+uPl2

Tjm-TPm=Wjm-WPmdWrdT

uTjm-TPm=ujm2+uPm2

Dij=Til-TPl-Tjm-TPm+δSl+δSm

uDij=uil2+ujm2+uAPl2+uAPm2+uSl2+uSm2

νij=uDij4uil4νil+ujm4νjm+uAPl4νAPl+uAPm4νAPm+uSl4νSl+uSm4νSm

UDij=kijuDij

In the model equation for Dij the terms δSl and δSm were added to take into account the drift of the transfer standards l and m over the period of the KC (δSl and δSm can be taken as zero but their uncertainty, uSl and uSm, respectively, and their number of degrees of freedom, νSl and νSm, respectively, must be taken into account when calculating the uncertainty in Dij).

In the expression for uDij, uAPl and uAPm are the Type A standard uncertainties of the pilot when measuring transfer standard l and m, respectively. The reason for using only Type A uncertainties of the pilot in this expression is that TPl and TPm are correlated, and the pilot Type B uncertainties were not included because they were assumed completely correlated.

The effective number of degrees of freedom νij of uDij is obtained by applying the Welch-Satterthwaite formula to the degrees of freedom νil, νjm, νAPl, νAPm, νSl and νSm of the corresponding standard uncertainties uil, ujm, uAPl, uAPm, uSl and uSm, respectively.

Finally, in the expression for UDij, the coverage factor kij is obtained from the Student’s distribution t0.95(υij) with the appropriate degrees of freedom υij.

7.2 Formal approach

The measurand Dij is obtained from four other quantities Til, TPl, Tjm, and TPm. To simplify the notations, we denote these variables x1, x2, x3 and x4, respectively. Dij is a function of x1, x2, x3 and x4:

Dij=fx1, x2, x3, x4=x1-x2- x3+x4

with correlation between the variables x2 and x4.

When the input quantities are correlated, the appropriate general expression for the combined uncertainty u2Dij is:

u2Dij=i=14∂f∂xi2u2xi+2i=13j=i+14∂f∂xi∂f∂xjuxi,xj

where uxi,xj is the estimated covariance associated with xi and xj.

Note that the first term of this expression is the usual term encountered when the input quantities are uncorrelated, and the second term arises when correlations among the input quantities are present.

In our specific example, there is correlation only between the input quantities x2 and x4. In this case the general expression above reduces to:

u2Dij=i=14∂f∂xi2u2xi+2∂f∂x2∂f∂x4ux2,x4

Moreover, ∂f∂x1=∂f∂x4=1 and ∂f∂x2=∂f∂x3=-1, so the expression reduces further to:

u2Dij=i=14u2xi-2ux2,x4

We only need to estimate the covariance associated with x2 and x4. If the estimates of  x2 and x4 (TPl and TPm, respectively) are obtained from n observations as:

x2=1nj=1nx2j

x4=1nj=1nx4j

The covariance ux2,x4 is obtained from:

ux2,x4=1n-1j=1nx2j-x2x4j-x4

8 Statistical models

Four kinds of statistical models are commonly used for meta-analysis and in interlaboratory studies, which can be fitted to sets of measurement results using either classical or Bayesian versions: common mean [28], common mean with inflated uncertainties [23], the fixed effects [29], and random effects [20].

Considering the features of CCT KC measurands (see discussion in Section 2), random effects models generally tend to be the most appropriate models for the measurement results of CCT KCs. This class of models comprises a great variety of specific models suitable for data sets like those illustrated in the four panels of Figure 2, as well as data sets exhibiting still other patterns, and to do so successfully even when all the results depicted should be included in the calculation of the KCRV. In this document we focus on random effects models, and on alternative techniques for estimating the parameters of these models, one of them being the KCRV.

A clarification is in order for the meaning of the term “random effects.” The effects in question describe whether participants tend to measure similar measurands high or low persistently. That is, these effects can be interpreted as biases. Modeling these effects as random variables (this being the reason why they are called “random”) does not really mean that they are variable: it means simply, and pragmatically, that their particular values are unknown to some extent, and that probability distributions are used to characterize this state of knowledge. In other words, the randomness here is epistemic, not stochastic.

As clarified in Section 3 (principle 5), the statistical model can detect dark uncertainty, which is evaluated as a variance in excess to that expected from the reported uncertainties. When the presence of dark uncertainty is observed, the CCT must decide if it should affect the DoEs. We believe the answer depends on the purpose that the DoEs serve. If the DoEs serve to express each participant’s belief about its capabilities, then they could exclude the uncertainty contribution from dark uncertainty. If the DoEs are intended to predict a future measurement result (either of one of the participants, or of a laboratory that did not participate in the KC but whose capabilities are judged to be similar to the participants), then the DoEs should recognize the dark uncertainty.

8.1 Random effects model

The random effects model is expressed by the model equation:

xi=μ+λi+εi

where xi is the value measured by participant i, μ is the true value of the measurand, λi is the laboratory effect of participant i, and εi is the laboratory-specific measurement error of participant.

The λi are modeled as non-observable outcomes of independent random variables with mean 0 and standard deviation τ: in some versions of the model their probability distribution is Gaussian, in others it can be Laplace (or double exponential), or some other still. Most often, and in the absence of compelling evidence to the contrary, the measurement errors εi are values of independent Gaussian random variables with mean 0 and standard deviations ui. These assumptions imply that the expected value (true mean) of the measured value is μ. This reflects the belief that the collective of the participants is focused on the true value of the measurand, even if some participants tend to measure high and others low (as described by the λi).

The λi can be distinguished from the εi because the reported uncertainties ui=u(xi) are an integral part of the measurement data. If the xi are more dispersed than the ui suggest that they should be, then the λi cannot all be zero, and a dark uncertainty τ is present. When the λi do not differ significantly from 0, there is no dark uncertainty (τ=0).

The model equation above is valid for the simplest comparison topology in which the same artifact is measured by all participants only one time. For the most general comparison topology, in which multiple artifacts are measured repeatedly by the participants, the model equation becomes [7]:

xi,j,k=μj+δi,j,k+λi+εi,j,k

where the indices i, j and k identify the participant, the artifact and the repeat of a participant’s measurement, respectively; μj is the true value of the j-th artifact; δi,j,k is the random error associated with the departure of the artifact value from the artifact model (e.g. variations due to transport) and is distributed about zero with a variance qj2;  λi is the bias in each participant’s measurement of the artifact value; εi,j,k is the random error associated with each participant’s measurement and is distributed about zero with a variance si2.

Thus, the variance in a participant’s single measurement of the artifact value is:

varxi,j,k=qj2+τ2+si2.

9 Statistical procedures for computing the Key Comparison Reference Value

Many different procedures have been used historically to compute the KCRV. The statistical model for the measured values alone does not stipulate the procedure that should be used to compute the KCRV, or the procedure to compute the DoEs. The optimal procedure depends both on the model and on the meaning of “optimal”. Next, we will review several of these procedures, and will explain the conditions under which they may be optimal. Refer to Borenstein et al [28], Hartung et al. [30], Koepke et al. [20] and Possolo et al. [21], for comprehensive reviews of estimation procedures for different kinds of models.

9.1 Simple average

The KCRV value xKCRV and its standard uncertainty uKCRV are given by:

xKCRV=1ni=1nxi

uKCRV2=1n(n-1)i=1nxi-xKCRV2

The simple average will be optimal in the sense of minimum mean squared error when the measured values are a sample from a single Gaussian distribution. In particular, there is no dark uncertainty (τ=0), all the xi are assumed to have the same standard uncertainty, but the standard uncertainties actually reported are tacitly assumed to be unreliable and all are disregarded.

9.2 Uncertainty-weighted average

The KCRV value xKCRV and its standard uncertainty uKCRV are given by:

xKCRV=i=1nwixii=1nwi

wi=1ui2

1uKCRV2=i=1n1ui2

The uncertainty-weighted average will be optimal in the sense of minimum mean squared error when the measured values are outcomes of Gaussian random variables with the same mean but possibly different standard deviations that are reliably estimated by the reported uncertainties. In particular, there is no dark uncertainty (τ=0). The formula for uKCRV also assumes that the reported uncertainties are based on infinitely many numbers of degrees of freedom.

9.3 Median

The KCRV value xKCRV and its standard uncertainty uKCRV are given by (when the participant values xi are arranged in increasing order):

xKCRV=12xm2+x(m2)+1,  &n evenxm+12,  &n odd

uKCRV2=π2nσ2

σ=1.483∙MAD

where MAD is the median of the absolute deviations: MAD=medianxi-xKCRV.

However, the formula given for the standard uncertainty σ is generally incorrect for two reasons: i) the measured values can have rather different standard uncertainties, hence a weighted median will generally be more appropriate than the simple median; ii) the formula applies to the median of large samples drawn from a Gaussian distribution. If the assumption in ii) is valid, then there is no reason to use the median to begin with.

The median will be optimal in the sense of minimum mean absolute error when the measured values are a sample from a single Laplace distribution. In particular, there is no dark uncertainty (τ=0), all the xi are assumed to have the same standard uncertainty, but the standard uncertainties actually reported are tacitly assumed to be unreliable and all are disregarded. If σ is the true, common standard deviation of the Laplace measurement errors, then the standard uncertainty of the median is σ2n, for large n.

In general, the formula for the standard uncertainty of the median involves the value that the probability density of the measured values takes at the median of their common distribution, and it provides no more than an approximation whose accuracy increases with the number, n, of participants.

Considering that the number of participants in a KC is generally small, the better alternative is to evaluate the standard uncertainty of the median as half the length of a 68% confidence interval for the median by analogy with the Gaussian case (Possolo and Bodnar [31], which can be computed in fair generality as described by Hettmansperger and Sheather [32]).

9.4 DerSimonian-Laird procedure (Adaptative Weighted Average, AWA)

The procedure suggested by DerSimonian and Laird [33] is suitable to fit the random effects model to the results from KCs both when τ>0 and when τ=0, and otherwise does not make specific distributional assumptions about either the λi or the εi introduced in Section 8.1. When τ=0, the procedure reduces to the conventional weighted average discussed in Section 9.2.

The KCRV value is a weighted average of the values xi measured by the participants, with weights wi=1ui2+τ2 (note the difference from the weights used in 9.2):

xKCRV=i=1nwixii=1nwi

wi=1ui2+τ2

An estimate of τ is obtained from:

τ=max0,τM

τM2=Q-n+1i=1nui-2-i=1nui-4i=1nui-2

Q=i=1nui-2xj-xKCRV2

The standard uncertainty of the KCRV is given by:

uKCRV=1i=1nwi

Neither this uncertainty evaluation for the KCRV, nor the uncertainty evaluations for the DoEs, take into account the fact that the estimate of τ above is based on the dispersion of the measured values, hence typically is supported by a fairly small number of degrees of freedom. The Adaptative Weighted Average implemented by the NIST Decision Tree [21] is a generalization of the DerSimonian-Laird procedure that addresses this shortcoming.

9.5 Maximum likelihood

This procedure is described by White [34] and by Koepke et al. [20] (subsection 5.1) in the metrological context, and in general by Searle et al. [35]. In this method the maximum likelihood estimate of the KCRV is obtained by maximizing the likelihood function that corresponds to a random effects model with Gaussian λi and εi, with respect to two parameters: the overall mean (μ in the model equation), and the dark uncertainty (the standard deviation τ of the laboratory effects in the model equation), assuming that the reported uncertainties are based on very large numbers of degrees of freedom. An approximation to the standard uncertainty associated with the KCRV is obtained as a side-effect, so that the larger the number of participants, the better the quality of the approximation.

R function “rma,” defined in package ‘metafor’ [36], can fit both the random effects and fixed effects variants of the model via maximum likelihood estimation (MLE). The random effects version of the model can also be fitted (again using function ‘rma’) via restricted maximum likelihood (REML) estimation, which is generally preferable to MLE, especially when the number of participating laboratories is small [37].

9.6 Linear pool

The linear pool is described by Koepke et al. [20]. It consists of representing each measurement result (measured value xi and associated uncertainty ui) by a probability density function fix (Gaussian or some other), and computing a weighted average of these functions gx=i=1nwifixi=1nwi. The KCRV is then obtained as the expected value of the distribution with probability density gx, which can easily be computed by Monte Carlo methods, and the standard uncertainty associated with the KCRV is the standard deviation of the same distribution.

The NIST Consensus Builder (available at https://consensus.nist.gov) can perform linear pooling, with or without weights, using Gaussian or (rescaled and shifted) Student’s t distributions, and produce the KCRV, associated uncertainty, and corresponding unilateral and bilateral degrees of equivalence.

Figure 3 shows the fix for CCT-K7.2021, chosen to be Gaussian with means equal to the values measured by the participants, and with standard deviations equal to their reported uncertainties. The linear pool has a mean -0.4 µK and standard deviation 41.4 µK.

Figure 3: CCT-K7.2021 Key Comparison of triple-point-of-water cells: individual participants’ distributions (thin curves) and linear pool distribution of results (thick black curve).

9.7 The NIST Decision Tree

The NIST Decision Tree (NDT) [2] is a useful tool that can be used in the analysis of KC results. The NDT is a web-based application that guides the user through a series of questions whose answers are informed by the results of statistical tests, to help the user decide on an appropriate statistical model and a data reduction procedure for the particular data set input by the user. It works as follows: 1) the user inputs the data set (the measured value and the uncertainty reported of each participant) into a spreadsheet of the NDT; 2) the NDT performs statistical tests (homogeneity, symmetry and normality) on the data set; 3) the user, based on the results of the statistical tests, on the recommendation that the NDT makes, and on their scientific knowledge and judgement, selects the statistical procedure; and 4) the NDT carries out the analysis for the selected statistical procedure, and displays the results of the data reductions, including the KCRV and the degrees of equivalence, in the form of plots, tables and a downloadable PDF report.

A flow diagram of the NDT is shown in Figure 4. The NDT is traversed by answering a question at each node (orange), and moving forward according to the answer given to the question: if the answer is yes, you move to the left (green branch), if the answer is no, you move to the right (red branch), until a leaf (blue) is reached, which is the procedure recommended by the NDT for the input data set.

Figure 4: Flow diagram of the NIST Decision Tree.

The NDT has been successfully used in the analysis of CCT-K7.2021 [38].

9.8 Degrees of freedom

The concept of number of degrees of freedom that is used in the statistical sciences conveys the number of independent pieces of information that underlie an estimate of a parameter of a probability distribution.

For example, the sample standard deviation is a function of the differences between the observations and their average. Since these differences add to zero, there are only n-1 independent such differences, where n denotes the number of observations. For this reason, we say that the sample standard deviation has n-1 degrees of freedom.

More generally, if one has n observations, and fits a statistical model to them that requires the estimation of k parameters, then the residuals (differences between observed and fitted values) have n-k degrees of freedom, and the calculation of the variance of these residuals should take this fact into account: the sum of the squared residuals should be divided by n-k.

In the context of metrology, the number of degrees of freedom associated with an uncertainty evaluation is an indicator of the reliability of the evaluation. This understanding harks back to the fact that n-1sσ2 has a chi-squared distribution with n-1 degrees of freedom where s is a sample standard deviation of a sample of size n drawn from a Gaussian distribution with standard deviation σ. This fact implies that, the larger the n, the closer to σ does s tend to be.

The numbers of degrees of freedom supporting the uncertainties that are part of the measurement results of a KC are similarly informative about the reliability of those uncertainties. If a weighted average is an appropriate KCRV, then the weights should reflect not only the values of the reported uncertainties, but also their numbers of degrees of freedom. This can be achieved using likelihood-based methods of data reduction: for example, (restricted) maximum likelihood estimation and Bayesian methods.

The dark uncertainty that affects many sets of measurement results obtained in KCs typically is estimated based on a very small number of degrees of freedom, yet plays an important role in the calculation of the KCRV and in the evaluation of its associated uncertainty. The DerSimonian-Laird method of data reduction uses an estimate of dark uncertainty but neglects the reliability of this estimate, hence tends to produce overoptimistic results.

10 The selection of the statistical procedure to be used for computing the Key Comparison Reference Value

Beyond the CIPM general statement about the KCRV given in [3-4], no guidance is given about how to determine the KCRV of a KC.

The selection of the statistical procedure for calculating the KCRV should be based on:

- Visual (graphical) inspection of the final data set (see section 3);

- The outcomes of the applied statistical tests (see section 7);

- The selected statistical model for the KC measurement (see section 8);

- The scientific knowledge and judgement of the participants.

1 and 2 allow the pilot to establish which of the 4 regions depicted in Figure 2 best represents the final data set. Based on the past history, the results of CCT KCs are all represented by one of the following three regions (see Table 2):

- Region A: mutually consistent results with no overdispersion and no outliers.

- Border region between A and B: some overdispersion (small excess variance) and no outliers.

- Region C: general consistency of results with a small number of outliers.

KC results represented by region D have not occurred, so in the following we will not consider this case.

| Key Comparison | Regions of Figure 1 | Statistical procedure used for computing the KCRV |
| --- | --- | --- |
| CCT-K1 | Border region between A and B | Weighted average |
| CCT-K2 | Region C | Weighted average |
| CCT-K3 | Border region between A and B | Not computed |
| CCT-K4 | Border region between A and B | Weighted average |
| CCT-K5 | Region C | Median |
| CCT-K6 | Border region between A and B | Weighted average |
| CCT-K7 | Border region between A and B | Simple average |
| CCT-K8 | Region A | Weighted average |
| CCT-K9 | Border region between A and B | Simple average |
| CCT-K10 | Region C | Median |
| CCT-K7.2021 | Border region between A and B | DerSimonian-Laird |

Table 2: Regions of Figure 1 representing the results of each CCT Key Comparison (second column) and corresponding adopted statistical procedure for calculating the KCRV (third column).

Moreover, as discussed in section 8, the most appropriate statistical model for CCT KCs is the random effects model, which allows laboratory bias λi having a non-zero standard deviation τ (dark uncertainty), which is evidence of some over-dispersion.

10.1 Results represented by region A

When the results are mutually consistent (region A, no outliers and no overdispersion), the uncertainty-weighted average is the appropriate method to be used. However, the DerSimonian-Laird procedure (adaptative weighted average) can also be used. The standard deviation τ will simply be close to zero and the KCRV value will coincide with the KCRV resulting from the application of the ordinary uncertainty-weighted average.

10.2 Results represented by border between region A and region B

When the results are generally consistent but some overdispersion is present (evidenced by τ≠0), the Der-Simonian Laird procedure is the appropriate method to be used. The uncertainty-weighted average should not be used when there is appreciable risk of significant over-dispersion. The Der-Simonian Laird procedure prevents small participant-specific uncertainties from having an excessive influence on the KCRV.

10.3 Results represented by region C

When the results are generally consistent but one or more outliers are present an estimation procedure that is robust (resistant) to the presence of outliers should be used. The median is a simple robust estimator, though there are more efficient robust estimators (M- and MM-estimators).

11 Conclusions

The main goal of this document is to provide guidance to CCT KC pilots and participants in the analysis of the results of CCT key comparisons. Although the analysis of KC results is traditionally addressed in the last stage of a KC lifetime, we recommend CCT KC pilots and participants to consult this document in the earliest phase of a KC, ideally when preparing the KC Technical Protocol.

There is no magic analysis method that is suitable for every possible KC data set (as there is no magic analysis method that can save poor measurement results). The KC pilots and participants, based on the results of the applied statistical tests and on their unique scientific knowledge of the measurement, have to select the most appropriate statistical model to be applied to the measurement. The reduction method (for the calculation of the KCRV and the unilateral degrees of equivalence) will follow naturally.

In spite of the unavoidable degree of discretion in the selection of the analysis method, this document made an effort harmonize all aspects of CCT KC analysis, trying to achieve, at least within the CCT community, a common understanding of the meaning to be attributed to the KCRV, and a common agreement on the general principles to be adopted when analyzing KC results.

Based on the past CCT KCs, we described only the analysis methods that are relevant for the CCT community. Obviously, by doing so, we limited the validity of this document to only CCT KCs.

Acknowledgements

The members of the CCT Working Group for Key Comparisons would like to express their deep gratitude to Dr. Antonio Possolo for his invaluable support.

References

[1] CCQM Guidance note “Estimation of a consensus KCRV and associated Degrees of Equivalence,”

[2] D. Newton, A. Koepke, A. Possolo, M. Winchester, “NIST Decision Tree User’s Manual”, 3rd edition, National Institute of Standards and Technology, 2024

[3] Bureau International des Poids et Mesures 1999 Mutual Recognition of National Measurement Standards and of Calibration and Measurement Certificates issued by National Metrology Institutes

[4] CIPM-MRA-G11 Measurement Comparisons in the CIPM MRA: Guidelines for organizing, participating and reporting (2021)

[5] C. Elster, B. Toman, Metrologia 50 (2013) 549-555

[6] U. Panne, W. Bremser, Accred Qual Assur (2009) 14: 281-287

[7] R. Rusby, D. Head, C. Meyer, W. Tew, O. Tamura, K.D. Hill, M. de Groot, A. Storm, A. Peruzzi, B. Fellmuth, “Final report on CCT-K1: Realizations of the ITS-90, 0.65 K to 24.5561 K, using rhodium-iron resistance thermometers,” 2006 Metrologia 43 03002

[8] A.G. Steele, B. Fellmuth, D.I. Head, Y. Hermier, K.H. Kang, P.P.M. Steur, W.L. Tew, “CCT-K2: key comparison of capsule-type standard platinum resistance thermometers from 13.8 K to 273.16 K,” 2002 Metrologia 39 551

[9] B.W. Mangum, G.F. Strouse, W.F. Guthrie, R. Pello, M. Stock, E. Renaot, Y. Hermier, G. Bonnier, P. Marcarino, K.S. Gam, K.H. Khang, Y.-G. Kim, J.V. Nicholas, D.R. White, T.D. Dransfield, Y. Duan, Y. Qu, J. Connolly, R.L. Rusby, J. Gray, G.J. Sutton, D.I. Head, K.D. Hill, A. Steele, K. Nara, E. Tegeler, U. Noatsch, D. Heyer, B. Fellmuth, B. Thiele-Krivoj, S. Duris, A.I. Pokhodun, N.P. Moiseeva, A.G. Ivanova, M.J. de Groot, J.F. Dubbeldam, 2002 Metrologia 39 179

[10] H.G. Nubbemeyer, J. Fischer, “Final report on key comparison CCT-K4 of local realizations of aluminium and silver freezing-point temperatures,” 2002 Metrologia 39 03001

[11] M.J. de Groot, E.W.M. van der Ham, R. Monshouwer, “CCT-K5: Comparison of local realizations of the ITS-90 between the silver point and 1700 °C using vacuum tungsten strip lamps as transfer standards”

[12] S. Bell, M. Stevens, H. Abe, R. Benyon, R. Bosma, V. Fernicola, M. Heinonen, P. Huang, H. Kitano, Z. Li, J. Nielsen, N. Ochi, O.A. Podmurnaya, G. Scace, D. Smorgon, T. Vicente, A.F. Vinge, L. Wang, H. Yi, “Final report to the CCT on key comparison CCT-K6 – Comparison of local realisations of dew-point temperature scales in the range -50 °C to +20 °C,” 2015 Metrologia 55 03005

[13] M. Stock, S. Solve, D. Del Campo, V. Chimenti, E. Mendez-Lango, H. Liedberg, P.P.M. Steur, P. Marcarino, R. Dematteis, E. Filipe, I. Lobo, K.H. Kang, K.S. Gam, Y.-G. Kim, E. Renaot, G. Bonnier, M. Valin, R. White, T.D. Dransfield, Y. Duan, Y. Xiaoke, G. Strouse, M. Ballico, D. Sukkar, M. Arai, A. Mans, M. de Groot, O. Kerkhof, R. Rusby, J. Gray, D. Head, K. Hill, E. Tegeler, U. Noatsch, S. Duris, H.Y. Kho, S. Ugur, A. Pokhodun, S.F. Gerasimov, “Final report on CCT-K7: Key comparison of water triple point cells”, 2006 Metrologia 43 03001

[14] J.G. Gallegos, H. Mitter, H. Abe, S. Bell, R. Benyon, P. Carroll, B.I. Choi, R. Derschermeir, V. Ebert, V. Fernicola, L. Wang, C. Meyer, D. Smorgon, T. Vicente, “Final Report to the CCT on key comparison CCT-K8 – Comparison of local realisations of dew-point temperature of humid gas in the range 30 °C to 95 °C” 2025 Metrologia 62 03001 DOI 10.1088/0026-1394/62/1A/03001

[15] T. Herman, M. Chojnacky, K. Hill, S. Rudtsch, I. Yang, P.P.M. Steur, R. Dematteis, G. Lopardo, F. Sparasci, C. Martin, L. RIsegari, J. Widiatmo, T. Nakano, I. Saito, K. Quelhas, P. Giorgio, J. Sun, J. Zhang, J. Pearce, J. Gray, “CCT-K9 Final Report – ITS-90 SPRT calibration from the Ar TP to the Zn FP,” 2023 Metrologia 60 03001

[16] H. McEvoy, D. Lowe, G. Machin, M.J. Martin, K. Anhalt, J. Hollandt, M. Sadli, F. Bourson, Y. Yamada, N. Sasajma, X. Lu, Y. Yoo, H. Yoon, C. Gibson, A. Todd, D. Woods, “Final report for CCT comparison of ITS-90 realisations above the silver point using two transfer radion thermometers and a set of high temperature fixed-point blackbody cells,” 2024 Metrologia 61 03003

[17] A. Peruzzi, S. Dedyulin, M. Levesque, D. Del Campo, B.C. Garcia Izquierdo, M.E. Gomez, K.N. Quelhas, M.A.P. Neto, B.M. Lozano, L. Eusebio, I. Yang, F. Sparasci, A. Martin, L. Risegari, P. Saunders, E. Molloy, X.K. Yan, J.P. Sun, X.J. Feng, J.T. Zhang, M.-K. Ho, T. Nakano, J.V. Widiatmo, I. Saito, E. Ejigu, J. Pearce, S. Rudtsch, L. Buenger, M. Kalemci, A. Uytun, C. Bruin-Barendregt, M. Panman, D.R. White, A. Possolo, “CCT-K7.2021: CIPM key comparison of water-triple-point cells,” 2023 Metrologia 60 03002.

[18] H. Preston-Thomas 1990 Metrologia 27 107

[19] C. Gaiser, B. Fellmuth, R.M. Gavioso, M. Kalemci, V. Kytin, T. Nakano, A. Pokhodun, P.M.C. Rourke, R. Rusby, F. Sparasci, P.P.M. Steur, W.L. Tew, R. Underwood, R. White, I. Yang, J. Zhang, Update for the differences between thermodynamic temperature and ITS-90 below 335 K, J. Phys. Chem. Ref. Data 51, 043105 (2022) https://doi.org/10.1063/5.0131026

[20] A. Koepke, T. Lafarge, A. Possolo, B. Toman, “Consensus building for interlaboratory studies, key comparisons, and meta-analysis,” Metrologia 54 (2017) S34-S62.

- [21] A. Possolo, A. Koepke, D. Newton, M.R. Winchester (2021) “Decision Tree for Key Comparisons,” J Res Natl Inst Stan 126:126007.

- [22] J. Mandel, The validation of measurement through interlaboratory studies, 1991 Chemometrics and Intelligent Laboratory Systems, 11, 109-119 https://doi.org/10.1016/0169-7439(91)80058-X

- [23] A. Possolo, Tracking Truth through Measurement and the Spyglass of Statistics, 2023 Statistical Science 38 (4): 655-671 https://doi.org/10.1214/23-STS899

- [24] W.G. Cochran (1954) “The combination of estimates from different experiments,” Biometrics 10(1):101-129.

- [25] S.S. Shapiro, M.B Wilk (1965) “An analysis of variance test for normality,” Biometrika 52(3,4): 591-611.

- [26] W. Miao, Y.R. Gel, J.L. Gastwirth (2006) “A new test of symmetry about an unknown median,” Random Walk, Sequential Analysis and Related Topics: A Festschrift in Honor of Yuan-Shih Chow, eds A.C. Hsiung, Z. Ying, C.H. Zhang, World Scientific Singapore, pp 199-214.

- [27] “Guide to the expression of uncertainty in measurement”, International Organization for Standardization, 1995.

- [28] M. Borenstein, L.V. Dedges, J.P.T. Higgins, H.R. Rothstein, 2010, A basic introduction to fixed-effect and random-effects models for meta-analysis, Research Synthesis Methods Vol. 1 pages 97-111.

- [29] B. Toman, A. Possolo, 2009, Laboratory effects models for interlaboratory comparisons, Accreditation and Quality Assurance, Vol. 14, Issue 10 553-563.

- [30] J. Hartung, G. Knapp, B.K. Sinha, (2008) Statistical Meta-Analysis with Applications, John Wiley & Son, Hoboken, NJ ISBN 978-0-470-29089-7

- [31] A. Possolo, O. Bodnar, Redefining Standard Measurement Uncertainty, Ukrainian Metrological Journal 2022(1): 34-37 DOI: 10.24027/2306-7039.1.2022.258815

- [32] T.P. Hettmansperger, S.J. Sheather, Confidence intervals based on interpolated order statistics, Statistics & Probability Letters, 1986 4(2) 75-79 DOI: 10.1016/0167-7152(86)90021-0

[33] R. DerSimonian, N. Laird (1986) “Meta-analysis in clinical trials,” Controlled Clinical Trials, 7 pp. 177-188

[34] D.R. White, On the analysis of measurement comparisons, Metrologia 41 (2004) 122-131

[35] S.R. Searle, G. Casella, C.E. McCulloch (1992) “Variance components,” John Wiley & Sons, New York, NY, ISBN 978-0-470-00959-8

- [36] W. Vlechtbauer (2010), “Conducting meta-analyses in R with the metafor package,” Journal of Statistical Software 36(3):1:48. DOI 10.18637/jss.v036.i03

- [37] A. Rukhin, B. Biggerstaff, M. Vanghel, Restricted maximum likelihood estimation of a common mean and the Mandel-Paul algorithm, Journal of Statistical Planning and Inference 83: 319-330 DOI: 10.1016/S0378-3758(99) 00098-1

- [38] A. Peruzzi et al. CCT-K7.2021: CIPM key comparison of water-triple-point cells, 2023 Metrologia 60 03002 DOI: 10.1088/0026-1394/60/1A/03002

Appendix I Overview of the analysis of the past CCT Key Comparisons

CCT-K1 [7]

In this KC, 7 participants copied their ITS-90 realizations in the temperature range from 0.65 K to 24.5561 K (3He/4He vapour pressure thermometers and 3He/4He interpolating constant-volume gas thermometers) to Rh-Fe resistance thermometers, delivered one or two calibrated Rh-Fe thermometers to the pilot, and all the Rh-Fe thermometers were intercompared in the pilot’s cryostat at 43 different temperatures.

For each Rh-Fe thermometer, the participants provided the interpolation equation Ti = Ti (Ri) and the calibration standard uncertainty ui. The experimental uncertainty on the measurement made in the comparison cryostat, estimated by the pilot, was uexp. The combined standard uncertainty of participant i for the comparison was then uCi=ui2+uexp2 .

The bilateral equivalence between participant i and j was simply Dij = Ti – Tj with standard uncertainty uDij given by uDij=ui2+uj2+2uexp2 , which is based on the assumption of no correlation between the uCi and uCj.

The KCRV was defined as the weighted mean with weights uCi=ui2+uexp2 :

TKCRV=iTiui2+uexp2i1ui2+uexp2

The participants agreed that no additional significance was attached to the KCRV with respect to ‘best estimates’ of T90. KCRV uncertainty was not computed.

CCT-K2 [8]

In this KC, each of the 7 participants calibrated 2 capsule-type SPRTs (CSPRTs) at the ITS-90 fixed points in the temperature range 13.8 K to 273.16 K, delivered the calibrated CSPRTs to the pilot, and all the CSPRTs were intercompared in the pilot’s cryostat at temperatures close to the 8 low-temperature defining fixed points of the ITS-90.

For each CSPRT, the participants provided the resistance ratios W(Ti) measured at the 8 calibration points, and the corresponding calibration standard uncertainty ui. The experimental uncertainty on the measurement made in the comparison cryostat, estimated by the pilot, was uexp. The combined standard uncertainty of participant i for the comparison was then uCi=ui2+uexp2 .

The bilateral equivalence between participant i and j was simply Dij = Ti – Tj with standard uncertainty uDij given by uDij=ui2+uj2+2uexp2 , which is based on the assumption of no correlation between the uCi and uCj.

The KCRV was defined as the weighted mean with weights uCi=ui2+uexp2 :

TKCRV=iTiui2+uexp2i1ui2+uexp2

The KCRV represents “a common baseline against which all laboratory values can be compared”. KCRV uncertainty was not computed.

CCT-K3 [9]

In this KC, 16 participants calibrated long-stem SPRTs (LSPRTs), provided by the pilot, at the ITS-90 defining fixed points in the range 83.8058 K to 933.473 K, and all SPRTs were measured at the pilot’s fixed points.

For each LSPRT, the participants provided the resistance ratios W(Ti) measured at the 7 fixed points, and the corresponding calibration combined standard uncertainty ui. The participants were asked to report the Type A uncertainty (SAini with SAi the repeatability for ni realizations), the Type B uncertainty SBi, and the corresponding degrees of freedom νAi and νBi. The combined standard uncertainty ui was thus given by: ui=SAi2ni+SBi2 and its degrees of freedom are obtained by applying the Welch-Satterthwaite formula (see [20]).

The bilateral equivalence between two participants i and j that measured only one and the same LSPRT was given by:

Dij=Ti-Tj=Wi-WjdWrdT

UDij=ki2SAi2ni+SBi2+kj2SAj2nj+SBj2+kSPRT2uSPRT2

where uSPRT is the uncertainty due to the drift of the LSPRT, and ki,j,SPRT are the coverage factors obtained from the t-distribution with νi,j,SPRT degrees of freedom: ki=t0.975(νi).

When participants i and j measured different LSPRTs, the bilateral equivalence was given by:

Dij=Ti-TPilot-Tj-TPilot

U2Dij=ki2SAi2ni+SBi2+kj2SAj2nj+SBj2+kNIST2SANIST2nNIST1+SANIST2nNIST2+kSPRT12uSPRT12+kSPRT22uSPRT22

It was decided that KCRV would not be given for this report. Although a minority of participants (5) thought that KCRV should be used, they withdrew their objections in order for the report to be completed.

CCT-K4 [10]

In this KC, the pilot circulated Al and Ag cells to the 3 co-pilots, and the co-pilots circulated their own Al and Ag cells to their loop participants (12 participants in total).

The analysis approach was similar to that of CCT-K3. The expanded uncertainties were calculated using a coverage factor ki=t0.95(νi) from t-distribution and the number of degrees of freedom for Type A uncertainties. Type B uncertainties were assumed to have an infinite number of degrees of freedom.

Given the comparison scheme, three different cases were considered. In the simplest case, in which the cell measured by participant i is the reference cell of participant j:

UTi-Tj=Ui2+Uj2-t0.952νjuAj2+UC2

where UC is the uncertainty associated to the transport of the cell.

In the case in which participant i and participant j belongs to the same loop but none of the two is co-pilot, the expanded uncertainty is:

UTi-Tj=U2Ti-Tcop+U2Tj-Tcop-2UBcop2

Where cop in the subscripts refers to co-pilot and U2Ti,j-Tcop are calculated using the previous equations.

In the most general case, in which participant i and participant j belong to two different loops and none of the two is co-pilot, the expanded uncertainty is:

U2Ti-Tj=U2Ti-Tcopi+U2Tpil-Tcopi+U2Tj-Tcopj+U2Tpil-Tcopj-2UBcopi2-2UBcopj2-2UBpil2

where pil in the subscripts refers to pilot.

The KCRV was calculated as the weighted mean of all the differences Ti-Tpil, with weights ui2+uC2:

TKCRV=iTi-Tpilui2+uC2i1ui2+uC2

The degrees of equivalence were defined by:

Di=Ti-TKCRV

uDi=ui2+uC2

CCT-K5 [11]

Two sets of two vacuum tungsten-strip lamps were circulated among the 14 participants in two geographical loops separated by pilot and assistant pilot measurements. The participants copied their ITS-90 local scales to the transfer lamps at 11 temperatures in the range 960 °C to 1700 °C.

The results, obtained in the two separate loops on two different set of lamps were linked through the pilot and assistant pilot measurements, who measured both sets of lamps.

For the calculation of the KCRV, some measurements were not included in the KCRV calculation due to problems with the measurement system. The median was chosen for calculating the KCRV.

CCT-K6 [12]

A pair of commercial chilled-mirror hygrometers (CMHs) was circulated among the 10 participants in two geographical loops separated by pilot measurements. The participants generated 5 dewpoint temperatures trefi from -50 °C to 20 °C using their primary dewpoint generators and measured them using the travelling CMHs. The dew-point value ti indicated by the travelling CMHs was the mirror PRT resistance, converted to a nominal temperature using IEC 60751 resistance-temperature characteristic. Each participant reported Ti=trefi-ti and the corresponding combined standard uncertainty ui.

The bilateral equivalence was given by:

Dij=Ti-Tj=trefi-ti-trefj-tj

uDij=ui2+uj2+udrift2

UDij=2uDij

The KCRV was calculated as the weighted mean of all results:

TKCRV=iTiui2i1ui2

1u2TKCRV=i1ui2

The results of one of the participants at two temperatures were excluded from the calculation of the KCRV. It was noted that the KCRV has no absolute significance outside the comparison, though it is essential to the use of comparison results for review of CMC claims.

CCT-K7 [13]

In this KC, each of the 21 participants calibrated a transfer cell against their national reference for the triple point of water and delivered it to the pilot. The pilot intercompared all transfer cells.

Indicating by Ttri the temperature realized by the transfer cell of participant i and by Trefi the temperature realized by the national reference of participant i, each participant measured Ttri-Trefi with standard uncertainty ui.

The pilot measured the differences Ttri-Trefpilot. From the participants’ measurements and the pilot’s measurements, the differences Trefi-Trefpilot could be calculated, with uncertainties:

u2Trefi-Trefpilot=u2Trefi-Trefpilot+u2Ttri-Trefi

The uncertainties in the two differences were considered totally uncorrelated.

The KCRV was calculated as the simple mean of the results from all the participants:

TKCRV-Trefpilot=1NiTrefi-Trefpilot

The uncertainty of the KCRV was calculated as the standard deviation of the mean:

uTKCRV-Trefpilot=iu2(Trefi-Trefpilot)N-1

CCT-K8 [14]

A pair of commercial chilled-mirror hygrometers (CMHs) was circulated among the 10 participants in two loops. The participants generated 7 dewpoint temperatures trefi from 30 °C to 95 °C using their primary dewpoint generators and measured them using the travelling CMHs. The dew-point value ti indicated by the travelling CMHs was the mirror PRT resistance, converted to a nominal temperature using IEC 60751 resistance-temperature characteristic. Each participant reported Ti=trefi-ti and the corresponding combined standard uncertainty ui.

The bilateral equivalence between participant i belonging to loop loop(i) and participant j belonging to loop loop(j) was obtained as:

Dij=ti-trefiloopi+tj-trefjloopj+Bij

The term Bij is a discrete function of i and j, used to link the two loops, obtained from the measurements at the linking labs.

Because the two loops measured different travelling standards, it was necessary to generate a virtual transfer standard having an intermediate behaviour between the two real travelling standards. This was done by applying a correction factor Cloopi to both loops (the fact that the linking labs measured all travelling artifacts 4 times made it possible to calculate the correction factors):

Cloop1=Bij2

Cloop2=-Bij2

u(Cloopi)=12u(Bij)

The KCRV was calculated as the weighted mean of all results with one exception: at the 3 highest temperatures, the results of one participant were excluded from the calculation of the KCRV, because its standard was metrologically traceable to another participant. However, the KCRV has no absolute significance.

CCT-K9 [15]

Fourteen NMI laboratories participated in this comparison. Each participating NMI selected two of their own SPRTs, calibrated them at every available ITS-90 fixed point within the range 83.8058 K to 692.677 K, and transported them to the pilot laboratory (NIST). Following the NIST calibration, the SPRTs were returned to their respective owners, and each of them recalibrated the SPRTs at their facility.

The calibrations were performed in terms of resistance ratio W at each fixed point. For each fixed point, each participant i measured the initial values of W for their two SPRTs: Wi1Ante and Wi2Ante. The SPRTs were then transported to, and measured at NIST: Wi1Pilot and Wi2Pilot. Finally, the SPRTs were returned to their respective NMIs for a final measurement: Wi1Post and Wi2Post.

The result of participant i was then (j identifies either SPRT 1 or SPRT 2):

Wij=12WijAnte+WijPost

The temperature difference ∆Tij between participant i and the pilot, calculated for SPRT j was:

∆Tij=Wij-WijPilotdWr/dT+Cij

where Cij accounts for the drift of the SPRT.

At each fixed point and for each participant i, the ∆Tij are combined to give ∆Ti:

∆Ti=12∆Ti1+∆Ti2

The KCRV ∆T was calculated as:

∆T=1ni=1n∆Ti

The SPRTs that failed the cut off criteria were excluded from the calculation of the KCRV.

The DoEs were calculated as:

Di=∆Ti-∆T

The uncertainty in the KCRV and in the DoEs were calculated in the same way of CCT-K3. Systematic uncertainties were treated as completely correlated and combined according to GUM section 5.2.2, the most conservative approach.

CCT-K10 [16]

CCT-K10 consisted of two parts:

1. Two transfer radiation thermometers and a transportable copper fixed-point blackbody source were circulated among the participants to compare their ITS-90 scale realizations at 14 temperatures over the range from the Ag point to 3000 °C.

2. Three high temperature fixed-point (HTFP) blackbody cells were circulated among the participants to probe their scale realizations at these particular temperatures.

The transfer instruments were circulated in the form of a semi-collapsed star (or flower), in which the artefacts were sent from the pilot laboratory to one RMO for a loop around that region, and then the artefacts were sent back to the pilot laboratory. Following check measurements at the pilot laboratory, the artefacts were then sent to the next RMO region for a further loop.

Initially the analysis of the results was carried out according to the method described in the protocol (weighted mean with cut off). However, due to the presence of some outlying data, and after further discussion amongst the participants, it was decided to use the median as the reference value (including the average pilot results) for the part of the comparison using radiation thermometers. For the comparison with HTFPs, the reference value was calculated as the weighted mean with cut off, as specified in the protocol.

The uncertainty u in the median value was calculated using the ‘median of absolute deviations’ (or ‘MAD’) as:

u≅1.858(n-1)MAD

where n is the number of results and

MAD=median{xi-med}

where xi is the result of laboratory i and med is the median value.

The DoEs were calculated for each participant for each thermometer and for each nominal temperature, using the difference between the result of each participant and the reference value (median) for that particular thermometer, and the uncertainty of the difference, which is the participant uncertainty (including a component for the thermometer drift) combined with the uncertainty of the reference value. The pair equivalence between each pair of participants was also determined.

Our definition deviates from that given in [3], which, although given only in words, seems to be Dij ≡ Di – Dj ≡ (xi – xKCRV) – (xj – xKCRV) and which involves the KCRV.