# Updated report on the evaluation of

# degrees of equivalence in regional dosimetry comparisons

# D T Burns (BIPM) and D Butler (ARPANSA)

This report is a revision of that presented in 2007 as CCRI(I)/07-04, with two main changes: (i) removal of the evaluation of pairs of degrees of equivalence, as these are no longer published in the KCDB, and (ii) advice on the evaluation of utr (formerly ustab) and uLINK, including the treatment of data for which the observed variations are larger than predicted from the relevant uncertainties.

The analysis assumes in the first instance a single linking laboratory, LINK, and a regional comparison of a quantity, which we take to be air-kerma rate, by measurement of the calibration coefficient NK for a single transfer chamber. The link to the BIPM is by means of an existing BIPM international comparison by the LINK for this quantity. The analysis is then generalized to multiple transfer chambers and linking laboratories.

The aim is a simplified analysis resembling the analysis that would be made for a comparison with the BIPM rather than linked through a regional comparison. To this uncertainty two components are added, utr arising from the transfer chambers and uLINK arising from the linking mechanism. The result is a relatively straightforward analysis that is appropriate for the level of uncertainty of the dosimetric input data.

# 1.  Evaluation of degrees of equivalence

The degree of equivalence for each of n participating laboratories i  = 1 to n (excluding the LINK, which by definition cannot obtain a new degree of equivalence) is defined as the difference Di = Ri – 1 and its expanded (k = 2) uncertainty Ui = 2uR,i, where

(1)

and we take

(2)

Here, NK,i is the transfer chamber calibration coefficient for laboratory i (expanded in terms of air-kerma rate and ionization current as Ki/Ii) and ui is its combined standard uncertainty. Similarly, for the LINK result in the regional comparison we have NK,LINK. The ratio RLINK,BIPM represents the result of the LINK in the corresponding BIPM international comparison (available in the KCDB), in which uBIPM is the combined standard uncertainty of the BIPM standard as described in the relevant comparison report.

The two expansions on the right-hand side of equation (1) express, respectively, a direct or indirect comparison with the BIPM, from which we anticipate cancellation of the non-statistical uncertainties for KLINK in both cases and, for an indirect comparison, cancellation also of those for ILINK at the expense of introducing those for IBIPM. The superscripts ‘reg’ and ‘inter’ for the LINK identify the regional and BIPM international comparison, respectively. The uncertainties utr and uLINK are discussed in Section 2.

The summation in equation (2) contains those components fj ui,j and fj uBIPM,j that are correlated between laboratory i and the BIPM, with correlation factor fj. The physical constants that enter in the air-kerma determinations are fully correlated (fj = 1); certain correction factors, for example kwall, might be considered partially correlated (0 < fj <1). When laboratory i is traceable to the BIPM, the summation contains all of the non-statistical components of uBIPM, each with correlation factor fj = 1. The case of a laboratory i traceable to the LINK might be noted; the LINK Type B uncertainties will then enter three times in equation (5). As always, they will cancel for NK,LINK and RLINK,BIPM, which means that they will remain for NK,i so that in the end this case is no different.

For a comparison of absorbed dose to water standards the equations are the same, but the correlated components j and the correlation factors fj are different, depending on whether a laboratory is traceable to a graphite calorimeter, to a water calorimeter, or to an ionometric standard. An example analysis, with choices for fj relating to mass-energy absorption coefficients and to the heat defect in water calorimetry, can be found in [1].

# 2.  The uncertainties utr and uLINK

The uncertainty ui for a given laboratory i will normally already include statistical components consistent with the repeatability of transfer chamber calibrations, assuming a ‘well-behaved’ transfer chamber. Under these conditions, the additional component utr in equation (2) is not required.

The behaviour of the transfer chamber should be characterized by the pilot laboratory through the ‘adjusted’ standard deviation of its m repeat calibrations NK,pilot,j (j = 1 to m);

(3)

where NK,pilot is the mean of the NK,pilot,j. If there is reason to believe that the transfer chamber is not ‘well-behaved’ (notably, if equation (3) yields a value that is greater than the normal repeatability at the pilot laboratory) then the value of utr from equation (3) should be included in equation (2). Furthermore, if the pilot laboratory measurements indicate a statistically-significant jump or drift in the response of the transfer chamber, then the ratios NK,i / NK,LINK entering in equation (1) must be corrected for this effect by a case-specific analysis (or the transfer chamber rejected if others have been used).

Regarding uLINK, as noted above the nature of the linking mechanism is such that many of the non-statistical components for the LINK cancel in the linking equation (1). However, the remaining non-statistical components are not always straightforward to estimate. This is among the reasons for using more than one LINK when possible, as this allows an alternative estimate of uLINK to be made from the difference between the results for different LINKs. Even so, the number of LINKs will rarely exceed two and therefore the latter estimate will never be robust.

Table 1 presents the components to be included in uLINK, depending on whether the BIPM comparison with the LINK was direct (both primary standards at one laboratory) or indirect (exchange of transfer instruments). Note that the statistical uncertainties for ILINK include chamber positioning and any other aspect of repeatability stated by the LINK. A value of 2 in the table indicates that the uncertainty should be included twice because the instrument was measured in both comparisons.

Not included in the table are uncertainties that might arise if reference conditions are not maintained for the LINK measurements during the two comparisons, for example if the LINK uses different values for distance or field size in the two comparisons. These can be difficult to estimate and are best avoided by using the same conditions in both.

Table 1. Uncertainty components to be included in uLINK. A value of 2 indicates that the component should be included twice.

|  | Direct BIPM comparison |  | Indirect BIPM comparison |  |
| --- | --- | --- | --- | --- |
| Parameter from equation (1) | statistical | non- statistical | statistical | non- statistical |
| KLINK | 2a | - | 2a | - |
| ILINK | 1a | 1 | 2a | - |
| IBIPM | - | - | 1b | 1b |

a In the unlikely event that utr is derived using equation (3) and the pilot laboratory is the LINK, each of these components should be reduced by 1 because they are already included in utr.

b These are required because uBIPM in equation (2) is for the air-kerma rate determination KBIPM only. Typically, the combined value will be around 0.03 % and can be found in the relevant comparison report for the LINK.

3.   Multiple transfer chambers

We consider a comparison using p transfer chambers (j = 1 to p), for which each of the n laboratories (i = 1 to n) has the p results NK,i,j and the LINK has the results NK,LINK,j. Relation (1) gives rise to the np values Ri,j with mean value for each laboratory

(4)

The standard uncertainty of each mean Ri gives information on the uncertainty arising from the transfer chambers,

(5)

This value for utr can be used in equation (2) for the case of multiple transfer chambers. Equation (5) reflects the improvement in uncertainty that can be achieved by the use of multiple transfer chambers (as well as the advantage of redundancy in the event of failure of a transfer chamber). However, this improvement is only realized if the spread of the transfer chamber results is consistent with the reproducibilities of the individual chambers.

4.   Multiple linking laboratories

A comparison can be made more robust, and in particular uLINK can be more reliably determined, by including more than one LINK. We consider a comparison with q linking laboratories (k = 1 to q) and a single transfer chamber. In this case relation (1) gives rise to the nq values Ri,k and Section 2 will yield the q estimates uLINK,k. Given the difficulty sometimes encountered in making a robust estimate of uLINK,k, an alternative estimate can be made from the spread of the Ri,k for different linking laboratories.

We evaluate Ri as the unweighted mean of the Ri,k

(6)

and make a first estimate of uLINK as

(7)

The best estimate of uLINK to be used in equation (2) is that derived from equation (7) or from

(8)

whichever is the larger. This approach prevents fortuitous agreement from resulting in an unrealistically low value for uLINK.

5.  General case

We consider the general case of n laboratories (i = 1 to n), p transfer chambers (j = 1 to p) and q linking laboratories (k = 1 to q), yielding the npq values Ri,j,k with q linking uncertainties uLINK,k. Treatment of these data can follow sequentially the analyses presented above. The method described in Section 3 can be used to derive the mean value

(9)

for each laboratory i, as determined using each linking laboratory k, along with the combined utr derived using

(10)

Following Section 4, we can then use these Ri,k in equation (6) to evaluate the Ri and use the uLINK,k to obtain the combined linking uncertainty uLINK as the maximum value from equations (7) and (8). Finally, utr and uLINK are inserted in equation (2) to yield the combined uR,i for each Ri.

## References

[1]	P. J. Allisy-Roberts and D. T. Burns, Summary of the BIPM.RI(I)-K4 comparison for absorbed dose to water in 60Co gamma radiation, Metrologia 42 (2005) Technical Supplement 06002.

To simplify the notation, air-kerma rate is denoted by K without a dot above.

For ease of presentation, equations are expressed as variances while the text refers to standard uncertainties.

The term ‘statistical’ is used to refer to those uncertainty components that result in variations for repeat measurements using a transfer chamber, and ‘non-statistical’ to those that are not revealed by repeat measurements. This classification does not necessarily correspond to that supplied by each laboratory as Type A and Type B components.

It can be shown empirically that the term (m–1.4) in the denominator is a better choice than (m–1) to estimate the standard deviation for low values of m, including m = 2.