## Introduction
This Excel workbook contains all the uncertainty budgets submitted by the participating laboratories in the EURAMET.AUV.A-K5 key comparison.

## NPL Level
NPL | Level |  | Note: green denotes uncertainties correlated across frequencies, orange denotes uncorrelated uncertainties.
|  | Valid for coupler UA1429 | [7 empty] | COMPONENTS OF TYPE B UNCERTAINTY EXPRESSED AS SEMI-RANGES (10-4 dB)
|  | SOURCE | [7 empty] | Frequency (Hz)
|  | Rectangular distribution applies unless stated | 2 | 4 | 8 | 16 | 20 | 25 | 31.5 | 63 | 125 | 250 | 500 | 1k | 2k | 2.5k | 3.15k | 4k | 5k | 6.3k | 8k | 10k
| 1 | Resistor Box Accuracy - Resistance | 0 | 0 | 0 | 0 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 1 | Resistor Box Accuracy - Capacitor | 238 | 242 | 243 | 244 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 1 | Stray Capacitance | 870 | 886 | 890 | 891 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| 1 | Non Linearity | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50
| 2 | Radius of Coupler | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 7 | 7 | 8 | 10 | 13
| 3 | Velocity of Sound (dry air) | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 39 | 38 | 37 | 35 | 32 | 27 | 18 | 1
| 4 | Ratio of Specific Heats | 15 | 10 | 7 | 5 | 5 | 4 | 4 | 3 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 5 | Ambient Pressure | 13 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 13 | 13 | 12
| 6 | Density of air | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 11 | 11 | 10
| 7 | Length of Coupler | 20 | 20 | 20 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 20 | 19 | 18 | 17 | 15 | 11 | 6 | 4 | 24
| 7 | Cavity Depth | 2 | 2 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 2 | 2 | 2 | 1
| 8 | Front Cavity Volume | 26 | 26 | 26 | 26 | 26 | 26 | 26 | 26 | 26 | 26 | 26 | 26 | 24 | 24 | 22 | 20 | 17 | 9 | 5 | 31
| 8 | Theory of Adding Volume * 0.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 6 | 9 | 14 | 23 | 36.5 | 60 | 105 | 193.5
| 9 | Compliance | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 9 | 40 | 66 | 111 | 187 | 268 | 232 | 118 | 593
| 9 | Mass | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 3 | 11 | 16 | 21 | 19 | 1 | 25 | 31 | 97
| 9 | Resistance | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 4 | 14 | 56 | 86 | 128 | 177 | 196 | 120 | 46 | 237
| 10 | Heat Conduction Theory (see note) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | Thermal Diffusivity | 154 | 104 | 73 | 52 | 46 | 41 | 37 | 26 | 18 | 13 | 9 | 6 | 5 | 4 | 4 | 3 | 3 | 2 | 2 | 2
| 10 | Thread Correction | 37 | 25 | 14 | 13 | 16 | 14 | 9 | 7 | 11 | 1 | 5 | 6 | 10 | 2 | 3 | 5 | 3 | 1 | 2 | 7
| 11 | Capillary Radius | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 11 | Air Viscosity | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 12 | Humidity Determination | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 2 | 3 | 3 | 5 | 7
| 13 | Polarising Voltage | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 21.712 | 21.712 | 21.712 | 21.712 | 21.712 | 21.712 | 21.712 | 21.712 | 21.712 | 21.712 | 21.712 | 21.712 | 21.712
| 14 | Temperature | 4 | 3 | 2 | 2 | 2 | 2 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 2 | 3 | 4 | 7 | 10 | 17 | 29
| 15 | Pressure Radial Non-Unif | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 6 | 22 | 34 | 54 | 86 | 132 | 203 | 308 | 424
| 16 | Mic. Temp. Dependence | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 24.0029 | 24.0201 | 24.1339 | 24.4318 | 25.4803 | 27.7894 | 28.9341 | 24.8489 | 38.0045 | 83.6444
| 16 | Mic. Press. Dependence | 32.2 | 32.2 | 32.2 | 32.2 | 32.2 | 32.2 | 32.2 | 32.2 | 32.2 | 32.2 | 32.2027 | 32.2439 | 32.758 | 33.6561 | 36.1245 | 41.001 | 42.9098 | 32.3372 | 88.8981 | 160.3714
| 20 | Transmitter Ground Shield | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22
| 20 | Receiver Ground Shield | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22
| 21 | Rounding Error | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50
| 22 | Frequency | 35 | 9 | 3 | 1 | 3 | 2 | 2 | 1 | 0 | 0 | 0 | 0 | 10 | 13 | 15 | 14 | 5 | 19 | 61 | 100
|  | ESTIMATE OF S.D. (sB) | 532.6493 | 537.3443 | 537.8233 | 537.7062 | 69.2768 | 68.0413 | 66.9797 | 65.0929 | 64.3564 | 63.635 | 63.5231 | 64.1302 | 76.0744 | 91.5788 | 121.4669 | 170.0968 | 216.6965 | 203.6047 | 219.3959 | 478.6857
sB = [S (a2/3)]1/2
COMPONENTS OF TYPE A UNCERTAINTY EXPRESSED AS STANDARD DEVIATIONS (10-4 dB)
|  | SOURCE | [7 empty] | Frequency (Hz)
|  | Normal distribution with n®¥ applies unless stated | 2 | 4 | 8 | 16 | 20 | 25 | 31.5 | 63 | 125 | 250 | 500 | 1k | 2k | 2.5k | 3.15k | 4k | 5k | 6.3k | 8k | 10k
| 1 | Allowed Repeatability | 1200 | 250 | 150 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 125 | 150 | 150
| 8 | Front Cavity Volume | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 36 | 34 | 32 | 28 | 24 | 12 | 6 | 44
|  | ESTIMATE OF S.D. (sA) | 1200.6015 | 252.8715 | 154.7385 | 106.9766 | 106.9766 | 106.9766 | 106.9766 | 106.9766 | 106.9766 | 106.9766 | 106.9766 | 106.9766 | 106.2826 | 105.622 | 104.9952 | 103.846 | 102.8397 | 125.5747 | 150.12 | 156.3202
sA = [S (a2)]1/2
OVERALL UNCERTAINTY (10-4 dB)
|  | SOURCE | [7 empty] | Frequency (Hz)
|  |  | 2 | 4 | 8 | 16 | 20 | 25 | 31.5 | 63 | 125 | 250 | 500 | 1k | 2k | 2.5k | 3.15k | 4k | 5k | 6.3k | 8k | 10k
|  | Type B at k=2 (2s B) | 1065.2986 | 1074.6887 | 1075.6467 | 1075.4124 | 138.5537 | 136.0825 | 133.9594 | 130.1858 | 128.7129 | 127.2701 | 127.0463 | 128.2604 | 152.1488 | 183.1577 | 242.9338 | 340.1936 | 433.393 | 407.2093 | 438.7917 | 957.3714
|  | Type A at k=2 (2s A) | 2401.203 | 505.743 | 309.477 | 213.9533 | 213.9533 | 213.9533 | 213.9533 | 213.9533 | 213.9533 | 213.9533 | 213.9533 | 213.9533 | 212.5653 | 211.2439 | 209.9905 | 207.6921 | 205.6794 | 251.1494 | 300.2399 | 312.6404
|  | Overall Uncertainty at k=2 | 2626.9064 | 1187.7423 | 1119.2818 | 1096.4888 | 254.8983 | 253.5635 | 252.4304 | 250.4483 | 249.6858 | 248.9451 | 248.8308 | 249.4529 | 261.4063 | 279.5903 | 321.1119 | 398.5821 | 479.7223 | 478.4302 | 531.6786 | 1007.1266
s = 2(sA2+sB2)1/2
|  | Quote (dB) | 0.27 | 0.12 | 0.12 | 0.11 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.04 | 0.05 | 0.05 | 0.06 | 0.11
Note on D7: JEMA: NPL uses a calibrated capacitor for frequencies below 20 Hz (19.95 Hz) and a set of calibrated resistors at 20 Hz and above. The uncertainty components associated with the resistance and capacitance are therefore correlated in two separate frequency ranges: 2 Hz - 16 Hz and 20 Hz - 10 kHz .
Note on W55: JEMA: The frequency uncertainty component considers the effect of the frequency being inaccuarte by a small amount. This uncertainty is considered uncorrelated for two reasons: There is no relationship between the frequency inaccuracy at different frequencies. The size of the effect on the microphone sensitivity also depends on the gradient of the microphone's sensitivity curve at the frequency in question. The gradient varies significantly with frequency. Therefore we can expect the size of the error due to this uncertainty component to vary randomly with frequency.
Note on W65: JEMA: A single value for Front Cavity Volume is used in the calculation of the microphone sensitivity at every frequency, so this uncertainty is correlated across frequencies even though it is a type A uncertainty.

## NPL Phase
NPL | Phase |  | Note: green denotes uncertainties correlated across frequencies, orange denotes uncorrelated uncertainties.
|  | Valid for coupler UA1429 | [7 empty] | COMPONENTS OF TYPE B UNCERTAINTY EXPRESSED AS SEMI-RANGES (10-4 degrees)
|  | SOURCE | [7 empty] | Frequency (Hz)
|  | Rectangular distribution applies unless stated | 2 | 4 | 8 | 16 | 20 | 25 | 31.5 | 63 | 125 | 250 | 500 | 1k | 2k | 2.5k | 3.15k | 4k | 5k | 6.3k | 8k | 10k
| 1 | Resistor Box Accuracy - Resistance | 9 | 4 | 2 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 2 | 2 | 1 | 1 | 2 | 1
| 1 | Resistor Box Accuracy - Capacitor | 491 | 250 | 125 | 63 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 1 | Stray Capacitance |  |  |  |  | 1900 | 1900 | 1900 | 1900 | 1900 | 1900 | 1900 | 1900 | 1900 | 1900 | 1900 | 1900 | 1900 | 1900 | 1900 | 1900
| 1 | Non Linearity | 2500 | 2500 | 2500 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000
| 2 | Radius of Coupler | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 2 | 2 | 2 | 1 | 7 | 16
| 3 | Velocity of Sound (dry air) | 4 | 3 | 3 | 2 | 2 | 1 | 1 | 1 | 0 | 1 | 2 | 5 | 12 | 15 | 20 | 28 | 39 | 53 | 67 | 78
| 4 | Ratio of Specific Heats | 161 | 123 | 92 | 67 | 61 | 55 | 49 | 35 | 26 | 18 | 13 | 9 | 6 | 5 | 4 | 4 | 3 | 2 | 0 | 2
| 5 | Ambient Pressure | 9 | 7 | 5 | 4 | 4 | 3 | 3 | 2 | 1 | 1 | 0 | 1 | 3 | 4 | 5 | 6 | 6 | 4 | 2 | 10
| 6 | Density of air | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 3 | 3 | 4 | 5 | 5 | 3 | 2 | 8
| 7 | Length of Coupler | 11 | 9 | 6 | 5 | 4 | 4 | 3 | 3 | 2 | 2 | 2 | 4 | 7 | 9 | 13 | 19 | 30 | 49 | 74 | 99
| 7 | Cavity Depth | 26 | 20 | 16 | 12 | 10 | 10 | 8 | 6 | 4 | 4 | 2 | 2 | 1 | 2 | 4 | 8 | 18 | 38 | 60 | 64
| 8 | Front Cavity Volume | 32 | 26 | 18 | 14 | 12 | 12 | 10 | 8 | 6 | 4 | 4 | 6 | 8 | 10 | 14 | 18 | 24 | 32 | 46 | 74
| 8 | Theory of Adding Volume * 0.5 | 0 | 0 | 0 | 0 | 0.5 | 0.5 | 0.5 | 1 | 1.5 | 3 | 6.5 | 12.5 | 31.5 | 38 | 47.5 | 61.5 | 81.5 | 113.5 | 164 | 261.5
| 9 | Compliance | 2 | 6 | 12 | 24 | 30 | 38 | 46 | 94 | 190 | 380 | 764 | 1528 | 3008 | 3684 | 4414 | 4918 | 4466 | 2460 | 1474 | 3828
| 9 | Mass | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 10 | 84 | 162 | 312 | 570 | 824 | 664 | 368 | 2166
| 9 | Resistance | 2 | 2 | 6 | 12 | 14 | 18 | 22 | 44 | 90 | 180 | 358 | 696 | 1220 | 1350 | 1332 | 928 | 12 | 756 | 836 | 2330
| 10 | Heat Conduction Theory (see note) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | Thermal Diffusivity | 1407 | 1072 | 800 | 588 | 531 | 479 | 430 | 310 | 223 | 159 | 113 | 79 | 53 | 46 | 39 | 31 | 23 | 14 | 1 | 19
| 10 | Thread Correction | 187 | 139 | 103 | 80 | 72 | 63 | 60 | 42 | 30 | 24 | 18 | 11 | 10 | 8 | 12 | 11 | 11 | 7 | 3 | 10
| 11 | Capillary Radius | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 11 | Air Viscosity | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 12 | Humidity Determination | 4 | 3 | 2 | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 6 | 13 | 20
| 13 | Polarising Voltage | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 14 | Temperature | 29 | 23 | 17 | 13 | 12 | 11 | 10 | 7 | 5 | 4 | 3 | 2 | 2 | 2 | 3 | 6 | 13 | 29 | 54 | 80
| 15 | Pressure Radial Non-Unif
| 16 | Mic. Temp. Dependence | 0 | 0 | 0 | 0 | 0 | 0 | 0.049504 | 0.11133 | 0.26696 | 0.70413 | 1.9746 | 5.4518 | 12.7756 | 15.5984 | 18.0556 | 19.2144 | 17.7835 | 10.8054 | 12.5506 | 76.6613
| 16 | Mic. Press. Dependence | 0.0061393 | 0.012577 | 0.02634 | 0.057353 | 0.07457 | 0.097662 | 0.13024 | 0.32697 | 0.87947 | 2.4822 | 6.5602 | 13.4297 | 13.7606 | 6.4015 | 16.9135 | 88.1912 | 248.1165 | 514.9051 | 585.2065 | 192.3884
| 20 | Transmitter Ground Shield | 20 | 10 | 5 | 3 | 2 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 20 | Receiver Ground Shield | 20 | 10 | 5 | 3 | 2 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 21 | Rounding Error | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500
| 22 | Frequency | 0 | 0 | 0 | 0 | 2 | 2 | 2 | 3 | 3 | 29 | 32 | 37 | 416 | 438 | 475 | 542 | 640 | 746 | 724 | 378
|  | ESTIMATE OF S.D. (sB) | 1711.2654 | 1607.0827 | 1546.5958 | 1239.879 | 1648.4742 | 1643.0266 | 1638.4744 | 1629.9323 | 1628.4554 | 1639.5012 | 1691.7721 | 1887.4646 | 2488.7179 | 2797.4287 | 3132.9421 | 3343.7124 | 3107.3362 | 2292.5608 | 1981.5423 | 3312.6696
sB = [S (a2/3)]1/2
COMPONENTS OF TYPE A UNCERTAINTY EXPRESSED AS SEMI-RANGES (10-4 degrees)
|  | SOURCE | [7 empty] | Frequency (Hz)
|  | Normal distribution with n®¥ applies unless stated | 2 | 4 | 8 | 16 | 20 | 25 | 31.5 | 63 | 125 | 250 | 500 | 1k | 2k | 2.5k | 3.15k | 4k | 5k | 6.3k | 8k | 10k
| 1 | Allowed Repeatability | 16000 | 7000 | 3500 | 2000 | 1500 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1500 | 2000
| 8 | Front Cavity Volume | 48 | 39 | 28 | 22 | 20 | 19 | 16 | 12 | 8 | 8 | 6 | 8 | 14 | 16 | 20 | 26 | 34 | 48 | 70 | 112
|  | ESTIMATE OF S.D. (sA) | 16000.072 | 7000.1086 | 3500.112 | 2000.121 | 1500.1333 | 1000.1805 | 1000.128 | 1000.072 | 1000.032 | 1000.032 | 1000.018 | 1000.032 | 1000.098 | 1000.128 | 1000.2 | 1000.3379 | 1000.5778 | 1001.1513 | 1501.6324 | 2003.1335
sA = [S (a2)]1/2
OVERALL UNCERTAINTY (10-4 dB)
|  | SOURCE | [7 empty] | Frequency (Hz)
|  |  | 2 | 4 | 8 | 16 | 20 | 25 | 31.5 | 63 | 125 | 250 | 500 | 1k | 2k | 2.5k | 3.15k | 4k | 5k | 6.3k | 8k | 10k
|  | Type B at k=2 (2s B) | 3422.5308 | 3214.1653 | 3093.1917 | 2479.7581 | 3296.9483 | 3286.0532 | 3276.9488 | 3259.8646 | 3256.9108 | 3279.0024 | 3383.5441 | 3774.9291 | 4977.4359 | 5594.8573 | 6265.8842 | 6687.4248 | 6214.6725 | 4585.1216 | 3963.0845 | 6625.3392
|  | Type A at k=2 (2s A) | 32000.144 | 14000.2173 | 7000.224 | 4000.242 | 3000.2667 | 2000.361 | 2000.256 | 2000.144 | 2000.064 | 2000.064 | 2000.036 | 2000.064 | 2000.196 | 2000.256 | 2000.4 | 2000.6759 | 2001.1557 | 2002.3027 | 3003.2649 | 4006.2671
|  | Overall Uncertainty at k=2 | 32182.6496 | 14364.4333 | 7653.1674 | 4706.4993 | 4457.7425 | 3847.0235 | 3839.1949 | 3824.5645 | 3822.0052 | 3840.8479 | 3930.4599 | 4272.0423 | 5364.2942 | 5941.6708 | 6577.4543 | 6980.2833 | 6528.9186 | 5003.2546 | 4972.4882 | 7742.4347
s = 2(sA2+sB2)1/2
|  | Quote (degrees) | 3.3 | 1.5 | 0.8 | 0.5 | 0.5 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.5 | 0.6 | 0.6 | 0.7 | 0.7 | 0.7 | 0.6 | 0.5 | 0.8
Note on D7: JEMA: NPL uses a calibrated capacitor for frequencies below 20 Hz (19.95 Hz) and a set of calibrated resistors at 20 Hz and above. The uncertainty components associated with the resistance and capacitance are therefore correlated in two separate frequency ranges: 2 Hz - 16 Hz and 20 Hz - 10 kHz .
Note on W55: JEMA: The frequency uncertainty component considers the effect of the frequency being inaccuarte by a small amount. This uncertainty is considered uncorrelated for two reasons: There is no relationship between the frequency inaccuracy at different frequencies. The size of the effect on the microphone sensitivity also depends on the gradient of the microphone's sensitivity curve at the frequency in question. The gradient varies significantly with frequency. Therefore we can expect the size of the error due to this uncertainty component to vary randomly with frequency.
Note on W65: JEMA: A single value for Front Cavity Volume is used in the calculation of the microphone sensitivity at every frequency, so this uncertainty is correlated across frequencies even though it is a type A uncertainty.

## LNE Level
LNE | Level
Level (dB*1000 re 1V/Pa)

## LNE Phase
LNE | Phase
Phase (degree*100)

## INRiM Level
INRiM | Level
Table I. Uncertainty for 4160.811012. Frequency range 20 Hz to 500 Hz.
| Frequency | Hz | 19,95 | 25,12 | 31,62 | 39,81 | 50,12 | 63,10 | 79,43 | 100,00 | 125,89 | 158,49 | 199,53 | 251,19 | 316,23 | 398,11 | 501,19
| Damping | dB 10-3 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,01 | 0,01 | 0,03 | 0,04
| Resonance Frequency | dB 10-3 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,01 | 0,01 | 0,01
| Equivalent Volume | dB 10-3 | 8,71 | 8,76 | 8,81 | 8,84 | 8,88 | 8,91 | 8,94 | 8,96 | 8,98 | 9,00 | 9,02 | 9,03 | 9,04 | 9,05 | 9,06
| Coupler Length | dB 10-3 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11
| Coupler diameter | dB 10-3 | 0,16 | 0,16 | 0,16 | 0,16 | 0,16 | 0,16 | 0,16 | 0,16 | 0,16 | 0,16 | 0,16 | 0,17 | 0,17 | 0,17 | 0,16
| Series impedance | dB 10-3 | 2,76 | 2,76 | 2,76 | 2,76 | 2,76 | 2,76 | 2,77 | 2,78 | 2,76 | 2,62 | 2,31 | 1,85 | 1,42 | 1,10 | 0,92
| Static Pressure | dB 10-3 | 0,32 | 0,32 | 0,32 | 0,32 | 0,32 | 0,32 | 0,31 | 0,31 | 0,31 | 0,31 | 0,31 | 0,31 | 0,31 | 0,31 | 0,31
| Temperature | dB 10-3 | 0,33 | 0,32 | 0,31 | 0,31 | 0,30 | 0,30 | 0,29 | 0,29 | 0,28 | 0,28 | 0,28 | 0,28 | 0,28 | 0,28 | 0,29
| Front Cavity Depth | dB 10-3 | 0,13 | 0,12 | 0,10 | 0,09 | 0,08 | 0,07 | 0,07 | 0,06 | 0,05 | 0,05 | 0,04 | 0,04 | 0,03 | 0,03 | 0,02
| Front cavity Volume | dB 10-3 | 3,56 | 3,58 | 3,60 | 3,61 | 3,63 | 3,64 | 3,65 | 3,66 | 3,67 | 3,67 | 3,68 | 3,69 | 3,69 | 3,69 | 3,70
| Correction to ref. Conditions | dB 10-3 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,39 | 3,39 | 3,39 | 3,39 | 3,39 | 3,38 | 3,38
| Polarisation Voltage | dB 10-3 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25
| Crosstalk&noise | dB 10-3 | 0,49 | 0,31 | 0,26 | 0,24 | 0,37 | 0,45 | 0,21 | 2,37 | 0,19 | 0,18 | 0,17 | 0,13 | 0,14 | 0,14 | 0,16
| Linearity | dB 10-3 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00
| Rounding Error | dB 10-3 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90
| Repeatability | dB 10-3 | 6,20 | 4,49 | 3,57 | 3,07 | 2,49 | 2,15 | 1,80 | 1,57 | 1,37 | 1,18 | 1,14 | 1,15 | 1,08 | 1,10 | 1,08
| U | dB 10-3 | 13,47 | 12,81 | 12,55 | 12,45 | 12,35 | 12,32 | 12,28 | 12,50 | 12,26 | 12,23 | 12,17 | 12,11 | 12,05 | 12,03 | 12,01
| Effective degrees of freedom |  | 67 | 198 | 458 | 808 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000
| k |  | 2,00 | 1,97 | 1,97 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96
| U95 | dB | 0,027 | 0,026 | 0,025 | 0,025 | 0,025 | 0,025 | 0,025 | 0,025 | 0,025 | 0,024 | 0,024 | 0,024 | 0,024 | 0,024 | 0,024
| U95 Rounded | dB | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03
Table II. Uncertainty for 4160.811012. Frequency range 630 Hz to 10 kHz.
| Frequency | Hz | 630,96 | 794,33 | 1000,00 | 1258,93 | 1584,89 | 1995,26 | 2511,89 | 3162,28 | 3981,07 | 5011,87 | 6309,57 | 7943,28 | 10000,00
| Damping | dB 10-3 | 0,07 | 0,11 | 0,18 | 0,28 | 0,45 | 0,73 | 1,16 | 1,83 | 2,81 | 4,04 | 4,43 | 2,12 | 0,65
| Resonance Frequency | dB 10-3 | 0,02 | 0,03 | 0,04 | 0,06 | 0,08 | 0,09 | 0,10 | 0,08 | 0,30 | 1,45 | 4,14 | 6,21 | 3,43
| Equivalent Volume | dB 10-3 | 9,06 | 9,05 | 9,03 | 8,99 | 8,92 | 8,81 | 8,61 | 8,26 | 7,63 | 6,58 | 4,61 | 1,95 | 0,90
| Coupler Length | dB 10-3 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11 | 0,11 | 0,10 | 0,10 | 0,09 | 0,08 | 0,06 | 0,04 | 0,02
| Coupler diameter | dB 10-3 | 0,16 | 0,16 | 0,16 | 0,16 | 0,16 | 0,16 | 0,16 | 0,15 | 0,14 | 0,11 | 0,08 | 0,03 | 0,05
| Series impedance | dB 10-3 | 0,84 | 0,85 | 0,92 | 0,94 | 0,88 | 0,92 | 1,37 | 1,91 | 1,84 | 1,84 | 1,83 | 1,83 | 1,84
| Static Pressure | dB 10-3 | 0,31 | 0,31 | 0,32 | 0,32 | 0,33 | 0,34 | 0,35 | 0,38 | 0,42 | 0,48 | 0,54 | 0,52 | 0,30
| Temperature | dB 10-3 | 0,30 | 0,32 | 0,34 | 0,38 | 0,42 | 0,49 | 0,60 | 0,80 | 1,15 | 1,66 | 2,22 | 2,37 | 1,40
| Front Cavity Depth | dB 10-3 | 0,02 | 0,02 | 0,01 | 0,00 | 0,00 | 0,02 | 0,04 | 0,07 | 0,11 | 0,23 | 0,31 | 0,32 | 0,27
| Front cavity Volume | dB 10-3 | 3,69 | 3,69 | 3,68 | 3,66 | 3,63 | 3,58 | 3,51 | 3,38 | 3,18 | 2,85 | 2,33 | 1,46 | 0,84
| Correction to ref. Conditions | dB 10-3 | 3,36 | 3,34 | 3,31 | 3,27 | 3,20 | 3,10 | 2,98 | 2,94 | 3,22 | 4,00 | 4,59 | 5,26 | 10,04
| Polarisation Voltage | dB 10-3 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25
| Crosstalk&noise | dB 10-3 | 0,18 | 0,20 | 0,22 | 0,27 | 0,33 | 0,40 | 0,47 | 0,59 | 0,77 | 0,96 | 1,20 | 1,50 | 1,95
| Linearity | dB 10-3 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00
| Rounding Error | dB 10-3 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90
| Repeatability | dB 10-3 | 1,06 | 1,05 | 0,85 | 1,07 | 1,09 | 1,21 | 1,75 | 3,21 | 5,77 | 2,87 | 4,01 | 14,91 | 15,45
| U | dB 10-3 | 12,00 | 11,99 | 11,95 | 11,93 | 11,86 | 11,76 | 11,72 | 11,90 | 12,67 | 11,68 | 12,07 | 18,59 | 19,93
| Effective degrees of freedom |  | 48819 | 50386 | 117676 | 46051 | 42169 | 26405 | 6019 | 570 | 70 | 825 | 247 | 7 | 8
| k |  | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,99 | 1,96 | 1,97 | 2,36 | 2,31
| U95 | dB | 0,024 | 0,024 | 0,024 | 0,024 | 0,024 | 0,024 | 0,023 | 0,024 | 0,025 | 0,023 | 0,024 | 0,044 | 0,046
| U95 Rounded | dB | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,05 | 0,06
Table V. Uncertainty for 4160. 2652754. Frequency range 20 Hz to 500 Hz.
| Frequency | Hz | 19,95 | 25,12 | 31,62 | 39,81 | 50,12 | 63,10 | 79,43 | 100,00 | 125,89 | 158,49 | 199,53 | 251,19 | 316,23 | 398,11 | 501,19
| Damping | dB 10-3 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,01 | 0,01 | 0,02 | 0,04
| Resonance Frequency | dB 10-3 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,01 | 0,01 | 0,01
| Equivalent Volume | dB 10-3 | 7,88 | 7,92 | 7,96 | 7,99 | 8,02 | 8,05 | 8,07 | 8,09 | 8,11 | 8,13 | 8,14 | 8,15 | 8,16 | 8,17 | 8,17
| Coupler Length | dB 10-3 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10
| Coupler diameter | dB 10-3 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15
| Series impedance | dB 10-3 | 2,76 | 2,76 | 2,76 | 2,76 | 2,76 | 2,76 | 2,77 | 2,78 | 2,76 | 2,61 | 2,30 | 1,85 | 1,42 | 1,10 | 0,92
| Static Pressure | dB 10-3 | 0,33 | 0,33 | 0,32 | 0,32 | 0,32 | 0,32 | 0,32 | 0,32 | 0,32 | 0,32 | 0,32 | 0,32 | 0,32 | 0,32 | 0,32
| Temperature | dB 10-3 | 0,33 | 0,32 | 0,31 | 0,31 | 0,30 | 0,30 | 0,29 | 0,29 | 0,28 | 0,28 | 0,28 | 0,28 | 0,28 | 0,28 | 0,29
| Front Cavity Depth | dB 10-3 | 0,12 | 0,11 | 0,09 | 0,08 | 0,07 | 0,07 | 0,06 | 0,05 | 0,05 | 0,04 | 0,04 | 0,03 | 0,03 | 0,03 | 0,02
| Front cavity Volume | dB 10-3 | 3,23 | 3,24 | 3,26 | 3,27 | 3,28 | 3,29 | 3,30 | 3,31 | 3,32 | 3,33 | 3,33 | 3,34 | 3,34 | 3,34 | 3,34
| Correction to ref. Conditions | dB 10-3 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,39 | 3,39 | 3,39 | 3,39 | 3,39 | 3,38 | 3,38
| Polarisation Voltage | dB 10-3 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25
| Crosstalk&noise | dB 10-3 | 0,49 | 0,31 | 0,26 | 0,24 | 0,37 | 0,45 | 0,21 | 2,37 | 0,19 | 0,18 | 0,17 | 0,13 | 0,14 | 0,14 | 0,16
| Linearity | dB 10-3 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00
| Rounding Error | dB 10-3 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90
| Repeatability | dB 10-3 | 1,55 | 1,27 | 1,20 | 1,09 | 1,09 | 1,03 | 1,03 | 0,93 | 0,98 | 0,95 | 0,92 | 0,94 | 0,90 | 0,85 | 0,83
| U | dB 10-3 | 11,37 | 11,37 | 11,39 | 11,40 | 11,43 | 11,45 | 11,47 | 11,72 | 11,49 | 11,47 | 11,41 | 11,34 | 11,28 | 11,25 | 11,23
| Effective degrees of freedom |  | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000
| k |  | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96
| U95 | dB | 0,023 | 0,023 | 0,023 | 0,023 | 0,023 | 0,023 | 0,023 | 0,023 | 0,023 | 0,023 | 0,023 | 0,023 | 0,023 | 0,022 | 0,022
| U95 Rounded | dB | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03
Table VI. Uncertainty for 4160. 2652754. Frequency range 630 Hz to 10 kHz.
| Frequency | Hz | 630,96 | 794,33 | 1000,00 | 1258,93 | 1584,89 | 1995,26 | 2511,89 | 3162,28 | 3981,07 | 5011,87 | 6309,57 | 7943,28 | 10000,00
| Damping | dB 10-3 | 0,06 | 0,10 | 0,16 | 0,26 | 0,41 | 0,65 | 1,03 | 1,60 | 2,41 | 3,28 | 3,22 | 1,63 | 2,77
| Resonance Frequency | dB 10-3 | 0,02 | 0,03 | 0,05 | 0,06 | 0,08 | 0,11 | 0,13 | 0,13 | 0,22 | 1,03 | 3,12 | 5,39 | 3,70
| Equivalent Volume | dB 10-3 | 8,16 | 8,15 | 8,13 | 8,08 | 8,00 | 7,86 | 7,62 | 7,22 | 6,51 | 5,36 | 3,41 | 1,44 | 2,69
| Coupler Length | dB 10-3 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,09 | 0,09 | 0,09 | 0,08 | 0,06 | 0,05 | 0,04 | 0,08
| Coupler diameter | dB 10-3 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,14 | 0,14 | 0,13 | 0,12 | 0,09 | 0,06 | 0,05 | 0,13
| Series impedance | dB 10-3 | 0,84 | 0,85 | 0,92 | 0,94 | 0,88 | 0,92 | 1,37 | 1,91 | 1,84 | 1,84 | 1,83 | 1,83 | 1,84
| Static Pressure | dB 10-3 | 0,32 | 0,32 | 0,32 | 0,33 | 0,33 | 0,34 | 0,36 | 0,38 | 0,43 | 0,49 | 0,55 | 0,52 | 0,29
| Temperature | dB 10-3 | 0,30 | 0,32 | 0,34 | 0,38 | 0,43 | 0,50 | 0,62 | 0,83 | 1,21 | 1,76 | 2,39 | 2,69 | 2,10
| Front Cavity Depth | dB 10-3 | 0,02 | 0,01 | 0,01 | 0,00 | 0,01 | 0,02 | 0,04 | 0,07 | 0,11 | 0,21 | 0,29 | 0,27 | 0,21
| Front cavity Volume | dB 10-3 | 3,34 | 3,33 | 3,32 | 3,30 | 3,27 | 3,21 | 3,11 | 2,96 | 2,72 | 2,31 | 1,73 | 1,59 | 2,85
| Correction to ref. Conditions | dB 10-3 | 3,36 | 3,34 | 3,31 | 3,27 | 3,20 | 3,10 | 2,98 | 2,94 | 3,22 | 4,00 | 4,59 | 5,26 | 10,04
| Polarisation Voltage | dB 10-3 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25
| Crosstalk&noise | dB 10-3 | 0,18 | 0,20 | 0,22 | 0,27 | 0,33 | 0,40 | 0,47 | 0,59 | 0,77 | 0,96 | 1,20 | 1,50 | 1,95
| Linearity | dB 10-3 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00 | 5,00
| Rounding Error | dB 10-3 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90
| Repeatability | dB 10-3 | 0,80 | 0,89 | 0,88 | 0,85 | 0,88 | 0,89 | 0,95 | 1,16 | 1,45 | 1,75 | 1,59 | 7,23 | 31,89
| U | dB 10-3 | 11,21 | 11,20 | 11,18 | 11,13 | 11,04 | 10,92 | 10,78 | 10,64 | 10,42 | 10,38 | 10,21 | 12,83 | 34,66
| Effective degrees of freedom |  | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | 30 | 4
| k |  | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 1,96 | 2,05 | 2,78
| U95 | dB | 0,022 | 0,022 | 0,022 | 0,022 | 0,022 | 0,022 | 0,022 | 0,021 | 0,021 | 0,021 | 0,020 | 0,026 | 0,096
| U95 Rounded | dB | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,03 | 0,1
The uncertainty at low frequency are reported for each microphone. However, as the repeatability data is common to both microphones (obtained from previous calibrations of LS1 microphones), the differences are negligible.
Table IX. Uncertainty for 4160.811012. Frequency range 2 Hz to 16 Hz.
| Frequency | Hz | 2,00 | 2,51 | 3,16 | 3,98 | 5,01 | 6,31 | 7,94 | 10,00 | 12,59 | 15,85
| Damping | dB 10-3 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00
| Resonance Frequency | dB 10-3 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00
| Equivalent Volume | dB 10-3 | 7,31 | 7,43 | 7,54 | 7,65 | 7,74 | 7,82 | 7,90 | 7,96 | 8,02 | 8,08
| Coupler Length | dB 10-3 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10
| Coupler diameter | dB 10-3 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15
| Series impedance | dB 10-3 | 5,52 | 5,52 | 5,52 | 5,52 | 5,52 | 5,52 | 5,52 | 5,52 | 5,52 | 5,52
| Static Pressure | dB 10-3 | 0,35 | 0,35 | 0,34 | 0,34 | 0,33 | 0,33 | 0,33 | 0,33 | 0,32 | 0,32
| Temperature | dB 10-3 | 0,46 | 0,44 | 0,42 | 0,40 | 0,39 | 0,38 | 0,36 | 0,35 | 0,34 | 0,33
| Front Cavity Depth | dB 10-3 | 0,36 | 0,33 | 0,29 | 0,26 | 0,24 | 0,21 | 0,19 | 0,17 | 0,15 | 0,13
| Front cavity Volume | dB 10-3 | 2,98 | 3,03 | 3,08 | 3,12 | 3,16 | 3,19 | 3,22 | 3,25 | 3,27 | 3,29
| Correction to ref. Conditions | dB 10-3 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38
| Polarisation Voltage | dB 10-3 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25
| Crosstalk&noise | dB 10-3 | 0,49 | 0,32 | 0,32 | 0,29 | 0,31 | 0,29 | 0,30 | 0,30 | 0,26 | 0,29
| Linearity | dB 10-3 | 10,00 | 10,00 | 10,00 | 10,00 | 10,00 | 10,00 | 10,00 | 10,00 | 5,00 | 5,00
| Rounding Error | dB 10-3 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90
| Repeatability | dB 10-3 | 43,64 | 32,22 | 23,74 | 17,71 | 13,35 | 10,16 | 8,51 | 7,05 | 5,65 | 5,12
| U | dB 10-3 | 46,03 | 35,42 | 27,97 | 23,12 | 20,01 | 18,08 | 17,25 | 16,61 | 13,57 | 13,40
| Effective degrees of freedom |  | 4 | 4 | 6 | 9 | 15 | 30 | 51 | 93 | 100 | 141
| k |  | 3,18 | 2,78 | 2,57 | 2,31 | 2,13 | 2,04 | 2,01 | 1,99 | 1,98 | 1,98
| U95 | dB | 0,146 | 0,098 | 0,072 | 0,053 | 0,043 | 0,037 | 0,035 | 0,033 | 0,027 | 0,027
| U95 Rounded | dB | 0,15 | 0,1 | 0,08 | 0,06 | 0,05 | 0,04 | 0,04 | 0,04 | 0,03 | 0,03
Table X. Uncertainty for 4160. 2652754. Frequency range 2 Hz to 16 Hz.
| Frequency | Hz | 2,00 | 2,51 | 3,16 | 3,98 | 5,01 | 6,31 | 7,94 | 10,00 | 12,59 | 15,85
| Damping | dB 10-3 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00
| Resonance Frequency | dB 10-3 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00 | 0,00
| Equivalent Volume | dB 10-3 | 7,31 | 7,43 | 7,54 | 7,65 | 7,74 | 7,82 | 7,90 | 7,96 | 8,02 | 8,08
| Coupler Length | dB 10-3 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10 | 0,10
| Coupler diameter | dB 10-3 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15 | 0,15
| Series impedance | dB 10-3 | 5,52 | 5,52 | 5,52 | 5,52 | 5,52 | 5,52 | 5,52 | 5,52 | 5,52 | 5,52
| Static Pressure | dB 10-3 | 0,35 | 0,35 | 0,34 | 0,34 | 0,33 | 0,33 | 0,33 | 0,33 | 0,32 | 0,32
| Temperature | dB 10-3 | 0,46 | 0,44 | 0,42 | 0,40 | 0,39 | 0,38 | 0,36 | 0,35 | 0,34 | 0,33
| Front Cavity Depth | dB 10-3 | 0,36 | 0,33 | 0,29 | 0,26 | 0,24 | 0,21 | 0,19 | 0,17 | 0,15 | 0,13
| Front cavity Volume | dB 10-3 | 2,98 | 3,03 | 3,08 | 3,12 | 3,16 | 3,19 | 3,22 | 3,25 | 3,27 | 3,29
| Correction to ref. Conditions | dB 10-3 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38 | 3,38
| Polarisation Voltage | dB 10-3 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25 | 1,25
| Crosstalk&noise | dB 10-3 | 0,49 | 0,32 | 0,32 | 0,29 | 0,31 | 0,29 | 0,30 | 0,30 | 0,26 | 0,29
| Linearity | dB 10-3 | 10,00 | 10,00 | 10,00 | 10,00 | 10,00 | 10,00 | 10,00 | 10,00 | 5,00 | 5,00
| Rounding Error | dB 10-3 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90 | 2,90
| Repeatability | dB 10-3 | 43,64 | 32,22 | 23,74 | 17,71 | 13,35 | 10,16 | 8,51 | 7,05 | 5,65 | 5,12
| U | dB 10-3 | 46,03 | 35,42 | 27,97 | 23,12 | 20,01 | 18,08 | 17,25 | 16,61 | 13,57 | 13,40
| Effective degrees of freedom |  | 4 | 4 | 6 | 9 | 15 | 30 | 51 | 93 | 100 | 141
| k |  | 3,18 | 2,78 | 2,57 | 2,31 | 2,13 | 2,04 | 2,01 | 1,99 | 1,98 | 1,98
| U95 | dB | 0,146 | 0,098 | 0,072 | 0,053 | 0,043 | 0,037 | 0,035 | 0,033 | 0,027 | 0,027
| U95 Rounded | dB | 0,15 | 0,1 | 0,08 | 0,06 | 0,05 | 0,04 | 0,04 | 0,04 | 0,03 | 0,03

## INRiM Phase
INRiM | Phase
Table III. Uncertainty for 4160.811012. Frequency range 20 Hz to 500 Hz: phase response.
| Frequency | Hz | 20,0 | 25,1 | 31,6 | 39,8 | 50,1 | 63,1 | 79,4 | 100,0 | 125,9 | 158,5 | 199,5 | 251,2 | 316,2 | 398,1 | 501,2
| Damping | deg 10-3 | 0,1 | 0,1 | 0,2 | 0,2 | 0,3 | 0,3 | 0,4 | 0,5 | 0,7 | 0,9 | 1,1 | 1,4 | 1,7 | 2,2 | 2,7
| Resonance Frequency | deg 10-3 | 0,1 | 0,1 | 0,1 | 0,2 | 0,2 | 0,2 | 0,3 | 0,4 | 0,5 | 0,6 | 0,8 | 1,0 | 1,2 | 1,6 | 2,0
| Equivalent Volume | deg 10-3 | 2,5 | 2,2 | 2,0 | 1,7 | 1,4 | 1,2 | 0,9 | 0,6 | 0,4 | 0,0 | 0,3 | 0,7 | 1,2 | 1,8 | 2,5
| Coupler Length | deg 10-3 | 0,015 | 0,013 | 0,012 | 0,011 | 0,010 | 0,009 | 0,008 | 0,008 | 0,007 | 0,007 | 0,007 | 0,006 | 0,007 | 0,007 | 0,007
| Coupler diameter | deg 10-3 | 0,019 | 0,018 | 0,016 | 0,014 | 0,013 | 0,012 | 0,011 | 0,010 | 0,009 | 0,009 | 0,009 | 0,009 | 0,009 | 0,009 | 0,010
| Series impedance | deg 10-3 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001
| Static Pressure | deg 10-3 | 0,087 | 0,006 | 0,006 | 0,005 | 0,004 | 0,004 | 0,003 | 0,003 | 0,002 | 0,002 | 0,001 | 0,000 | 0,001 | 0,002 | 0,003
| Temperature | deg 10-3 | 0,4 | 0,3 | 0,3 | 0,2 | 0,2 | 0,1 | 0,1 | 0,0 | 0,0 | 0,1 | 0,1 | 0,2 | 0,3 | 0,4 | 0,6
| Front Cavity Depth | deg 10-3 | 0,8 | 0,7 | 0,6 | 0,6 | 0,5 | 0,4 | 0,4 | 0,4 | 0,3 | 0,3 | 0,3 | 0,2 | 0,2 | 0,2 | 0,2
| Front cavity Volume | deg 10-3 | 1,1 | 1,0 | 0,9 | 0,8 | 0,7 | 0,7 | 0,6 | 0,5 | 0,5 | 0,5 | 0,4 | 0,4 | 0,4 | 0,4 | 0,4
| Repeatability | deg 10-3 | 184,5 | 147,1 | 117,0 | 93,8 | 74,4 | 60,0 | 47,0 | 37,9 | 29,2 | 23,1 | 18,5 | 14,6 | 11,4 | 8,4 | 7,0
| Correction to ref. conditions | deg 10-3 | 0,0 | 0,0 | 0,0 | 0,0 | 0,1 | 0,1 | 0,1 | 0,1 | 0,2 | 0,2 | 0,3 | 0,4 | 0,6 | 0,8 | 1,1
| Phase Linearity | deg 10-3 | 777,5 | 619,6 | 493,2 | 392,3 | 338,9 | 265,9 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0
| Rounding | deg 10-3 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0
| U | deg 10-3 | 799,6 | 637,5 | 507,7 | 404,4 | 348,2 | 274,1 | 256,0 | 254,5 | 253,4 | 252,7 | 252,4 | 252,1 | 251,9 | 251,8 | 251,8
| Effective degrees of freedom | deg 10-3 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000
| K | deg 10-3 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0
| U95 | deg | 1599 | 1275 | 1015 | 0,809 | 0,696 | 0,548 | 0,512 | 0,509 | 0,507 | 0,505 | 0,505 | 0,504 | 0,504 | 0,504 | 0,504
| U95 Rounded | deg | 1,6 | 1,3 | 1,1 | 0,9 | 0,7 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6
Table IV. Uncertainty for 4160.811012. Frequency range 630 Hz to 10 kHz: phase response.
| Frequency | Hz | 631,0 | 794,3 | 1000,0 | 1258,9 | 1584,9 | 1995,3 | 2511,9 | 3162,3 | 3981,1 | 5011,9 | 6309,6 | 7943,3 | 10000
| Damping | deg 10-3 | 3,4 | 4,3 | 5,4 | 6,7 | 8,3 | 10,1 | 11,9 | 13,2 | 12,5 | 6,9 | 7,2 | 17,8 | 7,7
| Resonance Frequency | deg 10-3 | 2,5 | 3,2 | 4,0 | 5,2 | 6,8 | 9,0 | 12,1 | 16,8 | 23,7 | 33,0 | 36,9 | 20,4 | 13,5
| Equivalent Volume | deg 10-3 | 3,4 | 4,5 | 5,8 | 7,5 | 9,6 | 12,2 | 15,6 | 19,8 | 24,8 | 30,5 | 33,5 | 25,8 | 11,1
| Coupler Length | deg 10-3 | 0,008 | 0,009 | 0,011 | 0,013 | 0,016 | 0,020 | 0,026 | 0,035 | 0,049 | 0,073 | 0,113 | 0,167 | 0,194
| Coupler diameter | deg 10-3 | 0,011 | 0,013 | 0,015 | 0,019 | 0,023 | 0,028 | 0,035 | 0,045 | 0,059 | 0,114 | 0,182 | 0,279 | 0,340
| Series impedance | deg 10-3 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001 | 0,001
| Static Pressure | deg 10-3 | 0,005 | 0,007 | 0,010 | 0,013 | 0,016 | 0,020 | 0,024 | 0,026 | 0,023 | 0,005 | 0,046 | 0,143 | 0,224
| Temperature | deg 10-3 | 0,8 | 1,0 | 1,2 | 1,5 | 1,8 | 2,2 | 2,6 | 3,1 | 3,2 | 2,1 | 1,9 | 9,7 | 16,8
| Front Cavity Depth | deg 10-3 | 0,1 | 0,1 | 0,1 | 0,1 | 0,1 | 0,0 | 0,0 | 0,2 | 0,4 | 0,9 | 1,9 | 3,5 | 4,3
| Front cavity Volume | deg 10-3 | 0,4 | 0,4 | 0,5 | 0,5 | 0,6 | 0,7 | 0,9 | 1,1 | 1,4 | 1,9 | 2,4 | 2,9 | 3,2
| Repeatability | deg 10-3 | 6,0 | 4,7 | 3,9 | 2,1 | 2,1 | 4,0 | 6,4 | 10,7 | 20,9 | 43,9 | 81,3 | 236,1 | 342,7
| Correction to ref. conditions | deg 10-3 | 1,6 | 2,1 | 2,8 | 3,5 | 4,4 | 5,2 | 5,7 | 5,6 | 4,9 | 12,2 | 35,3 | 66,6 | 75,1
| Phase Linearity | deg 10-3 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0
| Rounding | deg 10-3 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0
| U | deg 10-3 | 251,8 | 251,8 | 251,9 | 252,0 | 252,1 | 252,4 | 252,9 | 253,7 | 255,2 | 259,8 | 271,6 | 353,6 | 432,5
| Effective degrees of freedom | deg 10-3 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | 374 | 15 | 8
| k | deg 10-3 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,1 | 2,4
| U95 | deg | 0,504 | 0,504 | 0,504 | 0,504 | 0,504 | 0,505 | 0,506 | 0,507 | 0,510 | 0,520 | 0,543 | 0,754 | 1023
| U95 Rounded | deg | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,8 | 1,1
Table VII. Uncertainty for 4160. 2652754. Frequency range 20 Hz to 500 Hz: phase response.
| Frequency | Hz | 20,0 | 25,1 | 31,6 | 39,8 | 50,1 | 63,1 | 79,4 | 100,0 | 125,9 | 158,5 | 199,5 | 251,2 | 316,2 | 398,1 | 501,2
| Damping | deg 10-3 | 0,1 | 0,1 | 0,2 | 0,2 | 0,2 | 0,3 | 0,4 | 0,5 | 0,6 | 0,8 | 1,0 | 1,2 | 1,6 | 2,0 | 2,5
| Resonance Frequency | deg 10-3 | 0,1 | 0,1 | 0,1 | 0,1 | 0,2 | 0,2 | 0,3 | 0,4 | 0,4 | 0,6 | 0,7 | 0,9 | 1,1 | 1,4 | 1,7
| Equivalent Volume | deg 10-3 | 2,2 | 2,0 | 1,7 | 1,5 | 1,2 | 1,0 | 0,8 | 0,5 | 0,3 | 0,1 | 0,3 | 0,7 | 1,2 | 1,7 | 2,3
| Coupler Length | deg 10-3 | 0,012 | 0,011 | 0,010 | 0,009 | 0,008 | 0,008 | 0,007 | 0,006 | 0,006 | 0,006 | 0,005 | 0,005 | 0,005 | 0,006 | 0,006
| Coupler diameter | deg 10-3 | 0,014 | 0,013 | 0,012 | 0,010 | 0,010 | 0,009 | 0,008 | 0,007 | 0,007 | 0,006 | 0,006 | 0,006 | 0,007 | 0,007 | 0,008
| Series impedance | deg 10-3 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000
| Static Pressure | deg 10-3 | 0,083 | 0,006 | 0,005 | 0,005 | 0,004 | 0,004 | 0,003 | 0,003 | 0,002 | 0,002 | 0,001 | 0,000 | 0,001 | 0,002 | 0,003
| Temperature | deg 10-3 | 0,3 | 0,3 | 0,3 | 0,2 | 0,2 | 0,1 | 0,1 | 0,0 | 0,0 | 0,1 | 0,2 | 0,2 | 0,3 | 0,5 | 0,6
| Front Cavity Depth | deg 10-3 | 0,7 | 0,6 | 0,6 | 0,5 | 0,5 | 0,4 | 0,4 | 0,3 | 0,3 | 0,3 | 0,2 | 0,2 | 0,2 | 0,2 | 0,1
| Front cavity Volume | deg 10-3 | 1,0 | 0,9 | 0,8 | 0,7 | 0,6 | 0,6 | 0,5 | 0,5 | 0,4 | 0,4 | 0,4 | 0,3 | 0,3 | 0,3 | 0,3
| Repeatability | deg 10-3 | 184,5 | 147,1 | 117,0 | 93,8 | 74,4 | 60,0 | 47,0 | 37,9 | 29,2 | 23,1 | 18,5 | 14,6 | 11,4 | 8,4 | 7,0
| Correction To reference conditions | deg 10-3 | 0,0 | 0,0 | 0,0 | 0,0 | 0,1 | 0,1 | 0,1 | 0,1 | 0,2 | 0,2 | 0,3 | 0,4 | 0,6 | 0,8 | 1,1
| Phase Linearity | deg 10-3 | 777,5 | 619,6 | 493,2 | 392,3 | 338,9 | 265,9 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0
| Rounding | deg 10-3 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0
| U | deg 10-3 | 799,6 | 637,5 | 507,7 | 404,4 | 348,2 | 274,1 | 256,0 | 254,5 | 253,4 | 252,7 | 252,4 | 252,1 | 251,9 | 251,8 | 251,8
| Effective degrees of freedom | deg 10-3 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000
| k | deg 10-3 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0
| U95 | deg | 1,60 | 1,27 | 1,02 | 0,81 | 0,70 | 0,55 | 0,51 | 0,51 | 0,51 | 0,51 | 0,50 | 0,50 | 0,50 | 0,50 | 0,504
| U95 Rounded | deg | 1,6 | 1,3 | 1,1 | 0,9 | 0,7 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6
Table VIII. Uncertainty for 4160. 2652754. Frequency range 630 Hz to 10 kHz: phase response.
| Frequency | Hz | 631,0 | 794,3 | 1000,0 | 1258,9 | 1584,9 | 1995,3 | 2511,9 | 3162,3 | 3981,1 | 5011,9 | 6309,6 | 7943,3 | 10000
| Damping | deg 10-3 | 3,1 | 3,9 | 4,9 | 6,0 | 7,4 | 9,0 | 10,5 | 11,5 | 10,7 | 5,7 | 5,7 | 19,7 | 23,3
| Resonance Frequency | deg 10-3 | 2,2 | 2,8 | 3,6 | 4,6 | 6,0 | 7,9 | 10,7 | 14,7 | 20,6 | 28,3 | 31,3 | 25,0 | 50,0
| Equivalent Volume | deg 10-3 | 3,1 | 4,1 | 5,3 | 6,8 | 8,7 | 11,0 | 13,9 | 17,4 | 21,2 | 24,7 | 25,0 | 27,2 | 35,3
| Coupler Length | deg 10-3 | 0,007 | 0,008 | 0,009 | 0,011 | 0,013 | 0,017 | 0,022 | 0,030 | 0,043 | 0,065 | 0,106 | 0,172 | 0,265
| Coupler diameter | deg 10-3 | 0,009 | 0,011 | 0,013 | 0,015 | 0,019 | 0,023 | 0,030 | 0,038 | 0,050 | 0,103 | 0,172 | 0,287 | 0,448
| Series impedance | deg 10-3 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000 | 0,000
| Static Pressure | deg 10-3 | 0,005 | 0,007 | 0,009 | 0,012 | 0,016 | 0,020 | 0,023 | 0,025 | 0,021 | 0,004 | 0,051 | 0,151 | 0,237
| Temperature | deg 10-3 | 0,8 | 1,0 | 1,2 | 1,5 | 1,8 | 2,2 | 2,6 | 3,1 | 3,2 | 2,0 | 2,1 | 10,2 | 18,1
| Front Cavity Depth | deg 10-3 | 0,1 | 0,1 | 0,1 | 0,1 | 0,1 | 0,0 | 0,0 | 0,2 | 0,4 | 0,9 | 1,9 | 3,5 | 4,3
| Front cavity Volume | deg 10-3 | 0,3 | 0,4 | 0,4 | 0,4 | 0,5 | 0,6 | 0,8 | 0,9 | 1,2 | 1,6 | 2,2 | 3,1 | 5,5
| Repeatability | deg 10-3 | 6,0 | 4,7 | 3,9 | 2,1 | 2,1 | 4,0 | 6,4 | 10,7 | 20,9 | 43,9 | 81,3 | 236,1 | 342,7
| Correction To reference conditions | deg 10-3 | 1,6 | 2,1 | 2,8 | 3,5 | 4,4 | 5,2 | 5,7 | 5,6 | 4,9 | 12,2 | 35,3 | 66,6 | 75,1
| Phase Linearity | deg 10-3 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0 | 250,0
| Rounding | deg 10-3 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0 | 29,0
| U | deg 10-3 | 251,8 | 251,8 | 251,9 | 251,9 | 252,1 | 252,3 | 252,7 | 253,3 | 254,6 | 258,6 | 269,9 | 354,1 | 437,1
| Effective degrees of freedom | deg 10-3 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | >1000 | 365 | 15 | 8
| k | deg 10-3 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,0 | 2,1 | 2,4
| U95 | deg | 0,50 | 0,50 | 0,50 | 0,50 | 0,50 | 0,50 | 0,51 | 0,51 | 0,51 | 0,52 | 0,54 | 0,75 | 1,03
| U95 Rounded | deg | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,6 | 0,8 | 1,1

## METAS Level
METAS | Level
Uncertainty components for average sensitivity [dB]
| Parameter |  | nominal | Unit | Uncertainity |  |  | Frequencies [Hz]
|  |  | Value |  | relative | absolute
|  | frequency range (Min) | [5 empty] | 2 | 5.01 | 10 | 47.32 | 94.41 | 1995.26 | 4731.51 | 7943.28
|  | frequency range (Max) | [5 empty] | 5.01 | 10 | 47.32 | 94.41 | 1995.26 | 4731.51 | 7943.28 | 10000
| Ambi | Static pressure [Pa] | 101325 | Pa | 0.015% | 15 | U95 | 0.00034 | 0.00034 | 0.00034 | 0.00034 | 0.00038 | 0.00056 | 0.00063 | 0.00054
|  | Relative humidity [%] | 50 | % | 3% | 1.5 | U95 | 8e-05 | 8e-05 | 9e-05 | 0.0001 | 0.00012 | 0.00022 | 0.00045 | 0.00079
|  | Absolute temperature [K] | 296.15 | K | 0.101% | 0.3 | U95 | 0.00107 | 0.00098 | 0.00093 | 0.00087 | 0.0015 | 0.00474 | 0.00694 | 0.00678
|  | Pressure Coefficient | -0.015 | dB/kPa | 10% | 0 | U95 | 3e-05 | 3e-05 | 3e-05 | 3e-05 | 4e-05 | 4e-05 | 3e-05 | 7e-05
|  | Temperature Coefficient | -0.015 | dB/K | 10% | 0 | U95 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 4e-05 | 0.00013 | 0.00017 | 0.00014
| Microphone | Front cavity depth [m] | 0.0019 | m | 0.789% | 1.5e-05 | U95 | 0.00089 | 0.00058 | 0.00041 | 0.00019 | 0.00014 | 0.00051 | 0.00063 | 0.00042
|  | Front cavity volume [m3] | 5.35e-07 | m3 | 1.25% | 6.69e-09 | U95 | 0.01836 | 0.01884 | 0.0195 | 0.01966 | 0.01983 | 0.01897 | 0.01411 | 0.01827
|  | Diameter of diaphragm [m] | 0.0179 | m | 20% | 0.00358 | U95 | 0 | 0 | 0 | 2e-05 | 0.00796 | 0.04275 | 0.06333 | 0.08725
|  | Equivalent volume [m3] | 1.36e-07 | m3 | 10% | 1.36e-08 | U95 | 0.03568 | 0.03663 | 0.03791 | 0.03822 | 0.03857 | 0.03702 | 0.02478 | 0.01683
|  | Resonance frequency [kHz] | 7743.3 | kHz | 20% | 1550 | U95 | 0 | 1e-05 | 1e-05 | 2e-05 | 0.00117 | 0.03868 | 0.06618 | 0.02909
|  | Loss factor | 1.0132 |  | 20% | 0.203 | U95 | 0 | 1e-05 | 1e-05 | 1e-05 | 0.01011 | 0.04447 | 0.04554 | 0.04006
| Cpl | Coupler diameter [m] | 0.0186 | m | 0.003% | 5e-07 | U95 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 0.0001 | 0.00014 | 0.00019
|  | Coupler length [m] | 0.01 | m | 0.005% | 5e-07 | U95 | 4e-05 | 2e-05 | 2e-05 | 1e-05 | 2e-05 | 8e-05 | 0.00019 | 0.0003
| Electrical | Receiver voltage [V] | 0.5 | V | 0.1% | 0.0005 | U95 | 0.00324 | 0.00097 | 0.00265 | 0.00177 | 0.00256 | 0.0002 | 0.00019 | 0.00028
|  | Transmitter voltage [V] | 1.4 | V | 0.1% | 0.0014 | U95 | 0.00084 | 0.00038 | 0.00128 | 0.00075 | 0.00106 | 0.00012 | 0.00014 | 0.00013
|  | Receiver ref voltage [V] | 1 | V | 0.1% | 0.001 | U95 | 0.0068 | 0.00189 | 0.00587 | 0.00345 | 0.00485 | 0.00052 | 0.00059 | 0.00055
|  | Transmitter ref voltage [V] | 1 | V | 0.1% | 0.001 | U95 | 0.00398 | 0.00177 | 0.00584 | 0.00345 | 0.00483 | 0.00052 | 0.00059 | 0.00055
|  | Receiver voltage phase [°] | 180 | ° | 0% | 0.1 | U95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Receiver ref voltage phase [°] | 0 | ° | 0% | 0.01 | U95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Polarisation voltage [V] | 200 | V | 0.01% | 0.02 | U95 | 0.00087 | 0.00087 | 0.00087 | 0.00087 | 0.00087 | 0.00087 | 0.00087 | 0.00087
|  | Capacitance [nF] | 4.7 | nF | 0.02% | 0.00094 | U95 | 0.00087 | 0.00087 | 0.00087 | 0.00087 | 0.00087 | 0.00087 | 0.00087 | 0.00087
|  | Reproducibility | [5 empty] | 0.01123 | 0.00387 | 0.00825 | 0.02545 | 0.01769 | 0.01862 | 0.04679 | 0.04679
|  | Thermal model | [5 empty] | 0.12 | 0.048 | 0.024 | 0.005 | 0.003 | 0 | 0 | 0
|  | Max(Sum of squares) | [5 empty] | 0.127 | 0.063 | 0.048 | 0.05 | 0.047 | 0.078 | 0.102 | 0.11
|  | Total uncertainity | [5 empty] | 0.127 | 0.063 | 0.048 | 0.05 | 0.047 | 0.078 | 0.102 | 0.11

## METAS Phase
METAS | Phase
Uncertainty components for average phase [°]
| Parameter |  | nom. Value | Unit | Uncertainity |  |  | Frequencies [Hz]
[5 empty] | relative | absolute
|  | frequency range (Min) | [5 empty] | 2 | 5.01 | 10 | 47.32 | 94.41 | 1995.26 | 4731.51 | 7943.28
|  | frequency range (Max) | [5 empty] | 5.01 | 10 | 47.32 | 94.41 | 1995.26 | 4731.51 | 7943.28 | 10000
| Ambi | Static pressure [Pa] | 101325 | Pa | 0.015% | 15 | U95 | 8e-05 | 5e-05 | 4e-05 | 2e-05 | 9e-05 | 0.0002 | 0.00022 | 0.00035
|  | Relative humidity [%] | 50 | % | 3% | 1.5 | U95 | 0.00023 | 0.00015 | 0.00011 | 5e-05 | 4e-05 | 6e-05 | 0.00052 | 0.00117
|  | Absolute temperature [K] | 296.15 | K | 0.101% | 0.3 | U95 | 0.00109 | 0.00082 | 0.00064 | 0.00033 | 0.00024 | 0.00075 | 0.00416 | 0.00937
|  | Pressure Coefficient | -0.15 | °/kPa | 10% | 0 | U95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Temperature Coefficient | 0.15 | °/K | 10% | 0 | U95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Microphone | Front cavity depth [m] | 0.0019 | m | 0.789% | 1.5e-05 | U95 | 0.00424 | 0.00306 | 0.00232 | 0.00116 | 0.00084 | 0.00238 | 0.00868 | 0.00908
|  | Front cavity volume [m3] | 5.35e-07 | m3 | 1.25% | 6.69e-09 | U95 | 0.01323 | 0.00937 | 0.00704 | 0.00356 | 0.00387 | 0.00948 | 0.01984 | 0.03788
|  | Diameter of diaphragm [m] | 0.0179 | m | 20% | 0.00358 | U95 | 0 | 0 | 0 | 0 | 0.00805 | 0.13975 | 0.55133 | 0.60481
|  | Equivalent volume [m3] | 1.36e-07 | m3 | 10% | 1.36e-08 | U95 | 0.02565 | 0.01805 | 0.01337 | 0.00537 | 0.06023 | 0.13188 | 0.13429 | 0.10884
|  | Resonance frequency [kHz] | 7743.3 | kHz | 20% | 1550 | U95 | 0.00031 | 0.00063 | 0.0031 | 0.00623 | 0.15181 | 0.38659 | 0.38659 | 0.77673
|  | Loss factor | 1.0132 | 0 | 20% | 0.203 | U95 | 0.00031 | 0.00063 | 0.0031 | 0.00623 | 0.11795 | 0.14002 | 0.14841 | 0.08297
| Cpl | Coupler diameter [m] | 0.0186 | m | 0.003% | 5e-07 | U95 | 2e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 0.0001 | 0.0002
|  | Coupler length [m] | 0.01 | m | 0.005% | 5e-07 | U95 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 8e-05 | 0.00017
| Electrical | Receiver voltage [V] | 0.5 | V | 0.1% | 0.0005 | U95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Transmitter voltage [V] | 1.4 | V | 0.1% | 0.0014 | U95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Receiver ref voltage [V] | 1 | V | 0.1% | 0.001 | U95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Transmitter ref voltage [V] | 1 | V | 0.1% | 0.001 | U95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Receiver voltage phase [°] | 180 | ° | 0% | 0.1 | U95 | 0.00037 | 0.00011 | 0.00031 | 0.0002 | 0.00029 | 2e-05 | 2e-05 | 3e-05
|  | Receiver ref voltage phase [°] | 0 | ° | 0% | 0.01 | U95 | 0.00078 | 0.00022 | 0.00068 | 0.0004 | 0.00056 | 6e-05 | 7e-05 | 6e-05
|  | Polarisation voltage [V] | 200 | V | 0.01% | 0.02 | U95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Capacitance [nF] | 4.7 | nF | 0.02% | 0.00094 | U95 | 0.0001 | 4e-05 | 2e-05 | 0 | 0 | 0 | 0 | 0
|  | Reproducibility | [5 empty] | 0.39209 | 0.16714 | 0.08873 | 0.03421 | 0.07803 | 0.21246 | 0.24198 | 0.2895
|  | Thermal model | [5 empty] | 0.3 | 0.12 | 0.06 | 0.01 | 0.01 | 0 | 0 | 0
|  | Max(Sum of squares) | [5 empty] | 0.495 | 0.207 | 0.108 | 0.037 | 0.216 | 0.482 | 0.563 | 1.035
|  | Total uncertainity | [5 empty] | 0.49 | 0.21 | 0.11 | 0.04 | 0.22 | 0.48 | 0.56 | 1.04

## UME Level
Sources of Uncertainty | Probability distribution | Input Value | Unit | Standard Uncertainty (x 10-4 dB)
Frequency (Hz)
|  |  |  | 20 | 25 | 31.5 | 63 | 80 | 100 | 125 | 160 | 200 | 250 | 315 | 400 | 500 | 630 | 800 | 1000 | 1250 | 1600 | 2000 | 2500 | 3150 | 4000 | 5000 | 6300 | 8000 | 10000
microphone parameters
cavity depth | rectangular | 0.05 | mm | 7.6314 | 6.8275 | 6.0787 | 4.3277 | 3.8542 | 3.4468 | 3.0614 | 2.731 | 2.4227 | 2.1474 | 1.8941 | 1.6738 | 1.4426 | 1.2113 | 0.98008 | 0.70477 | 0.35239 | 0.41846 | 1.2113 | 2.3896 | 4.1956 | 6.7834 | 9.9219 | 11.3094 | 8.9528 | 3.8102
loss factor | rectangular | 0.05 |  | 0.011012 | 0.011012 | 0.011012 | 0.005506 | 0.005506 | 0.013765 | 0.02753 | 0.044048 | 0.082591 | 0.14316 | 0.23125 | 0.38542 | 0.60566 | 0.97732 | 1.5857 | 2.4777 | 3.9065 | 6.2439 | 9.5034 | 14.2056 | 20.0888 | 24.1522 | 15.6674 | 17.3275 | 10.3459 | 56.0075
resonance frequency | rectangular | 500 | Hz | -0.011012 | -0.022024 | -0.022024 | -0.022024 | -0.016518 | -0.02753 | -0.02753 | -0.030283 | -0.038542 | -0.038542 | -0.041295 | -0.038542 | 0.46801 | 0.56162 | 0.66623 | 1.104 | 1.371 | -1.6601 | -2.1143 | -3.2045 | -4.6471 | -5.3932 | -9.6934 | -81.4123 | 18.8912 | 8.2205
equivalent volume | rectangular | 10 | mm3 | 114.5697 | 114.9772 | 115.3516 | 115.7673 | 115.9958 | 116.194 | 116.3702 | 116.5106 | 116.6235 | 116.6895 | 116.7116 | 116.6482 | 116.4831 | 116.1444 | 115.5498 | 114.5367 | 112.8546 | 110.1016 | 105.5481 | 97.9607 | 85.1537 | 63.526 | 29.3362 | -6.3209 | 16.375 | 59.3551
front volume | rectangular | 5 | mm3 | 113.4465 | 114.1953 | 114.878 | 116.9153 | 117.3337 | 117.7192 | 118.0495 | 118.3248 | 118.5671 | 118.7323 | 118.8424 | 118.8644 | 118.7653 | 118.468 | 117.9174 | 116.9263 | 115.2745 | 112.5545 | 108.0726 | 100.7385 | 88.5041 | 67.449 | 28.9948 | 54.6309 | 23.8301 | 174.8498
pressure coefficient of the mic. | rectangular | 0.00395 | dB/kPa | 34.0714 | 34.0714 | 34.0714 | 42.1487 | 42.146 | 42.1487 | 42.1487 | 42.146 | 42.1487 | 42.1487 | 42.1543 | 42.1487 | 42.1487 | 42.1487 | 42.1515 | 42.146 | 42.1515 | 42.146 | 42.1487 | 42.146 | 42.1515 | 42.1515 | 42.1432 | 42.1515 | 42.1487 | 42.1487
temperature coefficient of the microphone | rectangular | 0.0011 | dB/°C | 5.1977 | 5.1977 | 5.1977 | 8.661 | 8.661 | 8.661 | 8.661 | 8.6638 | 8.6638 | 8.661 | 8.6638 | 8.661 | 8.661 | 8.661 | 8.6638 | 8.6665 | 8.6638 | 8.6638 | 8.661 | 8.6638 | 8.661 | 8.6638 | 8.6665 | 8.6638 | 8.661 | 8.661
electrical measurements
voltage ratio measurements | normal | 100 | ppm | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044 | 15.044
capacitance | normal | 1.1 | pF | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729 | 10.0729
cross-talk | rectangular | 66 | dB | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217 | 43.5217
polarization voltage | normal | 5 | mV | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714 | 2.1714
frequency of applied signal | rectangular | 5 | ppm | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429 | 0.43429
inherent noise | rectangular | 46 | dB | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909 | 1.0909
distortion | rectangular | 53 | dB | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766 | 0.21766
environmental corrections
static pressure | rectangular | 0.1 | kPa | 7.4772 | 7.4442 | 7.4332 | 7.4001 | 7.4001 | 7.3781 | 7.3671 | 7.3671 | 7.3561 | 7.3561 | 7.334 | 7.334 | 7.3561 | 7.3781 | 7.4111 | 7.4882 | 7.5873 | 7.7745 | 8.0278 | 8.4463 | 9.1621 | 10.3183 | 11.849 | 12.8731 | 10.8194 | 4.2672
temperature | rectangular | 0.3 | °C | 2.0593 | 2.0372 | 2.0152 | 1.9712 | 1.9602 | 1.9491 | 1.9381 | 1.9381 | 1.9381 | 1.9381 | 1.9491 | 1.9602 | 2.0482 | 2.1804 | 2.3566 | 2.6209 | 2.9623 | 3.4688 | 4.2066 | 5.462 | 7.7305 | 11.6177 | 17.4652 | 24.9203 | 18.9188 | 11.3314
relative humidity | rectangular | 1 | %rh | 0.17619 | 0.17619 | 0.17619 | 0.17619 | 0.17619 | 0.17619 | 0.17619 | 0.17619 | 0.17619 | 0.17619 | 0.17619 | 0.17619 | 0.17619 | 0.17619 | 0.18721 | 0.19822 | 0.19822 | 0.21474 | 0.24227 | 0.28631 | 0.3634 | 0.48453 | 0.7268 | 1.1783 | 0.84793 | 1.3545
coupler parameters
coupler length | rectangular | 1 | µm | 1.1976 | 1.1965 | 1.1976 | 1.2023 | 1.1994 | 1.1994 | 1.1976 | 1.2041 | 1.1965 | 1.1965 | 1.1965 | 1.1965 | 1.1918 | 1.1889 | 1.186 | 1.1662 | 1.1558 | 1.1248 | 1.0792 | 1.0096 | 0.90691 | 0.74955 | 0.61888 | 1.1432 | 0.20923 | 1.2003
coupler diameter | rectangular | 6 and 11,5 | µm | 11.4771 | 11.4497 | 11.4281 | 11.4248 | 11.4096 | 11.4003 | 11.388 | 11.3834 | 11.3758 | 11.3712 | 11.3695 | 11.3725 | 11.3817 | 11.3999 | 11.4228 | 11.4684 | 11.5478 | 11.6742 | 11.8771 | 12.2142 | 12.7681 | 13.6863 | 15.2451 | 18.0788 | 16.7604 | 21.3634
coupler surface | rectangular |  |  | 0.4042 | 0.36005 | 0.32105 | 0.22921 | 0.20245 | 0.18002 | 0.16773 | 0.14536 | 0.13306 | 0.12076 | 0.10846 | 0.096159 | 0.083866 | 0.07669 | 0.071579 | 0.061559 | 0.059302 | 0.049248 | 0.047044 | 0.036936 | 0.036936 | 0.034823 | 0.029651 | 0.024624 | 0.022024 | 0.022024
coupler volume | rectangular |  |  | 0.46057 | 0.40648 | 0.36473 | 0.26106 | 0.23644 | 0.20697 | 0.18721 | 0.16259 | 0.15029 | 0.13798 | 0.12312 | 0.10389 | 0.098495 | 0.086183 | 0.076492 | 0.06899 | 0.064211 | 0.056688 | 0.049248 | 0.049248 | 0.039705 | 0.039705 | 0.034823 | 0.034823 | 0.022024 | 0.022024
other
rounding | rectangular |  |  | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868 | 2.8868
repeatability | normal |  |  | 100 | 100 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 100 | 100 | 150 | 150 | 200 | 200 | 250 | 250
Combined Uncertainty | normal |  |  | 199.1797 | 199.8089 | 180.7091 | 184.0344 | 184.4319 | 184.7928 | 185.1061 | 185.364 | 185.5851 | 185.7284 | 185.8105 | 185.7818 | 185.6163 | 185.2173 | 184.5015 | 183.2537 | 181.193 | 177.8499 | 192.9474 | 185.2067 | 205.9678 | 190.4235 | 216.5768 | 235.1278 | 262.1539 | 323.278
Expanded Uncertainty | normal |  |  | 398.3593 | 399.6179 | 361.4182 | 368.0688 | 368.8638 | 369.5857 | 370.2121 | 370.7281 | 371.1701 | 371.4567 | 371.6209 | 371.5636 | 371.2326 | 370.4347 | 369.0029 | 366.5074 | 362.386 | 355.6999 | 385.8947 | 370.4135 | 411.9356 | 380.847 | 433.1535 | 470.2556 | 524.3078 | 646.5559
Reported Uncertainty (dB) |  |  |  | 0.039836 | 0.039962 | 0.036142 | 0.036807 | 0.036886 | 0.036959 | 0.037021 | 0.037073 | 0.037117 | 0.037146 | 0.037162 | 0.037156 | 0.037123 | 0.037043 | 0.0369 | 0.036651 | 0.036239 | 0.03557 | 0.038589 | 0.037041 | 0.041194 | 0.038085 | 0.043315 | 0.047026 | 0.052431 | 0.064656

## BIM Level
BIM | Level
System type 9699 NCM - Uncertainty Budget for Microphones LS1P type 4160
Input Parameter, Symbol | W. Unc.,Unit | 20 Hz | 25 Hz | 31.5 Hz | 40 Hz | 50 Hz | 63 Hz | 80 Hz | 100 Hz | 125 Hz | 160 Hz | 200 Hz | 250 Hz | 315 Hz | 400 Hz | 500 Hz | 630 Hz | 800 Hz | 1 kHz | 1.25 kHz | 1.6 kHz | 2 kHz | 2.5 kHz | 3.15 kHz | 4 kHz | 5 kHz | 6.3 kHz | 8 kHz | 10 kHz
Static Pressure, Psi, Ps2, Psa | 0.038 kPa | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.015 | 0.015 | 0.016 | 0.017 | 0.017 | 0.018
Temperature, Ti,T2,Ta | 0.869 K | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0003 | 0.0009 | 0.0012 | 0.0017 | 0.0028 | 0.0043 | 0.0065 | 0.0117
Relative Humidity, RHi, RHz, RH3 | 4.32 % | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0004 | 0.0005 | 0.0007 | 0.0009 | 0.0012 | 0.0018
Amb. Conditions |  | 0.014005 | 0.014005 | 0.014005 | 0.014005 | 0.014005 | 0.014004 | 0.014003 | 0.014003 | 0.014003 | 0.014003 | 0.014003 | 0.014003 | 0.014003 | 0.014003 | 0.014003 | 0.014004 | 0.014004 | 0.014004 | 0.014004 | 0.014005 | 0.014006 | 0.014032 | 0.015053 | 0.015104 | 0.016258 | 0.017558 | 0.01824 | 0.021544
Coupler Length, Icoup | 0.004 mm | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0002 | 0.0002 | 0.0003 | 0.0004 | 0.0007 | 0.0011 | 0.0016 | 0.003
Coupler Diameter, dcoup | 0.005 mm | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Coupler Volume, Vcoup | 1.54 mm3 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.0022 | 0.0022 | 0.0022 | 0.0021
Coup. Surf. Area, Scoup | 0.35 mm2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Coupler Leakage |  | 0.017 | 0.0095 | 0.0073 | 0.0041 | 0.0039 | 0.0037 | 0.0025 | 0.001 | 0.0003 | 0.0003 | 0.0002 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Coupler |  | 0.017118 | 0.0097088 | 0.0075697 | 0.0045629 | 0.0043841 | 0.0042071 | 0.0032016 | 0.0022361 | 0.0020224 | 0.0020224 | 0.00201 | 0.0020025 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.0020025 | 0.00201 | 0.00201 | 0.0020224 | 0.0020396 | 0.0023087 | 0.0024597 | 0.0027203 | 0.003662
Mic. Front Length (Nom.), La | 0.03 mm | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0002 | 0.0002 | 0.0002 | 0.0004 | 0.0005 | 0.0006 | 0.0011 | 0.0015 | 0.002 | 0.004 | 0.006 | 0.008 | 0.0146
Extra Fr. Length (thread), Lax | 0.35 mm | 0.0033 | 0.0028 | 0.0024 | 0.0021 | 0.0018 | 0.0015 | 0.0012 | 0.0009 | 0.0008 | 0.0007 | 0.0005 | 0.0004 | 0.0003 | 0.0002 | 0.0001 | 0.0002 | 0.0003 | 0.0004 | 0.0004 | 0.0005 | 0.0005 | 0.0006 | 0.0007 | 0.0008 | 0.0009 | 0.001 | 0.001 | 0.001
Eqv. and Front Vol., Veoa, Va (Vtot) | 2.4/0.8 mm3 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.018 | 0.015 | 0.012 | 0.011 | 0.01 | 0.01 | 0.00428 | 0.009 | 0.007
Diaphragm Res., foa | 550 Hz | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0002 | 0.0003 | 0.0004 | 0.001 | 0.003 | 0.0056 | 0.0087 | 0.014 | 0.017 | 0.018
Diaph. Damp. Fac., Da | 0.094 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0002 | 0.0005 | 0.0008 | 0.001 | 0.0015 | 0.0024 | 0.0036 | 0.0042 | 0.0055 | 0.0124 | 0.0152 | 0.0164 | 0.004 | 0.02
Microphones |  | 0.018301 | 0.018218 | 0.01816 | 0.018123 | 0.018091 | 0.018063 | 0.018041 | 0.018022 | 0.018018 | 0.018014 | 0.018007 | 0.018004 | 0.018003 | 0.018001 | 0.018001 | 0.018009 | 0.018021 | 0.018033 | 0.018072 | 0.018176 | 0.015451 | 0.012814 | 0.012767 | 0.017022 | 0.02058 | 0.02281 | 0.021237 | 0.031419
Series Cap., C | pF | 0.0046 | 0.0039 | 0.0032 | 0.0032 | 0.0032 | 0.0017 | 0.0015 | 0.0015 | 0.0015 | 0.0014 | 0.0013 | 0.0012 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0014 | 0.0015 | 0.0017 | 0.0019 | 0.0021 | 0.0022 | 0.0023 | 0.0025 | 0.0026 | 0.0027
Volt. Ratio, DVM VRab, VRac, VRbc |  | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0009 | 0.001 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012
Volt. Ratio, Cr-talk VRab, VRac, VRbc | < -66 dB - | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025
Volt. Ratio, Noise VRab, VRac, VRbc | < -46 dB - | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001
Volt. Ratio, Distort. VRab, VRac, VRbc | < -46 dB - | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001
Frequency, f | 12 ppm | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0002
Pol. Voltage, Uo | 0.08 V | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042 | 0.0042
Electrical Parameters |  | 0.006879 | 0.006432 | 0.0060332 | 0.0060332 | 0.0060332 | 0.0053141 | 0.0052536 | 0.0052536 | 0.0052536 | 0.0052259 | 0.0052 | 0.005098 | 0.0051225 | 0.0051225 | 0.0051225 | 0.0051225 | 0.0051225 | 0.0051225 | 0.0051488 | 0.0051769 | 0.0052383 | 0.0053226 | 0.0054148 | 0.0054945 | 0.0055362 | 0.0056223 | 0.0056675 | 0.0057166
Reproduc./Stabil., Mp | dB | 0.0123 | 0.0115 | 0.01 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0058 | 0.0068 | 0.008 | 0.009 | 0.011 | 0.012 | 0.028
Result Pounding, Mp | < 0.0005 dB | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006
Uncertainty at Measurement Conditions, dB |  | 0.031985 | 0.028218 | 0.026833 | 0.024703 | 0.024647 | 0.024429 | 0.024246 | 0.024124 | 0.024101 | 0.024092 | 0.02408 | 0.024056 | 0.02406 | 0.024059 | 0.024059 | 0.024065 | 0.024074 | 0.024083 | 0.024118 | 0.024202 | 0.022245 | 0.020676 | 0.02167 | 0.024832 | 0.028376 | 0.031426 | 0.031106 | 0.047768
Sensitivity Correction for Static Pressure, dB |  | 0.0045 | 0.0045 | 0.0045 | 0.0045 | 0.0045 | 0.0045 | 0.0045 | 0.0045 | 0.0045 | 0.0045 | 0.0045 | 0.0045 | 0.0045 | 0.0045 | 0.0045 | 0.0054 | 0.0061 | 0.007 | 0.008 | 0.01 | 0.011 | 0.012 | 0.014 | 0.019 | 0.021 | 0.024 | 0.027 | 0.031
Sensitivity Correction for Temperature, dB |  | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0036 | 0.0039 | 0.0039 | 0.0042 | 0.0046 | 0.0053 | 0.0097
Uncertainty at Reference Conditions, dB |  | 0.0325 | 0.0288 | 0.027444 | 0.025367 | 0.025312 | 0.0251 | 0.024921 | 0.024802 | 0.024781 | 0.024772 | 0.02476 | 0.024737 | 0.02474 | 0.024739 | 0.024739 | 0.024925 | 0.025094 | 0.025337 | 0.025664 | 0.026433 | 0.025076 | 0.024175 | 0.026092 | 0.031509 | 0.035551 | 0.039809 | 0.041529 | 0.057765
Specified Uncertainty (k=2), dB |  | 0.065 | 0.057601 | 0.054889 | 0.050733 | 0.050624 | 0.050199 | 0.049843 | 0.049605 | 0.049561 | 0.049544 | 0.049521 | 0.049474 | 0.049481 | 0.049479 | 0.049479 | 0.049849 | 0.050189 | 0.050673 | 0.051328 | 0.052866 | 0.050152 | 0.04835 | 0.052185 | 0.063018 | 0.071101 | 0.079618 | 0.083058 | 0.11553

## DMDM Level
DMDM | Level
| Input | Sym. | Unit | Unce. |  | 31.5Hz | 63Hz | 125Hz | 250Hz | 500Hz | 1kHz | 2kHz | 3.15 | 4kHz | 5kHz | 6.3 | 8kHz | 10kHz | 12.5
| parameter | [11 empty] | kHz |  |  | kHz |  |  | kHz
| Voltage ratio correction | CorR,n | dB | A | B | 0.0021 | 0.0021 | 0.0021 | 0.0017 | 0.0017 | 0.0017 | 0.0017 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.0023
|  |  |  | σ | 0.0017
| Polarization voltage | UPOL | V | 0.0008 V |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
norm.
| Coupler volume correction | CorCV,n | mm3 | 2.89 mm3 |  | 0.0037 | 0.0037 | 0.0037 | 0.0037 | 0.0037 | 0.0037 | 0.0037 | 0.0037 | 0.0037 | 0.0037 | 0.0037 | 0.0037 | 0.0037 | 0.0037
rect.
| Static pressure correction | CorPs | kPa | 0.035 kPa |  | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015
norm.
| Capacitance correction | CorC | nF | 0.0012 nF |  | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011
norm.
| Microphone front volume correct. | CorFV,n | mm3 | 5.78 mm3 |  | 0.0374 | 0.0374 | 0.0374 | 0.0374 | 0.0374 | 0.0374 | 0.0374 | 0.0374 | 0.0374 | 0.0374 | 0.0374 | 0.0374 | 0.0374 | 0.0374
rect.
| Static pressure correction | CP | kPa | 0.035 kPa |  | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0.0002
norm.
| Temperature correction | CT | K | 1.169 K |  | 0.0002 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0006 | 0.0016 | 0.0023 | 0.0039 | 0.0061 | 0.0097 | 0.0167 | 0.0393
norm.
| Humidity correction | CRH | % | 3.424% |  | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0003 | 0.0004 | 0.0005 | 0.0007 | 0.001 | 0.0015 | 0.0033
norm.
| Mic. front cavity length correct. | CL | mm | 0.058 mm |  | 0.0006 | 0.0004 | 0.0003 | 0.0002 | 0.0004 | 0.0009 | 0.003 | 0.0076 | 0.0122 | 0.0196 | 0.0311 | 0.0495 | 0.0864 | 0.2023
rect.
| Reproducibility | Mp | dB | σA |  | 0.02 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.015 | 0.015 | 0.015 | 0.02 | 0.025 | 0.03
| Result rounding | Mp | dB | 0.001 |  | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006
rect.
| Combined standard uncertainty |  |  |  |  | 0.0427 | 0.039 | 0.039 | 0.039 | 0.039 | 0.039 | 0.0391 | 0.0398 | 0.0425 | 0.0452 | 0.0515 | 0.0661 | 0.099 | 0.2117
| Expanded uncertainty (k=2) |  |  |  |  | 0.0854 | 0.078 | 0.078 | 0.078 | 0.078 | 0.078 | 0.0782 | 0.0795 | 0.0849 | 0.0904 | 0.103 | 0.1321 | 0.1979 | 0.4233
| Stated uncertainty (k=2) |  |  |  |  | 0.09 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.09 | 0.1 | 0.11 | 0.14 | 0.2 | 0.43

## BEV Level
BEV | Level
[24 empty] | s_result (dB) |  | u_result (dB) |  | U_result (dB)
f (Hz) | u_f (dB) | u_p (dB) | u_t (dB) | u_rH (dB) | u_Lfc (dB) | u_VF (dB) | u_Ve (dB) | u_f0 (dB) | u_D (dB) | u_dr (dB) | u_S0_add (dB) | u_c_d (dB) | u_c_h (dB) | u_c_sur (dB) | u_c_vol (dB) | u_p_corr (dB) | u_t_corr (dB) | u_C (dB) | u_R (dB) | u_f (dB) | u_PV (dB) | u_round (dB) | Sum (dB) | 811014 | 2036126 | 811014 | 2036126 | 811014 | 2036126
1.9953 | 2e-05 | 0.0031 | 0.00218 | 0.0001 | 0.00239 | 0.01497 | 0.0104 | 0 | 0 | 0 | 0.01079 | 0.0032 | 0.00272 | 0.00027 | 7e-05 | 8e-05 | 0.00048 | 0.00174 | 0.00524 | 0.00013 | 0.00029 | 0.00029 | 0.02275 | 0.01781 | 0.01927 | 0.02889 | 0.02981 | 0.05778 | 0.05962
2.5119 | 2e-05 | 0.00306 | 0.00202 | 0.00011 | 0.002 | 0.01452 | 0.01058 | 0 | 0 | 0 | 0.00916 | 0.00321 | 0.00278 | 0.00023 | 7e-05 | 0.00015 | 0.00051 | 0.00174 | 0.00524 | 0.00013 | 0.00029 | 0.00029 | 0.02175 | 0.01324 | 0.01043 | 0.02546 | 0.02412 | 0.05092 | 0.04824
3.1623 | 1e-05 | 0.00302 | 0.0019 | 0.00011 | 0.00168 | 0.01414 | 0.01074 | 0 | 0 | 0 | 0.00781 | 0.00322 | 0.00282 | 0.0002 | 7e-05 | 0.0002 | 0.00061 | 0.00174 | 0.00524 | 0.00013 | 0.00029 | 0.00029 | 0.02101 | 0.00945 | 0.00729 | 0.02303 | 0.02224 | 0.04607 | 0.04447
3.9811 | 1e-05 | 0.00299 | 0.0018 | 0.00012 | 0.00144 | 0.01386 | 0.01088 | 0 | 0 | 0 | 0.00677 | 0.00323 | 0.00284 | 0.00017 | 7e-05 | 0.0002 | 0.00066 | 0.00174 | 0.00524 | 0.00013 | 0.00029 | 0.00029 | 0.0205 | 0.00686 | 0.00566 | 0.02162 | 0.02127 | 0.04323 | 0.04253
5.0119 | 1e-05 | 0.00296 | 0.00172 | 0.00012 | 0.00127 | 0.01365 | 0.011 | 0 | 0 | 0 | 0.00598 | 0.00324 | 0.00286 | 0.00015 | 7e-05 | 0.00019 | 0.00058 | 0.00174 | 0.00524 | 0.00013 | 0.00029 | 0.00029 | 0.02015 | 0.00564 | 0.00482 | 0.02093 | 0.02072 | 0.04186 | 0.04144
6.3096 | 1e-05 | 0.00294 | 0.00166 | 0.00013 | 0.00114 | 0.01348 | 0.01111 | 0 | 0 | 0 | 0.00534 | 0.00325 | 0.00288 | 0.00013 | 7e-05 | 0.00019 | 0.00058 | 0.00174 | 0.00524 | 0.00013 | 0.00029 | 0.00029 | 0.01991 | 0.00482 | 0.00385 | 0.02048 | 0.02028 | 0.04096 | 0.04055
7.9433 | 1e-05 | 0.00292 | 0.00161 | 0.00013 | 0.00102 | 0.01333 | 0.01121 | 0 | 0 | 0 | 0.00478 | 0.00325 | 0.00289 | 0.00012 | 7e-05 | 0.00018 | 0.00059 | 0.00174 | 0.00524 | 0.00013 | 0.00029 | 0.00029 | 0.01971 | 0.00395 | 0.00321 | 0.0201 | 0.01997 | 0.0402 | 0.03993
10 | 1e-05 | 0.00291 | 0.00156 | 0.00013 | 0.00091 | 0.01319 | 0.0113 | 0 | 0 | 0 | 0.00426 | 0.00326 | 0.0029 | 0.00011 | 7e-05 | 0.00017 | 0.00059 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01953 | 0.00358 | 0.00271 | 0.01986 | 0.01972 | 0.03972 | 0.03944
12.5892 | 0 | 0.00289 | 0.00151 | 0.00014 | 0.00081 | 0.01306 | 0.01138 | 0 | 0 | 0 | 0.00379 | 0.00326 | 0.00291 | 0.0001 | 7e-05 | 0.00017 | 0.00059 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01939 | 0.00329 | 0.00242 | 0.01967 | 0.01954 | 0.03933 | 0.03908
15.8489 | 0 | 0.00288 | 0.00147 | 0.00014 | 0.00072 | 0.01295 | 0.01145 | 0 | 0 | 0 | 0.00336 | 0.00327 | 0.00292 | 8e-05 | 7e-05 | 0.00017 | 0.00063 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01927 | 0.00288 | 0.00212 | 0.01949 | 0.01939 | 0.03897 | 0.03878
19.9526 | 0 | 0.00286 | 0.00144 | 0.00014 | 0.00064 | 0.01285 | 0.01152 | 0 | 0 | 0 | 0.003 | 0.00327 | 0.00293 | 8e-05 | 8e-05 | 5e-05 | 0.00141 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01922 | 0.00209 | 0.00205 | 0.01933 | 0.01933 | 0.03866 | 0.03865
21.1349 | 0 | 0.00286 | 0.00143 | 0.00015 | 0.00062 | 0.01283 | 0.01153 | 0 | 0 | 0 | 0.00291 | 0.00327 | 0.00293 | 7e-05 | 8e-05 | 5e-05 | 0.00139 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0192 | 0.00184 | 0.00208 | 0.01928 | 0.01931 | 0.03857 | 0.03862
22.3872 | 0 | 0.00286 | 0.00142 | 0.00015 | 0.00061 | 0.0128 | 0.01155 | 0 | 0 | 0 | 0.00283 | 0.00327 | 0.00293 | 7e-05 | 8e-05 | 5e-05 | 0.00123 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01916 | 0.0019 | 0.00212 | 0.01926 | 0.01928 | 0.03852 | 0.03856
23.7137 | 0 | 0.00286 | 0.00141 | 0.00015 | 0.00059 | 0.01278 | 0.01156 | 0 | 0 | 0 | 0.00275 | 0.00327 | 0.00293 | 7e-05 | 8e-05 | 5e-05 | 0.00121 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01914 | 0.00221 | 0.0015 | 0.01927 | 0.0192 | 0.03854 | 0.03841
25.1189 | 0 | 0.00285 | 0.00141 | 0.00015 | 0.00057 | 0.01276 | 0.01157 | 0 | 0 | 0 | 0.00267 | 0.00327 | 0.00294 | 7e-05 | 8e-05 | 6e-05 | 0.00118 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01912 | 0.00217 | 0.00114 | 0.01925 | 0.01916 | 0.03849 | 0.03832
26.6073 | 0 | 0.00285 | 0.0014 | 0.00015 | 0.00056 | 0.01274 | 0.01159 | 0 | 0 | 0 | 0.00259 | 0.00327 | 0.00294 | 7e-05 | 8e-05 | 7e-05 | 0.00098 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01909 | 0.00236 | 0.00119 | 0.01924 | 0.01913 | 0.03848 | 0.03826
28.1838 | 0 | 0.00285 | 0.00139 | 0.00015 | 0.00054 | 0.01272 | 0.0116 | 0 | 0 | 0 | 0.00252 | 0.00327 | 0.00294 | 6e-05 | 8e-05 | 7e-05 | 0.00097 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01908 | 0.00213 | 0.00176 | 0.0192 | 0.01916 | 0.03839 | 0.03832
29.8538 | 0 | 0.00285 | 0.00138 | 0.00015 | 0.00052 | 0.0127 | 0.01161 | 0 | 0 | 0 | 0.00245 | 0.00328 | 0.00294 | 6e-05 | 8e-05 | 5e-05 | 0.00088 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01906 | 0.00234 | 0.00169 | 0.0192 | 0.01913 | 0.0384 | 0.03827
31.6228 | 0 | 0.00284 | 0.00138 | 0.00015 | 0.00051 | 0.01268 | 0.01162 | 0 | 0 | 0 | 0.00238 | 0.00328 | 0.00294 | 6e-05 | 8e-05 | 5e-05 | 0.0007 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01904 | 0.00208 | 0.00163 | 0.01915 | 0.01911 | 0.0383 | 0.03821
33.4965 | 0 | 0.00284 | 0.00137 | 0.00015 | 0.00049 | 0.01266 | 0.01164 | 0 | 0 | 0 | 0.00231 | 0.00328 | 0.00294 | 6e-05 | 8e-05 | 4e-05 | 0.0007 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01902 | 0.00224 | 0.00155 | 0.01915 | 0.01908 | 0.0383 | 0.03817
35.4813 | 0 | 0.00284 | 0.00136 | 0.00015 | 0.00048 | 0.01265 | 0.01165 | 0 | 0 | 0 | 0.00224 | 0.00328 | 0.00294 | 6e-05 | 8e-05 | 4e-05 | 0.00066 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01901 | 0.00216 | 0.00165 | 0.01913 | 0.01908 | 0.03826 | 0.03815
37.5837 | 0 | 0.00284 | 0.00135 | 0.00015 | 0.00047 | 0.01263 | 0.01166 | 0 | 0 | 0 | 0.00218 | 0.00328 | 0.00295 | 5e-05 | 8e-05 | 4e-05 | 0.00056 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01899 | 0.00206 | 0.00135 | 0.0191 | 0.01904 | 0.0382 | 0.03808
39.8107 | 0 | 0.00284 | 0.00135 | 0.00015 | 0.00045 | 0.01261 | 0.01167 | 0 | 0 | 0 | 0.00211 | 0.00328 | 0.00295 | 5e-05 | 8e-05 | 5e-05 | 0.00049 | 0.00174 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01898 | 0.00225 | 0.00142 | 0.01911 | 0.01903 | 0.03822 | 0.03806
42.1696 | 0 | 0.00283 | 0.00134 | 0.00015 | 0.00044 | 0.0126 | 0.01168 | 0 | 0 | 0 | 0.00205 | 0.00328 | 0.00295 | 5e-05 | 8e-05 | 5e-05 | 0.00049 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01892 | 0.00209 | 0.00142 | 0.01903 | 0.01897 | 0.03806 | 0.03794
44.6684 | 0 | 0.00283 | 0.00133 | 0.00015 | 0.00043 | 0.01258 | 0.01169 | 0 | 0 | 0 | 0.002 | 0.00328 | 0.00295 | 5e-05 | 8e-05 | 6e-05 | 0.00055 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01891 | 0.00212 | 0.00158 | 0.01902 | 0.01897 | 0.03805 | 0.03794
47.3151 | 0 | 0.00283 | 0.00133 | 0.00015 | 0.00042 | 0.01256 | 0.0117 | 0 | 0 | 0 | 0.00194 | 0.00328 | 0.00295 | 5e-05 | 8e-05 | 6e-05 | 0.00056 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0189 | 0.00229 | 0.00144 | 0.01903 | 0.01895 | 0.03807 | 0.0379
50.1187 | 0 | 0.00283 | 0.00132 | 0.00015 | 0.0004 | 0.01255 | 0.01171 | 0 | 0 | 0 | 0.00188 | 0.00328 | 0.00295 | 5e-05 | 8e-05 | 6e-05 | 0.0006 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01889 | 0.00206 | 0.00127 | 0.019 | 0.01893 | 0.038 | 0.03786
53.0885 | 0 | 0.00283 | 0.00132 | 0.00015 | 0.00039 | 0.01253 | 0.01172 | 0 | 0 | 0 | 0.00183 | 0.00328 | 0.00295 | 5e-05 | 8e-05 | 7e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01888 | 0.0022 | 0.00137 | 0.019 | 0.01892 | 0.03801 | 0.03785
56.2341 | 0 | 0.00283 | 0.00131 | 0.00015 | 0.00038 | 0.01252 | 0.01173 | 1e-05 | 0 | 0 | 0.00178 | 0.00328 | 0.00296 | 4e-05 | 8e-05 | 7e-05 | 0.0006 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01887 | 0.00208 | 0.00158 | 0.01898 | 0.01893 | 0.03796 | 0.03786
59.5662 | 0 | 0.00282 | 0.0013 | 0.00015 | 0.00037 | 0.01251 | 0.01174 | 1e-05 | 1e-05 | 0 | 0.00173 | 0.00328 | 0.00296 | 4e-05 | 8e-05 | 8e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01886 | 0.00221 | 0.00137 | 0.01899 | 0.01891 | 0.03797 | 0.03781
63.0957 | 0 | 0.00282 | 0.0013 | 0.00015 | 0.00036 | 0.01249 | 0.01174 | 1e-05 | 1e-05 | 0 | 0.00168 | 0.00328 | 0.00296 | 4e-05 | 8e-05 | 8e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01885 | 0.00215 | 0.0016 | 0.01897 | 0.01892 | 0.03794 | 0.03783
66.8344 | 0 | 0.00282 | 0.00129 | 0.00015 | 0.00035 | 0.01248 | 0.01175 | 1e-05 | 1e-05 | 0 | 0.00163 | 0.00328 | 0.00296 | 4e-05 | 8e-05 | 8e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01884 | 0.00233 | 0.00155 | 0.01898 | 0.0189 | 0.03797 | 0.03781
70.7946 | 0 | 0.00282 | 0.00129 | 0.00016 | 0.00034 | 0.01247 | 0.01176 | 1e-05 | 1e-05 | 0 | 0.00158 | 0.00328 | 0.00296 | 4e-05 | 8e-05 | 8e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01883 | 0.00213 | 0.00155 | 0.01895 | 0.0189 | 0.0379 | 0.03779
74.9894 | 0 | 0.00282 | 0.00128 | 0.00016 | 0.00033 | 0.01246 | 0.01177 | 1e-05 | 1e-05 | 0 | 0.00154 | 0.00329 | 0.00296 | 4e-05 | 8e-05 | 8e-05 | 0.00057 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01882 | 0.00216 | 0.00161 | 0.01895 | 0.01889 | 0.0379 | 0.03779
79.4328 | 0 | 0.00282 | 0.00128 | 0.00016 | 0.00032 | 0.01244 | 0.01178 | 1e-05 | 1e-05 | 0 | 0.00149 | 0.00329 | 0.00296 | 4e-05 | 8e-05 | 9e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01882 | 0.00221 | 0.0016 | 0.01895 | 0.01889 | 0.03789 | 0.03777
84.1395 | 0 | 0.00281 | 0.00127 | 0.00016 | 0.00031 | 0.01243 | 0.01178 | 1e-05 | 1e-05 | 0 | 0.00145 | 0.00329 | 0.00296 | 4e-05 | 8e-05 | 9e-05 | 0.00055 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01881 | 0.00241 | 0.0015 | 0.01896 | 0.01887 | 0.03793 | 0.03774
89.1251 | 0 | 0.00281 | 0.00127 | 0.00016 | 0.0003 | 0.01242 | 0.01179 | 1e-05 | 1e-05 | 0 | 0.00141 | 0.00329 | 0.00296 | 4e-05 | 8e-05 | 9e-05 | 0.00053 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0188 | 0.00207 | 0.00175 | 0.01892 | 0.01888 | 0.03783 | 0.03777
94.4061 | 0 | 0.00281 | 0.00126 | 0.00016 | 0.00029 | 0.01241 | 0.0118 | 1e-05 | 1e-05 | 0 | 0.00137 | 0.00329 | 0.00296 | 3e-05 | 8e-05 | 9e-05 | 0.00051 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0188 | 0.00239 | 0.00145 | 0.01895 | 0.01885 | 0.03789 | 0.0377
100 | 0 | 0.00281 | 0.00126 | 0.00016 | 0.00029 | 0.0124 | 0.0118 | 1e-05 | 1e-05 | 0 | 0.00133 | 0.00329 | 0.00297 | 3e-05 | 8e-05 | 9e-05 | 0.00045 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01879 | 0.00243 | 0.00156 | 0.01894 | 0.01885 | 0.03789 | 0.03771
105.9254 | 0 | 0.00281 | 0.00125 | 0.00016 | 0.00028 | 0.01239 | 0.01181 | 1e-05 | 1e-05 | 0 | 0.00129 | 0.00329 | 0.00297 | 3e-05 | 8e-05 | 0.0001 | 0.00046 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01878 | 0.00249 | 0.00151 | 0.01895 | 0.01884 | 0.03789 | 0.03769
112.2018 | 0 | 0.00281 | 0.00125 | 0.00016 | 0.00027 | 0.01238 | 0.01182 | 1e-05 | 1e-05 | 0 | 0.00126 | 0.00329 | 0.00297 | 3e-05 | 8e-05 | 0.0001 | 0.00047 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01878 | 0.0023 | 0.00165 | 0.01892 | 0.01885 | 0.03784 | 0.0377
118.8502 | 0 | 0.00281 | 0.00124 | 0.00016 | 0.00026 | 0.01237 | 0.01182 | 1e-05 | 1e-05 | 0 | 0.00122 | 0.00329 | 0.00297 | 3e-05 | 8e-05 | 0.0001 | 0.00057 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01878 | 0.00237 | 0.00181 | 0.01892 | 0.01886 | 0.03785 | 0.03772
125.8925 | 0 | 0.00281 | 0.00124 | 0.00016 | 0.00025 | 0.01236 | 0.01183 | 1e-05 | 1e-05 | 0 | 0.00118 | 0.00329 | 0.00297 | 3e-05 | 8e-05 | 0.0001 | 0.00054 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01877 | 0.00238 | 0.0019 | 0.01892 | 0.01886 | 0.03784 | 0.03773
133.3521 | 0 | 0.0028 | 0.00123 | 0.00016 | 0.00025 | 0.01235 | 0.01183 | 1e-05 | 1e-05 | 0 | 0.00115 | 0.00329 | 0.00297 | 3e-05 | 8e-05 | 0.0001 | 0.00056 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01877 | 0.00257 | 0.00177 | 0.01894 | 0.01885 | 0.03788 | 0.0377
141.2538 | 0 | 0.0028 | 0.00122 | 0.00016 | 0.00024 | 0.01234 | 0.01184 | 2e-05 | 1e-05 | 0 | 0.00112 | 0.00211 | 0.00297 | 3e-05 | 8e-05 | 6e-05 | 0.00085 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0186 | 0.00258 | 0.00204 | 0.01878 | 0.01871 | 0.03756 | 0.03742
149.6235 | 0 | 0.0028 | 0.00122 | 0.00016 | 0.00023 | 0.01234 | 0.01184 | 2e-05 | 1e-05 | 0 | 0.00109 | 0.00211 | 0.00297 | 3e-05 | 8e-05 | 6e-05 | 0.00084 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0186 | 0.00261 | 0.0021 | 0.01878 | 0.01871 | 0.03756 | 0.03743
158.4893 | 0 | 0.0028 | 0.00122 | 0.00016 | 0.00023 | 0.01233 | 0.01185 | 2e-05 | 1e-05 | 0 | 0.00106 | 0.00211 | 0.00297 | 3e-05 | 8e-05 | 6e-05 | 0.00084 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01859 | 0.00256 | 0.00207 | 0.01877 | 0.01871 | 0.03753 | 0.03741
167.8804 | 0 | 0.0028 | 0.00121 | 0.00016 | 0.00022 | 0.01232 | 0.01185 | 2e-05 | 1e-05 | 0 | 0.00103 | 0.00211 | 0.00297 | 3e-05 | 8e-05 | 7e-05 | 0.00084 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01859 | 0.00258 | 0.00216 | 0.01876 | 0.01871 | 0.03753 | 0.03742
177.8279 | 0 | 0.0028 | 0.00121 | 0.00016 | 0.00021 | 0.01231 | 0.01186 | 2e-05 | 1e-05 | 0 | 0.001 | 0.00211 | 0.00297 | 3e-05 | 8e-05 | 7e-05 | 0.00082 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01858 | 0.0024 | 0.00229 | 0.01874 | 0.01872 | 0.03747 | 0.03744
188.3649 | 0 | 0.0028 | 0.00121 | 0.00016 | 0.00021 | 0.0123 | 0.01186 | 3e-05 | 1e-05 | 0 | 0.00097 | 0.00211 | 0.00297 | 2e-05 | 8e-05 | 7e-05 | 0.00079 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01858 | 0.00259 | 0.00238 | 0.01876 | 0.01873 | 0.03751 | 0.03746
199.5262 | 0 | 0.00279 | 0.0012 | 0.00016 | 0.0002 | 0.0123 | 0.01187 | 3e-05 | 1e-05 | 0 | 0.00094 | 0.00211 | 0.00297 | 2e-05 | 8e-05 | 7e-05 | 0.0007 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01857 | 0.00249 | 0.00248 | 0.01873 | 0.01873 | 0.03747 | 0.03747
211.3489 | 0 | 0.00279 | 0.0012 | 0.00016 | 0.0002 | 0.01229 | 0.01187 | 3e-05 | 1e-05 | 0 | 0.00091 | 0.00211 | 0.00297 | 2e-05 | 8e-05 | 7e-05 | 0.00069 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01856 | 0.00277 | 0.0025 | 0.01877 | 0.01873 | 0.03754 | 0.03746
223.8721 | 0 | 0.00279 | 0.0012 | 0.00016 | 0.00019 | 0.01228 | 0.01187 | 3e-05 | 1e-05 | 0 | 0.00089 | 0.00211 | 0.00298 | 2e-05 | 8e-05 | 8e-05 | 0.00067 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01856 | 0.00285 | 0.00248 | 0.01878 | 0.01873 | 0.03755 | 0.03745
237.1374 | 0 | 0.00279 | 0.0012 | 0.00016 | 0.00019 | 0.01228 | 0.01188 | 4e-05 | 1e-05 | 1e-05 | 0.00086 | 0.00211 | 0.00298 | 2e-05 | 8e-05 | 8e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01855 | 0.00283 | 0.00263 | 0.01877 | 0.01874 | 0.03754 | 0.03748
251.1887 | 0 | 0.00279 | 0.00118 | 0.00015 | 0.0001 | 0.01211 | 0.01184 | 4e-05 | 1e-05 | 1e-05 | 0.00099 | 0.00209 | 0.00297 | 0 | 8e-05 | 8e-05 | 0.00059 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01843 | 0.00162 | 0.00118 | 0.0185 | 0.01846 | 0.03699 | 0.03693
266.0725 | 0 | 0.00279 | 0.00118 | 0.00015 | 0.0001 | 0.01211 | 0.01185 | 5e-05 | 1e-05 | 1e-05 | 0.00096 | 0.00209 | 0.00297 | 0 | 8e-05 | 8e-05 | 0.00056 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01842 | 0.00191 | 0.00143 | 0.01852 | 0.01848 | 0.03704 | 0.03696
281.8383 | 0 | 0.00279 | 0.00118 | 0.00015 | 0.0001 | 0.01211 | 0.01185 | 5e-05 | 1e-05 | 1e-05 | 0.00093 | 0.00209 | 0.00297 | 0 | 8e-05 | 8e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01842 | 0.002 | 0.00134 | 0.01853 | 0.01847 | 0.03706 | 0.03695
298.5382 | 0 | 0.00279 | 0.00118 | 0.00015 | 0.0001 | 0.01211 | 0.01185 | 6e-05 | 1e-05 | 1e-05 | 0.00091 | 0.00209 | 0.00297 | 0 | 8e-05 | 8e-05 | 0.00056 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01842 | 0.00202 | 0.00149 | 0.01853 | 0.01848 | 0.03707 | 0.03696
316.2278 | 0 | 0.00279 | 0.00118 | 0.00015 | 9e-05 | 0.01211 | 0.01185 | 6e-05 | 1e-05 | 1e-05 | 0.00088 | 0.00209 | 0.00297 | 0 | 8e-05 | 6e-05 | 0.00054 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01842 | 0.0021 | 0.00138 | 0.01854 | 0.01847 | 0.03708 | 0.03694
334.9655 | 0 | 0.00279 | 0.00118 | 0.00015 | 9e-05 | 0.01211 | 0.01185 | 7e-05 | 1e-05 | 1e-05 | 0.00086 | 0.0021 | 0.00297 | 0 | 8e-05 | 6e-05 | 0.00053 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01842 | 0.00206 | 0.00136 | 0.01853 | 0.01847 | 0.03707 | 0.03694
354.8134 | 0 | 0.00279 | 0.00119 | 0.00015 | 9e-05 | 0.0121 | 0.01185 | 8e-05 | 1e-05 | 1e-05 | 0.00083 | 0.0021 | 0.00297 | 0 | 8e-05 | 6e-05 | 0.00053 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01842 | 0.00217 | 0.00136 | 0.01855 | 0.01847 | 0.03709 | 0.03694
375.8374 | 0 | 0.00279 | 0.00119 | 0.00015 | 9e-05 | 0.0121 | 0.01185 | 9e-05 | 1e-05 | 1e-05 | 0.00081 | 0.0021 | 0.00297 | 0 | 8e-05 | 7e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01842 | 0.00196 | 0.00131 | 0.01852 | 0.01846 | 0.03705 | 0.03693
398.1071 | 0 | 0.00279 | 0.0012 | 0.00015 | 9e-05 | 0.0121 | 0.01185 | 0.0001 | 1e-05 | 1e-05 | 0.00078 | 0.0021 | 0.00297 | 0 | 8e-05 | 7e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01842 | 0.00198 | 0.00122 | 0.01852 | 0.01846 | 0.03705 | 0.03691
421.6965 | 0 | 0.00279 | 0.0012 | 0.00015 | 9e-05 | 0.0121 | 0.01185 | 0.00011 | 1e-05 | 2e-05 | 0.00076 | 0.0021 | 0.00297 | 0 | 8e-05 | 7e-05 | 0.00056 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01841 | 0.00196 | 0.00135 | 0.01852 | 0.01846 | 0.03704 | 0.03693
446.6837 | 0 | 0.00279 | 0.00121 | 0.00015 | 9e-05 | 0.0121 | 0.01185 | 0.00012 | 1e-05 | 2e-05 | 0.00074 | 0.0021 | 0.00297 | 0 | 8e-05 | 7e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01841 | 0.0021 | 0.00123 | 0.01853 | 0.01845 | 0.03707 | 0.03691
473.1512 | 0 | 0.00279 | 0.00122 | 0.00015 | 9e-05 | 0.0121 | 0.01185 | 0.00013 | 1e-05 | 2e-05 | 0.00072 | 0.0021 | 0.00297 | 0 | 8e-05 | 7e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01841 | 0.00227 | 0.00117 | 0.01855 | 0.01845 | 0.0371 | 0.0369
501.1873 | 0 | 0.00279 | 0.00123 | 0.00015 | 9e-05 | 0.0121 | 0.01184 | 0.00015 | 1e-05 | 2e-05 | 0.0007 | 0.0021 | 0.00297 | 0 | 8e-05 | 7e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01841 | 0.00212 | 0.00122 | 0.01853 | 0.01845 | 0.03706 | 0.0369
530.8843 | 0 | 0.00279 | 0.00125 | 0.00015 | 0.0001 | 0.0121 | 0.01184 | 0.00016 | 1e-05 | 2e-05 | 0.00068 | 0.0021 | 0.00297 | 0 | 8e-05 | 7e-05 | 0.00058 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01841 | 0.00202 | 0.00122 | 0.01852 | 0.01845 | 0.03704 | 0.0369
562.3413 | 0 | 0.00279 | 0.00126 | 0.00015 | 0.0001 | 0.0121 | 0.01183 | 0.00018 | 1e-05 | 3e-05 | 0.00066 | 0.0021 | 0.00297 | 0 | 8e-05 | 7e-05 | 0.00055 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0184 | 0.00198 | 0.00142 | 0.01851 | 0.01846 | 0.03702 | 0.03692
595.6623 | 0 | 0.0028 | 0.00128 | 0.00015 | 0.0001 | 0.0121 | 0.01183 | 0.0002 | 1e-05 | 3e-05 | 0.00064 | 0.0021 | 0.00297 | 0 | 8e-05 | 7e-05 | 0.00053 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0184 | 0.002 | 0.00128 | 0.01851 | 0.01844 | 0.03702 | 0.03689
630.9572 | 0 | 0.0028 | 0.00129 | 0.00015 | 0.0001 | 0.0121 | 0.01182 | 0.00022 | 1e-05 | 4e-05 | 0.00062 | 0.0021 | 0.00297 | 0 | 8e-05 | 7e-05 | 0.00049 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0184 | 0.00208 | 0.00122 | 0.01851 | 0.01844 | 0.03702 | 0.03687
668.344 | 0 | 0.0028 | 0.00131 | 0.00016 | 0.00011 | 0.0121 | 0.01181 | 0.00025 | 0 | 4e-05 | 0.0006 | 0.0021 | 0.00297 | 0 | 8e-05 | 8e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01839 | 0.00192 | 0.00156 | 0.01849 | 0.01846 | 0.03698 | 0.03691
707.9456 | 0 | 0.0028 | 0.00134 | 0.00016 | 0.00011 | 0.0121 | 0.0118 | 0.00028 | 0 | 4e-05 | 0.00058 | 0.0021 | 0.00296 | 0 | 8e-05 | 8e-05 | 0.00045 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01839 | 0.00205 | 0.00153 | 0.0185 | 0.01845 | 0.037 | 0.0369
749.8942 | 0 | 0.00281 | 0.00136 | 0.00016 | 0.00012 | 0.01211 | 0.01179 | 0.00031 | 0 | 5e-05 | 0.00057 | 0.0021 | 0.00296 | 0 | 8e-05 | 8e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01838 | 0.00221 | 0.00141 | 0.01851 | 0.01843 | 0.03703 | 0.03687
794.3285 | 0 | 0.00281 | 0.00139 | 0.00016 | 0.00012 | 0.01211 | 0.01177 | 0.00035 | 0 | 6e-05 | 0.00055 | 0.00211 | 0.00296 | 0 | 8e-05 | 8e-05 | 0.00045 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01838 | 0.00219 | 0.0015 | 0.01851 | 0.01844 | 0.03701 | 0.03687
841.395 | 0 | 0.00281 | 0.00142 | 0.00016 | 0.00013 | 0.01211 | 0.01176 | 0.00039 | 0 | 6e-05 | 0.00053 | 0.00211 | 0.00296 | 0 | 8e-05 | 8e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01837 | 0.00217 | 0.00171 | 0.0185 | 0.01845 | 0.03699 | 0.0369
891.251 | 0 | 0.00282 | 0.00145 | 0.00016 | 0.00015 | 0.01211 | 0.01174 | 0.00043 | 1e-05 | 7e-05 | 0.00052 | 0.00211 | 0.00296 | 0 | 8e-05 | 8e-05 | 0.00045 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01836 | 0.00215 | 0.00168 | 0.01849 | 0.01844 | 0.03698 | 0.03688
944.0607 | 0 | 0.00282 | 0.00149 | 0.00016 | 0.00016 | 0.01212 | 0.01171 | 0.00048 | 1e-05 | 8e-05 | 0.0005 | 0.00211 | 0.00295 | 0 | 8e-05 | 8e-05 | 0.00045 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01835 | 0.00221 | 0.00163 | 0.01849 | 0.01843 | 0.03697 | 0.03685
1000 | 0 | 0.00283 | 0.00152 | 0.00016 | 0.00017 | 0.01212 | 0.01169 | 0.00054 | 1e-05 | 9e-05 | 0.00049 | 0.00211 | 0.00295 | 0 | 8e-05 | 8e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01835 | 0.0024 | 0.00152 | 0.0185 | 0.01841 | 0.037 | 0.03682
1059.254 | 0 | 0.00284 | 0.00156 | 0.00016 | 0.00019 | 0.01212 | 0.01166 | 0.00061 | 2e-05 | 0.0001 | 0.00047 | 0.00211 | 0.00294 | 0 | 8e-05 | 8e-05 | 0.00045 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01834 | 0.00217 | 0.00213 | 0.01847 | 0.01846 | 0.03693 | 0.03692
1122.018 | 0 | 0.00284 | 0.00161 | 0.00016 | 0.00021 | 0.01213 | 0.01163 | 0.00068 | 2e-05 | 0.00011 | 0.00046 | 0.00211 | 0.00294 | 0 | 8e-05 | 9e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01833 | 0.00236 | 0.00174 | 0.01848 | 0.01841 | 0.03695 | 0.03682
1188.502 | 0 | 0.00285 | 0.00166 | 0.00017 | 0.00023 | 0.01214 | 0.01159 | 0.00076 | 3e-05 | 0.00012 | 0.00044 | 0.00211 | 0.00293 | 0 | 8e-05 | 9e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01831 | 0.00275 | 0.00214 | 0.01852 | 0.01844 | 0.03704 | 0.03688
1258.925 | 1e-05 | 0.00286 | 0.00171 | 0.00017 | 0.00025 | 0.01214 | 0.01155 | 0.00085 | 3e-05 | 0.00014 | 0.00043 | 0.00211 | 0.00293 | 0 | 8e-05 | 9e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0183 | 0.00248 | 0.00196 | 0.01847 | 0.01841 | 0.03694 | 0.03681
1333.521 | 1e-05 | 0.00287 | 0.00176 | 0.00017 | 0.00028 | 0.01215 | 0.01151 | 0.00095 | 4e-05 | 0.00016 | 0.00042 | 0.00212 | 0.00292 | 0 | 8e-05 | 9e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.01829 | 0.00235 | 0.00179 | 0.01844 | 0.01838 | 0.03688 | 0.03675
1412.538 | 0 | 0.00288 | 0.00174 | 0.00016 | 0.00026 | 0.014 | 0.01326 | 0.00121 | 7e-05 | 0.00021 | 0.00047 | 0.00153 | 0.00292 | 0 | 9e-05 | 3e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02061 | 0.00188 | 0.0015 | 0.0207 | 0.02067 | 0.0414 | 0.04134
1496.235 | 0 | 0.00289 | 0.00179 | 0.00017 | 0.00028 | 0.01401 | 0.01321 | 0.00136 | 8e-05 | 0.00024 | 0.00045 | 0.00153 | 0.00291 | 0 | 9e-05 | 3e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0206 | 0.00188 | 0.0014 | 0.02069 | 0.02065 | 0.04137 | 0.04129
1584.893 | 0 | 0.0029 | 0.00185 | 0.00017 | 0.00031 | 0.01403 | 0.01314 | 0.00152 | 9e-05 | 0.00026 | 0.00044 | 0.00153 | 0.0029 | 0 | 9e-05 | 3e-05 | 0.00042 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02059 | 0.00161 | 0.00113 | 0.02065 | 0.02062 | 0.0413 | 0.04123
1678.804 | 1e-05 | 0.00292 | 0.00191 | 0.00017 | 0.00035 | 0.01404 | 0.01307 | 0.0017 | 0.00011 | 0.0003 | 0.00042 | 0.00153 | 0.00289 | 0 | 9e-05 | 3e-05 | 0.00043 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02057 | 0.00188 | 0.00134 | 0.02066 | 0.02062 | 0.04132 | 0.04123
1778.279 | 1e-05 | 0.00294 | 0.00199 | 0.00017 | 0.00038 | 0.01406 | 0.01299 | 0.0019 | 0.00012 | 0.00033 | 0.00041 | 0.00153 | 0.00287 | 0 | 9e-05 | 3e-05 | 0.00039 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02056 | 0.00201 | 0.00149 | 0.02066 | 0.02061 | 0.04131 | 0.04122
1883.65 | 1e-05 | 0.00295 | 0.00206 | 0.00017 | 0.00042 | 0.01408 | 0.0129 | 0.00212 | 0.00014 | 0.00037 | 0.0004 | 0.00154 | 0.00286 | 0 | 9e-05 | 3e-05 | 0.00039 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02055 | 0.00261 | 0.00254 | 0.02071 | 0.0207 | 0.04142 | 0.04141
1995.262 | 1e-05 | 0.00298 | 0.00215 | 0.00017 | 0.00047 | 0.01411 | 0.0128 | 0.00237 | 0.00016 | 0.00042 | 0.00038 | 0.00154 | 0.00285 | 0 | 9e-05 | 4e-05 | 0.00041 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02054 | 0.00224 | 0.00176 | 0.02066 | 0.02061 | 0.04132 | 0.04123
2113.489 | 1e-05 | 0.003 | 0.00225 | 0.00018 | 0.00052 | 0.01413 | 0.01268 | 0.00265 | 0.00019 | 0.00047 | 0.00037 | 0.00154 | 0.00283 | 0 | 9e-05 | 4e-05 | 0.00042 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02053 | 0.00208 | 0.00134 | 0.02064 | 0.02058 | 0.04128 | 0.04115
2238.721 | 1e-05 | 0.00303 | 0.00237 | 0.00018 | 0.00058 | 0.01416 | 0.01255 | 0.00295 | 0.00022 | 0.00052 | 0.00036 | 0.00154 | 0.00281 | 0 | 9e-05 | 4e-05 | 0.00037 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02053 | 0.00206 | 0.00112 | 0.02064 | 0.02056 | 0.04127 | 0.04113
2371.374 | 1e-05 | 0.00306 | 0.0025 | 0.00018 | 0.00065 | 0.01419 | 0.01241 | 0.0033 | 0.00025 | 0.00058 | 0.00035 | 0.00155 | 0.00279 | 0 | 9e-05 | 4e-05 | 0.00039 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02054 | 0.00247 | 0.00148 | 0.02069 | 0.0206 | 0.04138 | 0.04119
2511.887 | 1e-05 | 0.00309 | 0.00265 | 0.00019 | 0.00072 | 0.01423 | 0.01225 | 0.00368 | 0.00028 | 0.00065 | 0.00033 | 0.00155 | 0.00276 | 0 | 9e-05 | 4e-05 | 0.00042 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02056 | 0.00276 | 0.00173 | 0.02075 | 0.02064 | 0.0415 | 0.04127
2660.725 | 1e-05 | 0.00313 | 0.00283 | 0.00019 | 0.0008 | 0.01427 | 0.01207 | 0.0041 | 0.00032 | 0.00073 | 0.00032 | 0.00156 | 0.00274 | 0 | 9e-05 | 4e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0206 | 0.00312 | 0.0022 | 0.02083 | 0.02072 | 0.04167 | 0.04143
2818.383 | 1e-05 | 0.00318 | 0.00304 | 0.00019 | 0.0009 | 0.01432 | 0.01187 | 0.00457 | 0.00037 | 0.00082 | 0.00031 | 0.00156 | 0.0027 | 0 | 9e-05 | 4e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02065 | 0.00434 | 0.00346 | 0.0211 | 0.02094 | 0.04221 | 0.04188
2985.382 | 2e-05 | 0.00323 | 0.00328 | 0.0002 | 0.001 | 0.01437 | 0.01165 | 0.00509 | 0.00042 | 0.00091 | 0.0003 | 0.00157 | 0.00267 | 0 | 9e-05 | 4e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02073 | 0.0044 | 0.00297 | 0.02119 | 0.02094 | 0.04239 | 0.04189
3162.278 | 2e-05 | 0.00328 | 0.00356 | 0.0002 | 0.00112 | 0.01443 | 0.0114 | 0.00566 | 0.00047 | 0.00102 | 0.00029 | 0.00157 | 0.00263 | 0 | 9e-05 | 4e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02084 | 0.00512 | 0.00334 | 0.02146 | 0.02111 | 0.04293 | 0.04222
3349.655 | 2e-05 | 0.00335 | 0.00388 | 0.00021 | 0.00125 | 0.01449 | 0.01112 | 0.00628 | 0.00054 | 0.00114 | 0.00028 | 0.00158 | 0.00259 | 0 | 9e-05 | 4e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02099 | 0.00582 | 0.00417 | 0.02179 | 0.0214 | 0.04358 | 0.04281
3548.134 | 2e-05 | 0.00342 | 0.00425 | 0.00022 | 0.0014 | 0.01456 | 0.01082 | 0.00696 | 0.00061 | 0.00127 | 0.00026 | 0.00159 | 0.00254 | 0 | 9e-05 | 5e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0212 | 0.00585 | 0.00434 | 0.02199 | 0.02164 | 0.04398 | 0.04327
3758.374 | 2e-05 | 0.00351 | 0.00467 | 0.00022 | 0.00157 | 0.01464 | 0.01048 | 0.00771 | 0.00069 | 0.00142 | 0.00025 | 0.00159 | 0.00248 | 0 | 0.0001 | 5e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02145 | 0.00681 | 0.00484 | 0.02251 | 0.02199 | 0.04502 | 0.04399
3981.071 | 3e-05 | 0.0036 | 0.00513 | 0.00023 | 0.00176 | 0.01473 | 0.01011 | 0.00851 | 0.00079 | 0.00158 | 0.00024 | 0.0016 | 0.00241 | 0 | 0.0001 | 5e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02178 | 0.00779 | 0.00544 | 0.02313 | 0.02245 | 0.04626 | 0.0449
4216.965 | 3e-05 | 0.0037 | 0.00565 | 0.00024 | 0.00197 | 0.01482 | 0.00969 | 0.00938 | 0.00089 | 0.00175 | 0.00023 | 0.00161 | 0.00235 | 0 | 0.0001 | 5e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02218 | 0.00853 | 0.00616 | 0.02377 | 0.02302 | 0.04754 | 0.04605
4466.837 | 3e-05 | 0.0038 | 0.0062 | 0.00025 | 0.00221 | 0.01493 | 0.00924 | 0.0103 | 0.00101 | 0.00195 | 0.00022 | 0.00162 | 0.00229 | 0 | 0.0001 | 5e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02267 | 0.00903 | 0.00813 | 0.0244 | 0.02409 | 0.04881 | 0.04817
4731.512 | 4e-05 | 0.00392 | 0.0068 | 0.00026 | 0.00247 | 0.01505 | 0.00875 | 0.01127 | 0.00115 | 0.00216 | 0.00021 | 0.00164 | 0.00223 | 0 | 0.0001 | 5e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02325 | 0.01088 | 0.00702 | 0.02567 | 0.02428 | 0.05133 | 0.04857
5011.873 | 4e-05 | 0.00403 | 0.00742 | 0.00028 | 0.00277 | 0.01517 | 0.00821 | 0.01228 | 0.0013 | 0.00239 | 0.00019 | 0.00165 | 0.00216 | 0 | 0.0001 | 6e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02391 | 0.01142 | 0.00756 | 0.02649 | 0.02507 | 0.05298 | 0.05015
5308.843 | 5e-05 | 0.00414 | 0.00806 | 0.00029 | 0.0031 | 0.01532 | 0.00763 | 0.01332 | 0.00148 | 0.00264 | 0.00018 | 0.00167 | 0.00208 | 0 | 0.0001 | 6e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02464 | 0.01242 | 0.0088 | 0.02759 | 0.02617 | 0.05519 | 0.05234
5623.413 | 5e-05 | 0.00424 | 0.00868 | 0.00031 | 0.00347 | 0.01547 | 0.00701 | 0.01436 | 0.00168 | 0.0029 | 0.00017 | 0.00169 | 0.002 | 0 | 0.0001 | 6e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02544 | 0.01023 | 0.00786 | 0.02742 | 0.02663 | 0.05483 | 0.05325
5956.623 | 6e-05 | 0.00432 | 0.00927 | 0.00033 | 0.00389 | 0.01564 | 0.00634 | 0.01535 | 0.00192 | 0.00319 | 0.00015 | 0.00171 | 0.0019 | 0 | 0.0001 | 6e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02626 | 0.00927 | 0.00751 | 0.02785 | 0.02731 | 0.0557 | 0.05463
6309.573 | 6e-05 | 0.00438 | 0.00979 | 0.00035 | 0.00436 | 0.01581 | 0.00564 | 0.01625 | 0.00219 | 0.00348 | 0.00014 | 0.00173 | 0.0018 | 0 | 0.0001 | 6e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02706 | 0.00858 | 0.00811 | 0.02838 | 0.02825 | 0.05677 | 0.05649
6683.44 | 7e-05 | 0.00438 | 0.0102 | 0.00037 | 0.00488 | 0.016 | 0.00493 | 0.01696 | 0.00252 | 0.00379 | 0.00013 | 0.00175 | 0.00169 | 0 | 0.0001 | 7e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02776 | 0.00667 | 0.00695 | 0.02855 | 0.02862 | 0.0571 | 0.05724
7079.456 | 7e-05 | 0.00433 | 0.01047 | 0.00039 | 0.00545 | 0.0162 | 0.0042 | 0.0174 | 0.0029 | 0.0041 | 0.00011 | 0.00177 | 0.00157 | 0 | 0.00011 | 7e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.0283 | 0.00429 | 0.00558 | 0.02862 | 0.02884 | 0.05725 | 0.05769
7498.942 | 8e-05 | 0.00421 | 0.01053 | 0.00042 | 0.00609 | 0.0164 | 0.00349 | 0.01746 | 0.00336 | 0.00439 | 0.0001 | 0.00179 | 0.00143 | 0 | 0.00011 | 7e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02858 | 0.00247 | 0.00465 | 0.02869 | 0.02895 | 0.05737 | 0.05791
7943.284 | 9e-05 | 0.004 | 0.01033 | 0.00045 | 0.0068 | 0.01661 | 0.00284 | 0.0169 | 0.00392 | 0.00466 | 8e-05 | 0.00182 | 0.00128 | 0 | 0.00011 | 6e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02846 | 0.00199 | 0.00355 | 0.02853 | 0.02868 | 0.05705 | 0.05735
8413.95 | 0.0001 | 0.00369 | 0.0098171 | 0.0004857 | 0.0075756 | 0.016803 | 0.002298 | 0.015469 | 0.0046001 | 0.0048722 | 6.08e-05 | 0.0018348 | 0.0011113 | 0 | 0.0001094 | 7.47e-05 | 0.0004392 | 0.0010857 | 0.0052362 | 6.5144e-05 | 0.00028868 | 0.00028868 | 0.027804 | 0.0026901 | 0.001766 | 0.027934 | 0.02786 | 0.055867 | 0.05572
8912.511 | 0.00011 | 0.0032801 | 0.0089226 | 0.0005225 | 0.0084187 | 0.016983 | 0.001929 | 0.012796 | 0.0054544 | 0.0049979 | 4.12e-05 | 0.0018558 | 0.0009278 | 0 | 0.0001107 | 7.67e-05 | 0.0004392 | 0.0010857 | 0.0052362 | 6.5144e-05 | 0.00028868 | 0.00028868 | 0.02656 | 0.0077702 | 0.0042175 | 0.027673 | 0.026893 | 0.055346 | 0.053785
9440.6
10000 | 0.00013 | 0.00221 | 0.00596 | 0.0006 | 0.01028 | 0.01723 | 0.00385 | 0.01179 | 0.00791 | 0.00476 | 8e-05 | 0.00189 | 0.0011 | 0 | 0.00011 | 7e-05 | 0.00044 | 0.00109 | 0.00524 | 7e-05 | 0.00029 | 0.00029 | 0.02676 | 0.04718 | 0.04325 | 0.05424 | 0.05086 | 0.10847 | 0.10171

## SP Level
SP | Level
| Frequency | 20 | 25 | 31.5 | 40 | 50 | 63 | 80 | 100 | 125 | 160 | 200 | 250 | 315 | 400 | 500 | 630 | 800 | 1k | 1,25k | 1,6k | 2k | 2,5k | 3,15k | 4k | 5k | 6,3k | 8k | 10k | 20 | 25 | 31.5 | 40 | 50 | 63 | 80 | 100 | 125 | 160 | 200 | 250 | 315 | 400 | 500 | 630 | 800 | 1k | 1,25k | 1,6k | 2k | 2,5k | 3,15k | 4k | 5k | 6,3k | 8k | 10k
| Type B uncertainties | Uncertaity components given as 1 standard deviation
Reference capacitor | Calibration | 0.003 | 0.002 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 9e-06 | 4e-06 | 2.56e-06 | 2.56e-06 | 2.56e-06 | 2.56e-06 | 2.56e-06 | 2.56e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06 | 1.96e-06
| Calibration method | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 0.0079 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05 | 6.241e-05
| Drift | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06 | 6.76e-06
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Voltage ratio measurement | DAQ unit | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06
| Microphone Calibration Apparatus and preamp | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06
| Noise | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06
| Crosstalk | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Geometry | Coupler diameter | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0005 | 0.0006 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 2.5e-07 | 3.6e-07
| Coupler length | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07
| Cavity depth | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 1e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08
| Total front cavity volume | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.013 | 0.012 | 0.012 | 0.012 | 0.011 | 0.011 | 0.01 | 0.008 | 0.005 | 0.005 | 0.015 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000169 | 0.000144 | 0.000144 | 0.000144 | 0.000121 | 0.000121 | 0.0001 | 6.4e-05 | 2.5e-05 | 2.5e-05 | 0.000225
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Equivalent volume | LF Equivalent volume | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0005 | 0.001 | 0.0015 | 0.002 | 0.004 | 0.005 | 0.008 | 0.013 | 0.019 | 0.016 | 0.012 | 0.053 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2.5e-07 | 1e-06 | 2.25e-06 | 4e-06 | 1.6e-05 | 2.5e-05 | 6.4e-05 | 0.000169 | 0.000361 | 0.000256 | 0.000144 | 0.002809
| Loss factor | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0003 | 0.0005 | 0.0007 | 0.0009 | 0.0011 | 0.0013 | 0.0015 | 0.002 | 0.003 | 0.004 | 0.004 | 0.003 | 0.002 | 0.006 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 9e-08 | 2.5e-07 | 4.9e-07 | 8.1e-07 | 1.21e-06 | 1.69e-06 | 2.25e-06 | 4e-06 | 9e-06 | 1.6e-05 | 1.6e-05 | 9e-06 | 4e-06 | 3.6e-05
| Resonance frequency | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0002 | 0.0003 | 0.0005 | 0.0007 | 0.001 | 0.0015 | 0.002 | 0.002 | 0.002 | 0.003 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1e-08 | 4e-08 | 9e-08 | 2.5e-07 | 4.9e-07 | 1e-06 | 2.25e-06 | 4e-06 | 4e-06 | 4e-06 | 9e-06
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Heat conduction and viscous losses | Theory imperfection | 0.00618 | 0.00549 | 0.00489 | 0.00438 | 0.0039 | 0.00348 | 0.00309 | 0.00276 | 0.00246 | 0.00222 | 0.00195 | 0.00174 | 0.00156 | 0.00138 | 0.00123 | 0.00108 | 0.00099 | 0.00084 | 0.00078 | 0.00066 | 0.0006 | 0.00048 | 0.00039 | 0.0003 | 0.00021 | 9e-05 | 9e-05 | 0.00036 | 3.8192e-05 | 3.014e-05 | 2.3912e-05 | 1.9184e-05 | 1.521e-05 | 1.211e-05 | 9.5481e-06 | 7.6176e-06 | 6.0516e-06 | 4.9284e-06 | 3.8025e-06 | 3.0276e-06 | 2.4336e-06 | 1.9044e-06 | 1.5129e-06 | 1.1664e-06 | 9.801e-07 | 7.056e-07 | 6.084e-07 | 4.356e-07 | 3.6e-07 | 2.304e-07 | 1.521e-07 | 9e-08 | 4.41e-08 | 8.1e-09 | 8.1e-09 | 1.296e-07
| Cavity surface area | 0.0029 | 0.0027 | 0.0024 | 0.0021 | 0.0019 | 0.0017 | 0.0015 | 0.0013 | 0.0012 | 0.001 | 0.001 | 0.0008 | 0.0007 | 0.0007 | 0.0006 | 0.0006 | 0.0004 | 0.0005 | 0.0004 | 0.0003 | 0.0003 | 0.0002 | 0.0002 | 0.0002 | 0.0001 | 0.0001 | 0 | 0 | 8.41e-06 | 7.29e-06 | 5.76e-06 | 4.41e-06 | 3.61e-06 | 2.89e-06 | 2.25e-06 | 1.69e-06 | 1.44e-06 | 1e-06 | 1e-06 | 6.4e-07 | 4.9e-07 | 4.9e-07 | 3.6e-07 | 3.6e-07 | 1.6e-07 | 2.5e-07 | 1.6e-07 | 9e-08 | 9e-08 | 4e-08 | 4e-08 | 4e-08 | 1e-08 | 1e-08 | 0 | 0
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Speed of sound | Theory | 0.0003 | 0.0002 | 0.0002 | 0.0002 | 0.0003 | 0.0002 | 0.0002 | 0.0003 | 0.0002 | 0.0003 | 0.0002 | 0.0002 | 0.0003 | 0.0002 | 0.0002 | 0.0002 | 0.0003 | 0.0002 | 0.0002 | 0.0002 | 0.0003 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0001 | 0 | 9e-08 | 4e-08 | 4e-08 | 4e-08 | 9e-08 | 4e-08 | 4e-08 | 9e-08 | 4e-08 | 9e-08 | 4e-08 | 4e-08 | 9e-08 | 4e-08 | 4e-08 | 4e-08 | 9e-08 | 4e-08 | 4e-08 | 4e-08 | 9e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 1e-08 | 0
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Radial wave motion |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0005 | 0.001 | 0.002 | 0.005 | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2.5e-07 | 1e-06 | 4e-06 | 2.5e-05 | 0.0001 | 0.0004 | 0.0009 | 0.0016 | 0.0025
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Climate determination | Air humidity | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0008 | 0.0008 | 0.001 | 0.001 | 0.001 | 0.0015 | 0.002 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 6.4e-07 | 6.4e-07 | 1e-06 | 1e-06 | 1e-06 | 2.25e-06 | 4e-06
| Temperature | 0.0005 | 0.0005 | 0.0005 | 0.0004 | 0.0003 | 0.0004 | 0.0004 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0002 | 0.0003 | 0.0003 | 0.0003 | 0.0002 | 0.0003 | 0.0003 | 0.0004 | 0.0005 | 0.0007 | 0.001 | 0.0015 | 0.0023 | 0.0036 | 0.0056 | 0.0096 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 1.6e-07 | 9e-08 | 1.6e-07 | 1.6e-07 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 4e-08 | 9e-08 | 9e-08 | 9e-08 | 4e-08 | 9e-08 | 9e-08 | 1.6e-07 | 2.5e-07 | 4.9e-07 | 1e-06 | 2.25e-06 | 5.29e-06 | 1.296e-05 | 3.136e-05 | 9.216e-05
| Static pressure | 0.0063 | 0.0062 | 0.0062 | 0.0062 | 0.0062 | 0.0061 | 0.0061 | 0.0061 | 0.0061 | 0.0061 | 0.006 | 0.006 | 0.0061 | 0.006 | 0.006 | 0.006 | 0.006 | 0.0059 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.0061 | 0.0062 | 0.0063 | 0.0064 | 0.0062 | 3.969e-05 | 3.844e-05 | 3.844e-05 | 3.844e-05 | 3.844e-05 | 3.721e-05 | 3.721e-05 | 3.721e-05 | 3.721e-05 | 3.721e-05 | 3.6e-05 | 3.6e-05 | 3.721e-05 | 3.6e-05 | 3.6e-05 | 3.6e-05 | 3.6e-05 | 3.481e-05 | 3.6e-05 | 3.6e-05 | 3.6e-05 | 3.6e-05 | 3.6e-05 | 3.721e-05 | 3.844e-05 | 3.969e-05 | 4.096e-05 | 3.844e-05
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ratio of specific heats K | Theory | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08 | 4e-08
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Polarising voltage | Voltmeter | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06 | 1.44e-06
| Voltage drop | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06 | 4.84e-06
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Environmental normalising | Correction coefficients | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0007 | 0.0008 | 0.0007 | 0.0007 | 0.0006 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.002 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 6.4e-07 | 4.9e-07 | 6.4e-07 | 4.9e-07 | 4.9e-07 | 3.6e-07 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 4e-06
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Rounding error |  | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Others |  | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05 | 2.5e-05
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Type A uncertainty |  | 0.0049212 | 0.0032349 | 0.0024745 | 0.0034278 | 0.0028569 | 0.0025518 | 0.0028299 | 0.003 | 0.0018817 | 0.0017245 | 0.002 | 0.0019734 | 0.0015177 | 0.0018609 | 0.002472 | 0.0024668 | 0.0019676 | 0.0021992 | 0.0019543 | 0.002577 | 0.0020937 | 0.0031843 | 0.0042246 | 0.0045833 | 0.0068323 | 0.0087651 | 0.01057 | 0.02473 | 2.4218e-05 | 1.0465e-05 | 6.1233e-06 | 1.175e-05 | 8.1621e-06 | 6.5117e-06 | 8.0083e-06 | 9e-06 | 3.5408e-06 | 2.9739e-06 | 4e-06 | 3.8944e-06 | 2.3035e-06 | 3.4631e-06 | 6.1106e-06 | 6.0853e-06 | 3.8714e-06 | 4.8365e-06 | 3.8194e-06 | 6.641e-06 | 4.3836e-06 | 1.014e-05 | 1.7848e-05 | 2.1007e-05 | 4.6681e-05 | 7.6828e-05 | 0.00011173 | 0.00061157
Expanded uncertainty (k=2) |  | 0.039786 | 0.038289 | 0.037575 | 0.037546 | 0.037097 | 0.036735 | 0.036642 | 0.036559 | 0.036123 | 0.036008 | 0.035932 | 0.035863 | 0.0358 | 0.035777 | 0.035901 | 0.035889 | 0.035771 | 0.0358 | 0.034481 | 0.034771 | 0.035414 | 0.035245 | 0.039196 | 0.047213 | 0.06465 | 0.075708 | 0.091036 | 0.16047
| Frequency | 20 | 25 | 31.5 | 40 | 50 | 63 | 80 | 100 | 125 | 160 | 200 | 250 | 315 | 400 | 500 | 630 | 800 | 1k | 1,25k | 1,6k | 2k | 2,5k | 3,15k | 4k | 5k | 6,3k | 8k | 10k | 20 | 25 | 31.5 | 40 | 50 | 63 | 80 | 100 | 125 | 160 | 200 | 250 | 315 | 400 | 500 | 630 | 800 | 1k | 1,25k | 1,6k | 2k | 2,5k | 3,15k | 4k | 5k | 6,3k | 8k | 10k

## SP Phase
SP | Phase
| Frequency | 20 | 25 | 31.5 | 40 | 50 | 63 | 80 | 100 | 125 | 160 | 200 | 250 | 315 | 400 | 500 | 630 | 800 | 1k | 1,25k | 1,6k | 2k | 2,5k | 3,15k | 4k | 5k | 6,3k | 8k | 10k | 20 | 25 | 31.5 | 40 | 50 | 63 | 80 | 100 | 125 | 160 | 200 | 250 | 315 | 400 | 500 | 630 | 800 | 1k | 1,25k | 1,6k | 2k | 2,5k | 3,15k | 4k | 5k | 6,3k | 8k | 10k
| Type B uncertainties | Uncertaity components given as 1 standard deviation
Microphone Calibration Apparatus and preamp | Serial resistance/stray capacitance | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225 | 0.0225
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Voltage ratio measurement | DAQ unit | 0.0046188 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0028868 | 0.0028868 | 0.0034641 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0023094 | 0.0023094 | 0.0017321 | 0.00057735 | 0.00057735 | 0.0023094 | 0.0057735 | 0.0086603 | 0.0086603 | 0.0086603 | 2.1333e-05 | 1.2e-05 | 1.2e-05 | 1.2e-05 | 1.2e-05 | 1.2e-05 | 1.2e-05 | 1.2e-05 | 1.2e-05 | 1.2e-05 | 1.2e-05 | 1.2e-05 | 8.3333e-06 | 8.3333e-06 | 1.2e-05 | 8.3333e-06 | 8.3333e-06 | 8.3333e-06 | 5.3333e-06 | 5.3333e-06 | 3e-06 | 3.3333e-07 | 3.3333e-07 | 5.3333e-06 | 3.3333e-05 | 7.5e-05 | 7.5e-05 | 7.5e-05
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Geometry | Coupler dimensions | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 0.0007 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07 | 4.9e-07
| Cavity depth | 0.0007 | 0.0007 | 0.0006 | 0.0005 | 0.0005 | 0.0004 | 0.0004 | 0.0004 | 0.0003 | 0.0003 | 0.0002 | 0.0002 | 0.0002 | 0.0001 | 0.0002 | 0.0001 | 0.0001 | 0.0001 | 0 | 0.0001 | 0 | 0.0001 | 0.0003 | 0.0005 | 0.0013 | 0.0025 | 0.004 | 0.0045 | 4.9e-07 | 4.9e-07 | 3.6e-07 | 2.5e-07 | 2.5e-07 | 1.6e-07 | 1.6e-07 | 1.6e-07 | 9e-08 | 9e-08 | 4e-08 | 4e-08 | 4e-08 | 1e-08 | 4e-08 | 1e-08 | 1e-08 | 1e-08 | 0 | 1e-08 | 0 | 1e-08 | 9e-08 | 2.5e-07 | 1.69e-06 | 6.25e-06 | 1.6e-05 | 2.025e-05
| Total front cavity volume | 0.004 | 0.003 | 0.003 | 0.003 | 0.003 | 0.002 | 0.002 | 0.002 | 0.001 | 0.001 | 0.002 | 0.001 | 0.002 | 0.002 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | 0.001 | 0.002 | 0.003 | 0.003 | 0.004 | 0.006 | 0.008 | 0.011 | 0.019 | 1.6e-05 | 9e-06 | 9e-06 | 9e-06 | 9e-06 | 4e-06 | 4e-06 | 4e-06 | 1e-06 | 1e-06 | 4e-06 | 1e-06 | 4e-06 | 4e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 4e-06 | 1e-06 | 4e-06 | 9e-06 | 9e-06 | 1.6e-05 | 3.6e-05 | 6.4e-05 | 0.000121 | 0.000361
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Equivalent volume | LF Equivalent volume | 0.001 | 0.002 | 0.002 | 0.002 | 0.003 | 0.004 | 0.005 | 0.006 | 0.008 | 0.01 | 0.012 | 0.015 | 0.019 | 0.024 | 0.03 | 0.038 | 0.047 | 0.06 | 0.075 | 0.094 | 0.117 | 0.143 | 0.17 | 0.189 | 0.173 | 0.099 | 0.1 | 0.149 | 1e-06 | 4e-06 | 4e-06 | 4e-06 | 9e-06 | 1.6e-05 | 2.5e-05 | 3.6e-05 | 6.4e-05 | 0.0001 | 0.000144 | 0.000225 | 0.000361 | 0.000576 | 0.0009 | 0.001444 | 0.002209 | 0.0036 | 0.005625 | 0.008836 | 0.013689 | 0.020449 | 0.0289 | 0.035721 | 0.029929 | 0.009801 | 0.01 | 0.022201
| Loss factor | 0 | 0 | 0 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | 0.002 | 0.003 | 0.003 | 0.003 | 0.005 | 0.006 | 0.007 | 0.01 | 0.011 | 0.013 | 0.016 | 0.016 | 0.012 | 0.003 | 0.006 | 0.011 | 0.033 | 0 | 0 | 0 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 4e-06 | 4e-06 | 9e-06 | 9e-06 | 9e-06 | 2.5e-05 | 3.6e-05 | 4.9e-05 | 0.0001 | 0.000121 | 0.000169 | 0.000256 | 0.000256 | 0.000144 | 9e-06 | 3.6e-05 | 0.000121 | 0.001089
| Resonance frequency | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.001 | 0.001 | 0.001 | 0.002 | 0.001 | 0.002 | 0.003 | 0.003 | 0.005 | 0.005 | 0.007 | 0.01 | 0.012 | 0.014 | 0.015 | 0.008 | 0.009 | 0.051 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1e-06 | 1e-06 | 1e-06 | 4e-06 | 1e-06 | 4e-06 | 9e-06 | 9e-06 | 2.5e-05 | 2.5e-05 | 4.9e-05 | 0.0001 | 0.000144 | 0.000196 | 0.000225 | 6.4e-05 | 8.1e-05 | 0.002601
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Heat conduction and viscous losses | Theory imperfection | 0.03837 | 0.03438 | 0.03084 | 0.02763 | 0.02475 | 0.02214 | 0.0198 | 0.0177 | 0.01581 | 0.01413 | 0.01263 | 0.01125 | 0.01005 | 0.00894 | 0.00795 | 0.00705 | 0.00627 | 0.00552 | 0.00486 | 0.0042 | 0.0036 | 0.00303 | 0.0024 | 0.00168 | 0.0009 | 0.00015 | 0.00156 | 0.00393 | 0.0014723 | 0.001182 | 0.00095111 | 0.00076342 | 0.00061256 | 0.00049018 | 0.00039204 | 0.00031329 | 0.00024996 | 0.00019966 | 0.00015952 | 0.00012656 | 0.000101 | 7.9924e-05 | 6.3203e-05 | 4.9702e-05 | 3.9313e-05 | 3.047e-05 | 2.362e-05 | 1.764e-05 | 1.296e-05 | 9.1809e-06 | 5.76e-06 | 2.8224e-06 | 8.1e-07 | 2.25e-08 | 2.4336e-06 | 1.5445e-05
| Cavity surface area | 0.0179 | 0.0162 | 0.0146 | 0.013 | 0.0117 | 0.0105 | 0.0094 | 0.0085 | 0.0076 | 0.0068 | 0.006 | 0.0054 | 0.0048 | 0.0043 | 0.0039 | 0.0034 | 0.003 | 0.0027 | 0.0023 | 0.0021 | 0.0018 | 0.0015 | 0.0012 | 0.001 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.00032041 | 0.00026244 | 0.00021316 | 0.000169 | 0.00013689 | 0.00011025 | 8.836e-05 | 7.225e-05 | 5.776e-05 | 4.624e-05 | 3.6e-05 | 2.916e-05 | 2.304e-05 | 1.849e-05 | 1.521e-05 | 1.156e-05 | 9e-06 | 7.29e-06 | 5.29e-06 | 4.41e-06 | 3.24e-06 | 2.25e-06 | 1.44e-06 | 1e-06 | 3.6e-07 | 3.6e-07 | 3.6e-07 | 3.6e-07
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Radial wave motion |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.01 | 0.01 | 0.02 | 0.03 | 0.05 | 0.01 | 0.15 | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0004 | 0.0009 | 0.0025 | 0.0001 | 0.0225 | 0.04
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Climate determination | Air humidity | 0.0005 | 0 | 0 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0 | 0 | 0 | 0.0005 | 0 | 0.0005 | 0.0005 | 0 | 0 | 0 | 0 | 0.0005 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0005 | 0.0015 | 0.0025 | 2.5e-07 | 0 | 0 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 0 | 0 | 0 | 2.5e-07 | 0 | 2.5e-07 | 2.5e-07 | 0 | 0 | 0 | 0 | 2.5e-07 | 0 | 0 | 0 | 0 | 0 | 0 | 2.5e-07 | 2.25e-06 | 6.25e-06
| Temperature | 0.002 | 0.0019 | 0.0017 | 0.0015 | 0.0013 | 0.0012 | 0.0011 | 0.001 | 0.0009 | 0.0008 | 0.0007 | 0.0006 | 0.0005 | 0.0005 | 0.0005 | 0.0004 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0005 | 0.0009 | 0.0019 | 0.004 | 0.0075 | 0.012 | 4e-06 | 3.61e-06 | 2.89e-06 | 2.25e-06 | 1.69e-06 | 1.44e-06 | 1.21e-06 | 1e-06 | 8.1e-07 | 6.4e-07 | 4.9e-07 | 3.6e-07 | 2.5e-07 | 2.5e-07 | 2.5e-07 | 1.6e-07 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 9e-08 | 2.5e-07 | 8.1e-07 | 3.61e-06 | 1.6e-05 | 5.625e-05 | 0.000144
| Static pressure | 0.002 | 0.0017 | 0.0015 | 0.0014 | 0.0013 | 0.0011 | 0.001 | 0.0009 | 0.0008 | 0.0007 | 0.0006 | 0.0005 | 0.0004 | 0.0003 | 0.0002 | 0.0001 | 0 | 0.0001 | 0.0002 | 0.0004 | 0.0006 | 0.0008 | 0.001 | 0.0013 | 0.0013 | 0.0009 | 0.0005 | 0.0024 | 4e-06 | 2.89e-06 | 2.25e-06 | 1.96e-06 | 1.69e-06 | 1.21e-06 | 1e-06 | 8.1e-07 | 6.4e-07 | 4.9e-07 | 3.6e-07 | 2.5e-07 | 1.6e-07 | 9e-08 | 4e-08 | 1e-08 | 0 | 1e-08 | 4e-08 | 1.6e-07 | 3.6e-07 | 6.4e-07 | 1e-06 | 1.69e-06 | 1.69e-06 | 8.1e-07 | 2.5e-07 | 5.76e-06
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ratio of specific heats K | Theory | 0.006 | 0.006 | 0.005 | 0.005 | 0.004 | 0.004 | 0.003 | 0.003 | 0.002 | 0.002 | 0.002 | 0.002 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 3.6e-05 | 3.6e-05 | 2.5e-05 | 2.5e-05 | 1.6e-05 | 1.6e-05 | 9e-06 | 9e-06 | 4e-06 | 4e-06 | 4e-06 | 4e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06 | 1e-06
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Environmental normalising | Correction coefficients | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.04 | 0.06 | 0.08 | 0.1 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0004 | 0.0016 | 0.0036 | 0.0064 | 0.01
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Rounding error |  | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841 | 0.000841
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Others |  | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01
[30 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Type A uncertainty |  | 0.11129 | 0.082082 | 0.063058 | 0.048092 | 0.038763 | 0.027765 | 0.027787 | 0.02 | 0.017802 | 0.016038 | 0.012 | 0.0086108 | 0.009 | 0.0092192 | 0.0020394 | 0.0025547 | 0.0067243 | 0.0038321 | 0.0053275 | 0.0097395 | 0.011694 | 0.01062 | 0.011088 | 0.021164 | 0.014048 | 0.01642 | 0.042492 | 0.10435 | 0.012386 | 0.0067374 | 0.0039763 | 0.0023128 | 0.0015025 | 0.00077089 | 0.00077212 | 0.0004 | 0.00031691 | 0.00025723 | 0.000144 | 7.4146e-05 | 8.1e-05 | 8.4995e-05 | 4.1593e-06 | 6.5267e-06 | 4.5216e-05 | 1.4685e-05 | 2.8383e-05 | 9.4859e-05 | 0.00013675 | 0.00011278 | 0.00012295 | 0.00044792 | 0.00019734 | 0.00026963 | 0.0018056 | 0.010888
Expanded uncertainty (k=2) |  | 0.43682 | 0.40837 | 0.39313 | 0.38337 | 0.37812 | 0.37344 | 0.37281 | 0.37036 | 0.36959 | 0.36913 | 0.36852 | 0.36834 | 0.36895 | 0.37002 | 0.37121 | 0.37413 | 0.37841 | 0.38555 | 0.39628 | 0.41254 | 0.43639 | 0.46683 | 0.50312 | 0.53359 | 0.52108 | 0.43532 | 0.54598 | 0.69498
| Frequency | 20 | 25 | 31.5 | 40 | 50 | 63 | 80 | 100 | 125 | 160 | 200 | 250 | 315 | 400 | 500 | 630 | 800 | 1k | 1,25k | 1,6k | 2k | 2,5k | 3,15k | 4k | 5k | 6,3k | 8k | 10k | 20 | 25 | 31.5 | 40 | 50 | 63 | 80 | 100 | 125 | 160 | 200 | 250 | 315 | 400 | 500 | 630 | 800 | 1k | 1,25k | 1,6k | 2k | 2,5k | 3,15k | 4k | 5k | 6,3k | 8k | 10k

## PTB Level
PTB | Level
LS1P microphones in 3 plane-wave couplers | [10 empty] | Frequency in Hz
|  |  |  | 2 | 4 | 8 | 16 | 31.5 | 63 | 125 | 250 | 500 | 1k | 2k | 4k | 8k | 10k
Type A uncertainty, as standard deviation (10-4 dB)
Source of uncertainty |  |  | Allowed repeatibility
Normal distribution
|  |  |  | 1000 | 500 | 300 | 150 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 150 | 180
Estimate of a type A uncertainty (S.D.), k=1 |  |  |  | 1000 | 500 | 300 | 150 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 150 | 180
Type B uncertainty, as semi-ranges (10-4 dB)
Source of uncertainty
Rectangular distribution
Measurement |  |  | Measurement Capacitor | 100 | 50 | 100 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
|  |  | Voltage ratio | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
|  |  | Polarization Voltage | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22
Microphone parameters |  |  | Acoustic impedance incl. volume (fit) | 600 | 300 | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 350 | 600
|  |  | Cavity depth | 2 | 2 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 2
Corrections |  |  | Heat conduction | 400 | 300 | 200 | 150 | 100 | 70 | 50 | 40 | 40 | 40 | 40 | 40 | 40 | 70
|  |  | Thread influence neglected | 700 | 400 | 200 | 100 | 50 | 50 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Couplers |  |  | Diameter | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 20
|  |  | Length | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 15 | 15 | 15 | 20
|  |  | Leakage | 100 | 75 | 50 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Correction of results to normal environmental conditions |  |  | Static pressure | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 100 | 150
|  |  | Temperature | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 50 | 100
Environmental conditions |  |  | Static pressure | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 30 | 150
|  |  | Temperature | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 25 | 30
|  |  | Humidity | 2 | 2 | 2 | 2 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 10 | 15
Humid air properties |  |  |  | 200 | 200 | 200 | 100 | 100 | 100 | 100 | 100 | 100 | 50 | 50 | 50 | 50 | 50
Rounding error |  |  |  | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50
Estimate of type B uncertainty (S.D.), k=1 |  |  |  | 598.7604 | 362.2462 | 243.6131 | 171.5022 | 150.7315 | 144.9828 | 139.236 | 138.1545 | 138.1545 | 128.7892 | 128.5626 | 128.5626 | 219.4979 | 377.6299
Overall uncertainty (10-4 dB)
|  | Type A, k=2 |  | 2000 | 1000 | 600 | 300 | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 300 | 360
|  | Type B, k=2 |  | 1197.5208 | 724.4925 | 487.2262 | 343.0044 | 301.4631 | 289.9655 | 278.472 | 276.309 | 276.309 | 257.5785 | 257.1251 | 257.1251 | 438.9958 | 755.2598
|  | Overall uncertainty, k=2 |  | 2331.1062 | 1234.8641 | 772.9097 | 455.6885 | 361.7734 | 352.2499 | 342.8508 | 341.0963 | 341.0963 | 326.1084 | 325.7504 | 325.7504 | 531.7117 | 836.6704
|  | Overall value (dB) |  | 0.23311 | 0.12349 | 0.077291 | 0.045569 | 0.036177 | 0.035225 | 0.034285 | 0.03411 | 0.03411 | 0.032611 | 0.032575 | 0.032575 | 0.053171 | 0.083667
|  | Frequency /Hz) |  | 2 | 4 | 8 | 16 | 31.5 | 63 | 125 | 250 | 500 | 1k | 2k | 4k | 8k | 10k

## PTB Phase
PTB | Phase
LS1P microphones in 3 plane-wave couplers | [10 empty] | Frequency in Hz
|  |  |  | 2 | 4 | 8 | 16 | 31.5 | 63 | 125 | 250 | 500 | 1k | 2k | 4k | 8k | 10k
Type A uncertainty, as standard deviation (10-4 degrees)
Source of uncertainty
Normal distribution
|  |  | Allowed repeatibility | 20000 | 10000 | 5000 | 2000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 2000 | 2000
Estimate of a type A uncertainty (S.D.), k=1 |  |  |  | 20000 | 10000 | 5000 | 2000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 2000 | 2000
Type B uncertainty, as semi-ranges (10-4 degrees)
Source of uncertainty
Rectangular distribution
Measurement |  |  | Measurement Capacitor | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  |  | Phase ratio | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000
|  |  | Polarization Voltage | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Microphone parameters |  |  | Acoustic impedance incl. volume(fit) | 10000 | 5000 | 2000 | 1500 | 800 | 500 | 300 | 200 | 150 | 500 | 1000 | 2000 | 3000 | 5000
|  |  | Cavity depth | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 100
Corrections |  |  | Heat conduction | 400 | 300 | 200 | 150 | 100 | 70 | 50 | 40 | 40 | 40 | 40 | 40 | 40 | 70
|  |  | Thread influence neglected | 5000 | 3000 | 2000 | 1500 | 1000 | 1000 | 600 | 400 | 300 | 300 | 200 | 100 | 100 | 100
Couplers |  |  | Diameter | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 75 | 100
|  |  | Length | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 75 | 100
|  |  | Leakage | 2000 | 1000 | 600 | 300 | 100 | 75 | 50 | 20 | 10 | 0 | 0 | 0 | 0 | 0
Correction of results to normal environmental conditions |  |  | Static pressure | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 300 | 500 | 500
|  |  | Temperature | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 100
Environmental conditions |  |  | Static pressure | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 30 | 150
|  |  | Temperature | 50 | 50 | 50 | 50 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 80 | 100
|  |  | Humidity | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 20 | 30
Humid air properties |  |  |  | 2000 | 1500 | 1000 | 1000 | 500 | 300 | 300 | 300 | 100 | 100 | 100 | 100 | 100 | 100
Rounding error |  |  |  | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500
Estimate of type B uncertainty (S.D.), k=1 |  |  |  | 6768.0672 | 3723.8063 | 2133.8697 | 1814.0838 | 1433.8875 | 1367.3118 | 1265.3195 | 1231.5437 | 1208.6218 | 1239.5833 | 1330.3759 | 1670.0499 | 2124.8961 | 3141.0349
Overall uncertainty (10-4 degrees)
|  | Type A, k=2 |  | 40000 | 20000 | 10000 | 4000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 4000 | 4000
|  | Type B, k=2 |  | 13536.1344 | 7447.6126 | 4267.7394 | 3628.1676 | 2867.775 | 2734.6237 | 2530.6389 | 2463.0875 | 2417.2436 | 2479.1665 | 2660.7518 | 3340.0998 | 4249.7922 | 6282.0697
|  | Overall uncertainty, k=2 |  | 42228.2717 | 21341.6713 | 10872.6078 | 5400.3333 | 3496.3028 | 3387.9443 | 3225.5439 | 3172.8221 | 3137.3662 | 3185.3205 | 3328.6033 | 3893.105 | 5836.1574 | 7447.4425
|  | Overall value (degrees) |  | 4.2228 | 2.1342 | 1.0873 | 0.54003 | 0.34963 | 0.33879 | 0.32255 | 0.31728 | 0.31374 | 0.31853 | 0.33286 | 0.38931 | 0.58362 | 0.74474
|  |  |  | 2 | 4 | 8 | 16 | 31.5 | 63 | 125 | 250 | 500 | 1k | 2k | 4k | 8k | 10k
|  | Quoted value (degrees) |  | 4.5 | 2.5 | 1.5 | 1 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 1 | 1

## CEM Level
CEM | Level
Frequency /Hz |  | 10 | 12.5893 | 15.8489 | 19.9526 | 25.1189 | 31.6228 | 39.8107 | 50.1187 | 63.0957 | 79.4328 | 100 | 125.8925 | 158.4893 | 199.5262 | 251.1886 | 316.2278 | 398.1072 | 501.1872 | 630.9573 | 794.3282 | 1000 | 1258.9254 | 1584.8932 | 1995.2623 | 2511.8864 | 3162.2777 | 3981.0717 | 5011.8723 | 6309.5734 | 7943.2823 | 10000
Type B (rectangular) | Electrical parameters /(10-4 dB) | 25.1936 | 11.8048 | 9.4186 | 7.7745 | 8.0538 | 7.7525 | 8.2352 | 9.6705 | 4.6082 | 4.6623 | 5.2215 | 4.435 | 5.3642 | 10.8353 | 7.7035 | 6.3586 | 5.0337 | 4.9532 | 5.4619 | 5.6485 | 4.5363 | 5.3793 | 4.453 | 4.4111 | 4.4845 | 4.4738 | 4.4464 | 4.4199 | 4.4234 | 4.4254 | 4.4237
| Microphone parameters /(10-4 dB) | 376.4664 | 377.1067 | 377.1721 | 377.2086 | 218.9675 | 218.9779 | 218.961 | 218.9049 | 166.9857 | 161.1678 | 160.2836 | 161.2339 | 160.6152 | 160.457 | 160.6772 | 160.581 | 160.7268 | 159.9913 | 160.1085 | 154.715 | 156.2681 | 154.123 | 155.4533 | 150.4651 | 144.3488 | 151.2074 | 151.1901 | 226.7492 | 324.4854 | 323.8829 | 469.6512
| Couplers volume /(10-4 dB) | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
| Static pressure /(10-4 dB) | -2.3182 | -2.3182 | -2.3182 | -2.3182 | -2.3182 | -2.3182 | -2.3182 | -2.3182 | -2.3182 | -2.3182 | -2.3181 | -2.318 | -2.3179 | -2.3177 | -2.3175 | -2.3172 | -2.3169 | -2.3091 | -2.2949 | -2.2771 | -2.2457 | -2.2006 | -2.128 | -2.0194 | -1.8476 | -1.5579 | -1.0918 | -0.44119 | 0.087094 | -0.47321 | -3.2529
| Humidity /(10-4 dB) | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434 | 1.4434
| Temperature /(10-4 dB) | -1.2503 | -1.2503 | -1.2503 | -1.2503 | -1.2503 | -1.2503 | -1.2503 | -1.2503 | -1.2503 | -1.2503 | -1.2512 | -1.253 | -1.2552 | -1.258 | -1.2615 | -1.266 | -1.2716 | -1.3111 | -1.38 | -1.4668 | -1.5854 | -1.7406 | -1.9399 | -2.221 | -2.6862 | -3.5579 | -5.0962 | -7.3091 | -9.3201 | -8.7352 | -1.5797
Type A (normal) | Repeteability and reproducibility (811014) /(10-4 dB) | 87.346 | 87.4193 | 87.4268 | 87.431 | 66.9279 | 66.9295 | 66.927 | 66.9186 | 122.4154 | 146.9495 | 147.8948 | 146.8841 | 139.2478 | 147.4012 | 139.0663 | 147.5271 | 147.4195 | 139.979 | 139.8255 | 145.7633 | 144.1355 | 154.2959 | 152.9839 | 165.3926 | 177.85 | 199.9979 | 219.8125 | 168.8477 | 84.2657 | 86.6121 | 97.4322
| Repeteability and reproducibility (2036126) /(10-4 dB) | 107.9552 | 108.0145 | 108.0206 | 108.0239 | 110.4733 | 110.4742 | 97.8735 | 110.4676 | 64.1134 | 63.1994 | 65.3673 | 63.047 | 64.5364 | 64.2425 | 64.1439 | 64.5311 | 64.2845 | 66.0993 | 65.7736 | 63.6156 | 59.7917 | 65.0555 | 61.8795 | 58.5637 | 57.9277 | 56.561 | 56.5024 | 78.1636 | 133.513 | 256.177 | 576.1016
U (k=2) /dB (811014) |  | 0.08 | 0.08 | 0.08 | 0.08 | 0.05 | 0.05 | 0.05 | 0.05 | 0.046 | 0.048 | 0.048 | 0.048 | 0.047 | 0.048 | 0.047 | 0.048 | 0.048 | 0.047 | 0.047 | 0.047 | 0.047 | 0.048 | 0.048 | 0.049 | 0.05 | 0.054 | 0.057 | 0.06 | 0.07 | 0.07 | 0.098
U (k=2) /dB (2036126) |  | 0.081 | 0.081 | 0.081 | 0.081 | 0.053 | 0.053 | 0.052 | 0.053 | 0.041 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.039 | 0.039 | 0.039 | 0.039 | 0.038 | 0.037 | 0.038 | 0.038 | 0.052 | 0.073 | 0.085 | 0.15

## NIS Level
NIS | Level |  |  | Uncertainity Budget For LS1P Microphone (dB)
Microphone Reciprocity Calibration System Type 9699
S.No. 811014 | Source of Uncertainity | Symbol | Value | Unit | Probabilty Distribution Sensitivity | [14 empty] | Frequency (Hz)
[6 empty] | 20 Hz | 25 Hz | 31.6 Hz | 40 Hz | 50 Hz | 63 Hz | 80 Hz | 100 Hz | 125 Hz | 160 Hz | 200 Hz | 250 Hz | 316 Hz | 400 Hz | 500 Hz | 630 Hz | 800 Hz | 1 kHz | 1.25 kHz | 1 .6kHz | 2 kHz | 2 5kHz | 3.16 kHz | 4 kHz | 5 kHz | 6.3 kHz | 8 kHz | 10 kHz
| Electrical Transfer Impedance |  | S | dB |  | 0.011021 | 0.011021 | 0.010976 | 0.010976 | 0.010976 | 0.010931 | 0.010931 | 0.010931 | 0.010845 | 0.010845 | 0.010845 | 0.010762 | 0.010762 | 0.010762 | 0.010803 | 0.010803 | 0.010803 | 0.010845 | 0.010845 | 0.010845 | 0.010888 | 0.010888 | 0.010888 | 0.010931 | 0.010931 | 0.010976 | 0.011021 | 0.011021
1 | Series Capacitor | c | 1 | nf | rectangular (type B) | 0.005 | 0.005 | 0.0049 | 0.0049 | 0.0049 | 0.0048 | 0.0048 | 0.0048 | 0.0046 | 0.0046 | 0.0046 | 0.0044 | 0.0044 | 0.0044 | 0.0045 | 0.0045 | 0.0045 | 0.0046 | 0.0046 | 0.0046 | 0.0047 | 0.0047 | 0.0047 | 0.0048 | 0.0048 | 0.0049 | 0.005 | 0.005
2 | Votage Ratios | vr | 1 | - | normal(type B) | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071
3 | Cross Talk |  | 1 | - |  | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709
4 | Inherent and Ambient Noise |  | 1 | - |  | 0.00010857 | 0.0001 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857
5 | Distortion |  | 1 | - |  | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002
6 | Frequency | f | 1 | Hz |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
7 | Reciever Ground Shielding |  | 1 | - | rectangular(type B) | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006
8 | Reciever Ground Shielding(Guard Length) |  | 0.15 | mm | rectangular(type B) | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016
9 | Transmitter Ground Shielding |  | 1 | - | rectangular(type B) | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004
10 | Transmitter Ground Shielding(Guard Length) |  | 0.15 | mm | rectangular(type B) | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016
| Acoustic Transfer Impedance(Coupler and Gas) |  |  | dB |  | 0.021164 | 0.019443 | 0.018642 | 0.018642 | 0.018642 | 0.018172 | 0.018172 | 0.018172 | 0.01814 | 0.01814 | 0.01814 | 0.018138 | 0.018138 | 0.018138 | 0.018138 | 0.018138 | 0.018138 | 0.018139 | 0.018139 | 0.018139 | 0.018148 | 0.018148 | 0.018148 | 0.018293 | 0.01852 | 0.019065 | 0.020397 | 0.024386
11 | Coupler Length (normalized at 500Hz) | Lcoup | 0.01 | mm | rectangular(type B) | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0003 | 0.0003 | 0.0003 | 0.001 | 0.0016 | 0.0025 | 0.004 | 0.0073
12 | Coupler Diameter (normalized at 500Hz) | d coup | 0.01 | mm | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0002
13 | Coupler Voulme (SHORT7.5 MM)+Volume fitting | v coup | 5 | mm3 | rectangular(type B) | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065
14 | Coupler Volume (Long 15 mm)+Volume fitting |  | 7 | mm3 | rectangular(type B) | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065
15 | Coupler Surface Area (Short7.5 mm) | scoup | 0.2 | % | rectangular(type B) | 0.00012 | 0.0001 | 8e-05 | 8e-05 | 8e-05 | 6e-05 | 6e-05 | 6e-05 | 4e-05 | 4e-05 | 4e-05 | 4e-05 | 4e-05 | 4e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 0 | 0 | 0 | 0 | 0
16 | Coupler Surface Area (Long15 mm) |  | 0.2 | % | rectangular(type B) | 0.00014 | 0.00012 | 0.0001 | 0.0001 | 0.0001 | 8e-05 | 8e-05 | 8e-05 | 6e-05 | 6e-05 | 6e-05 | 4e-05 | 4e-05 | 4e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | - | - | - | - | - | - | - | -
17 | Capillary Tube Length |  | - | - | rectangular(type B) | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
18 | Capillary Tube Diameter |  | - | - | rectangular(type B) | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
19 | Coupler Leakage |  | 1 | - | rectangular(type B) | 0.0109 | 0.007 | 0.0043 | 0.0043 | 0.0043 | 0.0011 | 0.0011 | 0.0011 | 0.0003 | 0.0003 | 0.0003 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 | Static Pressure of air in Coupler | p | 0.04 | kPa | rectangular(type B) | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.00164 | 0.00168 | 0.00168 | 0.00168
21 | Temperature of air in Coupler | T | 1 | C | rectangular(type B) | 0.0003 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.00015 | 0.00015 | 0.00015 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0002 | 0.0005 | 0.0005 | 0.0005 | 0.0021 | 0.0033 | 0.0052 | 0.0083 | 0.0144
22 | Relative Humidity of air in Coupler | RH | 5 | % | rectangular(type B) | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.0003 | 0.0003 | 0.0003 | 0.00035 | 0.00035 | 0.00035 | 0.00055 | 0.0007 | 0.001 | 0.0014 | 0.0022
| Acoustic Transfer Impedance(Microphone) |  |  | dB |  | 0.0090826 | 0.0083102 | 0.0076522 | 0.0076522 | 0.0076522 | 0.0061471 | 0.0061471 | 0.0061471 | 0.0052997 | 0.0052997 | 0.0052997 | 0.0047635 | 0.0047635 | 0.0047635 | 0.0044563 | 0.0044563 | 0.0044563 | 0.0044686 | 0.0044686 | 0.0044686 | 0.006256 | 0.006256 | 0.006256 | 0.017215 | 0.023116 | 0.024617 | 0.028146 | 0.051859
23 | Microphone Front Cavity Depth | Lf | 0.03 | mm | rectangular(type B) | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0 | 0 | 0 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0015 | 0.0015 | 0.0015 | 0.0063 | 0.0102 | 0.0162 | 0.0258 | 0.0447
24 | Microphone Front Cavity Volume | Vf | 3 | mm3 | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
25 | Microphone Equivalent Volume | Veq | 3.96 | mm3 | rectangular(type B) | 0.000297 | 0.000297 | 0.000198 | 0.000198 | 0.000198 | 0.000198 | 0.000198 | 0.000198 | 9.9e-05 | 9.9e-05 | 9.9e-05 | 0 | 0 | 0 | 0 | 0 | 0 | 9.9e-05 | 9.9e-05 | 9.9e-05 | 0.000594 | 0.000594 | 0.000594 | 0.002871 | 0.004851 | 0.007524 | 0.00891 | 0.005742
26 | Microphone Diaphragm Resonance Frequency | F0 | 550 | Hz | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0005 | 0.0005 | 0.0005 | 0.0054 | 0.0105 | 0.012 | 0.0039 | 0.0127
27 | Microphone Diaphragm Loss Factor | D | 10 | % | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0002 | 0.0002 | 0.0002 | 0.0011 | 0.0011 | 0.0011 | 0.0043 | 0.0043 | 0.0043 | 0.0142 | 0.0167 | 0.0112 | 0.0038 | 0.0219
28 | Microphone Diaphragm Compliance |  | not appl. | m3/Pa | rectangular(type B) | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
29 | Microphone Diaphragm Mass |  | not appl. | kg/m4 | rectangular(type B) | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
30 | Microphone Diaphragm Resistance |  | not appl. | Pa s/m3 | rectangular(type B) | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
31 | Mic.Cavity Thread- Short Coup.(heat cond.) |  | 33 | % | rectangular(type B) | 0.006864 | 0.006072 | 0.005412 | 0.005412 | 0.005412 | 0.003828 | 0.003828 | 0.003828 | 0.002772 | 0.002772 | 0.002772 | 0.00198 | 0.00198 | 0.00198 | 0.00132 | 0.00132 | 0.00132 | 0.000924 | 0.000924 | 0.000924 | 0.00066 | 0.00066 | 0.00066 | 0.000528 | 0.000528 | 0.000396 | 0.000396 | 0.000264
32 | Mic.Cavity Thread- Long Coup.(heat cond.) |  | 33 | % | rectangular(type B) | 0.004224 | 0.003828 | 0.003432 | 0.003432 | 0.003432 | 0.002376 | 0.002376 | 0.002376 | 0.001716 | 0.001716 | 0.001716 | 0.001188 | 0.001188 | 0.001188 | 0.000792 | 0.000792 | 0.000792 | 0.00066 | 0.00066 | 0.00066 | - | - | - | - | - | - | - | -
33 | Microphone Polarization Voltage |  | 0.08 | V | rectangular(type B) | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664
| Imperfection of Theory |  |  | dB |  | 0.01535 | 0.01375 | 0.0123 | 0.0123 | 0.0123 | 0.0088 | 0.0088 | 0.0088 | 0.00625 | 0.00625 | 0.00625 | 0.0044 | 0.0044 | 0.0044 | 0.0030504 | 0.0030504 | 0.0030504 | 0.001966 | 0.001966 | 0.001966 | 0.0012104 | 0.0012104 | 0.0012104 | 0.0040162 | 0.0063885 | 0.0093489 | 0.012615 | 0.018491
34 | Heat Conduction |  | 100 | % | rectangular(type B) | 0.01535 | 0.01375 | 0.0123 | 0.0123 | 0.0123 | 0.0088 | 0.0088 | 0.0088 | 0.00625 | 0.00625 | 0.00625 | 0.0044 | 0.0044 | 0.0044 | 0.00305 | 0.00305 | 0.00305 | 0.00195 | 0.00195 | 0.00195 | 0.00075 | 0.00075 | 0.00075 | 0.0013 | 0.00245 | 0.004 | 0.00635 | 0.01095
35 | Adding of Excess Voulme |  | - | - | rectangular(type B) | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
36 | Viscosity Losses |  | 1 | - | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
37 | Radial Wave-motion |  | 100 | % | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5e-05 | 5e-05 | 5e-05 | 0.00025 | 0.00025 | 0.00025 | 0.00095 | 0.00095 | 0.00095 | 0.0038 | 0.0059 | 0.00845 | 0.0109 | 0.0149
| Processing of Results |  |  | dB |  | 0.011712 | 0.011378 | 0.01289 | 0.012402 | 0.01118 | 0.012558 | 0.011793 | 0.011793 | 0.012255 | 0.012033 | 0.012073 | 0.012519 | 0.011519 | 0.011558 | 0.012277 | 0.01152 | 0.011482 | 0.012298 | 0.011187 | 0.010751 | 0.011959 | 0.010449 | 0.012242 | 0.021802 | 0.025994 | 0.035883 | 0.032065 | 0.046658
38 | Rounding to 2decimals |  | 3 | - |  | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735
39 | Repeatability of Measurements |  |  | - | normal(type A) | 0.0079654 | 0.0074657 | 0.0096138 | 0.0089489 | 0.0071591 | 0.0091636 | 0.0080829 | 0.0080829 | 0.0087444 | 0.0084293 | 0.008487 | 0.00911 | 0.0076788 | 0.0077365 | 0.0087738 | 0.0076788 | 0.007621 | 0.0084691 | 0.006755 | 0.0060044 | 0.0076478 | 0.0049652 | 0.0080829 | 0.017597 | 0.020805 | 0.031786 | 0.019494 | 0.027648
40 | Static Pressure Cor.∆P and Coef.(500Hz) |  | 0.04 | kPa | rectangular(type B) | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632
41 | Static Pressure Cor.∆P and Ucoef.(500Hz) |  | 5 | kPa | rectangular(type B) | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008
42 | Static Pressure Cor.∆P and normalized Coef(f) |  | 0.04 | kPa | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8e-06 | 8e-06 | 8e-06 | 2.4e-05 | 2.4e-05 | 2.4e-05 | 9.6e-05 | 9.6e-05 | 9.6e-05 | 0.000368 | 0.000544 | 0.000656 | 0.000368 | 0.000472
43 | Static Pressure Cor.∆P and Unormalized Coef(f) |  | 5 | kPa | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0005 | 0.0005 | 0.0005 | 0.0015 | 0.0015 | 0.0015 | 0.0055 | 0.006 | 0.001 | 0.02 | 0.033
44 | Temperature Cor.∆T and Coef(500Hz) |  | 1 | K | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023
45 | Temperature Cor.∆T and Ucoef(500Hz) |  | 3 | K | rectangular(type B) | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003
46 | Temperature Cor.∆T and normalized Coef(f) |  | 1 | K | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0004 | 0.0004 | 0.0004 | 0.0017 | 0.0017 | 0.0017 | 0.007 | 0.0108 | 0.014 | 0.0112 | 0.0032
47 | Temperature Cor.∆T and Unormalized Coef(f) |  | 3 | K | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0003 | 0.0003 | 0.0003 | 0.0006 | 0.0006 | 0.0006 | 0.0027 | 0.0033 | 0.0009 | 0.0066 | 0.0153
| Combined uncertainity |  |  |  | normal | 0.03201 | 0.029784 | 0.029051 | 0.028838 | 0.028334 | 0.026882 | 0.026533 | 0.026533 | 0.025769 | 0.025664 | 0.025683 | 0.025369 | 0.024891 | 0.024909 | 0.025011 | 0.024649 | 0.024631 | 0.024935 | 0.024406 | 0.024209 | 0.02513 | 0.024448 | 0.025266 | 0.035241 | 0.041392 | 0.049649 | 0.05017 | 0.07697
| Expanded uncertainity |  |  |  |  | 0.064021 | 0.059568 | 0.058103 | 0.057677 | 0.056668 | 0.053764 | 0.053066 | 0.053066 | 0.051538 | 0.051327 | 0.051365 | 0.050738 | 0.049781 | 0.049817 | 0.050023 | 0.049297 | 0.049262 | 0.049869 | 0.048812 | 0.048418 | 0.050261 | 0.048896 | 0.050532 | 0.070483 | 0.082785 | 0.099297 | 0.10034 | 0.15394
Uncertainity Budget For LS1P Microphone (dB)
Microphone Reciprocity Calibration System Type 9699
S.No. 236126 | Source of Uncertainity | Symbol | Value | Unit | Probabilty Distribution Sensitivity | [14 empty] | Frequency (Hz)
[6 empty] | 20 Hz | 25 Hz | 31.6 Hz | 40 Hz | 50 Hz | 63 Hz | 80 Hz | 100 Hz | 125 Hz | 160 Hz | 200 Hz | 250 Hz | 316 Hz | 400 Hz | 500 Hz | 630 Hz | 800 Hz | 1 kHz | 1.25 kHz | 1.6 kHz | 2 kHz | 2.5 kHz | 3.16 kHz | 4 kHz | 5 kHz | 6.3 kHz | 8 kHz | 10 kHz
| Electrical Transfer Impedance |  | S | dB |  | 0.011021 | 0.011021 | 0.010976 | 0.010976 | 0.010976 | 0.010931 | 0.010931 | 0.010931 | 0.010845 | 0.010845 | 0.010845 | 0.010762 | 0.010762 | 0.010762 | 0.010803 | 0.010803 | 0.010803 | 0.010845 | 0.010845 | 0.010845 | 0.010888 | 0.010888 | 0.010888 | 0.010931 | 0.010931 | 0.010976 | 0.011021 | 0.011021
1 | Series Capacitor | c | 1 | nf | rectangular (type B) | 0.005 | 0.005 | 0.0049 | 0.0049 | 0.0049 | 0.0048 | 0.0048 | 0.0048 | 0.0046 | 0.0046 | 0.0046 | 0.0044 | 0.0044 | 0.0044 | 0.0045 | 0.0045 | 0.0045 | 0.0046 | 0.0046 | 0.0046 | 0.0047 | 0.0047 | 0.0047 | 0.0048 | 0.0048 | 0.0049 | 0.005 | 0.005
2 | Votage Ratios | vr | 1 | - | normal(type B) | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071 | 0.0071
3 | Cross Talk |  | 1 | - |  | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709 | 0.0021709
4 | Inherent and Ambient Noise |  | 1 | - |  | 0.00010857 | 0.0001 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857 | 0.00010857
5 | Distortion |  | 1 | - |  | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002
6 | Frequency | f | 1 | Hz |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
7 | Reciever Ground Shielding |  | 1 | - | rectangular(type B) | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006
8 | Reciever Ground Shielding(Guard Length) |  | 0.15 | mm | rectangular(type B) | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016
9 | Transmitter Ground Shielding |  | 1 | - | rectangular(type B) | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004
10 | Transmitter Ground Shielding(Guard Length) |  | 0.15 | mm | rectangular(type B) | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016
| Acoustic Transfer Impedance(Coupler and Gas) |  |  | dB |  | 0.021164 | 0.019443 | 0.018642 | 0.018642 | 0.018642 | 0.018172 | 0.018172 | 0.018172 | 0.01814 | 0.01814 | 0.01814 | 0.018138 | 0.018138 | 0.018138 | 0.018138 | 0.018138 | 0.018138 | 0.018139 | 0.018139 | 0.018139 | 0.018148 | 0.018148 | 0.018148 | 0.018293 | 0.01852 | 0.019065 | 0.020397 | 0.024386
11 | Coupler Length (normalized at 500Hz) | Lcoup | 0.01 | mm | rectangular(type B) | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0003 | 0.0003 | 0.0003 | 0.001 | 0.0016 | 0.0025 | 0.004 | 0.0073
12 | Coupler Diameter (normalized at 500Hz) | d coup | 0.01 | mm | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0002
13 | Coupler Voulme (SHORT7.5 MM)+Volume fitting | v coup | 5 | mm3 | rectangular(type B) | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065
14 | Coupler Volume (Long 15mm)+Volume fitting |  | 7 | mm3 | rectangular(type B) | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065 | 0.018065
15 | Coupler Surface Area (Short7.5 mm) | scoup | 0.2 | % | rectangular(type B) | 0.00012 | 0.0001 | 8e-05 | 8e-05 | 8e-05 | 6e-05 | 6e-05 | 6e-05 | 4e-05 | 4e-05 | 4e-05 | 4e-05 | 4e-05 | 4e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 0 | 0 | 0 | 0 | 0
16 | Coupler Surface Area (Long15 mm) |  | 0.2 | % | rectangular(type B) | 0.00014 | 0.00012 | 0.0001 | 0.0001 | 0.0001 | 8e-05 | 8e-05 | 8e-05 | 6e-05 | 6e-05 | 6e-05 | 4e-05 | 4e-05 | 4e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | 2e-05 | - | - | - | - | - | - | - | -
17 | Capillary Tube Length |  | - | - | rectangular(type B) | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
18 | Capillary Tube Diameter |  | - | - | rectangular(type B) | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
19 | Coupler Leakage |  | 1 | - | rectangular(type B) | 0.0109 | 0.007 | 0.0043 | 0.0043 | 0.0043 | 0.0011 | 0.0011 | 0.0011 | 0.0003 | 0.0003 | 0.0003 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 | Static Pressure of air in Coupler | p | 0.04 | kPa | rectangular(type B) | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0.00164 | 0.00168 | 0.00168 | 0.00168
21 | Temperature of air in Coupler | T | 1 | C | rectangular(type B) | 0.0003 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.00015 | 0.00015 | 0.00015 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0002 | 0.0005 | 0.0005 | 0.0005 | 0.0021 | 0.0033 | 0.0052 | 0.0083 | 0.0144
22 | Relative Humidity of air in Coupler | RH | 5 | % | rectangular(type B) | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.0003 | 0.0003 | 0.0003 | 0.00035 | 0.00035 | 0.00035 | 0.00055 | 0.0007 | 0.001 | 0.0014 | 0.0022
| Acoustic Transfer Impedance(Microphone) |  |  | dB |  | 0.0090826 | 0.0083102 | 0.0076522 | 0.0076522 | 0.0076522 | 0.0061471 | 0.0061471 | 0.0061471 | 0.0052997 | 0.0052997 | 0.0052997 | 0.0047635 | 0.0047635 | 0.0047635 | 0.0044563 | 0.0044563 | 0.0044563 | 0.0044686 | 0.0044686 | 0.0044686 | 0.006256 | 0.006256 | 0.006256 | 0.017215 | 0.023116 | 0.024617 | 0.028146 | 0.051859
23 | Microphone Front Cavity Depth | Lf | 0.03 | mm | rectangular(type B) | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0 | 0 | 0 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0015 | 0.0015 | 0.0015 | 0.0063 | 0.0102 | 0.0162 | 0.0258 | 0.0447
24 | Microphone Front Cavity Volume | Vf | 3 | mm3 | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
25 | Microphone Equivalent Volume | Veq | 3.96 | mm3 | rectangular(type B) | 0.000297 | 0.000297 | 0.000198 | 0.000198 | 0.000198 | 0.000198 | 0.000198 | 0.000198 | 9.9e-05 | 9.9e-05 | 9.9e-05 | 0 | 0 | 0 | 0 | 0 | 0 | 9.9e-05 | 9.9e-05 | 9.9e-05 | 0.000594 | 0.000594 | 0.000594 | 0.002871 | 0.004851 | 0.007524 | 0.00891 | 0.005742
26 | Microphone Diaphragm Resonance Frequency | F0 | 550 | Hz | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0005 | 0.0005 | 0.0005 | 0.0054 | 0.0105 | 0.012 | 0.0039 | 0.0127
27 | Microphone Diaphragm Loss Factor | D | 10 | % | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0002 | 0.0002 | 0.0002 | 0.0011 | 0.0011 | 0.0011 | 0.0043 | 0.0043 | 0.0043 | 0.0142 | 0.0167 | 0.0112 | 0.0038 | 0.0219
28 | Microphone Diaphragm Compliance |  | not appl. | m3/Pa | rectangular(type B) | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
29 | Microphone Diaphragm Mass |  | not appl. | kg/m4 | rectangular(type B) | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
30 | Microphone Diaphragm Resistance |  | not appl. | Pa s/m3 | rectangular(type B) | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
31 | Mic.Cavity Thread- Short Coup.(heat cond.) |  | 33 | % | rectangular(type B) | 0.006864 | 0.006072 | 0.005412 | 0.005412 | 0.005412 | 0.003828 | 0.003828 | 0.003828 | 0.002772 | 0.002772 | 0.002772 | 0.00198 | 0.00198 | 0.00198 | 0.00132 | 0.00132 | 0.00132 | 0.000924 | 0.000924 | 0.000924 | 0.00066 | 0.00066 | 0.00066 | 0.000528 | 0.000528 | 0.000396 | 0.000396 | 0.000264
32 | Mic.Cavity Thread- Long Coup.(heat cond.) |  | 33 | % | rectangular(type B) | 0.004224 | 0.003828 | 0.003432 | 0.003432 | 0.003432 | 0.002376 | 0.002376 | 0.002376 | 0.001716 | 0.001716 | 0.001716 | 0.001188 | 0.001188 | 0.001188 | 0.000792 | 0.000792 | 0.000792 | 0.00066 | 0.00066 | 0.00066 | - | - | - | - | - | - | - | -
33 | Microphone Polarization Voltage |  | 0.08 | V | rectangular(type B) | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664 | 0.0041664
| Imperfection of Theory |  |  | dB |  | 0.01535 | 0.01375 | 0.0123 | 0.0123 | 0.0123 | 0.0088 | 0.0088 | 0.0088 | 0.00625 | 0.00625 | 0.00625 | 0.0044 | 0.0044 | 0.0044 | 0.0030504 | 0.0030504 | 0.0030504 | 0.001966 | 0.001966 | 0.001966 | 0.0012104 | 0.0012104 | 0.0012104 | 0.0040162 | 0.0063885 | 0.0093489 | 0.012615 | 0.018491
34 | Heat Conduction |  | 100 | % | rectangular(type B) | 0.01535 | 0.01375 | 0.0123 | 0.0123 | 0.0123 | 0.0088 | 0.0088 | 0.0088 | 0.00625 | 0.00625 | 0.00625 | 0.0044 | 0.0044 | 0.0044 | 0.00305 | 0.00305 | 0.00305 | 0.00195 | 0.00195 | 0.00195 | 0.00075 | 0.00075 | 0.00075 | 0.0013 | 0.00245 | 0.004 | 0.00635 | 0.01095
35 | Adding of Excess Voulme |  | - | - | rectangular(type B) | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
36 | Viscosity Losses |  | 1 | - | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
37 | Radial Wave-motion |  | 100 | % | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5e-05 | 5e-05 | 5e-05 | 0.00025 | 0.00025 | 0.00025 | 0.00095 | 0.00095 | 0.00095 | 0.0038 | 0.0059 | 0.00845 | 0.0109 | 0.0149
| Processing of Results |  |  | dB |  | 0.010728 | 0.0098546 | 0.0096492 | 0.0094879 | 0.0095401 | 0.0095215 | 0.0093254 | 0.009406 | 0.0093762 | 0.0094131 | 0.0093618 | 0.0093379 | 0.009371 | 0.0093873 | 0.009428 | 0.0093814 | 0.0093837 | 0.0095906 | 0.009684 | 0.0097273 | 0.010078 | 0.010263 | 0.010854 | 0.014488 | 0.017382 | 0.0192 | 0.027306 | 0.042304
38 | Rounding to 2decimals |  | 3 | - |  | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735
39 | Repeatability of Measurements |  |  | - | normal(type A) | 0.0064304 | 0.0048354 | 0.0044016 | 0.0040357 | 0.0041569 | 0.0041142 | 0.0036373 | 0.0038394 | 0.0037657 | 0.0038567 | 0.0037297 | 0.0036692 | 0.0037528 | 0.0037932 | 0.0038904 | 0.0037759 | 0.0037816 | 0.0035293 | 0.0037759 | 0.0038856 | 0.0041262 | 0.0045611 | 0.0057677 | 0.006648 | 0.0076987 | 0.00956 | 0.0098712 | 0.019417
40 | Static Pressure Cor.∆P and Coef.(500Hz) |  | 0.04 | kPa | rectangular(type B) | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632 | 0.000632
41 | Static Pressure Cor.∆P and Ucoef.(500Hz) |  | 5 | kPa | rectangular(type B) | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008
42 | Static Pressure Cor.∆P and normalized Coef(f) |  | 0.04 | kPa | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8e-06 | 8e-06 | 8e-06 | 2.4e-05 | 2.4e-05 | 2.4e-05 | 9.6e-05 | 9.6e-05 | 9.6e-05 | 0.000368 | 0.000544 | 0.000656 | 0.000368 | 0.000472
43 | Static Pressure Cor.∆P and Unormalized Coef(f) |  | 5 | kPa | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0005 | 0.0005 | 0.0005 | 0.0015 | 0.0015 | 0.0015 | 0.0055 | 0.006 | 0.001 | 0.02 | 0.033
44 | Temperature Cor.∆T and Coef(500Hz) |  | 1 | K | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023 | 0.0023
45 | Temperature Cor.∆T and Ucoef(500Hz) |  | 3 | K | rectangular(type B) | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003
46 | Temperature Cor.∆T and normalized Coef(f) |  | 1 | K | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0004 | 0.0004 | 0.0004 | 0.0017 | 0.0017 | 0.0017 | 0.007 | 0.0108 | 0.014 | 0.0112 | 0.0032
47 | Temperature Cor.∆T and Unormalized Coef(f) |  | 3 | K | rectangular(type B) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0003 | 0.0003 | 0.0003 | 0.0006 | 0.0006 | 0.0006 | 0.0027 | 0.0033 | 0.0009 | 0.0066 | 0.0153
| Combined uncertainity |  |  |  | normal | 0.031663 | 0.029236 | 0.027766 | 0.02771 | 0.027728 | 0.025604 | 0.025532 | 0.025562 | 0.02453 | 0.024545 | 0.024525 | 0.023959 | 0.023972 | 0.023979 | 0.023743 | 0.023724 | 0.023725 | 0.023716 | 0.023754 | 0.023772 | 0.024291 | 0.024369 | 0.024623 | 0.031249 | 0.036603 | 0.039319 | 0.04727 | 0.074411
| Expanded uncertainity |  |  |  |  | 0.063326 | 0.058471 | 0.055531 | 0.05542 | 0.055456 | 0.051209 | 0.051064 | 0.051123 | 0.049061 | 0.049089 | 0.04905 | 0.047919 | 0.047944 | 0.047957 | 0.047486 | 0.047449 | 0.047451 | 0.047433 | 0.047509 | 0.047544 | 0.048583 | 0.048738 | 0.049247 | 0.062498 | 0.073205 | 0.078639 | 0.09454 | 0.14882