Measurement Results

Laboratory: FGUP VNIIOFI                            Filter Identifier: A

Measurement results for Step 4. All uncertainties should be reported as absolute uncertainties.

Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance (a) | 0,9131376 | 0,9147526 | 0,9171418 | 0,9186229 | 0,9196406 | 0,9203963 | 0,9209704 | 0,9218300 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Type A Uncertainty(b) | 2,8156 E-05 | 4,0052 E-05 | 2,8151 E-05 | 2,4720 E-05 | 1,6485 E-05 | 1,2403 E-05 | 7,8057 E-06 | 1,6309 E-05 |
| Type B Uncertainty(c) | 4,9937 E-04 | 1,2933 E-03 | 1,3255 E-03 | 1,0052 E-03 | 7,6411 E-04 | 7,8985 E-04 | 7,1547 E-04 | 1,1462 E-03 |
| Total Uncertainty(d) | 5,0017 E-04 | 1,2940 E-03 | 1,3258 E-03 | 1,0055 E-03 | 7,6429 E-04 | 7,8995 E-04 | 7,1552 E-04 | 1,1463 E-03 |
| Degrees of Freedom | 8,9620 E+05 | 9,8044 E+06 | 4,4270 E+07 | 2,4634 E+07 | 4,1582 E+07 | 1,4810 E+08 | 6,3545 E+08 | 2,1965 E+08 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Type B Uncertainty Budget

Measurement results for Step 4. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: А

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 7,2171 E-06 | 8,7833 E-06 | 1,3860 E-05 | 1,3983 E-05 | 8,7377 E-06 | 1,2001 E-05 | 1,0165 E-05 | 1,3118 E-05 | 0,99 |
| Wavelength | 8,6271 E-06 | 2,9978 E-06 | 1,4648 E-06 | 1,1898 E-06 | 8,8534 E-07 | 7,3765 E-07 | 8,9227 E-07 | 0 | 0,8 |
| Stray Light | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Beam Size & Position | 1,2702 E-04 | 9,4685 E-04 | 1,2471 E-03 | 9,2953 E-04 | 5,2539 E-04 | 1,1547 E-04 | 2,3094 E-05 | 5,7735 E-06 | 0,9 |
| Inter-reflection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Obliquity | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Polarization | 4,8282 E-04 | 8,8096 E-04 | 4,4883 E-04 | 3,8224 E-04 | 5,5476 E-04 | 7,8127 E-04 | 7,1503 E-04 | 1,1461 E-03 | 0,99 |
| Source Drift & Fluctuation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Bandwidth | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Total Type B Uncertainty | 4,9937 E-04 | 1,2933 E-03 | 1,3255 E-03 | 1,0052 E-03 | 7,6411 E-04 | 7,8985 E-04 | 7,1547 E-04 | 1,1462 E-03 |  |
| Degrees of Freedom | 8,9620 E+05 | 9,8044 E+06 | 4,4270 E+07 | 2,4634 E+07 | 4,1582 E+07 | 1,4810 E+08 | 6,3545 E+08 | 2,1965 E+08 |  |

Measurement Results

Measurement results for Step 4. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: B

Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance (a) | 0,4085577 | 0,6069410 | 0,6211547 | 0,6094001 | 0,6368400 | 0,5762352 | 0,5012691 | 0,4529791 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Type A Uncertainty(b) | 8,7716 E-05 | 2,1979 E-05 | 1,8460 E-05 | 2,7258 E-05 | 1,8769 E-05 | 1,4744 E-05 | 6,7638 E-06 | 5,6853 E-06 |
| Type B Uncertainty(c) | 1,2370 E-03 | 6,6428 E-04 | 5,1607 E-04 | 4,1801 E-04 | 4,3858 E-04 | 4,9914 E-04 | 4,5284 E-04 | 5,5023 E-04 |
| Total Uncertainty(d) | 1,2401 E-03 | 6,6465 E-04 | 5,1640 E-04 | 4,1890 E-04 | 4,3899 E-04 | 4,9935 E-04 | 4,5289 E-04 | 5,5026 E-04 |
| Degrees of Freedom | 3,5952 E+05 | 7,5264 E+06 | 5,5112 E+06 | 5,0202 E+05 | 2,6931 E+06 | 1,1842 E+07 | 1,8091 E+08 | 7,8975 E+08 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Type B Uncertainty Budget

Measurement results for Step 4. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: В

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 1,2825 E-05 | 1,1692 E-05 | 9,9724 E-06 | 9,1111 E-06 | 1,6413 E-06 | 1,7521 E-05 | 1,0161 E-05 | 5,6834 E-06 | 0,99 |
| Wavelength | 1,1453 E-03 | 1,6461 E-05 | 1,3738 E-05 | 3,1704 E-05 | 6,9918 E-05 | 8,6643 E-05 | 5,5641 E-05 | 5,5641 E-05 | 0,8 |
| Stray Light | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Beam Size & Position | 3,7528 E-04 | 5,3116 E-04 | 4,3301 E-04 | 3,8105 E-04 | 3,2909 E-04 | 3,2332 E-04 | 2,7713 E-04 | 3,4064 E-04 | 0,9 |
| Inter-reflection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Obliquity | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Polarization | 2,7836 E-04 | 3,9841 E-04 | 2,8025 E-04 | 1,6866 E-04 | 2,8136 E-04 | 3,6985 E-04 | 3,5365 E-04 | 4,2847 E-04 | 0,99 |
| Source Drift & Fluctuation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Bandwidth | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Total Type B Uncertainty | 1,2370 E-03 | 6,6428 E-04 | 5,1607 E-04 | 4,1801 E-04 | 4,3858 E-04 | 4,9914 E-04 | 4,5284 E-04 | 5,5023 E-04 |  |
| Degrees of Freedom | 3,5952 E+05 | 7,5264 E+06 | 5,5112 E+06 | 5,0202 E+05 | 2,6931 E+06 | 1,1842 E+07 | 1,8091 E+08 | 7,8975 E+08 |  |

Measurement Results

Measurement results for Step 4. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: C

Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance (a) | 0,0210100 | 0,0940007 | 0,0907814 | 0,0761198 | 0,1599051 | 0,1484756 | 0,1015163 | 0,0754628 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Type A Uncertainty(b) | 2,1930 E-05 | 1,6604 E-05 | 4,4444 E-06 | 1,3507 E-05 | 1,2081 E-05 | 4,2939 E-06 | 3,0904 E-06 | 3,6539 E-06 |
| Type B Uncertainty(c) | 1,1403 E-03 | 1,0381 E-04 | 5,5074 E-05 | 1,0818 E-04 | 1,5653 E-04 | 1,2157 E-04 | 9,7219 E-05 | 1,1790 E-04 |
| Total Uncertainty(d) | 1,1405 E-03 | 1,0513 E-04 | 5,5253 E-05 | 1,0902 E-04 | 1,5699 E-04 | 1,2164 E-04 | 9,7268 E-05 | 1,1796 E-04 |
| Degrees of Freedom | 6,5847 E+07 | 1,4466 E+04 | 2,1499 E+05 | 3,8204 E+04 | 2,5663 E+05 | 5,7968 E+06 | 8,8326 E+06 | 9,7750 E+06 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Type B Uncertainty Budget

Measurement results for Step 4. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: C

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 2,2910 E-06 | 2,8574 E-06 | 6,5889 E-06 | 7,4669 E-06 | 4,4322 E-06 | 8,3053 E-06 | 7,1397 E-06 | 7,2453 E-06 | 0,99 |
| Wavelength | 4,2106 E-04 | 3,8198 E-06 | 1,6956 E-05 | 9,6754 E-05 | 1,3168 E-05 | 5,4218 E-05 | 3,0073 E-05 | 8,7106 E-05 | 0,8 |
| Stray Light | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Beam Size & Position | 8,6603 E-05 | 9,8150 E-05 | 5,1962 E-05 | 4,6188 E-05 | 1,5588 E-04 | 9,8150 E-05 | 8,6603 E-05 | 6,9282 E-05 | 0,9 |
| Inter-reflection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Obliquity | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Polarization | 1,0562 E-03 | 3,3486 E-05 | 1,5011 E-06 | 1,2394 E-05 | 2,8098 E-06 | 4,6227 E-05 | 3,1562 E-05 | 3,8221 E-05 | 0,99 |
| Source Drift & Fluctuation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Bandwidth | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Total Type B Uncertainty | 1,1403 E-03 | 1,0381 E-04 | 5,5074 E-05 | 1,0818 E-04 | 1,5653 E-04 | 1,2157 E-04 | 9,7219 E-05 | 1,1790 E-04 |  |
| Degrees of Freedom | 6,5847 E+07 | 1,4466 E+04 | 2,1499 E+05 | 3,8204 E+04 | 2,5663 E+05 | 5,7968 E+06 | 8,8326 E+06 | 9,7750 E+06 |  |

Measurement Results

Measurement results for Step 4. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: D

Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance (a) |  | 0,0051655 | 0,0084360 | 0,0083642 | 0,0265792 | 0,0321917 | 0,0224820 | 0,0166342 |
| Number of Measurements |  | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Type A Uncertainty(b) |  | 3,8126 E-06 | 1,2996 E-06 | 2,3561 E-06 | 3,0140 E-06 | 1,2894 E-06 | 9,8883 E-07 | 1,2806 E-06 |
| Type B Uncertainty(c) |  | 1,2583 E-04 | 9,3142 E-06 | 2,3010 E-05 | 3,1560 E-05 | 2,2706 E-05 | 1,3137 E-05 | 2,2533 E-05 |
| Total Uncertainty(d) |  | 1,2589 E-04 | 9,4044 E-06 | 2,3131 E-05 | 3,1704 E-05 | 2,2743 E-05 | 1,3175 E-05 | 2,2569 E-05 |
| Degrees of Freedom |  | 1,0697 E+07 | 2,4682 E+04 | 8,3603 E+04 | 1,1018 E+05 | 8,7115 E+05 | 2,8360 E+05 | 8,6822 E+05 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Type B Uncertainty Budget

Measurement results for Step 4. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: D

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity |  | 5,0863 E-06 | 7,2938 E-06 | 7,2350 E-06 | 5,9134 E-06 | 3,1843 E-06 | 2,6306 E-06 | 6,0397 E-06 | 0,99 |
| Wavelength |  | 3,7364 E-06 | 9,4351 E-08 | 2,1018 E-05 | 6,5124 E-06 | 1,1202 E-05 | 6,7509 E-06 | 1,9221 E-05 | 0,8 |
| Stray Light |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Beam Size & Position |  | 0 | 5,7735 E-06 | 5,7735 E-06 | 2,8868 E-05 | 1,1547 E-05 | 5,7735 E-06 | 5,7735 E-06 | 0,9 |
| Inter-reflection |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Obliquity |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Polarization |  | 1,2567 E-04 | 4,6188 E-07 | 1,4241 E-06 | 9,2376 E-06 | 1,5704 E-05 | 9,3146 E-06 | 8,2754 E-06 | 0,99 |
| Source Drift & Fluctuation |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Bandwidth |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Total Type B Uncertainty |  | 1,2583 E-04 | 9,3142 E-06 | 2,3010 E-05 | 3,1560 E-05 | 2,2706 E-05 | 1,3137 E-05 | 2,2533 E-05 |  |
| Degrees of Freedom |  | 1,0697 E+07 | 2,4682 E+04 | 8,3603 E+04 | 1,1018 E+05 | 8,7115 E+05 | 2,8360 E+05 | 8,6822 E+05 |  |

Measurement Results

Measurement results for Step 4. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: E

Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance (a) |  |  | 0,0009567 | 0,0010284 | 0,0051548 | 0,0098266 | 0,0085558 | 0,0072300 |
| Number of Measurements |  |  | 10 | 10 | 10 | 10 | 10 | 10 |
| Type A Uncertainty(b) |  |  | 8,9505 E-07 | 1,0667 E-06 | 2,0483 E-06 | 9,4516 E-07 | 8,6667 E-07 | 1,1547 E-06 |
| Type B Uncertainty(c) |  |  | 8,1914 E-06 | 1,0497 E-05 | 8,2371 E-06 | 1,8227 E-05 | 1,9007 E-05 | 1,5653 E-05 |
| Total Uncertainty(d) |  |  | 8,2402 E-06 | 1,0551 E-05 | 8,4880 E-06 | 1,8252 E-05 | 1,9026 E-05 | 1,5695 E-05 |
| Degrees of Freedom |  |  | 6,4655 E+04 | 8,6147 E+04 | 2,6539 E+03 | 1,2516 E+06 | 2,0905 E+06 | 3,0723 E+05 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Type B Uncertainty Budget

Measurement results for Step 4. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: E

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity |  |  | 7,7485 E-06 | 4,5164 E-06 | 2,1536 E-06 | 5,1232 E-06 | 7,2818 E-06 | 6,1808 E-06 | 0,99 |
| Wavelength |  |  | 7,9151 E-08 | 4,7611 E-06 | 5,4001 E-06 | 1,4631 E-06 | 1,5319 E-06 | 8,3531 E-06 | 0,8 |
| Stray Light |  |  | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Beam Size & Position |  |  | 0 | 5,7735 E-06 | 5,7735 E-06 | 1,7321 E-05 | 1,7321 E-05 | 1,1547 E-05 | 0,9 |
| Inter-reflection |  |  | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Obliquity |  |  | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Polarization |  |  | 2,6558 E-06 | 5,8120 E-06 | 8,4678 E-07 | 1,9630 E-06 | 2,4249 E-06 | 1,9245 E-06 | 0,99 |
| Source Drift & Fluctuation |  |  | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Bandwidth |  |  | 0 | 0 | 0 | 0 | 0 | 0 |  |
| Total Type B Uncertainty |  |  | 8,1914 E-06 | 1,0497 E-05 | 8,2371 E-06 | 1,8227 E-05 | 1,9007 E-05 | 1,5653 E-05 |  |
| Degrees of Freedom |  |  | 6,4655 E+04 | 8,6147 E+04 | 2,6539 E+03 | 1,2516 E+06 | 2,0905 E+06 | 3,0723 E+05 |  |