Appendix A.2 Measurement Results (1)

Red: Revised at the pre-Draft A stage for consistency check

Blue: Revised at the Draft A stage to hide identification of the DUT filter set

Please reproduce the following tables for each of the five filters measured at the completion of step 2 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: National Metrology Institute of Japan (NMIJ, AIST)

Filter Type: A

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 9.135E-01 | 9.147E-01 | 9.169E-01 | 9.180E-01 | 9.192E-01 | 9.197E-01 | 9.204E-01 | 9.208E-01 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Temperature | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 |
| Type A Uncertainty(b) | 1.4E-04 | 1.4E-04 | 1.2E-04 | 1.2E-04 | 1.2E-04 | 1.2E-04 | 1.2E-04 | 1.2E-04 |
| Type B Uncertainty(c) | 5.2E-04 | 5.1E-04 | 3.1E-04 | 3.0E-04 | 3.0E-04 | 3.2E-04 | 3.0E-04 | 3.0E-04 |
| Total Uncertainty(d) | 5.4E-04 | 5.3E-04 | 3.3E-04 | 3.3E-04 | 3.3E-04 | 3.4E-04 | 3.3E-04 | 3.2E-04 |
| Degrees of Freedom | 1981 | 1819 | 503 | 479 | 473 | 556 | 464 | 432 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

Filter Type: A

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 1.6E-04 | 1.6E-04 | 1.6E-04 | 1.6E-04 | 1.6E-04 | 1.6E-04 | 1.6E-04 | 1.6E-04 |
| Temperature | 2.5E-04 | 3.0E-04 | 9.0E-05 | 7.4E-05 | 6.9E-05 | 1.2E-04 | 7.9E-06 | 8.0E-06 |
| Wavelength | 2.4E-05 | 5.7E-07 | 1.6E-06 | 6.4E-06 | 3.0E-06 | 3.9E-06 | 5.2E-05 | 1.3E-06 |
| Stray Light | 1.3E-06 | 1.1E-06 | 6.6E-07 | 4.2E-07 | 2.3E-07 | 1.0E-07 | 3.9E-09 | 9.6E-08 |
| Beam Size & Position | 1.4E-05 | 1.4E-05 | 1.4E-05 | 1.4E-05 | 1.4E-05 | 1.4E-05 | 1.4E-05 | 1.4E-05 |
| Inter-reflection | 9.9E-05 | 9.9E-05 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 |
| Obliquity | 7.6E-05 | 7.6E-05 | 7.6E-05 | 7.6E-05 | 7.6E-05 | 7.6E-05 | 7.6E-05 | 7.6E-05 |
| Polarization | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 |
| Source Drift & Fluctuation * | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A |
| Bandwidth | 1.6E-05 | 3.7E-07 | 1.0E-06 | 4.2E-06 | 2.0E-06 | 2.5E-06 | 3.4E-05 | 8.3E-07 |
| Filter Stability | 3.9E-04 | 3.4E-04 | 1.8E-04 | 1.8E-04 | 1.8E-04 | 1.8E-04 | 1.8E-04 | 1.8E-04 |
| Total Type B Uncertainty | 5.2E-04 | 5.1E-04 | 3.1E-04 | 3.0E-04 | 3.0E-04 | 3.2E-04 | 3.0E-04 | 3.0E-04 |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

*) “Source Drift & Fluctuation” is included in the Type A uncertainty.

Signature : 	Date :

Hiroshi Shitomi			2016-02-29

* Original data submitted:  2014-12-05

* Revision at the pre-Draft A stage: 2015-03-31

Appendix A.2 Measurement Results (2)

Please reproduce the following tables for each of the five filters measured at the completion of step 2 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: National Metrology Institute of Japan (NMIJ, AIST)

Filter Type: B

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 4.125E-01 | 6.083E-01 | 6.220E-01 | 6.099E-01 | 6.374E-01 | 5.769E-01 | 5.017E-01 | 4.532E-01 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Temperature | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 |
| Type A Uncertainty(b) | 1.9E-04 | 2.8E-04 | 1.7E-04 | 1.7E-04 | 1.7E-04 | 1.6E-04 | 1.4E-04 | 1.2E-04 |
| Type B Uncertainty(c) | 9.3E-04 | 2.3E-04 | 2.4E-04 | 2.4E-04 | 2.5E-04 | 2.5E-04 | 2.1E-04 | 1.8E-04 |
| Total Uncertainty(d) | 9.5E-04 | 3.6E-04 | 2.9E-04 | 2.9E-04 | 3.0E-04 | 3.0E-04 | 2.5E-04 | 2.2E-04 |
| Degrees of Freedom | 5773 | 26 | 84 | 85 | 86 | 116 | 110 | 87 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

Filter Type: B

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 1.0E-04 | 1.5E-04 | 1.6E-04 | 1.5E-04 | 1.6E-04 | 1.5E-04 | 1.3E-04 | 1.1E-04 |
| Temperature | 1.6E-05 | 2.8E-05 | 1.3E-05 | 2.9E-05 | 1.6E-05 | 1.1E-05 | 9.4E-06 | 1.3E-05 |
| Wavelength | 7.5E-04 | 5.9E-05 | 1.4E-05 | 7.7E-07 | 2.6E-05 | 9.8E-05 | 7.8E-05 | 2.6E-05 |
| Stray Light | 2.0E-05 | 2.0E-05 | 2.2E-05 | 2.0E-05 | 2.5E-05 | 1.3E-05 | 1.8E-06 | 1.1E-05 |
| Beam Size & Position | 3.0E-05 | 4.4E-05 | 4.5E-05 | 4.5E-05 | 4.7E-05 | 4.2E-05 | 3.7E-05 | 3.3E-05 |
| Inter-reflection | 4.5E-05 | 6.6E-05 | 6.8E-05 | 6.6E-05 | 6.9E-05 | 6.3E-05 | 5.4E-05 | 4.9E-05 |
| Obliquity | 3.4E-05 | 5.0E-05 | 5.1E-05 | 5.0E-05 | 5.3E-05 | 4.8E-05 | 4.2E-05 | 3.8E-05 |
| Polarization | 7.1E-05 | 1.0E-04 | 1.1E-04 | 1.0E-04 | 1.1E-04 | 9.9E-05 | 8.6E-05 | 7.8E-05 |
| Source Drift & Fluctuation * | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A |
| Bandwidth | 4.9E-04 | 3.9E-05 | 9.4E-06 | 5.0E-07 | 1.7E-05 | 6.4E-05 | 5.1E-05 | 1.7E-05 |
| Filter Stability | 1.9E-04 | 5.9E-05 | 1.1E-04 | 1.1E-04 | 1.1E-04 | 1.0E-04 | 9.0E-05 | 8.1E-05 |
| Total Type B Uncertainty | 9.3E-04 | 2.3E-04 | 2.4E-04 | 2.4E-04 | 2.5E-04 | 2.5E-04 | 2.1E-04 | 1.8E-04 |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

*) “Source Drift & Fluctuation” is included in the Type A uncertainty.

Signature : 	Date :

Hiroshi Shitomi			2016-02-29

* Original data submitted:  2014-12-05

Appendix A.2 Measurement Results (3)

Please reproduce the following tables for each of the five filters measured at the completion of step 2 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: National Metrology Institute of Japan (NMIJ, AIST)

Filter Type: C

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 2.244E-02 | 9.581E-02 | 9.263E-02 | 7.761E-02 | 1.621E-01 | 1.509E-01 | 1.033E-01 | 7.699E-02 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Temperature | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 |
| Type A Uncertainty(b) | 4.6E-05 | 1.1E-04 | 6.9E-05 | 5.8E-05 | 1.2E-04 | 1.1E-04 | 7.7E-05 | 6.3E-05 |
| Type B Uncertainty(c) | 2.4E-04 | 9.1E-05 | 7.2E-05 | 5.4E-05 | 2.2E-04 | 1.1E-04 | 6.2E-05 | 5.0E-05 |
| Total Uncertainty(d) | 2.5E-04 | 1.4E-04 | 1.0E-04 | 7.9E-05 | 2.5E-04 | 1.6E-04 | 9.9E-05 | 8.0E-05 |
| Degrees of Freedom | 7673 | 25 | 39 | 32 | 174 | 34 | 24 | 24 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

Filter Type: C

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 3.9E-06 | 1.7E-05 | 1.6E-05 | 1.3E-05 | 2.8E-05 | 2.6E-05 | 1.8E-05 | 1.3E-05 |
| Temperature | 3.9E-06 | 2.3E-05 | 1.1E-05 | 2.1E-05 | 1.9E-05 | 1.1E-05 | 7.1E-06 | 7.6E-06 |
| Wavelength | 2.0E-04 | 6.4E-05 | 4.3E-05 | 2.3E-05 | 1.7E-04 | 6.1E-05 | 2.4E-05 | 2.3E-05 |
| Stray Light | 1.8E-05 | 2.9E-06 | 3.6E-06 | 6.6E-06 | 1.0E-05 | 8.1E-06 | 1.4E-06 | 6.7E-06 |
| Beam Size & Position | 1.5E-06 | 6.4E-06 | 6.2E-06 | 5.2E-06 | 1.1E-05 | 1.0E-05 | 6.9E-06 | 5.1E-06 |
| Inter-reflection | 2.4E-06 | 1.0E-05 | 1.0E-05 | 8.4E-06 | 1.8E-05 | 1.6E-05 | 1.1E-05 | 8.4E-06 |
| Obliquity | 1.9E-06 | 7.9E-06 | 7.7E-06 | 6.4E-06 | 1.3E-05 | 1.2E-05 | 8.6E-06 | 6.4E-06 |
| Polarization | 5.0E-06 | 2.1E-05 | 2.1E-05 | 1.7E-05 | 3.6E-05 | 3.4E-05 | 2.3E-05 | 1.7E-05 |
| Source Drift & Fluctuation * | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A |
| Bandwidth | 1.3E-04 | 4.2E-05 | 2.8E-05 | 1.5E-05 | 1.1E-04 | 4.0E-05 | 1.6E-05 | 1.5E-05 |
| Filter Stability | 3.6E-05 | 3.1E-05 | 3.9E-05 | 3.3E-05 | 6.8E-05 | 6.3E-05 | 4.3E-05 | 3.2E-05 |
| Total Type B Uncertainty | 2.4E-04 | 9.1E-05 | 7.2E-05 | 5.4E-05 | 2.2E-04 | 1.1E-04 | 6.2E-05 | 5.0E-05 |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

*) “Source Drift & Fluctuation” is included in the Type A uncertainty.

Signature : 	Date :

Hiroshi Shitomi			2016-02-29

* Original data submitted:  2014-12-05

Appendix A.2 Measurement Results (4)

Please reproduce the following tables for each of the five filters measured at the completion of step 2 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: National Metrology Institute of Japan (NMIJ, AIST)

Filter Type: D

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 4.044E-04 | 5.454E-03 | 8.842E-03 | 8.747E-03 | 2.763E-02 | 3.346E-02 | 2.354E-02 | 1.740E-02 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Temperature | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 |
| Type A Uncertainty(b) | 8.8E-06 | 2.4E-05 | 3.8E-05 | 3.2E-05 | 8.2E-05 | 1.0E-04 | 7.0E-05 | 3.6E-05 |
| Type B Uncertainty(c) | 1.6E-05 | 3.4E-05 | 2.2E-05 | 1.7E-05 | 6.6E-05 | 5.6E-05 | 6.6E-05 | 2.9E-05 |
| Total Uncertainty(d) | 1.8E-05 | 4.2E-05 | 4.4E-05 | 3.6E-05 | 1.1E-04 | 1.1E-04 | 9.6E-05 | 4.6E-05 |
| Degrees of Freedom | 162 | 77 | 16 | 14 | 24 | 16 | 31 | 25 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

Filter Type: D

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 1.5E-07 | 2.0E-06 | 3.2E-06 | 3.2E-06 | 1.0E-05 | 1.2E-05 | 8.5E-06 | 6.3E-06 |
| Temperature | 1.5E-07 | 3.0E-06 | 2.1E-06 | 4.6E-06 | 6.5E-06 | 4.3E-06 | 2.7E-06 | 2.8E-06 |
| Wavelength | 1.0E-05 | 2.0E-05 | 2.8E-06 | 5.3E-06 | 4.1E-05 | 1.0E-05 | 4.5E-05 | 3.5E-06 |
| Stray Light | 4.9E-06 | 3.9E-06 | 3.2E-06 | 3.2E-06 | 5.7E-07 | 1.7E-06 | 2.5E-07 | 1.5E-06 |
| Beam Size & Position | 1.0E-07 | 1.4E-06 | 2.3E-06 | 2.2E-06 | 7.0E-06 | 8.5E-06 | 6.0E-06 | 4.4E-06 |
| Inter-reflection | 4.4E-08 | 5.9E-07 | 9.6E-07 | 9.5E-07 | 3.0E-06 | 3.6E-06 | 2.6E-06 | 1.9E-06 |
| Obliquity | 3.3E-08 | 4.5E-07 | 7.3E-07 | 7.2E-07 | 2.3E-06 | 2.8E-06 | 1.9E-06 | 1.4E-06 |
| Polarization | 4.1E-07 | 5.5E-06 | 9.0E-06 | 8.9E-06 | 2.8E-05 | 3.4E-05 | 2.4E-05 | 1.8E-05 |
| Source Drift & Fluctuation * | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A |
| Bandwidth | 6.7E-06 | 1.3E-05 | 1.8E-06 | 3.5E-06 | 2.6E-05 | 6.6E-06 | 2.9E-05 | 2.3E-06 |
| Filter Stability | 8.8E-06 | 2.3E-05 | 1.9E-05 | 1.0E-05 | 3.3E-05 | 4.0E-05 | 2.8E-05 | 2.1E-05 |
| Total Type B Uncertainty | 1.6E-05 | 3.4E-05 | 2.2E-05 | 1.7E-05 | 6.6E-05 | 5.6E-05 | 6.6E-05 | 2.9E-05 |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

*) “Source Drift & Fluctuation” is included in the Type A uncertainty.

Signature : 	Date :

Hiroshi Shitomi			2016-02-29

* Original data submitted:  2014-12-05

Appendix A.2 Measurement Results (5)

Red: Revised at the pre-Draft A stage for consistency check

Blue: Revised at the Draft A stage to hide identification of the DUT filter set

Please reproduce the following tables for each of the five filters measured at the completion of step 2 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: National Metrology Institute of Japan (NMIJ, AIST)

Filter Type: E

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 1.247E-05 | 3.257E-04 | 9.073E-04 | 9.883E-04 | 4.956E-03 | 9.525E-03 | 8.370E-03 | 6.953E-03 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Temperature | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 |
| Type A Uncertainty(b) | 6.2E-06 | 7.4E-06 | 5.8E-06 | 6.4E-06 | 5.2E-05 | 4.4E-05 | 3.8E-05 | 8.7E-05 |
| Type B Uncertainty(c) | 2.3E-06 | 2.9E-06 | 4.9E-06 | 5.4E-06 | 2.8E-05 | 4.8E-05 | 4.8E-05 | 3.4E-05 |
| Total Uncertainty(d) | 6.6E-06 | 7.9E-06 | 7.7E-06 | 8.3E-06 | 5.9E-05 | 6.5E-05 | 6.2E-05 | 9.4E-05 |
| Degrees of Freedom | 12 | 12 | 27 | 27 | 15 | 44 | 60 | 12 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

Filter Type: E

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 4.3E-09 | 1.1E-07 | 3.2E-07 | 3.4E-07 | 1.7E-06 | 3.3E-06 | 2.9E-06 | 2.4E-06 |
| Temperature | 6.7E-09 | 2.8E-07 | 3.2E-07 | 7.6E-07 | 1.7E-06 | 1.7E-06 | 1.2E-06 | 1.7E-06 |
| Wavelength | 1.3E-06 | 1.7E-06 | 2.6E-07 | 4.9E-07 | 1.0E-05 | 2.2E-06 | 2.0E-05 | 2.3E-06 |
| Stray Light | 1.7E-06 | 1.6E-06 | 1.5E-06 | 1.5E-06 | 6.6E-07 | 2.7E-07 | 1.4E-08 | 2.8E-07 |
| Beam Size & Position | 1.4E-09 | 3.5E-08 | 9.8E-08 | 1.1E-07 | 5.3E-07 | 1.0E-06 | 9.1E-07 | 7.3E-07 |
| Inter-reflection | 4.0E-09 | 1.1E-07 | 2.9E-07 | 3.2E-07 | 1.6E-06 | 3.1E-06 | 2.7E-06 | 2.2E-06 |
| Obliquity | 1.0E-09 | 2.7E-08 | 7.5E-08 | 8.2E-08 | 4.1E-07 | 7.9E-07 | 6.9E-07 | 5.6E-07 |
| Polarization | 3.8E-08 | 9.9E-07 | 2.8E-06 | 3.0E-06 | 1.5E-05 | 2.9E-05 | 2.6E-05 | 2.1E-05 |
| Source Drift & Fluctuation * | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A |
| Bandwidth | 8.7E-07 | 1.1E-06 | 1.7E-07 | 3.2E-07 | 6.7E-06 | 1.4E-06 | 1.3E-05 | 1.5E-06 |
| Filter Stability | 3.9E-07 | 9.0E-07 | 3.8E-06 | 4.1E-06 | 2.0E-05 | 3.8E-05 | 3.3E-05 | 2.7E-05 |
| Total Type B Uncertainty | 2.3E-06 | 2.9E-06 | 4.9E-06 | 5.4E-06 | 2.8E-05 | 4.8E-05 | 4.8E-05 | 3.4E-05 |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

*) “Source Drift & Fluctuation” is included in the Type A uncertainty.

Signature : 	Date :

Hiroshi Shitomi			2016-02-29

* Original data submitted:  2014-12-05

* Revision at the pre-Draft A stage: 2015-03-31