Description of the measurement facility and measurement method

Laboratory: PTB AG 4.24 Reflection and Transmission

General remark:

The spectral transmittance of the designated filter set was, for each step of the comparison, determined in two different set-ups of the German national standard facility for Regular Transmittance as described below. In the first set-up, called NaNoRef, a parallel wide diameter measurement beam according to the technical protocol was realized. The second set-up (NaNoTrans) uses a small spot slightly converging beam, as it is applied in the majority of commercial spectrophotometers. The main difference in both set-ups therefore is the geometry of the measurement beam. Effects of the converging beam are calculated for the second set-up and a correction is applied on transmittance. Where appropriate, the different effects of uncertainty are determined individually for the different set-ups. The mean transmittance scale results from the average of the results of both set-ups and results are given in the following five tables for each step of the process. For sake of simplicity the resulting uncertainties are calculated such that they account for the maximum observed effects for both set-ups.

Laboratory: PTB AG 4.24 Reflection and Transmission, Set-up NaNoRef

Indicate whether this table relates to Step 2 [] or Step 4 []

The following table and other information relates to PTB measurements during Step 2 and Step 4.

Table A-1 Details of Measurement Setup

| Make and Type of Spectrophotometer | Set-up NaNoRef Double monochromator system with quasi parallel measurement beam |
| --- | --- |
| Additional Stray Light Rejection | by double monochromator, higher order rejection filters, shielding baffles in set-up |
| Source Drift Monitoring | None, residual linear drift elimination by measuring technique |
| Source | Temperature controlled Halogen lamp |
| Detector | Photomultiplier with diffuser, Si-Photodiode with integrating sphere, InGaAs diode with integrating sphere |
| Temperature(a) | Nominal temperature 23°C, permanent temperature monitoring by calibrated PT-100 system, maximum observed temperature range covering all measurements 22,5°C – 23,0°C, max. drift during individual measuring cycles ≤ 0,2°C, in most cases smaller, correction for individual measurements made |
| Humidity | Not exceeding 60% |
| Beam Size | Quasi circular beam, full beam h 17,4 mm x v 20 mm |
| Beam Collimation | Quasi parallel beam, full divergence h 0,25° v 1,8° |
| Measurement Sequence(b) | General sequence called cycle is H F1 F2 ...Fx..Fx ..F2 F1 H, denominating the 100% m. and individual filter sequence which is run in opposite order on the way back, dark currents measured at each step; sequence usually repeated 10 times at given wavelength, repeated 4 times |
| Bandwidth | 1 nm nominal, in few cases modified for intensity reasons |

Make and type of the spectrophotometer

A self-built reference spectrophotometer (NaNoRef) was used consisting of lamp section (1), single pass double monochromator (2; S&I Spectroscopy &Imaging GmbH, TriVista-System focal length f = 0,5 m), beam forming chamber (4 upstream), and detection chamber (see 4 in figure 1).

multiplier

integr. sphere

detect. arm position

ART = 180°

Sample rotation

AGR = 45°

detection arm without weight for balancing

beam preparation with

f = 250 mm / plano mirror set-up

meas. diaphragm 15 mm

sep. shield, diameter 27 mm

coll. mirror f = 300 mm

diaphragm at entrance of chamber: maximum

Fig.1	Sketch of the set-up NaNoRef set-up for the measurement of

regular transmittance, on top: overview of complete system,

below: set-up details in sample chamber

The beam leaving the monochromator unit was formed by means of reflective mirrors configured in z-shape to be parallel close to the desired diameter.

## Characterisation of the spectrophotometer

In this study a temperature controlled tungsten halogen lamp was used for the transmittance measurements, a HeNe-Laser 543 nm for alignment, and for wavelength calibration the laser and the lines of a mercury spectral lamp. Depending on the wavelength, different diffraction gratings were chosen for the double monochromator and the width of the three monochromator slits was set to such values that the spectral bandwidth as listed in table A-2 was resulting. Also listed in this table are cut off filters which were used for higher order elimination.

The beam was formed to be parallel having a slightly elliptical shape at the sample with a diameters of h 17,4 mm and v 20 mm as horizontal and vertical dimensions. The residual full divergence was in horizontal direction h = 0,25° and for the vertical direction v = 1,8°. Normal incidence was used.

Measurements were taken at 380 nm, 400 nm, 500 nm, 600 nm, 700 nm, 800 nm, 900 nm, and 1000 nm with the gratings and filters according to table 1.

In the spectral range up to 800 nm a photomultiplier detector (EMI 9558) and Si detector (Hamamatsu 1010 BR) with integrating sphere, above 800 nm a cooled InGaAs detector (Polytec IGA-050-TE2-H/SEL) was used.

## Description of the measuring technique

The samples were fed into a computer controlled sample-holder which is able to carry six filters and has a seventh position for a reference filter or for the determination of the 100% signal, respectively. For the current measurements one cycle at a specific wavelength consists of the reading of the 100% value, the reading for up to three filters at the measurement position and the reverse process. The mean of both values for each of different position is calculated in order to cope for a drift of the lamp radiation. A complete measurement consists of up to 10 individual cycles. The samples were not cleaned.

## Description of the calibration laboratory conditions

The temperature during the measurements was controlled by the climatic system of the laboratory, able to stabilize the temperature in the laboratory to be 23 °C  0,5 °C. Measurements of the temperature inside the measurement chamber were performed for each cycle. The observed maximum overall variation within a measurement consisting of 10 cycles was less than 0,2°C in all cases, most often considerably smaller. A correction of the filter transmittance was performed for each individual measurement by using the mean temperature observed within 10 cycles. The relative humidity was also controlled by the climatic system to lie below 60%.

Table A-2	Spectral bandwidth of the spectrophotometer and cut off filters

used for the measurements

|  | Wavelength | Geometrical slitwidth | Spectral bandwidth | Grating | Cut off filter |
| --- | --- | --- | --- | --- | --- |
|  | nm | mm | nom. / used nm | lines/mm |  |
|  | 380 | 1,90 | 1,0 / 1,0 | 1800 | none |
|  | 400 | 1,93 | 1,0 / 1,0 | 1800 | KV389 |
|  | 500 | 1,98 | 1,0 / 1,0 | 1800 | KV389 |
|  | 600 | 1,56 | 1,0 /0,8 | 1800 | KV389 |
|  | 700 | 1,77 | 1,0 /0,9 | 1800 | KV389 |
|  | 800 | 2,06 | 1,0 / 1,0 | 1800 | RG780 |
|  | 900 | 2,02 | 1,0 / 1,5 | 1200 | RG780 |
|  | 1000 | 1,37 | 1,0 / 1,0 | 1200 | RG780 |

Table A-3: Correlation coefficients for individual contributions of uncertainty between

Step 2 and 4

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient(b) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity |  |  |  |  |  |  |  |  | 1 |
| Temperature1 |  |  |  |  |  |  |  |  | 0,5 |
| Wavelength |  |  |  |  |  |  |  |  | 0 |
| Stray Light |  |  |  |  |  |  |  |  | 1 |
| Beam Size & Position |  |  |  |  |  |  |  |  | 0 |
| Inter-reflection |  |  |  |  |  |  |  |  | 1 |
| Obliquity |  |  |  |  |  |  |  |  | 1 |
| Polarization |  |  |  |  |  |  |  |  | 1 |
| Bandwidth |  |  |  |  |  |  |  |  | 0 |
| Resolution |  |  |  |  |  |  |  |  | 0 |
| Prismatic effect2 |  |  |  |  |  |  |  |  | 1 |
| Total Type B Uncertainty |  |  |  |  |  |  |  |  |  |
| Degrees of Freedom |  |  |  |  |  |  |  |  |  |

Remark to Table A-3ii Type B Uncertainty Budget(a) for individual filters

The last column for correlation coefficient is not added to the corresponding tables following which contain the details for the estimation of uncertainty contributions. Instead the entries listed in the table may be used for all tables concerning the Type B contributions.

1For filter A__ no temperature correction has been performed.

2For Filter A__ a distinct prismatic effect was observed and a corresponding uncertainty component was derived with a correlation coefficient of 1.

Description of measurement facility and measurement method

Laboratory: PTB AG 4.24 Reflection and Transmission, Set-up NaNoTrans

Indicate whether this table relates to Step 2 [] or Step 4 [x]

The following table and other information relates to PTB measurements during Step 2 and Step 4.

Table A-1.2 Details of Measurement Setup

| Make and Type of Spectrophotometer | Set-up NaNoTrans Double monochromator system with slightly convergent beam |
| --- | --- |
| Additional Stray Light Rejection | by double monochromator, higher order rejection filters, shielding baffles in set-up |
| Source Drift Monitoring | None, linear drift elimination by measuring technique |
| Source | Temperature controlled Halogen lamp |
| Detector | Photomultiplier with diffuser, Si-Photodiode with integrating sphere, InGaAs diode with integrating sphere |
| Temperature(a) | Nominal temperature 23°C, permanent temperature monitoring by calibrated PT-100 system, maximum observed temperature range covering all measurements 22,72°C – 23,3°C, drift during individual measuring cycles ≤ 0,2°C, in most cases smaller, correction for individual measurements made |
| Humidity | Not exceeding 60% |
| Beam Size | Rectangular beam h 2,8 mm x v 12 mm |
| Beam Collimation | Slightly convergent beam, full divergence h 2,9° v 5,2° |
| Measurement Sequence(b) | General sequence called cycle is H F1 F2 ...Fx..Fx ..F2 F1 H, denominating the 100% m. and individual filter sequence which is run in opposite order on the way back, dark currents measured at each step; sequence usually repeated 10 times at given wavelength, repeated 4 times |
| Bandwidth | 1 nm nominal, in few cases modified for intensity reasons |

Make and type of the spectrophotometer

A self-built reference spectrophotometer (NaNoTrans) was used consisting of lamp section (1), single pass double monochromator (2; S&I Spectroscopy &Imaging GmbH, TriVista-System focal length f = 0,5 m), and set-up NaNoTrans (see 3 in figure 1.2).

|  | Fig. A1.2 Sketch of the set-up NaNoTrans on top: system overview left: details of set-up in part 3 - monochromator exit slit 2 sample plane, enlarged image of exit slit - photomultiplier - integrating sphere equipped with diode type detectors - optional beam condenser optics - monitor diode |
| --- | --- |

The set-up NaNoTrans is a permanent set-up for the measurement of transmission. It uses a convergent beam similar to the vast majority of commercial spectrophotometers. The beam cross section at the sample has a rectangular shape and is for the current measurements approx. 2,8 mm x 12 mm.

## Characterisation of the spectrophotometer

In this study a temperature controlled tungsten halogen lamp was used for the transmittance measurements, a HeNe-Laser 543 nm for alignment, and for the wavelength calibration the laser and the lines of a mercury spectral lamp. Depending on the wavelength, different diffraction gratings were chosen for the double monochromator and the width of the three monochromator slits was set to such values that the spectral bandwidth as listed in table 2 was resulting. Also listed in this table are cut off filters which were used for higher order elimination.

The beam was formed to be slightly convergent having a rectangular shape at the sample with dimension and divergence as given above.

Measurements were taken at 380 nm, 400 nm, 500 nm, 600 nm, 700 nm, 800 nm, 900 nm, and 1000 nm with the gratings according to table 2.2.

In the spectral range up to 800 nm a photomultiplier detector (EMI 9558) and Si detector (Hamamatsu 1010 BR) with integrating sphere, above 800 nm a cooled InGaAs detector (Polytec IGA-050-TE2-H/SEL) was used.

## Description of the measuring technique

The samples were fed into a computer controlled sample-holder which is able to carry up to six filters and has a seventh position for a reference filter or for the determination of the 100% signal, respectively. For the current measurements one cycle at a specific wavelength consists of the reading of the 100% value, the reading for up to six filters and the reverse process. The mean of both values for each of different positions is calculated in order to cope for a drift of the lamp radiation. A complete measurement consists of up to 10 individual cycles. The samples were not cleaned.

## Description of the calibration laboratory conditions

The temperature during the measurements was controlled by the climatic system of the laboratory, able to stabilize the temperature in the laboratory to be 23 °C  0,5 °C. Measurements of the temperature inside the measurement chamber were performed for each cycle. The observed maximum overall variation within a measurement consisting of 10 cycles was less than 0,2°C in all cases, most often considerably smaller. A correction of the filter transmittance was performed for each individual measurement by using the mean temperature observed within 10 cycles. The relative humidity was also controlled by the climatic system to lie below 60%.

Table A2.2	Spectral bandwidth of the spectrophotometer and cut off filters

used for the measurements

|  | Wavelength | Geometrical slitwidth | Spectral bandwidth nom. / used | Grating | Cut off filter |
| --- | --- | --- | --- | --- | --- |
|  | nm | mm | nm | lines/mm |  |
|  | 380 | 1,90 | 1,0 / 1,0 | 1800 | none |
|  | 400 | 1,93 | 1,0 | 1800 | KV389 |
|  | 500 | 1,98 | 1,0 | 1800 | KV389 |
|  | 600 | 1,77 | 1,0 / 0,9 | 1800 | KV389 |
|  | 700 | 1,97 | 1,0 / 1,0 | 1800 | KV389 |
|  | 800 | 2,06 | 1,0 / 1,0 | 1800 | RG780 |
|  | 900 | 2,02 | 1,0 / 1,5 | 1200 | RG780 |
|  | 1000 | 1,37 | 1,0 / 1,0 | 1200 | RG780 |