Appendix A.1 Description of measurement facility and measurement method

Laboratory: MKEH

Indicate whether this table relates to Step 2 [   X   ] or Step 4 [      ]

If the measurement setup has not changed from Step 2, check here [     ] and the following table does not need to be completed. Otherwise, please fill out the whole table.

Table A-1 Details of Measurement Setup

| Make and Type of Spectrophotometer | OMH custom design with double grating monochromator |
| --- | --- |
| Additional Stray Light Rejection | PMT with integrating sphere |
| Source Drift Monitoring | Monitor detector, test measurement |
| Source | Incandescent band lamp |
| Detector | Si photodetector, PMT tube |
| Temperature(a) | 22-23 °C |
| Humidity | 35-55% |
| Beam Size | 8 mm diameter |
| Beam Collimation | 3,2 mrad |
| Measurement Sequence(b) | 3-10/filter, filter orientation was not changed |
| Bandwidth | 1 nm |

(a) i.e. describe method of temperature monitoring of filters and range of temperatures  (b) i.e. describe number of measurements and whether filter orientation with respect to beam changes between measurements

Description of measuring technique (please include a diagram)

If any damage, contamination or cleaning of the filters was carried out, please give details:

Two corners are slightly damaged in the filter holder in case of filters: A24, B26, E25.

Signature :	Date :	2014-11-03

Appendix A.2 Measurement Results

Please reproduce the following tables for each of the five filters measured at the completion of step 2 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory:               MKEH               Filter Identifier: A--

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0,9145 | 0,9156 | 0,9189 | 0,9211 | 0,9227 | 0,9240 | 0,9244 | 0,9248 |
| Number of Measurements | 39 | 43 | 29 | 29 | 35 | 31 | 1 | 1 |
| Temperature | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C |
| Type A Uncertainty(b) | 0,0005 | 0,0004 | 0,0006 | 0,0005 | 0,0005 | 0,0006 | 0,0009 | 0,0011 |
| Type B Uncertainty(c) | 0,0009 | 0,0008 | 0,0008 | 0,0008 | 0,0007 | 0,0008 | 0,0008 | 0,0009 |
| Total Uncertainty(d) | 0,0010 | 0,0009 | 0,0010 | 0,0009 | 0,0009 | 0,0010 | 0,0012 | 0,0014 |
| Degrees of Freedom | 1049 | 836 | 216 | 239 | 396 | 231 | 4 | 3 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 0,00029 | 0,00029 | 0,00029 | 0,00029 | 0,00029 | 0,00029 | 0,00029 | 0,00029 |
| Temperature | 0,00005 | 0,00005 | 0,00005 | 0,00005 | 0,00005 | 0,00005 | 0,00005 | 0,00005 |
| Wavelength | 0,00003 | 0,00003 | 0,00003 | 0,00003 | 0,00003 | 0,00003 | 0,00003 | 0,00003 |
| Stray Light | 0,00004 | 0,00004 | 0,00004 | 0,00004 | 0,00004 | 0,00004 | 0,00004 | 0,00004 |
| Beam Size & Position | 0,00085 | 0,00074 | 0,00067 | 0,00065 | 0,00064 | 0,00072 | 0,00076 | 0,00082 |
| Inter-reflection | 0,00009 | 0,00009 | 0,00009 | 0,00009 | 0,00009 | 0,00009 | 0,00009 | 0,00009 |
| Obliquity | 0,00006 | 0,00006 | 0,00006 | 0,00006 | 0,00006 | 0,00006 | 0,00006 | 0,00006 |
| Polarization | 0,00001 | 0,00001 | 0,00001 | 0,00001 | 0,00001 | 0,00001 | 0,00001 | 0,00001 |
| Source Drift & Fluctuation | 0,00018 | 0,00018 | 0,00018 | 0,00018 | 0,00018 | 0,00018 | 0,00018 | 0,00018 |
| Bandwidth | 0,00005 | 0,00003 | 0,00002 | 0,00002 | 0,00002 | 0,00002 | 0,00002 | 0,00002 |
| Other(b) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 0,00093 | 0,00083 | 0,00076 | 0,00075 | 0,00074 | 0,00081 | 0,00084 | 0,00090 |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : 	Date :	2014-11-03

Laboratory:               MKEH               Filter Identifier: B--

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0,4088 | 0,6085 | 0,6228 | 0,6111 | 0,6390 | 0,5786 | 0,5027 | 0,4536 |
| Number of Measurements | 66 | 71 | 45 | 45 | 51 | 44 | 2 | 1 |
| Temperature | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C |
| Type A Uncertainty(b) | 0,00014 | 0,00007 | 0,00003 | 0,00008 | 0,00006 | 0,00009 | 0,00020 | 0,00020 |
| Type B Uncertainty(c) | 0,00139 | 0,00045 | 0,00042 | 0,00041 | 0,00041 | 0,00039 | 0,00036 | 0,00037 |
| Total Uncertainty(d) | 0,00140 | 0,00046 | 0,00042 | 0,00042 | 0,00041 | 0,00040 | 0,00041 | 0,00042 |
| Degrees of Freedom | 671229 | 150059 | 1115501 | 41755 | 104103 | 21787 | 35 | 19 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 0,000191 | 0,000192 | 0,000196 | 0,000193 | 0,000201 | 0,000182 | 0,000158 | 0,000212 |
| Temperature | 0,000047 | 0,000102 | 0,000040 | 0,000105 | 0,000055 | 0,000026 | 0,000010 | 0,000033 |
| Wavelength | 0,001349 | 0,000107 | 0,000043 | 0,000022 | 0,000050 | 0,000058 | 0,000050 | 0,000045 |
| Stray Light | 0,000016 | 0,000024 | 0,000025 | 0,000024 | 0,000026 | 0,000023 | 0,000020 | 0,000018 |
| Beam Size & Position | 0,000264 | 0,000352 | 0,000330 | 0,000312 | 0,000313 | 0,000315 | 0,000291 | 0,000275 |
| Inter-reflection | 0,000041 | 0,000061 | 0,000062 | 0,000061 | 0,000064 | 0,000058 | 0,000050 | 0,000045 |
| Obliquity | 0,000029 | 0,000043 | 0,000044 | 0,000043 | 0,000045 | 0,000041 | 0,000035 | 0,000032 |
| Polarization | 0,000004 | 0,000006 | 0,000006 | 0,000006 | 0,000006 | 0,000006 | 0,000005 | 0,000005 |
| Source Drift & Fluctuation | 0,000082 | 0,000122 | 0,000125 | 0,000122 | 0,000128 | 0,000116 | 0,000101 | 0,000091 |
| Bandwidth | 0,000020 | 0,000018 | 0,000012 | 0,000012 | 0,000013 | 0,000012 | 0,000010 | 0,000009 |
| Other(b) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 0,001392 | 0,000451 | 0,000416 | 0,000409 | 0,000409 | 0,000394 | 0,000356 | 0,000368 |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : 	Date :	2014-11-03

Laboratory:               MKEH               Filter Identifier: C--

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0,0209 | 0,0939 | 0,0907 | 0,0759 | 0,1592 | 0,1481 | 0,1011 | 0,0751 |
| Number of Measurements | 61 | 65 | 36 | 36 | 37 | 34 | 2 | 3 |
| Temperature | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C |
| Type A Uncertainty(b) | 0,000044 | 0,000008 | 0,000007 | 0,000018 | 0,000030 | 0,000013 | 0,000040 | 0,000050 |
| Type B Uncertainty(c) | 0,000439 | 0,000167 | 0,000103 | 0,000112 | 0,000224 | 0,000168 | 0,000182 | 0,000153 |
| Total Uncertainty(d) | 0,000441 | 0,000168 | 0,000104 | 0,000113 | 0,000226 | 0,000168 | 0,000186 | 0,000161 |
| Degrees of Freedom | 597797 | 12785265 | 2219533 | 56247 | 112544 | 894416 | 940 | 322 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 0,000021 | 0,000069 | 0,000066 | 0,000055 | 0,000096 | 0,000089 | 0,000061 | 0,000055 |
| Temperature | 0,000011 | 0,000084 | 0,000034 | 0,000077 | 0,000064 | 0,000025 | 0,000007 | 0,000019 |
| Wavelength | 0,000438 | 0,000103 | 0,000034 | 0,000028 | 0,000156 | 0,000075 | 0,000152 | 0,000128 |
| Stray Light | 0,000002 | 0,000009 | 0,000009 | 0,000008 | 0,000016 | 0,000015 | 0,000010 | 0,000008 |
| Beam Size & Position | 0,000018 | 0,000071 | 0,000059 | 0,000048 | 0,000104 | 0,000111 | 0,000075 | 0,000058 |
| Inter-reflection | 0,000002 | 0,000009 | 0,000009 | 0,000008 | 0,000016 | 0,000015 | 0,000010 | 0,000008 |
| Obliquity | 0,000001 | 0,000007 | 0,000006 | 0,000005 | 0,000011 | 0,000010 | 0,000007 | 0,000005 |
| Polarization | 0,000000 | 0,000001 | 0,000001 | 0,000001 | 0,000002 | 0,000001 | 0,000001 | 0,000001 |
| Source Drift & Fluctuation | 0,000004 | 0,000019 | 0,000018 | 0,000015 | 0,000032 | 0,000030 | 0,000020 | 0,000015 |
| Bandwidth | 0,000001 | 0,000003 | 0,000002 | 0,000002 | 0,000003 | 0,000003 | 0,000002 | 0,000002 |
| Other(b) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 0,000439 | 0,000167 | 0,000103 | 0,000112 | 0,000224 | 0,000168 | 0,000182 | 0,000153 |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : 	Date :	2014-11-03

Laboratory:               MKEH               Filter Identifier: D--

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0,00035 | 0,00515 | 0,00843 | 0,00835 | 0,02658 | 0,03230 | 0,02248 | 0,01656 |
| Number of Measurements | 90 | 47 | 9 | 9 | 9 | 120 | 1 | 1 |
| Temperature | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C |
| Type A Uncertainty(b) | 0,0000089 | 0,0000170 | 0,0000035 | 0,0000047 | 0,0000082 | 0,0000692 | 0,0000400 | 0,0000400 |
| Type B Uncertainty(c) | 0,0000061 | 0,0000223 | 0,0000157 | 0,0000213 | 0,0000604 | 0,0000515 | 0,0000404 | 0,0000292 |
| Total Uncertainty(d) | 0,0000108 | 0,0000280 | 0,0000161 | 0,0000218 | 0,0000609 | 0,0000863 | 0,0000569 | 0,0000496 |
| Degrees of Freedom | 196 | 347 | 4037 | 4364 | 27143 | 290 | 4 | 2 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 0,0000006 | 0,0000058 | 0,0000095 | 0,0000094 | 0,0000228 | 0,0000278 | 0,0000222 | 0,0000164 |
| Temperature | 0,0000004 | 0,0000105 | 0,0000064 | 0,0000162 | 0,0000216 | 0,0000102 | 0,0000026 | 0,0000068 |
| Wavelength | 0,0000060 | 0,0000178 | 0,0000070 | 0,0000070 | 0,0000463 | 0,0000297 | 0,0000252 | 0,0000152 |
| Stray Light | 0,0000006 | 0,0000005 | 0,0000008 | 0,0000008 | 0,0000027 | 0,0000032 | 0,0000022 | 0,0000017 |
| Beam Size & Position | 0,0000004 | 0,0000055 | 0,0000078 | 0,0000072 | 0,0000215 | 0,0000288 | 0,0000216 | 0,0000171 |
| Inter-reflection | 0,00000004 | 0,0000005 | 0,0000008 | 0,0000008 | 0,0000027 | 0,0000032 | 0,0000022 | 0,0000017 |
| Obliquity | 0,00000002 | 0,0000004 | 0,0000006 | 0,0000006 | 0,0000019 | 0,0000023 | 0,0000016 | 0,0000012 |
| Polarization | 0,00000000 | 0,0000001 | 0,0000001 | 0,0000001 | 0,0000003 | 0,0000003 | 0,0000002 | 0,0000002 |
| Source Drift & Fluctuation | 0,00000007 | 0,0000010 | 0,0000017 | 0,0000017 | 0,0000053 | 0,0000065 | 0,0000045 | 0,0000033 |
| Bandwidth | 0,00000002 | 0,0000002 | 0,0000002 | 0,0000002 | 0,0000005 | 0,0000006 | 0,0000004 | 0,0000003 |
| Other(b) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 0,0000061 | 0,0000223 | 0,0000157 | 0,0000213 | 0,0000604 | 0,0000515 | 0,0000404 | 0,0000292 |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : 	Date :	2014-11-03

Laboratory:               MKEH               Filter Identifier: E--

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0,000013 | 0,000344 | 0,000962 | 0,001031 | 0,005169 | 0,009895 | 0,008607 | 0,007274 |
| Number of Measurements | 90 | 47 | 9 | 9 | 9 | 120 | 2 | 1 |
| Temperature | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C | 22-23 °C |
| Type A Uncertainty(b) | 0,0000080 | 0,0000083 | 0,0000019 | 0,0000038 | 0,0000145 | 0,0000404 | 0,0000150 | 0,0000190 |
| Type B Uncertainty(c) | 0,0000007 | 0,0000068 | 0,0000068 | 0,0000040 | 0,0000150 | 0,0000159 | 0,0000144 | 0,0000140 |
| Total Uncertainty(d) | 0,0000081 | 0,0000107 | 0,0000071 | 0,0000055 | 0,0000209 | 0,0000434 | 0,0000208 | 0,0000236 |
| Degrees of Freedom | 92 | 131 | 1778 | 39 | 39 | 160 | 7 | 2 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 0,00000003 | 0,0000006 | 0,0000015 | 0,0000014 | 0,0000058 | 0,0000111 | 0,0000097 | 0,0000082 |
| Temperature | 0,00000002 | 0,0000011 | 0,0000011 | 0,0000029 | 0,0000063 | 0,0000043 | 0,0000013 | 0,0000058 |
| Wavelength | 0,00000036 | 0,0000066 | 0,0000064 | 0,0000020 | 0,0000113 | 0,0000030 | 0,0000052 | 0,0000051 |
| Stray Light | 0,00000065 | 0,0000005 | 0,0000005 | 0,0000005 | 0,0000005 | 0,0000010 | 0,0000009 | 0,0000007 |
| Beam Size & Position | 0,00000002 | 0,0000004 | 0,0000010 | 0,0000010 | 0,0000048 | 0,0000097 | 0,0000090 | 0,0000081 |
| Inter-reflection | 0,000000001 | 0,00000003 | 0,0000001 | 0,0000001 | 0,0000005 | 0,0000010 | 0,0000009 | 0,0000007 |
| Obliquity | 0,000000001 | 0,00000002 | 0,0000001 | 0,0000001 | 0,0000004 | 0,0000007 | 0,0000006 | 0,0000005 |
| Polarization | 0,0000000001 | 0,000000003 | 0,00000001 | 0,00000001 | 0,0000001 | 0,0000001 | 0,0000001 | 0,0000001 |
| Source Drift & Fluctuation | 0,000000003 | 0,0000001 | 0,0000002 | 0,0000002 | 0,0000010 | 0,0000020 | 0,0000017 | 0,0000015 |
| Bandwidth | 0,000000001 | 0,00000001 | 0,00000002 | 0,00000002 | 0,0000001 | 0,0000002 | 0,0000002 | 0,0000001 |
| Other(b) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 0,0000007 | 0,0000068 | 0,0000068 | 0,0000040 | 0,0000150 | 0,0000159 | 0,0000144 | 0,0000140 |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : 	Date :	2014-11-03