Appendix A.1 Description of measurement facility and measurement method

Laboratory: National Metrology Institute of Japan (NMIJ, AIST)

Indicate whether this table relates to Step 2 [ * ] or Step 4 [      ]

If the measurement setup has not changed from Step 2, check here [     ] and the following table does not need to be completed. Otherwise, please fill out the whole table.

Table A-1 Details of Measurement Setup

| Make and Type of Spectrophotometer | Spectrophotometer used for the spectral transmittance measurement is based on a JASCO double-grating subtractive monochromator (f=250 mm), which is designed for spectral measurements from 200 nm to 2500 nm. Four sources (400 W tungsten halogen lamp, 300 W Xe-arc lamp, 30 W deuterium lamp and low pressure Hg lamp) are installed in conjunction with five kinds of order-sorting filters, which are selectable according to the wavelength range. It is equipped with three gratings of 600 lines/mm with blazed wavelength at 300 nm, 500 nm and 1250 nm. Radiation exiting from the monochromator is collimated by a toroidal mirror and reduced in diameter by an iris placed at the beam exit. A DUT, which is placed on a transition table with 6-axis alignment stage, is irradiated with normal angle of incidence. Transmitted radiation is detected by a 3-elements reflection-type Si-trap detector comprised of one S1337-21 and two S1337-11 photodiodes, aiming at preventing the inter-reflection effect between the surface of the detector and the backside of the DUT. |
| --- | --- |
| Additional Stray Light Rejection | - Aperture screens coated with black material whose nominal reflectance is below 0.01 for all the wavelength points of this comparison - A set of order-sorting filters |
| Source Drift Monitoring | Monitored by applying a time-symmetrical sequence |
| Source | 400 W tungsten halogen lamp( > 700 nm) 300 W Xe arc lamp (380 nm ≤  ≤ 700 nm) |
| Detector | Si-trap detector |
| Temperature(a) | - Monitored by a thermometer places in the vicinity of the filters - Range of Temperatures during the Step 2: 22.8 ºC to 23.5 ºC |
| Humidity | 40 % ± 5 % |
| Beam Size | 10 mm in diameter |
| Beam Collimation | Nearly collimated |
| Measurement Sequence(b) | - Acquired twelve data in one measurement, that consist of six data acquired in a time-symmetrical sequence (Dark - Reference - DUT - DUT - Reference - Dark) with two orthogonal orientations - Number of measurement sets: 10 |
| Bandwidth | 1.0 nm |

(a) i.e. describe method of temperature monitoring of filters and range of temperatures  (b) i.e. describe number of measurements and whether filter orientation with respect to beam changes between measurements

Description of measuring technique (please include a diagram):

The spectral transmittance of five filters (A, B, C, D, and E) were measured relative to air as 100 % reference reading, which is obtained by taking the ratio of measured signals from a detector with and without the filter in the beam path. Each filter was placed on a motorized linear table with a 6-axis alignment stage. The filter was moved on the motorized linear stage perpendicular to the optical axis to change the filter position within and outside the beam path. The measurement sequences at each wavelength comprised the following data acquisition, Si,j: a filter in the beam, Ri,j: reference (air) in the beam and Di,j: dark (with a shutter closed), where i =1 or 2 and j= v (vertical) or h (horizontal), respectively. One measurement set consisted on the acquisition of in total 12 data, D1,v, R1,v, S1,v, S2,v, R1,v, D1,v, D1,h, R1,h, S1,h, S2,h, R1,h and D1,h. First, six data, at a fixed orientation along with the optical axis, were measured in a time-symmetrical sequence and then, another 6 data were measured at the orthogonalized orientation. Each data consisted of 20 readings from the 8.5-digits DMM that corresponded to the current-to-voltage converted photo-signals from the 3-elements reflection type Si-trap detector. The gain of the amplifier used for the current-to-voltage conversion was 107 ohm. By using the 12 data above, the spectral transmittance of the filter was obtained as follows,

T(λ)=12⋅(S1,v(λ)+S2,v(λ))−(D1,v(λ)+D2,v(λ))(R1,v(λ)+R2,v(λ))−(D1,v(λ)+D2,v(λ))+(S1,h(λ)+S2,h(λ))−(D1,h(λ)+D2,h(λ))(R1,h(λ)+R2,h(λ))−(D1,h(λ)+D2,h(λ))

where  is the wavelength (380 nm, 400 nm, 500 nm, 600 nm, 700 nm, 800 nm, 900 nm and 1000 nm). For each measurement set, the 5 filters were measured in sequence, starting with the filter A and ending with the filter E. In total 10 measurement sets were performed and their average was taken as a final result.

The schematic diagram of the measurement system is shown below.

If any damage, contamination or cleaning of the filters was carried out, please give details:

N/A

Signature :	Date :

Hiroshi Shitomi			2015-1-11