Appendix A.1 Description of measurement facility and measurement method

Laboratory: National Metrology Centre – Agency for Science, Technology and Research

(NMC/A*STAR)

Indicate whether this table relates to Step 2 [  √  ] or Step 4 [      ]

If the measurement setup has not changed from Step 2, check here [     ] and the following table does not need to be completed. Otherwise, please fill out the whole table.

Table A-1 Details of Measurement Setup

| Make and Type of Spectrophotometer | The reference spectrophotometer used for this comparison is a single beam spectrometer developed at NMC-A*STAR. A Czerny-Turner double monochromator with focal length 240mm and an effective aperture of f/4 is used in this system. Order-sorting filters are placed in front of the entrance slit to provide high order rejection. |
| --- | --- |
| Additional Stray Light Rejection | No |
| Source Drift Monitoring | Source drift effect was mostly removed by the measurement sequence |
| Source | 24V 250W tungsten halogen lamp |
| Detector | photomultiplier and Si-photodiode |
| Temperature(a) | 25 ˚C measured in the enclosed sample compartment |
| Humidity | 60 % ± 10 % |
| Beam Size | 20 mm diameter |
| Beam Collimation | 0.013 rad |
| Measurement Sequence(b) | Four repeat measurements without changing the filter orientation |
| Bandwidth | 1 nm |

(a) i.e. describe method of temperature monitoring of filters and range of temperatures  (b) i.e. describe number of measurements and whether filter orientation with respect to beam changes between measurements

Description of measuring technique (please include a diagram)

A 24V 250W tungsten halogen lamp is used as the radiation source. Spherical mirrors are used to focus the source radiation on the entrance slit and to collimate the exit beam from the monochromator. The collimation of the beam so obtained is about 0.013 rad. The detection unit comprises an averaging sphere and a 100 mm² silicon photodiode (or a photomultiplier) with a low noise transimpedance amplifier. The photomultiplier is used only for filters D, E at wavelengths 380 nm, 400 nm where the signal becomes too noisy for Si-photodiode. DC mode is used for the signal amplification and measurement. The solid angle of collection of the system subtends about 0.004 sr at the filter position.

The reference spectrometer at NMC-A*STAR

In this comparison a circular aperture with 20-mm diameter is used. The bandwidth of the

monochromator used is 1 nm. For filters of low optical densities, the transmittance against air is obtained directly by measuring the ratio of signals with and without the filter in the beam path. For filters with higher optical densities the step-down method is used to minimise uncertainty induced by detector nonlinearity. In this method, the relative transmittance of the test filter relative to a reference filter (of low optical density) is measured first. This is done by measuring the ratio of the two signals obtained with the two filters in the beam path separately. The transmittance value of the test filter against the air is then calculated by multiplying the relative transmittance value with the known (or measured) transmittance value of the reference filter. In this comparison filters A, B and C are measured directly. D is measured relative to C. Likewise, E is measured relative to D. The data acquisition is performed automatically following the sequence

Id Io Id Ii Id Io Id Ii Id Io Id Ii Id Io Id Ii Id Io Id

where Id is the dark current when shutter is closed, Io the signal current with the test filter moved

out of the beam and Ii the signal current with the test filter moved into the beam. Totally 19

current readings are obtained from one measurement. Each current reading is the average of 15

readings repeatedly taken within 10 seconds. There is a 15 seconds time interval between

adjacent current readings to eliminate the fatigue effect of the detector system.

The dark currents on both sides of the signal currents are averaged and subtracted from these

signal currents, yielding a sequence of 9 adjusted signal currents. The transmittance of the filter

is thus

= Ii / Io-ave

where Io-ave is the average of Io's before and after each Ii . Four values are obtained and the final

value is the average of these four.

If any damage, contamination or cleaning of the filters was carried out, please give details

Signature : Liu Yuanjie	Date :	19 Feb 2016

Appendix A.2 Measurement Results

Please reproduce the following tables for each of the five filters measured at the completion of step 2 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory:   NMC/A*STAR                           Filter Identifier: A

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0.913330 | 0.914527 | 0.917587 | 0.918829 | 0.919644 | 0.920514 | 0.920813 | 0.921620 |
| Number of Measurements | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 |
| Temperature | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 |
| Type A Uncertainty(b) | 0.000098 | 0.000162 | 0.000044 | 0.000131 | 0.000078 | 0.000159 | 0.000136 | 0.000184 |
| Type B Uncertainty(c) | 0.000299 | 0.000299 | 0.000295 | 0.000289 | 0.000286 | 0.000283 | 0.000284 | 0.000284 |
| Total Uncertainty(d) | 0.000315 | 0.000341 | 0.000298 | 0.000317 | 0.000296 | 0.000325 | 0.000315 | 0.000339 |
| Degrees of Freedom | 210 | 39 | 4140 | 68 | 410 | 35 | 57 | 23 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory:   NMC/A*STAR                           Filter Identifier: A

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 0.000030 | 0.000030 | 0.000030 | 0.000030 | 0.000030 | 0.000030 | 0.000030 | 0.000030 |
| Temperature | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Wavelength | 0.000010 | 0.000005 | 0.000003 | 0.000001 | 0.000005 | 0.000003 | 0.000003 | 0.000001 |
| Stray Light | 0.000003 | 0.000003 | 0.000003 | 0.000003 | 0.000003 | 0.000003 | 0.000003 | 0.000003 |
| Beam Size & Position | 0.000276 | 0.000276 | 0.000276 | 0.000276 | 0.000276 | 0.000276 | 0.000276 | 0.000276 |
| Filter Instability | 0.000110 | 0.000112 | 0.000099 | 0.000079 | 0.000068 | 0.000056 | 0.000061 | 0.000062 |
| Inter-reflection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Obliquity | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Polarization | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Source Drift & Fluctuation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Bandwidth | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Other(b) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Total Type B Uncertainty | 0.000299 | 0.000299 | 0.000295 | 0.000289 | 0.000286 | 0.000283 | 0.000284 | 0.000284 |
| Degrees of Freedom | infinite | infinite | infinite | infinite | infinite | infinite | infinite | infinite |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : Liu Yuanjie	Date :	19 Feb 2016

Appendix A.2 Measurement Results

Please reproduce the following tables for each of the five filters measured at the completion of step 2 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory:   NMC/A*STAR                           Filter Identifier: B

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0.406658 | 0.606022 | 0.620344 | 0.608278 | 0.635798 | 0.575124 | 0.499974 | 0.451556 |
| Number of Measurements | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 |
| Temperature | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 |
| Type A Uncertainty(b) | 0.000016 | 0.000093 | 0.000041 | 0.000077 | 0.000117 | 0.000084 | 0.000077 | 0.000166 |
| Type B Uncertainty(c) | 0.001304 | 0.000245 | 0.000236 | 0.000218 | 0.000228 | 0.000228 | 0.000186 | 0.000152 |
| Total Uncertainty(d) | 0.001304 | 0.000262 | 0.000240 | 0.000231 | 0.000256 | 0.000243 | 0.000201 | 0.000225 |
| Degrees of Freedom | 98212627 | 127 | 2362 | 160 | 46 | 137 | 94 | 7 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory:   NMC/A*STAR                           Filter Identifier: B

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 0.000030 | 0.000030 | 0.000030 | 0.000030 | 0.000030 | 0.000030 | 0.000030 | 0.000030 |
| Temperature | 0.000031 | 0.000053 | 0.000025 | 0.000055 | 0.000031 | 0.000021 | 0.000019 | 0.000026 |
| Wavelength | 0.001277 | 0.000078 | 0.000044 | 0.000070 | 0.000095 | 0.000088 | 0.000064 | 0.000033 |
| Stray Light | 0.000003 | 0.000003 | 0.000003 | 0.000003 | 0.000003 | 0.000003 | 0.000003 | 0.000003 |
| Beam Size & Position | 0.000123 | 0.000183 | 0.000186 | 0.000183 | 0.000192 | 0.000180 | 0.000150 | 0.000135 |
| Filter Instability | 0.000231 | 0.000129 | 0.000133 | 0.000072 | 0.000065 | 0.000102 | 0.000082 | 0.000048 |
| Inter-reflection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Obliquity | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Polarization | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Source Drift & Fluctuation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Bandwidth | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Other(b) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Total Type B Uncertainty | 0.001304 | 0.000245 | 0.000236 | 0.000218 | 0.000228 | 0.000228 | 0.000186 | 0.000152 |
| Degrees of Freedom | infinite | infinite | infinite | infinite | infinite | infinite | infinite | infinite |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : Liu Yuanjie	Date :	19 Feb 2016

Appendix A.2 Measurement Results

Please reproduce the following tables for each of the five filters measured at the completion of step 2 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory:   NMC/A*STAR                           Filter Identifier: C

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0.021442 | 0.094841 | 0.091554 | 0.076700 | 0.160524 | 0.149657 | 0.102531 | 0.076266 |
| Number of Measurements | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 |
| Temperature | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 |
| Type A Uncertainty(b) | 0.000014 | 0.000024 | 0.000007 | 0.000010 | 0.000013 | 0.000027 | 0.000016 | 0.000035 |
| Type B Uncertainty(c) | 0.000261 | 0.000124 | 0.000061 | 0.000061 | 0.000162 | 0.000082 | 0.000064 | 0.000046 |
| Total Uncertainty(d) | 0.000261 | 0.000126 | 0.000061 | 0.000061 | 0.000163 | 0.000086 | 0.000066 | 0.000058 |
| Degrees of Freedom | 272475 | 1444 | 14860 | 2692 | 45689 | 215 | 520 | 15 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory:   NMC/A*STAR                           Filter Identifier: C

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 0.000050 | 0.000030 | 0.000030 | 0.000030 | 0.000030 | 0.000030 | 0.000030 | 0.000030 |
| Temperature | 0.000007 | 0.000045 | 0.000021 | 0.000041 | 0.000036 | 0.000021 | 0.000014 | 0.000015 |
| Wavelength | 0.000254 | 0.000104 | 0.000037 | 0.000021 | 0.000147 | 0.000052 | 0.000037 | 0.000016 |
| Stray Light | 0.000003 | 0.000003 | 0.000003 | 0.000003 | 0.000003 | 0.000003 | 0.000003 | 0.000003 |
| Beam Size & Position | 0.000009 | 0.000030 | 0.000030 | 0.000024 | 0.000051 | 0.000045 | 0.000033 | 0.000024 |
| Filter Instability | 0.000033 | 0.000026 | 0.000010 | 0.000007 | 0.000000 | 0.000026 | 0.000023 | 0.000012 |
| Inter-reflection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Obliquity | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Polarization | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Source Drift & Fluctuation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Bandwidth | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Other(b) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Total Type B Uncertainty | 0.000261 | 0.000124 | 0.000061 | 0.000061 | 0.000162 | 0.000082 | 0.000064 | 0.000046 |
| Degrees of Freedom | infinite | infinite | infinite | infinite | infinite | infinite | infinite | infinite |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : Liu Yuanjie	Date :	19 Feb 2016

Appendix A.2 Measurement Results

Please reproduce the following tables for each of the five filters measured at the completion of step 2 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory:   NMC/A*STAR                           Filter Identifier: D

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0.0003596 | 0.0051736 | 0.0084305 | 0.0083527 | 0.0264825 | 0.0322337 | 0.0225231 | 0.0166760 |
| Number of Measurements | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 |
| Temperature | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 |
| Type A Uncertainty(b) | 0.0000074 | 0.0000027 | 0.0000002 | 0.0000002 | 0.0000020 | 0.0000097 | 0.0000088 | 0.0000063 |
| Type B Uncertainty(c) | 0.0000086 | 0.0000153 | 0.0000099 | 0.0000112 | 0.0000517 | 0.0000283 | 0.0000197 | 0.0000129 |
| Total Uncertainty(d) | 0.0000113 | 0.0000155 | 0.0000099 | 0.0000112 | 0.0000518 | 0.0000299 | 0.0000216 | 0.0000144 |
| Degrees of Freedom | 11 | 2209 | 5669167 | 13379462 | 911618 | 184 | 72 | 53 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory:   NMC/A*STAR                           Filter Identifier: D

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 0.0000014 | 0.0000050 | 0.0000039 | 0.0000040 | 0.0000069 | 0.0000079 | 0.0000073 | 0.0000069 |
| Temperature | 0.0000003 | 0.0000056 | 0.0000040 | 0.0000085 | 0.0000122 | 0.0000082 | 0.0000051 | 0.0000053 |
| Wavelength | 0.0000085 | 0.0000126 | 0.0000068 | 0.0000043 | 0.0000480 | 0.0000201 | 0.0000135 | 0.0000054 |
| Stray Light | 0.0000001 | 0.0000003 | 0.0000004 | 0.0000004 | 0.0000007 | 0.0000008 | 0.0000007 | 0.0000007 |
| Beam Size & Position | 0.0000002 | 0.0000023 | 0.0000039 | 0.0000038 | 0.0000117 | 0.0000142 | 0.0000101 | 0.0000074 |
| Filter Instability | 0.0000000 | 0.0000036 | 0.0000021 | 0.0000020 | 0.0000056 | 0.0000083 | 0.0000049 | 0.0000024 |
| Inter-reflection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Obliquity | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Polarization | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Source Drift & Fluctuation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Bandwidth | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Other(b) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Total Type B Uncertainty | 0.0000086 | 0.0000153 | 0.0000099 | 0.0000112 | 0.0000517 | 0.0000283 | 0.0000197 | 0.0000129 |
| Degrees of Freedom | infinite | infinite | infinite | infinite | infinite | infinite | infinite | infinite |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : Liu Yuanjie	Date :	19 Feb 2016

Appendix A.2 Measurement Results

Please reproduce the following tables for each of the five filters measured at the completion of step 2 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory:   NMC/A*STAR                           Filter Identifier: E

Table A-2i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0.00001173 | 0.00032304 | 0.00090833 | 0.00097493 | 0.00493977 | 0.00950047 | 0.00828975 | 0.00700148 |
| Number of Measurements | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 |
| Temperature | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 |
| Type A Uncertainty(b) | 0.00000024 | 0.00000026 | 0.00000078 | 0.00000031 | 0.00000071 | 0.00000267 | 0.00000306 | 0.00000593 |
| Type B Uncertainty(c) | 0.00000039 | 0.00000142 | 0.00000142 | 0.00000176 | 0.00001376 | 0.00000977 | 0.00000774 | 0.00000552 |
| Total Uncertainty(d) | 0.00000046 | 0.00000144 | 0.00000161 | 0.00000179 | 0.00001378 | 0.00001013 | 0.00000832 | 0.00000810 |
| Degrees of Freedom | 27 | 1957 | 37 | 2318 | 276910 | 413 | 110 | 7 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory:   NMC/A*STAR                           Filter Identifier: E

Table A-2ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 2.    All uncertainties should be reported as absolute uncertainties.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 0.00000004 | 0.00000039 | 0.00000052 | 0.00000050 | 0.00000120 | 0.00000164 | 0.00000132 | 0.00000108 |
| Temperature | 0.00000001 | 0.00000053 | 0.00000064 | 0.00000146 | 0.00000339 | 0.00000337 | 0.00000240 | 0.00000350 |
| Wavelength | 0.00000039 | 0.00000120 | 0.00000108 | 0.00000073 | 0.00001312 | 0.00000802 | 0.00000629 | 0.00000274 |
| Stray Light | 0.00000000 | 0.00000002 | 0.00000004 | 0.00000004 | 0.00000012 | 0.00000016 | 0.00000013 | 0.00000011 |
| Beam Size & Position | 0.00000001 | 0.00000014 | 0.00000041 | 0.00000044 | 0.00000209 | 0.00000414 | 0.00000356 | 0.00000302 |
| Filter Instability | 0.00000000 | 0.00000034 | 0.00000000 | 0.00000007 | 0.00000000 | 0.00000000 | 0.00000000 | 0.00000069 |
| Inter-reflection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Obliquity | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Polarization | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Source Drift & Fluctuation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Bandwidth | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Other(b) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Total Type B Uncertainty | 0.00000039 | 0.00000142 | 0.00000142 | 0.00000176 | 0.00001376 | 0.00000977 | 0.00000774 | 0.00000552 |
| Degrees of Freedom | infinite | infinite | infinite | infinite | infinite | infinite | infinite | infinite |

(a) Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : Liu Yuanjie	Date :	19 Feb 2016

A.3 Measurement Results

Please reproduce the following two tables for each of the five filters measured at the completion of step 4 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: NMC/A*STAR                           Filter Identifier: A

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0.913631 | 0.915116 | 0.917758 | 0.919090 | 0.920093 | 0.920789 | 0.921286 | 0.922547 |
| Number of Measurements | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 |
| Temperature | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C |
| Type A Uncertainty(b) | 0.000062 | 0.000080 | 0.000228 | 0.000308 | 0.000384 | 0.000138 | 0.000212 | 0.000410 |
| Type B Uncertainty(c) | 0.000299 | 0.000299 | 0.000295 | 0.000289 | 0.000286 | 0.000283 | 0.000284 | 0.000284 |
| Total Uncertainty(d) | 0.000305 | 0.000310 | 0.000373 | 0.000422 | 0.000479 | 0.000315 | 0.000355 | 0.000500 |
| Degrees of Freedom | 1169 | 449 | 14 | 7 | 5 | 54 | 16 | 4 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory: NMC/A*STAR                           Filter Identifier: A

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [ √  ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient(b) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity |  |  |  |  |  |  |  |  | 1 |
| Temperature |  |  |  |  |  |  |  |  | 1 |
| Wavelength |  |  |  |  |  |  |  |  | 0.5 |
| Stray Light |  |  |  |  |  |  |  |  | 1 |
| Beam Size & Position |  |  |  |  |  |  |  |  | 0.5 |
| Filter Instability |  |  |  |  |  |  |  |  | 0 |
| Inter-reflection |  |  |  |  |  |  |  |  | NA |
| Obliquity |  |  |  |  |  |  |  |  | NA |
| Polarization |  |  |  |  |  |  |  |  | NA |
| Source Drift & Fluctuation |  |  |  |  |  |  |  |  | NA |
| Bandwidth |  |  |  |  |  |  |  |  | NA |
| Other(c) |  |  |  |  |  |  |  |  | NA |
| Total Correlated(b) |  |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty |  |  |  |  |  |  |  |  |  |
| Degrees of Freedom |  |  |  |  |  |  |  |  |  |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : Liu Yuanjie	Date :	19 Feb 2016

A.3 Measurement Results

Please reproduce the following two tables for each of the five filters measured at the completion of step 4 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: NMC/A*STAR                           Filter Identifier: B

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0.406985 | 0.606586 | 0.620862 | 0.608838 | 0.636383 | 0.575550 | 0.500300 | 0.451719 |
| Number of Measurements | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 |
| Temperature | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C |
| Type A Uncertainty(b) | 0.000017 | 0.000094 | 0.000044 | 0.000045 | 0.000134 | 0.000068 | 0.000320 | 0.000059 |
| Type B Uncertainty(c) | 0.001304 | 0.000245 | 0.000236 | 0.000218 | 0.000228 | 0.000228 | 0.000186 | 0.000152 |
| Total Uncertainty(d) | 0.001304 | 0.000262 | 0.000240 | 0.000222 | 0.000265 | 0.000238 | 0.000370 | 0.000163 |
| Degrees of Freedom | 64996624 | 122 | 1834 | 1173 | 30 | 304 | 4 | 117 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory: NMC/A*STAR                           Filter Identifier: B

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [     ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient(b) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity |  |  |  |  |  |  |  |  | 1 |
| Temperature |  |  |  |  |  |  |  |  | 1 |
| Wavelength |  |  |  |  |  |  |  |  | 0.5 |
| Stray Light |  |  |  |  |  |  |  |  | 1 |
| Beam Size & Position |  |  |  |  |  |  |  |  | 0.5 |
| Filter Instability |  |  |  |  |  |  |  |  | 0 |
| Inter-reflection |  |  |  |  |  |  |  |  | NA |
| Obliquity |  |  |  |  |  |  |  |  | NA |
| Polarization |  |  |  |  |  |  |  |  | NA |
| Source Drift & Fluctuation |  |  |  |  |  |  |  |  | NA |
| Bandwidth |  |  |  |  |  |  |  |  | NA |
| Other(c) |  |  |  |  |  |  |  |  | NA |
| Total Correlated(b) |  |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty |  |  |  |  |  |  |  |  |  |
| Degrees of Freedom |  |  |  |  |  |  |  |  |  |

((a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : Liu Yuanjie	Date :	19 Feb 2016

A.3 Measurement Results

Please reproduce the following two tables for each of the five filters measured at the completion of step 4 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: NMC/A*STAR                           Filter Identifier: C

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0.021387 | 0.094808 | 0.091527 | 0.076699 | 0.160534 | 0.149612 | 0.102478 | 0.076237 |
| Number of Measurements | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 |
| Temperature | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C |
| Type A Uncertainty(b) | 0.000016 | 0.000024 | 0.000022 | 0.000042 | 0.000045 | 0.000049 | 0.000054 | 0.000048 |
| Type B Uncertainty(c) | 0.000261 | 0.000124 | 0.000061 | 0.000061 | 0.000162 | 0.000082 | 0.000064 | 0.000046 |
| Total Uncertainty(d) | 0.000261 | 0.000126 | 0.000065 | 0.000074 | 0.000168 | 0.000095 | 0.000084 | 0.000066 |
| Degrees of Freedom | 132214 | 1492 | 147 | 18 | 383 | 29 | 12 | 7 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory: NMC/A*STAR                           Filter Identifier: C

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [     ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient(b) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity |  |  |  |  |  |  |  |  | 1 |
| Temperature |  |  |  |  |  |  |  |  | 1 |
| Wavelength |  |  |  |  |  |  |  |  | 0.5 |
| Stray Light |  |  |  |  |  |  |  |  | 1 |
| Beam Size & Position |  |  |  |  |  |  |  |  | 0.5 |
| Filter Instability |  |  |  |  |  |  |  |  | 0 |
| Inter-reflection |  |  |  |  |  |  |  |  | NA |
| Obliquity |  |  |  |  |  |  |  |  | NA |
| Polarization |  |  |  |  |  |  |  |  | NA |
| Source Drift & Fluctuation |  |  |  |  |  |  |  |  | NA |
| Bandwidth |  |  |  |  |  |  |  |  | NA |
| Other(c) |  |  |  |  |  |  |  |  | NA |
| Total Correlated(b) |  |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty |  |  |  |  |  |  |  |  |  |
| Degrees of Freedom |  |  |  |  |  |  |  |  |  |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : Liu Yuanjie	Date :	19 Feb 2016

A.3 Measurement Results

Please reproduce the following two tables for each of the five filters measured at the completion of step 4 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: NMC/A*STAR                           Filter Identifier: D

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0.0003590 | 0.0051745 | 0.0084350 | 0.0083583 | 0.0265033 | 0.0322262 | 0.0225018 | 0.0166627 |
| Number of Measurements | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 |
| Temperature | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C |
| Type A Uncertainty(b) | 0.0000001 | 0.0000042 | 0.0000025 | 0.0000107 | 0.0000092 | 0.0000089 | 0.0000160 | 0.0000174 |
| Type B Uncertainty(c) | 0.0000086 | 0.0000153 | 0.0000099 | 0.0000112 | 0.0000517 | 0.0000283 | 0.0000197 | 0.0000129 |
| Total Uncertainty(d) | 0.0000086 | 0.0000158 | 0.0000102 | 0.0000155 | 0.0000525 | 0.0000297 | 0.0000254 | 0.0000216 |
| Degrees of Freedom | 1476172744 | 406 | 589 | 9 | 2110 | 252 | 13 | 5 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory: NMC/A*STAR                           Filter Identifier: D

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [     ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient(b) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity |  |  |  |  |  |  |  |  | 1 |
| Temperature |  |  |  |  |  |  |  |  | 1 |
| Wavelength |  |  |  |  |  |  |  |  | 0.5 |
| Stray Light |  |  |  |  |  |  |  |  | 1 |
| Beam Size & Position |  |  |  |  |  |  |  |  | 0.5 |
| Filter Instability |  |  |  |  |  |  |  |  | 0 |
| Inter-reflection |  |  |  |  |  |  |  |  | NA |
| Obliquity |  |  |  |  |  |  |  |  | NA |
| Polarization |  |  |  |  |  |  |  |  | NA |
| Source Drift & Fluctuation |  |  |  |  |  |  |  |  | NA |
| Bandwidth |  |  |  |  |  |  |  |  | NA |
| Other(c) |  |  |  |  |  |  |  |  | NA |
| Total Correlated(b) |  |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty |  |  |  |  |  |  |  |  |  |
| Degrees of Freedom |  |  |  |  |  |  |  |  |  |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : Liu Yuanjie	Date :	19 Feb 2016

A.3 Measurement Results

Please reproduce the following two tables for each of the five filters measured at the completion of step 4 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: NMC/A*STAR                           Filter Identifier: E

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 0.00001185 | 0.00032435 | 0.00090756 | 0.00097391 | 0.00493243 | 0.00948687 | 0.00825632 | 0.00698392 |
| Number of Measurements | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 |
| Temperature | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C | 25 ˚C |
| Type A Uncertainty(b) | 0.00000004 | 0.00000013 | 0.00000083 | 0.00000086 | 0.00000365 | 0.00000299 | 0.00000820 | 0.00000886 |
| Type B Uncertainty(c) | 0.00000039 | 0.00000142 | 0.00000142 | 0.00000176 | 0.00001376 | 0.00000977 | 0.00000774 | 0.00000552 |
| Total Uncertainty(d) | 0.00000039 | 0.00000142 | 0.00000164 | 0.00000196 | 0.00001424 | 0.00001022 | 0.00001127 | 0.00001044 |
| Degrees of Freedom | 19369 | 29780 | 31 | 55 | 461 | 274 | 7 | 4 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory: NMC/A*STAR                           Filter Identifier: E

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [     ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient(b) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity |  |  |  |  |  |  |  |  | 1 |
| Temperature |  |  |  |  |  |  |  |  | 1 |
| Wavelength |  |  |  |  |  |  |  |  | 0.5 |
| Stray Light |  |  |  |  |  |  |  |  | 1 |
| Beam Size & Position |  |  |  |  |  |  |  |  | 0.5 |
| Filter Instability |  |  |  |  |  |  |  |  | 0 |
| Inter-reflection |  |  |  |  |  |  |  |  | NA |
| Obliquity |  |  |  |  |  |  |  |  | NA |
| Polarization |  |  |  |  |  |  |  |  | NA |
| Source Drift & Fluctuation |  |  |  |  |  |  |  |  | NA |
| Bandwidth |  |  |  |  |  |  |  |  | NA |
| Other(c) |  |  |  |  |  |  |  |  | NA |
| Total Correlated(b) |  |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty |  |  |  |  |  |  |  |  |  |
| Degrees of Freedom |  |  |  |  |  |  |  |  |  |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : Liu Yuanjie	Date :	19 Feb 2016