Appendix A.1 Description of measurement facility and measurement method

Laboratory:

Indicate whether this table relates to Step 2 [      ] or Step 4 [   ]

If the measurement setup has not changed from Step 2, check here [   ] and the following table does not need to be completed. Otherwise, please fill out the whole table.

Table A-1 Details of Measurement Setup

| Make and Type of Spectrophotometer |  |
| --- | --- |
| Additional Stray Light Rejection |  |
| Source Drift Monitoring |  |
| Source |  |
| Detector |  |
| Temperature(a) | temperature range: (24.9 ~ 29.1)C |
| Humidity |  |
| Beam Size |  |
| Beam Collimation |  |
| Measurement Sequence(b) |  |
| Bandwidth |  |

(a) i.e. describe method of temperature monitoring of filters and range of temperatures  (b) i.e. describe number of measurements and whether filter orientation with respect to beam changes between measurements

Description of measuring technique (please include a diagram)

If any damage, contamination or cleaning of the filters was carried out, please give details

Signature :	Date :

A.3 Measurement Results

Please reproduce the following two tables for each of the five filters measured at the completion of step 4 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory:     KRISS                        Filter Identifier: A

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 9.1317E-01 | 9.1450E-01 | 9.1713E-01 | 9.1868E-01 | 9.1977E-01 | 9.2035E-01 | 9.2079E-01 | 9.2072E-01 |
| Number of Measurements | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 |
| Temperature | (27.9 2.4) C | (27.9 2.4) C | (27.9 2.4) C | (27.9 2.4) C | (27.9 2.4) C | (27.9 2.4) C | (27.9 2.4) C | (27.9 2.4) C |
| Type A Uncertainty(b) | 1.53E-04 | 1.43E-04 | 7.70E-05 | 1.31E-04 | 7.10E-05 | 1.33E-04 | 9.97E-05 | 2.30E-04 |
| Type B Uncertainty(c) | 3.08E-03 | 1.47E-03 | 2.05E-03 | 1.36E-03 | 1.27E-03 | 1.25E-03 | 1.39E-03 | 1.96E-03 |
| Total Uncertainty(d) | 3.08E-03 | 1.48E-03 | 2.05E-03 | 1.37E-03 | 1.27E-03 | 1.26E-03 | 1.40E-03 | 1.97E-03 |
| Degrees of Freedom |  |  |  |  |  |  |  |  |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory: KRISS                        Filter Identifier: A

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [     ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

| Wavelength (nm) | 380 |  | 400 |  | 500 |  | 600 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) |
| Nonlinearity | 9.13E-06 | 1 | 9.14E-06 | 1 | 9.17E-06 | 1 | 9.19E-06 | 1 |
| Temperature | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Wavelength | 1.87E-05 | 0 | 1.11E-05 | 0 | 8.29E-06 | 0 | 5.26E-06 | 0 |
| Stray Light | 9.13E-07 | 1 | 9.14E-07 | 1 | 9.17E-07 | 1 | 9.19E-07 | 1 |
| Beam Size & Position | 3.66E-04 | 1 | 3.98E-04 | 1 | 2.91E-04 | 1 | 2.24E-04 | 1 |
| Inter-reflection | 1.10E-03 | 1 | 1.10E-03 | 1 | 1.10E-03 | 1 | 1.10E-03 | 1 |
| Obliquity | 1.79E-04 | 1 | 3.74E-04 | 1 | 5.73E-05 | 1 | 2.43E-04 | 1 |
| Polarization | 5.52E-04 | 1 | 3.45E-05 | 1 | 1.42E-05 | 1 | 1.20E-05 | 1 |
| Source Drift & Fluctuation | 2.79E-03 | 0 | 7.97E-04 | 0 | 1.71E-03 | 0 | 7.32E-04 | 0 |
| Bandwidth | 3.00E-05 | 1 | 1.40E-04 | 1 | 9.99E-06 | 1 | 3.00E-05 | 1 |
| Other(c) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 3.08E-03 | 0.4 | 1.47E-03 | 0.7 | 2.05E-03 | 0.4 | 1.36E-03 | 0.7 |
| Degrees of Freedom |  |  |  |  |  |  |  |  |

| Wavelength (nm) | 700 |  | 800 |  | 900 |  | 1000 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) |
| Nonlinearity | 9.20E-06 | 1 | 9.20E-06 | 1 | 9.21E-06 | 1 | 9.21E-06 | 1 |
| Temperature | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Wavelength | 2.91E-06 | 0 | 4.56E-06 | 0 | 4.99E-06 | 0 | 8.54E-06 | 0 |
| Stray Light | 9.20E-07 | 1 | 9.20E-07 | 1 | 9.21E-07 | 1 | 9.21E-07 | 1 |
| Beam Size & Position | 1.83E-04 | 1 | 1.62E-04 | 1 | 1.01E-04 | 1 | 1.23E-04 | 1 |
| Inter-reflection | 1.10E-03 | 1 | 1.10E-03 | 1 | 1.10E-03 | 1 | 1.10E-03 | 1 |
| Obliquity | 2.96E-05 | 1 | 1.73E-04 | 1 | 3.21E-04 | 1 | 1.98E-04 | 1 |
| Polarization | 5.01E-06 | 1 | 9.75E-06 | 1 | 4.12E-06 | 1 | 3.93E-06 | 1 |
| Source Drift & Fluctuation | 5.77E-04 | 0 | 5.34E-04 | 0 | 7.82E-04 | 0 | 1.60E-03 | 0 |
| Bandwidth | 1.60E-04 | 1 | 6.00E-05 | 1 | 9.99E-06 | 1 | 9.98E-06 | 1 |
| Other(c) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 1.27E-03 | 0.7 | 1.25E-03 | 0.9 | 1.39E-03 | 0.7 | 1.96E-03 | 0.5 |
| Degrees of Freedom |  |  |  |  |  |  |  |  |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : 	Date :

Laboratory:     KRISS                        Filter Identifier: B

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 4.1273E-01 | 6.0740E-01 | 6.2144E-01 | 6.0919E-01 | 6.3671E-01 | 5.7595E-01 | 5.0052E-01 | 4.5230E-01 |
| Number of Measurements | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 |
| Temperature | (28.5 0.1) C | (26.7 0.9) C | (26.7 0.9) C | (26.7 0.9) C | (26.7 0.9) C | (26.7 0.9) C | (26.7 0.9) C | (26.7 0.9) C |
| Type A Uncertainty(b) | 5.85E-04 | 7.04E-05 | 1.61E-04 | 1.09E-04 | 9.90E-05 | 1.73E-04 | 2.85E-04 | 2.05E-05 |
| Type B Uncertainty(c) | 2.59E-03 | 1.77E-03 | 1.74E-03 | 1.42E-03 | 1.74E-03 | 2.68E-03 | 9.77E-04 | 2.49E-03 |
| Total Uncertainty(d) | 2.66E-03 | 1.77E-03 | 1.74E-03 | 1.42E-03 | 1.74E-03 | 2.68E-03 | 1.02E-03 | 2.49E-03 |
| Degrees of Freedom | 1.70E+03 |  |  |  |  |  | 651.0 |  |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory: KRISS                        Filter Identifier: B

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [     ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

| Wavelength (nm) | 380 |  | 400 |  | 500 |  | 600 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) |
| Nonlinearity | 4.13E-06 | 1 | 6.08E-06 | 1 | 6.21E-06 | 1 | 6.09E-06 | 1 |
| Temperature | 7.41E-05 | 0.7 | 1.40E-04 | 0.6 | 6.03E-05 | 0.7 | 1.45E-04 | 0.6 |
| Wavelength | 7.78E-04 | 0 | 9.83E-04 | 0 | 1.99E-05 | 0 | 1.40E-05 | 0 |
| Stray Light | 1.05E-05 | 1 | 3.25E-06 | 1 | 4.28E-06 | 1 | 3.42E-06 | 1 |
| Beam Size & Position | 2.92E-04 | 1 | 1.22E-04 | 1 | 2.56E-04 | 1 | 2.10E-04 | 1 |
| Inter-reflection | 4.96E-04 | 1 | 7.29E-04 | 1 | 7.46E-04 | 1 | 7.31E-04 | 1 |
| Obliquity | 3.62E-05 | 1 | 1.81E-04 | 1 | 1.94E-04 | 1 | 3.25E-05 | 1 |
| Polarization | 2.49E-04 | 1 | 2.30E-05 | 1 | 9.63E-06 | 1 | 7.98E-06 | 1 |
| Source Drift & Fluctuation | 1.60E-03 | 0 | 1.22E-03 | 0 | 1.53E-03 | 0 | 1.19E-03 | 0 |
| Bandwidth | 1.77E-03 | 1 | 3.20E-04 | 1 | 0 | 1 | 4.99E-05 | 1 |
| Other(c) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 2.59E-03 | 0.6 | 1.77E-03 | 0.3 | 1.74E-03 | 0.4 | 1.42E-03 | 0.5 |
| Degrees of Freedom |  |  |  |  |  |  |  |  |

| Wavelength (nm) | 700 |  | 800 |  | 900 |  | 1000 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) |
| Nonlinearity | 6.37E-06 | 1 | 5.76E-06 | 1 | 5.00E-06 | 1 | 4.52E-06 | 1 |
| Temperature | 7.84E-05 | 0.7 | 4.57E-05 | 0.8 | 3.57E-05 | 1.0 | 5.77E-05 | 0.8 |
| Wavelength | 4.86E-05 | 0 | 1.53E-04 | 0 | 1.11E-04 | 0 | 5.70E-05 | 0 |
| Stray Light | 5.28E-06 | 1 | 1.05E-06 | 1 | 4.26E-06 | 1 | 7.63E-06 | 1 |
| Beam Size & Position | 9.74E-05 | 1 | 1.21E-04 | 1 | 2.71E-04 | 1 | 8.97E-05 | 1 |
| Inter-reflection | 7.64E-04 | 1 | 6.91E-04 | 1 | 6.01E-04 | 1 | 5.43E-04 | 1 |
| Obliquity | 2.46E-04 | 1 | 1.76E-04 | 1 | 2.07E-04 | 1 | 2.58E-04 | 1 |
| Polarization | 3.47E-06 | 1 | 6.10E-06 | 1 | 2.24E-06 | 1 | 1.93E-06 | 1 |
| Source Drift & Fluctuation | 1.54E-03 | 0 | 2.57E-03 | 0 | 6.82E-04 | 0 | 2.42E-03 | 0 |
| Bandwidth | 7.99E-05 | 1 | 2.00E-05 | 1 | 0 | 1 | 1.00E-05 | 1 |
| Other(c) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 1.74E-03 | 0.4 | 2.68E-03 | 0.2 | 9.77E-04 | 0.6 | 2.49E-03 | 0.2 |
| Degrees of Freedom |  |  |  |  |  |  |  |  |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : 	Date :

Laboratory:     KRISS                        Filter Identifier: C

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 2.2523E-02 | 9.5545E-02 | 9.2075E-02 | 7.7168E-02 | 1.6162E-01 | 1.5023E-01 | 1.0301E-01 | 7.6828E-02 |
| Number of Measurements | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 |
| Temperature | (28.5 0.3) C | (26.9 0.5) C | (26.9 0.5) C | (26.9 0.5) C | (26.9 0.5) C | (26.9 0.5) C | (26.9 0.5) C | (26.9 0.5) C |
| Type A Uncertainty(b) | 8.18E-05 | 6.60E-05 | 3.59E-05 | 3.37E-05 | 7.26E-05 | 3.36E-05 | 1.35E-05 | 2.18E-05 |
| Type B Uncertainty(c) | 4.55E-04 | 5.96E-04 | 2.41E-04 | 2.01E-04 | 6.07E-04 | 6.23E-04 | 3.20E-04 | 2.75E-04 |
| Total Uncertainty(d) | 4.62E-04 | 6.00E-04 | 2.44E-04 | 2.04E-04 | 6.11E-04 | 6.24E-04 | 3.21E-04 | 2.75E-04 |
| Degrees of Freedom | 4.06E+03 |  | 7.67E+03 | 4.75E+03 |  |  |  |  |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory: KRISS                        Filter Identifier: C

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [     ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

| Wavelength (nm) | 380 |  | 400 |  | 500 |  | 600 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) |
| Nonlinearity | 2.26E-07 | 1 | 9.59E-07 | 1 | 9.19E-07 | 1 | 7.69E-07 | 1 |
| Temperature | 1.83E-05 | 0.6 | 9.95E-05 | 0.6 | 4.59E-05 | 0.6 | 9.16E-05 | 0.6 |
| Wavelength | 1.52E-04 | 0 | 5.24E-04 | 0 | 4.93E-05 | 0 | 8.96E-06 | 0 |
| Stray Light | 5.82E-06 | 1 | 6.35E-07 | 1 | 8.05E-07 | 1 | 1.83E-06 | 1 |
| Beam Size & Position | 1.09E-05 | 1 | 6.82E-05 | 1 | 3.77E-05 | 1 | 3.38E-05 | 1 |
| Inter-reflection | 2.71E-05 | 1 | 1.15E-04 | 1 | 1.10E-04 | 1 | 9.22E-05 | 1 |
| Obliquity | 9.81E-06 | 1 | 5.05E-05 | 1 | 4.02E-05 | 1 | 5.20E-05 | 1 |
| Polarization | 1.36E-05 | 1 | 3.62E-06 | 1 | 1.43E-06 | 1 | 1.01E-06 | 1 |
| Source Drift & Fluctuation | 9.97E-05 | 0 | 1.44E-04 | 0 | 1.95E-04 | 0 | 1.39E-04 | 0 |
| Bandwidth | 4.15E-04 | 1 | 1.72E-04 | 1 | 1.99E-05 | 1 | 9.92E-06 | 1 |
| Other(c) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 4.55E-04 | 0.9 | 5.96E-04 | 0.2 | 2.41E-04 | 0.4 | 2.01E-04 | 0.5 |
| Degrees of Freedom |  |  |  |  |  |  |  |  |

| Wavelength (nm) | 700 |  | 800 |  | 900 |  | 1000 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) |
| Nonlinearity | 1.61E-06 | 1 | 1.50E-06 | 1 | 1.03E-06 | 1 | 7.69E-07 | 1 |
| Temperature | 8.01E-05 | 0.6 | 4.34E-05 | 0.6 | 2.78E-05 | 0.8 | 3.14E-05 | 0.6 |
| Wavelength | 2.76E-04 | 0 | 9.15E-05 | 0 | 6.52E-05 | 0 | 2.74E-05 | 0 |
| Stray Light | 4.02E-06 | 1 | 3.26E-06 | 1 | 7.02E-08 | 1 | 1.91E-06 | 1 |
| Beam Size & Position | 3.40E-05 | 1 | 5.26E-05 | 1 | 2.64E-05 | 1 | 1.40E-05 | 1 |
| Inter-reflection | 1.94E-04 | 1 | 1.80E-04 | 1 | 1.24E-04 | 1 | 9.23E-05 | 1 |
| Obliquity | 7.28E-05 | 1 | 6.36E-05 | 1 | 5.85E-05 | 1 | 7.40E-05 | 1 |
| Polarization | 8.79E-07 | 1 | 1.59E-06 | 1 | 4.61E-07 | 1 | 3.28E-07 | 1 |
| Source Drift & Fluctuation | 4.92E-04 | 0 | 5.82E-04 | 0 | 2.80E-04 | 0 | 2.44E-04 | 0 |
| Bandwidth | 2.00E-05 | 1 | 9.97E-06 | 1 | 0 | 1 | 1.00E-05 | 1 |
| Other(c) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 6.07E-04 | 0.2 | 6.23E-04 | 0.3 | 3.20E-04 | 0.4 | 2.75E-04 | 0.4 |
| Degrees of Freedom |  |  |  |  |  |  |  |  |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : 	Date :

Laboratory:     KRISS                        Filter Identifier: D

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 3.8669E-04 | 5.1718E-03 | 8.3909E-03 | 8.3166E-03 | 2.6514E-02 | 3.2148E-02 | 2.2488E-02 | 1.6654E-02 |
| Number of Measurements | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 |
| Temperature | (28.4 0.6) C | (26.2 1.3) C | (26.2 1.3) C | (26.2 1.3) C | (26.2 1.3) C | (26.2 1.3) C | (26.2 1.3) C | (26.2 1.3) C |
| Type A Uncertainty(b) | 7.48E-07 | 7.36E-06 | 7.44E-06 | 1.08E-05 | 1.53E-05 | 1.35E-05 | 1.81E-06 | 2.70E-06 |
| Type B Uncertainty(c) | 2.24E-05 | 5.74E-05 | 2.43E-05 | 3.45E-05 | 1.59E-04 | 6.37E-05 | 9.49E-05 | 5.04E-05 |
| Total Uncertainty(d) | 2.24E-05 | 5.79E-05 | 2.55E-05 | 3.61E-05 | 1.59E-04 | 6.51E-05 | 9.49E-05 | 5.05E-05 |
| Degrees of Freedom |  |  | 532.0 | 496.0 |  | 1.89E+03 |  |  |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory: KRISS                        Filter Identifier: D

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [     ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

| Wavelength (nm) | 380 |  | 400 |  | 500 |  | 600 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) |
| Nonlinearity | 3.89E-09 | 1 | 5.21E-08 | 1 | 8.37E-08 | 1 | 8.26E-08 | 1 |
| Temperature | 6.98E-07 | 0.6 | 1.75E-05 | 0.4 | 1.11E-05 | 0.4 | 2.69E-05 | 0.4 |
| Wavelength | 3.28E-06 | 0 | 4.01E-05 | 0 | 3.26E-06 | 0 | 5.56E-06 | 0 |
| Stray Light | 1.19E-06 | 1 | 8.42E-07 | 1 | 5.85E-07 | 1 | 5.85E-07 | 1 |
| Beam Size & Position | 3.28E-07 | 1 | 5.80E-06 | 1 | 1.46E-06 | 1 | 2.63E-06 | 1 |
| Inter-reflection | 4.67E-07 | 1 | 6.25E-06 | 1 | 1.00E-05 | 1 | 9.92E-06 | 1 |
| Obliquity | 4.09E-07 | 1 | 1.74E-05 | 1 | 2.36E-07 | 1 | 1.97E-06 | 1 |
| Polarization | 2.35E-07 | 1 | 1.97E-07 | 1 | 1.30E-07 | 1 | 1.08E-07 | 1 |
| Source Drift & Fluctuation | 2.32E-06 | 0 | 3.17E-05 | 0 | 1.60E-05 | 0 | 1.50E-05 | 0 |
| Bandwidth | 2.20E-05 | 1 | 0 | 1 | 9.95E-06 | 1 | 9.89E-06 | 1 |
| Other(c) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 2.24E-05 | 1.0 | 5.74E-05 | 0.2 | 2.43E-05 | 0.6 | 3.45E-05 | 0.5 |
| Degrees of Freedom |  |  |  |  |  |  |  |  |

| Wavelength (nm) | 700 |  | 800 |  | 900 |  | 1000 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) |
| Nonlinearity | 2.64E-07 | 1 | 3.21E-07 | 1 | 2.25E-07 | 1 | 1.67E-07 | 1 |
| Temperature | 3.67E-05 | 0.4 | 1.97E-05 | 0.5 | 9.04E-06 | 0.7 | 1.29E-05 | 0.5 |
| Wavelength | 6.15E-05 | 0 | 1.33E-05 | 0 | 1.54E-05 | 0 | 4.83E-06 | 0 |
| Stray Light | 6.75E-07 | 1 | 1.07E-06 | 1 | 3.96E-07 | 1 | 1.40E-08 | 1 |
| Beam Size & Position | 1.63E-06 | 1 | 3.70E-06 | 1 | 3.38E-06 | 1 | 5.76E-06 | 1 |
| Inter-reflection | 3.17E-05 | 1 | 3.85E-05 | 1 | 2.70E-05 | 1 | 2.00E-05 | 1 |
| Obliquity | 1.54E-05 | 1 | 2.04E-05 | 1 | 6.12E-06 | 1 | 1.01E-05 | 1 |
| Polarization | 1.44E-07 | 1 | 3.40E-07 | 1 | 1.01E-07 | 1 | 7.12E-08 | 1 |
| Source Drift & Fluctuation | 1.37E-04 | 0 | 3.96E-05 | 0 | 8.89E-05 | 0 | 4.14E-05 | 0 |
| Bandwidth | 0 | 1 | 0 | 1 | 0 | 1 | 1.00E-05 | 1 |
| Other(c) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 1.59E-04 | 0.2 | 6.37E-05 | 0.6 | 9.49E-05 | 0.3 | 5.04E-05 | 0.5 |
| Degrees of Freedom |  |  |  |  |  |  |  |  |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : 	Date :

Laboratory:     KRISS                        Filter Identifier: E

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 1.2404E-05 | 3.1929E-04 | 8.8299E-04 | 9.4752E-04 | 4.8538E-03 | 9.3227E-03 | 8.1086E-03 | 6.8541E-03 |
| Number of Measurements | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 |
| Temperature | (28.6 0.4) C | (28.6 0.4) C | (26.3 0.6) C | (26.3 0.6) C | (26.3 0.6) C | (26.3 0.6) C | (26.3 0.6) C | (26.3 0.6) C |
| Type A Uncertainty(b) | 2.10E-07 | 4.37E-07 | 6.70E-07 | 6.47E-07 | 6.53E-06 | 4.40E-06 | 2.35E-06 | 1.61E-06 |
| Type B Uncertainty(c) | 2.09E-06 | 8.74E-06 | 5.90E-06 | 4.60E-06 | 2.11E-05 | 1.98E-05 | 3.77E-05 | 3.34E-05 |
| Total Uncertainty(d) | 2.10E-06 | 8.75E-06 | 5.94E-06 | 4.65E-06 | 2.21E-05 | 2.03E-05 | 3.78E-05 | 3.35E-05 |
| Degrees of Freedom |  |  |  | 6.22E+03 | 506.0 | 1.28E+03 |  |  |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Laboratory: KRISS                        Filter Identifier: E

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [     ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

| Wavelength (nm) | 380 |  | 400 |  | 500 |  | 600 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) |
| Nonlinearity | 1.25E-10 | 1 | 3.25E-09 | 1 | 8.80E-09 | 1 | 9.39E-09 | 1 |
| Temperature | 3.18E-08 | 0.7 | 1.16E-06 | 0.9 | 1.31E-06 | 0.7 | 3.25E-06 | 0.7 |
| Wavelength | 1.07E-07 | 0 | 2.69E-06 | 0 | 4.23E-07 | 0 | 1.13E-06 | 0 |
| Stray Light | 4.68E-07 | 1 | 3.09E-07 | 1 | 2.58E-07 | 1 | 2.42E-07 | 1 |
| Beam Size & Position | 1.21E-07 | 1 | 7.31E-06 | 1 | 7.16E-07 | 1 | 8.57E-07 | 1 |
| Inter-reflection | 1.50E-08 | 1 | 3.90E-07 | 1 | 1.06E-06 | 1 | 1.13E-06 | 1 |
| Obliquity | 1.22E-07 | 1 | 1.82E-06 | 1 | 5.48E-07 | 1 | 2.24E-07 | 1 |
| Polarization | 7.56E-09 | 1 | 1.23E-08 | 1 | 1.36E-08 | 1 | 1.23E-08 | 1 |
| Source Drift & Fluctuation | 1.87E-06 | 0 | 3.22E-06 | 0 | 5.49E-06 | 0 | 2.69E-06 | 0 |
| Bandwidth | 7.80E-07 | 1 | 5.44E-07 | 1 | 9.27E-07 | 1 | 7.04E-08 | 1 |
| Other(c) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 2.09E-06 | 0.4 | 8.74E-06 | 0.7 | 5.90E-06 | 0.2 | 4.60E-06 | 0.5 |
| Degrees of Freedom |  |  |  |  |  |  |  |  |

| Wavelength (nm) | 700 |  | 800 |  | 900 |  | 1000 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) | uncertainty | Correlation Coefficient(b) |
| Nonlinearity | 4.83E-08 | 1 | 9.31E-08 | 1 | 8.10E-08 | 1 | 6.87E-08 | 1 |
| Temperature | 7.34E-06 | 0.7 | 6.41E-06 | 0.7 | 4.01E-06 | 0.8 | 7.22E-06 | 0.7 |
| Wavelength | 1.54E-05 | 0 | 1.72E-06 | 0 | 3.52E-06 | 0 | 4.92E-07 | 0 |
| Stray Light | 2.66E-08 | 1 | 3.40E-07 | 1 | 2.56E-07 | 1 | 1.68E-07 | 1 |
| Beam Size & Position | 2.87E-06 | 1 | 2.67E-06 | 1 | 4.85E-06 | 1 | 4.58E-06 | 1 |
| Inter-reflection | 5.80E-06 | 1 | 1.12E-05 | 1 | 9.73E-06 | 1 | 8.25E-06 | 1 |
| Obliquity | 1.10E-06 | 1 | 2.46E-06 | 1 | 2.01E-06 | 1 | 1.96E-06 | 1 |
| Polarization | 2.63E-08 | 1 | 9.86E-08 | 1 | 3.63E-08 | 1 | 2.93E-08 | 1 |
| Source Drift & Fluctuation | 1.05E-05 | 0 | 1.45E-05 | 0 | 3.57E-05 | 0 | 3.12E-05 | 0 |
| Bandwidth | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 |
| Other(c) |  |  |  |  |  |  |  |  |
| Total Type B Uncertainty | 2.11E-05 | 0.3 | 1.98E-05 | 0.5 | 3.77E-05 | 0.3 | 3.34E-05 | 0.3 |
| Degrees of Freedom |  |  | 4.02E+03 |  |  |  |  |  |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

Signature : 	Date :