A.3 Measurement Results (1)

Blue: Revised at the Draft A stage to hide identification of the DUT filter set

Please reproduce the following two tables for each of the five filters measured at the completion of step 4 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: National Metrology Institute of Japan (NMIJ, AIST)

Filter Type: A

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 9.142E-01 | 9.156E-01 | 9.175E-01 | 9.187E-01 | 9.197E-01 | 9.202E-01 | 9.208E-01 | 9.213E-01 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Temperature | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 |
| Type A Uncertainty(b) | 1.1E-04 | 1.1E-04 | 1.4E-04 | 1.4E-04 | 1.4E-04 | 1.4E-04 | 1.4E-04 | 1.4E-04 |
| Type B Uncertainty(c) | 5.2E-04 | 5.1E-04 | 3.1E-04 | 3.0E-04 | 3.0E-04 | 3.2E-04 | 3.0E-04 | 3.0E-04 |
| Total Uncertainty(d) | 5.3E-04 | 5.2E-04 | 3.4E-04 | 3.3E-04 | 3.3E-04 | 3.5E-04 | 3.3E-04 | 3.3E-04 |
| Degrees of Freedom | 4550 | 4148 | 319 | 305 | 301 | 350 | 297 | 277 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [  ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

Filter Type: A

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient(b) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 1.6E-04 | 1.6E-04 | 1.6E-04 | 1.6E-04 | 1.6E-04 | 1.6E-04 | 1.6E-04 | 1.6E-04 | +1 |
| Temperature | 2.5E-04 | 2.9E-04 | 8.9E-05 | 7.3E-05 | 6.8E-05 | 1.2E-04 | 7.7E-06 | 7.9E-06 | 0 |
| Wavelength | 2.4E-05 | 5.7E-07 | 1.6E-06 | 6.4E-06 | 3.0E-06 | 3.9E-06 | 5.2E-05 | 1.3E-06 | +1 |
| Stray Light | 1.3E-06 | 1.1E-06 | 6.6E-07 | 4.2E-07 | 2.3E-07 | 1.0E-07 | 3.9E-09 | 9.6E-08 | +1 |
| Beam Size & Position | 1.4E-05 | 1.4E-05 | 1.4E-05 | 1.4E-05 | 1.4E-05 | 1.4E-05 | 1.4E-05 | 1.4E-05 | +1 |
| Inter-reflection | 9.9E-05 | 9.9E-05 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | +1 |
| Obliquity | 7.6E-05 | 7.6E-05 | 7.6E-05 | 7.6E-05 | 7.6E-05 | 7.6E-05 | 7.6E-05 | 7.6E-05 | 0 |
| Polarization | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | 1.0E-04 | +1 |
| Source Drift & Fluctuation * | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A |
| Bandwidth | 1.6E-05 | 3.7E-07 | 1.0E-06 | 4.2E-06 | 2.0E-06 | 2.5E-06 | 3.4E-05 | 8.3E-07 | +1 |
| Filter Stability | 3.9E-04 | 3.4E-04 | 1.8E-04 | 1.8E-04 | 1.8E-04 | 1.8E-04 | 1.8E-04 | 1.8E-04 | 0 |
| Total Type B Uncertainty | 5.2E-04 | 5.1E-04 | 3.1E-04 | 3.0E-04 | 3.0E-04 | 3.2E-04 | 3.0E-04 | 3.0E-04 | N.A |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | N.A |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

*) “Source Drift & Fluctuation” is included in the Type A uncertainty.

Signature : 	Date :

Hiroshi Shitomi			2016-02-29

* Original data submitted:  2014-12-05

A.3 Measurement Results (2)

Blue: Revised at the Draft A stage to hide identification of the DUT filter set

Please reproduce the following two tables for each of the five filters measured at the completion of step 4 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: National Metrology Institute of Japan (NMIJ, AIST)

Filter Type: B

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 4.132E-01 | 6.085E-01 | 6.223E-01 | 6.102E-01 | 6.376E-01 | 5.770E-01 | 5.017E-01 | 4.534E-01 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Temperature | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 |
| Type A Uncertainty(b) | 2.0E-04 | 3.0E-04 | 1.3E-04 | 1.3E-04 | 1.3E-04 | 1.2E-04 | 1.1E-04 | 9.6E-05 |
| Type B Uncertainty(c) | 9.3E-04 | 2.3E-04 | 2.4E-04 | 2.4E-04 | 2.5E-04 | 2.5E-04 | 2.1E-04 | 1.8E-04 |
| Total Uncertainty(d) | 9.5E-04 | 3.8E-04 | 2.8E-04 | 2.7E-04 | 2.8E-04 | 2.8E-04 | 2.4E-04 | 2.0E-04 |
| Degrees of Freedom | 4370 | 23 | 175 | 177 | 178 | 253 | 237 | 181 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [  ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

Filter Type: B

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient(b) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 1.0E-04 | 1.5E-04 | 1.6E-04 | 1.5E-04 | 1.6E-04 | 1.5E-04 | 1.3E-04 | 1.1E-04 | +1 |
| Temperature | 1.7E-05 | 3.1E-05 | 1.4E-05 | 3.1E-05 | 1.7E-05 | 1.1E-05 | 9.5E-06 | 1.4E-05 | 0 |
| Wavelength | 7.5E-04 | 5.9E-05 | 1.4E-05 | 7.7E-07 | 2.6E-05 | 9.8E-05 | 7.8E-05 | 2.6E-05 | +1 |
| Stray Light | 2.0E-05 | 2.0E-05 | 2.2E-05 | 2.0E-05 | 2.5E-05 | 1.3E-05 | 1.8E-06 | 1.1E-05 | +1 |
| Beam Size & Position | 3.0E-05 | 4.4E-05 | 4.5E-05 | 4.5E-05 | 4.7E-05 | 4.2E-05 | 3.7E-05 | 3.3E-05 | +1 |
| Inter-reflection | 4.5E-05 | 6.6E-05 | 6.8E-05 | 6.6E-05 | 6.9E-05 | 6.3E-05 | 5.4E-05 | 4.9E-05 | +1 |
| Obliquity | 3.4E-05 | 5.0E-05 | 5.2E-05 | 5.1E-05 | 5.3E-05 | 4.8E-05 | 4.2E-05 | 3.8E-05 | 0 |
| Polarization | 7.1E-05 | 1.0E-04 | 1.1E-04 | 1.0E-04 | 1.1E-04 | 9.9E-05 | 8.6E-05 | 7.8E-05 | +1 |
| Source Drift & Fluctuation * | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A |
| Bandwidth | 4.9E-04 | 3.9E-05 | 9.4E-06 | 5.0E-07 | 1.7E-05 | 6.4E-05 | 5.1E-05 | 1.7E-05 | +1 |
| Filter Stability | 1.9E-04 | 5.9E-05 | 1.1E-04 | 1.1E-04 | 1.1E-04 | 1.0E-04 | 9.0E-05 | 8.1E-05 | 0 |
| Total Type B Uncertainty | 9.3E-04 | 2.3E-04 | 2.4E-04 | 2.4E-04 | 2.5E-04 | 2.5E-04 | 2.1E-04 | 1.8E-04 | N.A |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | N.A |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

*) “Source Drift & Fluctuation” is included in the Type A uncertainty.

Signature : 	Date :

Hiroshi Shitomi			2016-02-29

* Original data submitted:  2014-12-05

A.3 Measurement Results (3)

Blue: Revised at the Draft A stage to hide identification of the DUT filter set

Please reproduce the following two tables for each of the five filters measured at the completion of step 4 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: National Metrology Institute of Japan (NMIJ, AIST)

Filter Type: C

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 2.257E-02 | 9.594E-02 | 9.250E-02 | 7.755E-02 | 1.623E-01 | 1.509E-01 | 1.033E-01 | 7.695E-02 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Temperature | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 |
| Type A Uncertainty(b) | 3.9E-05 | 3.9E-05 | 1.0E-04 | 8.7E-05 | 1.8E-04 | 1.7E-04 | 1.2E-04 | 4.3E-05 |
| Type B Uncertainty(c) | 2.4E-04 | 9.2E-05 | 7.2E-05 | 5.6E-05 | 2.2E-04 | 1.1E-04 | 6.2E-05 | 5.0E-05 |
| Total Uncertainty(d) | 2.5E-04 | 1.0E-04 | 1.3E-04 | 1.0E-04 | 2.9E-04 | 2.0E-04 | 1.3E-04 | 6.6E-05 |
| Degrees of Freedom | 14615 | 408 | 20 | 18 | 56 | 18 | 15 | 51 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [  ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

Filter Type: C

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient(b) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 3.9E-06 | 1.7E-05 | 1.6E-05 | 1.3E-05 | 2.8E-05 | 2.6E-05 | 1.8E-05 | 1.3E-05 | +1 |
| Temperature | 4.5E-06 | 2.8E-05 | 1.3E-05 | 2.5E-05 | 2.2E-05 | 1.2E-05 | 7.2E-06 | 8.4E-06 | 0 |
| Wavelength | 2.0E-04 | 6.4E-05 | 4.3E-05 | 2.3E-05 | 1.7E-04 | 6.1E-05 | 2.4E-05 | 2.3E-05 | +1 |
| Stray Light | 1.8E-05 | 2.9E-06 | 3.6E-06 | 6.6E-06 | 1.0E-05 | 8.1E-06 | 1.4E-06 | 6.7E-06 | +1 |
| Beam Size & Position | 1.5E-06 | 6.4E-06 | 6.2E-06 | 5.2E-06 | 1.1E-05 | 1.0E-05 | 6.9E-06 | 5.1E-06 | +1 |
| Inter-reflection | 2.4E-06 | 1.0E-05 | 1.0E-05 | 8.4E-06 | 1.8E-05 | 1.6E-05 | 1.1E-05 | 8.4E-06 | +1 |
| Obliquity | 1.9E-06 | 7.9E-06 | 7.7E-06 | 6.4E-06 | 1.3E-05 | 1.2E-05 | 8.6E-06 | 6.4E-06 | 0 |
| Polarization | 5.0E-06 | 2.1E-05 | 2.1E-05 | 1.7E-05 | 3.6E-05 | 3.4E-05 | 2.3E-05 | 1.7E-05 | +1 |
| Source Drift & Fluctuation * | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A |
| Bandwidth | 1.3E-04 | 4.2E-05 | 2.8E-05 | 1.5E-05 | 1.1E-04 | 4.0E-05 | 1.6E-05 | 1.5E-05 | +1 |
| Filter Stability | 3.6E-05 | 3.1E-05 | 3.9E-05 | 3.3E-05 | 6.8E-05 | 6.3E-05 | 4.3E-05 | 3.2E-05 | 0 |
| Total Type B Uncertainty | 2.4E-04 | 9.2E-05 | 7.2E-05 | 5.6E-05 | 2.2E-04 | 1.1E-04 | 6.2E-05 | 5.0E-05 | N.A |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | N.A |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

*) “Source Drift & Fluctuation” is included in the Type A uncertainty.

Signature : 	Date :

Hiroshi Shitomi			2016-02-29

* Original data submitted:  2014-12-05

A.3 Measurement Results (4)

Blue: Revised at the Draft A stage to hide identification of the DUT filter set

Please reproduce the following two tables for each of the five filters measured at the completion of step 4 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: National Metrology Institute of Japan (NMIJ, AIST)

Filter Type: D

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 4.348E-04 | 5.516E-03 | 8.907E-03 | 8.811E-03 | 2.766E-02 | 3.343E-02 | 2.359E-02 | 1.750E-02 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Temperature | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 |
| Type A Uncertainty(b) | 7.0E-06 | 2.6E-05 | 1.3E-05 | 1.3E-05 | 5.6E-05 | 6.7E-05 | 4.7E-05 | 3.6E-05 |
| Type B Uncertainty(c) | 1.6E-05 | 3.4E-05 | 2.2E-05 | 1.7E-05 | 6.6E-05 | 5.6E-05 | 6.6E-05 | 2.9E-05 |
| Total Uncertainty(d) | 1.7E-05 | 4.3E-05 | 2.6E-05 | 2.1E-05 | 8.7E-05 | 8.7E-05 | 8.1E-05 | 4.6E-05 |
| Degrees of Freedom | 327 | 63 | 126 | 60 | 53 | 26 | 77 | 24 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [  ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

Filter Type: D

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient(b) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 1.6E-07 | 2.0E-06 | 3.2E-06 | 3.2E-06 | 1.0E-05 | 1.2E-05 | 8.6E-06 | 6.4E-06 | +1 |
| Temperature | 1.7E-07 | 3.4E-06 | 2.3E-06 | 5.1E-06 | 7.2E-06 | 4.5E-06 | 2.7E-06 | 3.0E-06 | 0 |
| Wavelength | 1.0E-05 | 2.0E-05 | 2.8E-06 | 5.3E-06 | 4.1E-05 | 1.0E-05 | 4.5E-05 | 3.5E-06 | +1 |
| Stray Light | 4.9E-06 | 3.9E-06 | 3.2E-06 | 3.2E-06 | 5.7E-07 | 1.7E-06 | 2.5E-07 | 1.5E-06 | +1 |
| Beam Size & Position | 1.1E-07 | 1.4E-06 | 2.3E-06 | 2.2E-06 | 7.0E-06 | 8.5E-06 | 6.0E-06 | 4.5E-06 | +1 |
| Inter-reflection | 4.7E-08 | 6.0E-07 | 9.7E-07 | 9.6E-07 | 3.0E-06 | 3.6E-06 | 2.6E-06 | 1.9E-06 | +1 |
| Obliquity | 3.6E-08 | 4.6E-07 | 7.4E-07 | 7.3E-07 | 2.3E-06 | 2.8E-06 | 2.0E-06 | 1.4E-06 | 0 |
| Polarization | 4.4E-07 | 5.6E-06 | 9.0E-06 | 8.9E-06 | 2.8E-05 | 3.4E-05 | 2.4E-05 | 1.8E-05 | +1 |
| Source Drift & Fluctuation * | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A |
| Bandwidth | 6.7E-06 | 1.3E-05 | 1.8E-06 | 3.5E-06 | 2.6E-05 | 6.6E-06 | 2.9E-05 | 2.3E-06 | +1 |
| Filter Stability | 8.8E-06 | 2.3E-05 | 1.9E-05 | 1.0E-05 | 3.3E-05 | 4.0E-05 | 2.8E-05 | 2.1E-05 | 0 |
| Total Type B Uncertainty | 1.6E-05 | 3.4E-05 | 2.2E-05 | 1.7E-05 | 6.6E-05 | 5.6E-05 | 6.6E-05 | 2.9E-05 | N.A |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | N.A |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

*) “Source Drift & Fluctuation” is included in the Type A uncertainty.

Signature : 	Date :

Hiroshi Shitomi			2016-02-29

* Original data submitted:  2014-12-05

A.3 Measurement Results (5)

Blue: Revised at the Draft A stage to hide identification of the DUT filter set

Please reproduce the following two tables for each of the five filters measured at the completion of step 4 of the comparison. All uncertainties should be reported as absolute uncertainties.

Laboratory: National Metrology Institute of Japan (NMIJ, AIST)

Filter Type: E

Table A-3i Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance(a) | 1.382E-05 | 3.292E-04 | 9.200E-04 | 9.956E-04 | 5.050E-03 | 9.602E-03 | 8.345E-03 | 7.055E-03 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Temperature | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 | 23.0 |
| Type A Uncertainty(b) | 2.6E-06 | 4.9E-06 | 5.0E-06 | 5.4E-06 | 7.7E-06 | 1.5E-05 | 1.3E-05 | 7.9E-06 |
| Type B Uncertainty(c) | 2.3E-06 | 2.9E-06 | 5.0E-06 | 5.4E-06 | 2.8E-05 | 4.8E-05 | 4.8E-05 | 3.5E-05 |
| Total Uncertainty(d) | 3.5E-06 | 5.7E-06 | 7.0E-06 | 7.6E-06 | 2.9E-05 | 5.0E-05 | 5.0E-05 | 3.6E-05 |
| Degrees of Freedom | 30 | 16 | 36 | 37 | 1860 | 1272 | 2114 | 3818 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Table A-3ii Type B Uncertainty Budget(a)

Fill out the table below for the uncertainty contributions in measurement at Step 4. All uncertainties should be reported as absolute uncertainties.

If the uncertainty budget has not changed from Step 2 (Table A-2ii), check here [ ] and only the Correlation Coefficient column of the following table needs to be filled out. Otherwise, please fill out the whole table.

Filter Type: E

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | Correlation Coefficient(b) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 4.8E-09 | 1.1E-07 | 3.2E-07 | 3.5E-07 | 1.8E-06 | 3.3E-06 | 2.9E-06 | 2.5E-06 | +1 |
| Temperature | 7.8E-09 | 3.0E-07 | 3.5E-07 | 8.1E-07 | 1.9E-06 | 1.8E-06 | 1.2E-06 | 1.9E-06 | 0 |
| Wavelength | 1.3E-06 | 1.7E-06 | 2.6E-07 | 4.9E-07 | 1.0E-05 | 2.2E-06 | 2.0E-05 | 2.3E-06 | +1 |
| Stray Light | 1.7E-06 | 1.6E-06 | 1.5E-06 | 1.5E-06 | 6.6E-07 | 2.7E-07 | 1.4E-08 | 2.8E-07 | +1 |
| Beam Size & Position | 4.5E-09 | 1.1E-07 | 3.0E-07 | 3.2E-07 | 1.6E-06 | 3.1E-06 | 2.7E-06 | 2.3E-06 | +1 |
| Inter-reflection | 1.5E-09 | 3.6E-08 | 1.0E-07 | 1.1E-07 | 5.5E-07 | 1.0E-06 | 9.1E-07 | 7.7E-07 | +1 |
| Obliquity | 1.1E-09 | 2.7E-08 | 7.6E-08 | 8.2E-08 | 4.2E-07 | 7.9E-07 | 6.9E-07 | 5.8E-07 | 0 |
| Polarization | 4.2E-08 | 1.0E-06 | 2.8E-06 | 3.0E-06 | 1.5E-05 | 2.9E-05 | 2.5E-05 | 2.2E-05 | +1 |
| Source Drift & Fluctuation * | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A | N.A |
| Bandwidth | 8.7E-07 | 1.1E-06 | 1.7E-07 | 3.2E-07 | 6.7E-06 | 1.4E-06 | 1.3E-05 | 1.5E-06 | +1 |
| Filter Stability | 3.9E-07 | 9.0E-07 | 3.8E-06 | 4.1E-06 | 2.0E-05 | 3.8E-05 | 3.3E-05 | 2.7E-05 | 0 |
| Total Type B Uncertainty | 2.3E-06 | 2.9E-06 | 5.0E-06 | 5.4E-06 | 2.8E-05 | 4.8E-05 | 4.8E-05 | 3.5E-05 | N.A |
| Degrees of Freedom | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ | N.A |

(a)Please record any uncertainties considered negligible as zero (rather than e.g. < some value) (b)Correlation coefficient of uncertainty component relative to measurements made during step 2, expressed as a number between -1 and 1.  The covariance between the step 2 and step 4 measurements will be calculated according to section 5.2 of the GUM [5]. (c)Add lines to the table as necessary, itemising other components of uncertainty considered

*) “Source Drift & Fluctuation” is included in the Type A uncertainty.

Signature : 	Date :

Hiroshi Shitomi			2016-02-29

* Original data submitted:  2014-12-05