Description of measurement facility and measurement method

Laboratory: All-Russian Research Institute for Optical and Physical Measurements

FGUP VNIIOFI

Step 2

Table A-1 Details of Measurement Setup

| Make and Type of Spectrophotometer | Double monochromator DTMc300 |
| --- | --- |
| Additional Stray Light Rejection | 10-5 |
| Source Drift Monitoring | No |
| Source | Tungsten Strip Lamp |
| Detector | HAMAMATSU Si photodiode with stabilisation of temperature at 25C |
| Temperature(a) | Room temperature 23 ± 1 C |
| Humidity | 50 ± 10 % |
| Beam Size | Diameter 17 mm |
| Beam Collimation | Full divergence angle of the beam is 48 at horizontal and 88 at vertical. |
| Measurement Sequence(b) | Set of 10 measurements with realignment and repositioning of the filter between each individual measurement |
| Bandwidth | 1,35 nm |

(a) i.e. describe method of temperature monitoring of filters and range of temperatures  (b) i.e. describe number of measurements and whether filter orientation with respect to beam changes between measurements

Description of measuring technique

Diagram of the measuring equipment is shown in Fig. 1

The tungsten strip lamp is used as the radiation source. The spectral lamp is used to calibrate the monochromator wavelength. Radiation beam coming out of the monochromator is collimated using an off-axis parabolic mirror. A collimated beam of radiation passes through the measurement sample and focused onto a detector using a second with parabolic mirror.

Measured sample is introduced into and removed from the beam using a linear positioner, which is controlled by computer. Adjusting the position of the sample perpendicular to the incident radiation beam is performed using a HeNe laser. Detector dark current is measured when the shatter 1 and shatter 2 are closed. All equipment is covered by an opaque cover and measurements are performed automatically under computer control.

Fig.1

Diagram of the measuring technique