Measurement Results

Laboratory: FGUP VNIIOFI                            Filter Identifier: A

Measurement results for Step 2. All uncertainties should be reported as absolute uncertainties.

Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance (a) | 0,9128892 | 0,9143505 | 0,9169858 | 0,9182640 | 0,9192960 | 0,9201133 | 0,9207546 | 0,9215064 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Type A Uncertainty(b) | 9,3433 E-05 | 1,2213 E-04 | 1,1255 E-04 | 1,0040 E-04 | 8,2904 E-05 | 7,5316 E-05 | 4,9124 E-05 | 6,8812 E-05 |
| Type B Uncertainty(c) | 4,9937 E-04 | 1,2933 E-03 | 1,3255 E-03 | 1,0052 E-03 | 7,6411 E-04 | 7,8985 E-04 | 7,1547 E-04 | 1,1462 E-03 |
| Total Uncertainty(d) | 5,0804 E-04 | 1,2991E -03 | 1,3302E-03 | 1,0102 E-03 | 7,6860 E-04 | 7,9343 E-04 | 7,1716 E-04 | 1,1483 E-03 |
| Degrees of Freedom | 7,8673 E+03 | 1,1521 E+05 | 1,7561 E+05 | 9,2225 E+04 | 6,6486 E+04 | 1,1085 E+05 | 4,0882 E+05 | 6,9785 E+05 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Type B Uncertainty Budget

Measurement results for Step 2. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: А

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 7,2171E-06 | 8,7833E-06 | 1,3860E-05 | 1,3983E-05 | 8,7377E-06 | 1,2001E-05 | 1,0165E-05 | 1,3118E-05 |
| Wavelength | 8,6271E-06 | 2,9978E-06 | 1,4648E-06 | 1,1898E-06 | 8,8534E-07 | 7,3765E-07 | 8,9227E-07 | 0 |
| Stray Light | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Beam Size & Position | 1,2702E-04 | 9,4685E-04 | 1,2471E-03 | 9,2953E-04 | 5,2539E-04 | 1,1547E-04 | 2,3094E-05 | 5,7735E-06 |
| Inter-reflection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Obliquity | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Polarization | 4,8282E-04 | 8,8096E-04 | 4,4883E-04 | 3,8224E-04 | 5,5476E-04 | 7,8127E-04 | 7,1503E-04 | 1,1461E-03 |
| Source Drift & Fluctuation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Bandwidth | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Total Type B Uncertainty | 4,9937E-04 | 1,2933E-03 | 1,3255E-03 | 1,0052E-03 | 7,6411E-04 | 7,8985E-04 | 7,1547E-04 | 1,1462E-03 |
| Degrees of Freedom | 7,8673 E+03 | 1,1521 E+05 | 1,7561 E+05 | 9,2225 E+04 | 6,6486 E+04 | 1,1085 E+05 | 4,0882 E+05 | 6,9785 E+05 |

Measurement Results

Measurement results for Step 2. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: B

Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance (a) | 0,4090793 | 0,6074100 | 0,6215563 | 0,6097502 | 0,6371600 | - | 0,5016254 | - |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | - | 10 | - |
| Type A Uncertainty(b) | 5,0961 E-05 | 3,3860 E-05 | 2,8999 E-05 | 1,2861 E-05 | 2,2239 E-05 | - | 2,3165 E-05 | - |
| Type B Uncertainty(c) | 1,2370 E-03 | 6,6428 E-04 | 5,1607 E-04 | 4,1801 E-04 | 4,3858 E-04 | - | 4,5284 E-04 | - |
| Total Uncertainty(d) | 1,2380 E-03 | 6,6515 E-04 | 5,1688 E-04 | 4,1821 E-04 | 4,3915 E-04 | - | 4,5343 E-04 | - |
| Degrees of Freedom | 3,1348 E+06 | 1,3402 E+06 | 9,0845 E+05 | 1,0062 E+07 | 1,3683 E+06 | - | 1,3211 E+06 | - |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Type B Uncertainty Budget

Measurement results for Step 2. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: В

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 1,2825E-05 | 1,1692E-05 | 9,9724E-06 | 9,1111E-06 | 1,6413E-06 | - | 1,0161E-05 | - |
| Wavelength | 1,1453E-03 | 1,6461E-05 | 1,3738E-05 | 3,1704E-05 | 6,9918E-05 | - | 5,5641E-05 | - |
| Stray Light | 0 | 0 | 0 | 0 | 0 | - | 0 | - |
| Beam Size & Position | 3,7528E-04 | 5,3116E-04 | 4,3301E-04 | 3,8105E-04 | 3,2909E-04 | - | 2,7713E-04 | - |
| Inter-reflection | 0 | 0 | 0 | 0 | 0 | - | 0 | - |
| Obliquity | 0 | 0 | 0 | 0 | 0 | - | 0 | - |
| Polarization | 2,7836E-04 | 3,9841E-04 | 2,8025E-04 | 1,6866E-04 | 2,8136E-04 | - | 3,5365E-04 | - |
| Source Drift & Fluctuation | 0 | 0 | 0 | 0 | 0 | - | 0 | - |
| Bandwidth | 0 | 0 | 0 | 0 | 0 | - | 0 | - |
| Total Type B Uncertainty | 1,2370E-03 | 6,6428E-04 | 5,1607E-04 | 4,1801E-04 | 4,3858E-04 | - | 4,5284E-04 | - |
| Degrees of Freedom | 3,1348 E+06 | 1,3402 E+06 | 9,0845 E+05 | 1,0062 E+07 | 1,3683 E+06 | - | 1,3211 E+06 | - |

Measurement Results

Measurement results for Step 2. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: C

Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance (a) | 0,0211483 | 0,0940485 | 0,0907277 | 0,0760497 | 0,1598392 | 0,1484337 | 0,1014803 | 0,0754358 |
| Number of Measurements | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Type A Uncertainty(b) | 4,5538 E-05 | 2,3188 E-05 | 1,1637 E-05 | 1,5709 E-05 | 1,5456 E-05 | 5,2144 E-06 | 1,5937 E-06 | 8,7464 E-07 |
| Type B Uncertainty(c) | 1,1403 E-03 | 1,0381 E-04 | 5,5074 E-05 | 1,0818 E-04 | 1,5653 E-04 | 1,2157 E-04 | 9,7219 E-05 | 1,1790 E-04 |
| Total Uncertainty(d) | 1,1412 E-03 | 1,0637 E-04 | 5,6290 E-05 | 1,0932 E-04 | 1,5729 E-04 | 1,2168 E-04 | 9,7232 E-05 | 1,1790 E-04 |
| Degrees of Freedom | 3,5502 E+06 | 3,9859 E+03 | 4,9277 E+03 | 2,1107 E+04 | 9,6522 E+04 | 2,6687 E+06 | 1,2468 E+08 | 2,9720 E+09 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Type B Uncertainty Budget

Measurement results for Step 2. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: C

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity | 2,2910E-06 | 2,8574E-06 | 6,5889E-06 | 7,4669E-06 | 4,4322E-06 | 8,3053E-06 | 7,1397E-06 | 7,2453E-06 |
| Wavelength | 4,2106E-04 | 3,8198E-06 | 1,6956E-05 | 9,6754E-05 | 1,3168E-05 | 5,4218E-05 | 3,0073E-05 | 8,7106E-05 |
| Stray Light | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Beam Size & Position | 8,6603E-05 | 9,8150E-05 | 5,1962E-05 | 4,6188E-05 | 1,5588E-04 | 9,8150E-05 | 8,6603E-05 | 6,9282E-05 |
| Inter-reflection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Obliquity | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Polarization | 1,0562E-03 | 3,3486E-05 | 1,5011E-06 | 1,2394E-05 | 2,8098E-06 | 4,6227E-05 | 3,1562E-05 | 3,8221E-05 |
| Source Drift & Fluctuation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Bandwidth | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Total Type B Uncertainty | 1,1403E-03 | 1,0381E-04 | 5,5074E-05 | 1,0818E-04 | 1,5653E-04 | 1,2157E-04 | 9,7219E-05 | 1,1790E-04 |
| Degrees of Freedom | 3,5502 E+06 | 3,9859 E+03 | 4,9277 E+03 | 2,1107 E+04 | 9,6522 E+04 | 2,6687 E+06 | 1,2468 E+08 | 2,9720 E+09 |

Measurement Results

Measurement results for Step 2. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: D

Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance (a) |  | 0,0052079 | 0,0084320 | 0,0083527 | 0,0265527 | 0,0321930 | 0,0224921 | 0,0166450 |
| Number of Measurements |  | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Type A Uncertainty(b) |  | 1,0683 E-05 | 1,8692 E-06 | 4,0983 E-06 | 1,1828 E-05 | 1,8993 E-06 | 1,2677 E-06 | 3,4157 E-06 |
| Type B Uncertainty(c) |  | 1,2583 E-04 | 9,3142 E-06 | 2,3010 E-05 | 3,1560 E-05 | 2,2706 E-05 | 1,3137 E-05 | 2,2533 E-05 |
| Total Uncertainty(d) |  | 1,2628 E-04 | 9,4999 E-06 | 2,3372 E-05 | 3,3704 E-05 | 2,2786 E-05 | 1,3198 E-05 | 2,2790 E-05 |
| Degrees of Freedom |  | 1,7573 E+05 | 6,0053 E+03 | 9,5201 E+03 | 5,9328 E+02 | 1,8642 E+05 | 1,0576 E+05 | 1,7838 E+04 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Type B Uncertainty Budget

Measurement results for Step 2. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: D

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity |  | 5,0863E-06 | 7,2938E-06 | 7,2350E-06 | 5,9134E-06 | 3,1843E-06 | 2,6306E-06 | 6,0397E-06 |
| Wavelength |  | 3,7364E-06 | 9,4351E-08 | 2,1018E-05 | 6,5124E-06 | 1,1202E-05 | 6,7509E-06 | 1,9221E-05 |
| Stray Light |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Beam Size & Position |  | 0 | 5,7735E-06 | 5,7735E-06 | 2,8868E-05 | 1,1547E-05 | 5,7735E-06 | 5,7735E-06 |
| Inter-reflection |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Obliquity |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Polarization |  | 1,2567E-04 | 4,6188E-07 | 1,4241E-06 | 9,2376E-06 | 1,5704E-05 | 9,3146E-06 | 8,2754E-06 |
| Source Drift & Fluctuation |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Bandwidth |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Total Type B Uncertainty |  | 1,2583E-04 | 9,3142E-06 | 2,3010E-05 | 3,1560E-05 | 2,2706E-05 | 1,3137E-05 | 2,2533E-05 |
| Degrees of Freedom |  | 1,7573 E+05 | 6,0053 E+03 | 9,5201 E+03 | 5,9328 E+02 | 1,8642 E+05 | 1,0576 E+05 | 1,7838 E+04 |

Measurement Results

Measurement results for Step 2. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: E

Measurement Results

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Transmittance (a) |  |  | 0,0009569 | 0,0010272 | 0,0051476 | 0,0098284 | 0,0085614 | 0,0072348 |
| Number of Measurements |  |  | 10 | 10 | 10 | 10 | 10 | 10 |
| Type A Uncertainty(b) |  |  | 1,0714 E-06 | 1,5581 E-06 | 3,6551 E-06 | 1,9391E-06 | 2,3152 E-06 | 3,8781 E-06 |
| Type B Uncertainty(c) |  |  | 8,1914 E-06 | 1,0497 E-05 | 8,2371 E-06 | 1,8227E-05 | 1,9007 E-05 | 1,5653 E-05 |
| Total Uncertainty(d) |  |  | 8,2612 E-06 | 1,0612 E-05 | 9,0117 E-06 | 1,8330E-05 | 1,9147 E-05 | 1,6126 E-05 |
| Degrees of Freedom |  |  | 3,1810 E+04 | 1,9362 E+04 | 3,3254 E+02 | 7,1870E+04 | 4,2104 E+04 | 2,6907 E+03 |

(a)Spectral transmittance. The value of the spectral transmittance of the central 17 mm diameter of the filter as measured by the participant laboratory. (b)Type A Uncertainty. The uncertainties associated with the spectral transmittance values attributed to reproducibility of the measurement. (c)Type B Uncertainty. The uncertainties associated with the spectral transmittance values attributed to all type B sources. (d)Total Uncertainty. The total uncertainty of the measurement of spectral transmittance for a coverage factor of k=1.

Type B Uncertainty Budget

Measurement results for Step 2. All uncertainties should be reported as absolute uncertainties.

Laboratory:  FGUP VNIIOFI                            Filter Identifier: E

| Wavelength (nm) | 380 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nonlinearity |  |  | 7,7485E-06 | 4,5164E-06 | 2,1536E-06 | 5,1232E-06 | 7,2818E-06 | 6,1808E-06 |
| Wavelength |  |  | 7,9151E-08 | 4,7611E-06 | 5,4001E-06 | 1,4631E-06 | 1,5319E-06 | 8,3531E-06 |
| Stray Light |  |  | 0 | 0 | 0 | 0 | 0 | 0 |
| Beam Size & Position |  |  | 0 | 5,7735E-06 | 5,7735E-06 | 1,7321E-05 | 1,7321E-05 | 1,1547E-05 |
| Inter-reflection |  |  | 0 | 0 | 0 | 0 | 0 | 0 |
| Obliquity |  |  | 0 | 0 | 0 | 0 | 0 | 0 |
| Polarization |  |  | 2,6558E-06 | 5,8120E-06 | 8,4678E-07 | 1,9630E-06 | 2,4249E-06 | 1,9245E-06 |
| Source Drift & Fluctuation |  |  | 0 | 0 | 0 | 0 | 0 | 0 |
| Bandwidth |  |  | 0 | 0 | 0 | 0 | 0 | 0 |
| Total Type B Uncertainty |  |  | 8,1914E-06 | 1,0497E-05 | 8,2371E-06 | 1,8227E-05 | 1,9007E-05 | 1,5653E-05 |
| Degrees of Freedom |  |  | 3,1810 E+04 | 1,9362 E+04 | 3,3254 E+02 | 7,1870 E+04 | 4,2104 E+04 | 2,6907 E+03 |