## eps
Estd-Limit(e) |  | 1 | [80 empty] | last rev. 23.1.2018 by JH-CMI
5 VA | I1 (%IN) | CMI |  |  |  |  | PTB |  |  |  |  | METAS |  |  |  |  | NPL |  |  |  |  | LCOE |  |  |  |  | MIKES |  |  |  |  | BEV |  |  |  |  | VSL |  |  |  |  | SP |  |  |  |  | DMDM |  |  |  |  | GUM |  |  |  |  | BIM |  |  |  |  | weighted mean er (ppm) | u(er) (ppm) | U(er) (ppm) | u(estd) (ppm) | UME |  |  |  |  | LNE |  |  |  |  | INRIM
kI (kA/A) |  | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) |  |  |  |  | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-)
4/5 | 120 | -92.525 | 8.5 | -0.35875 | 9.516 | 0.01885 | -92 | 4 | 0.16625 | 5.857 | 0.014192 | -86 | 11 | 6.1662 | 11.8027 | 0.26122 | -94 | 15 | -1.8338 | 15.5982 | 0.058781 | -91 | 21 | 1.1662 | 21.4314 | 0.027209 | -105 | 13 | -12.8338 | 13.6859 | 0.46887 | -88 | 18.5 | 4.1662 | 18.9883 | 0.10971 | -98 | 7 | -5.8338 | 8.2039 | 0.35555 | [5 empty] | -84.8333 | 36.0258 | 7.3329 | 36.279 | 0.10106 | -87 | 19.4 | 5.1662 | 19.8662 | 0.13003 | -84 | 11 | 8.1662 | 11.8027 | 0.34595 | -92.1662 | 2.5876 | 5.1753 | 5 | -89 | 10.3 | 3.1662 | 11.1532 | 0.14194 | -91.8063 | 17.6245 | 0.35992 | 18.1363 | 0.0099226 | -89 | 15 | 3.1662 | 15.5982 | 0.10149
| 100 | -90.1833 | 8.5 | 0.50235 | 9.5281 | 0.026361 | -91 | 4 | -0.31432 | 5.8767 | 0.026743 | -84 | 11 | 6.6857 | 11.8125 | 0.28299 | -98 | 15 | -7.3143 | 15.6056 | 0.23435 | -89 | 21 | 1.6857 | 21.4368 | 0.039318 | -101 | 12 | -10.3143 | 12.7489 | 0.40452 | -85 | 18.5 | 5.6857 | 18.9943 | 0.14967 | -96 | 7 | -5.3143 | 8.218 | 0.32333 | -92 | 20 | -1.3143 | 20.4581 | 0.032122 | -82.5 | 36.0295 | 8.1857 | 36.2858 | 0.11279 | -85 | 19.4 | 5.6857 | 19.872 | 0.14306 | -84 | 10 | 6.6857 | 10.8874 | 0.30704 | -90.6857 | 2.5426 | 5.0851 | 5 | -86 | 10.3 | 4.6857 | 11.1636 | 0.20986 | -89.4537 | 17.4147 | 1.232 | 17.939 | 0.034338 | -87 | 15 | 3.6857 | 15.6056 | 0.11809
| 50 | -80.5 | 8.5 | 1.4482 | 9.036 | 0.080135 | -83 | 4 | -1.0518 | 5.0398 | 0.10435 | -76 | 11 | 5.9482 | 11.4192 | 0.26045 | -90 | 15 | -8.0518 | 15.3101 | 0.26296 | -77 | 20.5 | 4.9482 | 20.728 | 0.11936 | -89 | 11 | -7.0518 | 11.4192 | 0.30877 | -78 | 18.5 | 3.9482 | 18.7523 | 0.10527 | -85 | 8 | -3.0518 | 8.5673 | 0.17811 | -86 | 20 | -4.0518 | 20.2336 | 0.10013 | -75.1667 | 36.0196 | 6.7815 | 36.1499 | 0.093797 | -74 | 19.4 | 7.9482 | 19.6408 | 0.20234 | -78 | 10 | 3.9482 | 10.4594 | 0.18874 | -81.9482 | 2.5692 | 5.1384 | 4 | -78 | 10.2 | 3.9482 | 10.6508 | 0.18535 | -79.4339 | 17.484 | 2.5143 | 17.7508 | 0.070823 | -79 | 15 | 2.9482 | 15.3101 | 0.096283
| 20 | -70.95 | 8.5 | 0.78441 | 9.0472 | 0.043351 | -73 | 4 | -1.2656 | 5.0597 | 0.12506 | -67 | 11 | 4.7344 | 11.4281 | 0.20714 | -72 | 15 | -0.26559 | 15.3167 | 0.0086699 | -74 | 20.5 | -2.2656 | 20.7328 | 0.054638 | -76 | 11 | -4.2656 | 11.4281 | 0.18663 | -63 | 18.5 | 8.7344 | 18.7577 | 0.23282 | -73 | 7 | -1.2656 | 7.6551 | 0.082663 | -85 | 20 | -13.2656 | 20.2386 | 0.32773 | -62.8333 | 36.0184 | 8.9011 | 36.1514 | 0.12311 | -59 | 19.4 | 12.7344 | 19.6459 | 0.3241 | -68 | 10 | 3.7344 | 10.4691 | 0.17835 | -71.7344 | 2.5296 | 5.0592 | 4 | -70 | 10.2 | 1.7344 | 10.6603 | 0.081349 | -72.4854 | 17.4951 | -0.75102 | 17.7674 | 0.021135 | -69 | 15 | 2.7344 | 15.3167 | 0.089263
| 10 | -65.1167 | 8.5 | 1.8579 | 9.039 | 0.10277 | -68 | 4 | -1.0255 | 5.0452 | 0.10163 | -63 | 12 | 3.9745 | 12.3877 | 0.16042 | -57 | 15 | 9.9745 | 15.3119 | 0.32571 | -70 | 37 | -3.0255 | 37.1275 | 0.040744 | -77 | 11 | -10.0255 | 11.4217 | 0.43888 | -62 | 18.5 | 4.9745 | 18.7538 | 0.13263 | -68 | 7 | -1.0255 | 7.6455 | 0.067062 | -92 | 22 | -25.0255 | 22.2138 | 0.56329 | -44 | 36.0153 | 22.9745 | 36.1463 | 0.3178 | -48 | 19.4 | 18.9745 | 19.6422 | 0.48301 | -63 | 10 | 3.9745 | 10.462 | 0.18995 | -66.9745 | 2.5585 | 5.1169 | 4 | -67 | 10.2 | -0.025454 | 10.6534 | 0.0011947 | -69.1213 | 17.7714 | -2.1468 | 18.0354 | 0.059516 | -65 | 15 | 1.9745 | 15.3119 | 0.064477
| 5 | -61.475 | 8.5 | 0.92872 | 8.9988 | 0.051603 | -64 | 4 | -1.5963 | 4.9727 | 0.1605 | -59 | 18 | 3.4037 | 18.2408 | 0.093299 | -39 | 15 | 23.4037 | 15.2882 | 0.76542 | -68 | 37 | -5.5963 | 37.1178 | 0.075385 | -68 | 11 | -5.5963 | 11.3898 | 0.24567 | -52 | 27.3 | 10.4037 | 27.4594 | 0.18944 | -64 | 7 | -1.5963 | 7.5979 | 0.10505 | -111 | 30 | -48.5963 | 30.1451 | 0.80604 | -16 | 44.7547 | 46.4037 | 44.8521 | 0.5173 | -32 | 19.4 | 30.4037 | 19.6237 | 0.77467 | -60 | 15 | 2.4037 | 15.2882 | 0.078614 | -62.4037 | 2.6967 | 5.3934 | 4 | -64 | 12.3 | -1.5963 | 12.6498 | 0.063095 | -65.7714 | 17.8798 | -3.3676 | 18.1223 | 0.092914 | -62 | 15 | 0.40372 | 15.2882 | 0.013204
| 2 | -58.1 | 8.5 | 0.63029 | 9.4549 | 0.033332 | -60 | 4 | -1.2697 | 5.7571 | 0.11027 | -55 | 18 | 3.7303 | 18.4701 | 0.10098 | -10 | 26 | 48.7303 | 26.3276 | 0.92546 | -71 | 38 | -12.2697 | 38.2249 | 0.16049 | -64 | 12 | -5.2697 | 12.6943 | 0.20756 | -50 | 41.5 | 8.7303 | 41.706 | 0.10466 | -58 | 7 | 0.73029 | 8.1329 | 0.044897 | -193 | 50 | -134.2697 | 50.1712 | 1.3381 | [5 empty] | 5 | 28.4 | 63.7303 | 28.7003 | 1.1103 | -58 | 15 | 0.73029 | 15.561 | 0.023465 | -58.7303 | 2.8028 | 5.6055 | 5 | -60 | 13.3 | -1.2697 | 13.9296 | 0.045576 | -61.5803 | 18.4313 | -2.85 | 18.8907 | 0.075435 | -59 | 18 | -0.26971 | 18.4701 | 0.0073013
| 1 | -56.2833 | 8.5 | -0.45451 | 9.4333 | 0.024091 | -57 | 4 | -1.1712 | 5.7216 | 0.10235 | -51 | 18 | 4.8288 | 18.4591 | 0.1308 | 26 | 35 | 81.8288 | 35.2383 | 1.1611 | -81 | 38 | -25.1712 | 38.2196 | 0.3293 | -53 | 11 | 2.8288 | 11.7361 | 0.12052 | -46 | 51.2 | 9.8288 | 51.3632 | 0.09568 | -54 | 7 | 1.8288 | 8.1078 | 0.11278 | -329 | 100 | -273.1712 | 100.0836 | 1.3647 | [10 empty] | -60 | 15 | -4.1712 | 15.5479 | 0.13414 | -55.8288 | 2.8746 | 5.7491 | 5 | -56 | 15.4 | -0.17117 | 15.9341 | 0.0053713 | [5 empty] | -57 | 21 | -1.1712 | 21.3948 | 0.027371
10/5 | 120 | [5 empty] | -34 | 3.5 | -0.85164 | 4.2997 | 0.099035 | -33 | 11 | 0.14836 | 11.28 | 0.0065764 | [15 empty] | -30 | 18.5 | 3.1484 | 18.6678 | 0.084326 | [25 empty] | -33.1484 | 3.1245 | 6.2491 | 4 | -27 | 10.2 | 6.1484 | 10.5013 | 0.29274
| 100 | -30.75 | 8.5 | 0.90801 | 8.9634 | 0.050651 | -33 | 3.5 | -1.342 | 4.5102 | 0.14877 | -32 | 11 | -0.34199 | 11.3619 | 0.01505 | -33 | 15 | -1.342 | 15.2674 | 0.04395 | [10 empty] | -28 | 18.5 | 3.658 | 18.7174 | 0.097717 | [15 empty] | -18 | 19.4 | 13.658 | 19.6075 | 0.34829 | [5 empty] | -31.658 | 2.8121 | 5.6241 | 4 | -28 | 10.2 | 3.658 | 10.5893 | 0.17272 | [5 empty] | -24 | 18 | 7.658 | 18.2234 | 0.21011
| 50 | -21.3 | 8.5 | 0.62389 | 9.0589 | 0.034435 | -24 | 3.5 | -2.0761 | 4.6972 | 0.22099 | -24 | 11 | -2.0761 | 11.4374 | 0.09076 | -22 | 15 | -0.076107 | 15.3236 | 0.0024833 | -24 | 26.5 | -2.0761 | 26.6845 | 0.038901 | -22 | 13 | -0.076107 | 13.3721 | 0.0028457 | -9 | 18.5 | 12.9239 | 18.7634 | 0.34439 | -19 | 7 | 2.9239 | 7.669 | 0.19063 | [5 empty] | 12.3333 | 43.2655 | 34.2572 | 43.3787 | 0.39486 | -14 | 19.4 | 7.9239 | 19.6513 | 0.20161 | -14 | 15 | 7.9239 | 15.3236 | 0.25855 | -21.9239 | 2.4873 | 4.9745 | 4 | -21 | 10.1 | 0.92389 | 10.5747 | 0.043684 | -28.3301 | 21.0096 | -6.4062 | 21.2419 | 0.15079 | -17 | 18 | 4.9239 | 18.2706 | 0.13475
| 20 | -11.1 | 8.5 | 1.0857 | 9.0689 | 0.059857 | -14 | 3.5 | -1.8143 | 4.7165 | 0.19234 | -9 | 11 | 3.1857 | 11.4453 | 0.13917 | -18 | 15 | -5.8143 | 15.3296 | 0.18964 | -7 | 26.5 | 5.1857 | 26.6879 | 0.097154 | -12 | 11 | 0.18567 | 11.4453 | 0.0081113 | -3 | 18.5 | 9.1857 | 18.7682 | 0.24471 | -9 | 7 | 3.1857 | 7.6809 | 0.20738 | -27 | 20 | -14.8143 | 20.2484 | 0.36582 | 54.6667 | 42.3705 | 66.8523 | 42.4883 | 0.78671 | -2 | 19.4 | 10.1857 | 19.6559 | 0.2591 | -8.7 | 15 | 3.4857 | 15.3296 | 0.11369 | -12.1857 | 2.4504 | 4.9007 | 4 | -10 | 10.1 | 2.1857 | 10.5833 | 0.10326 | -23.8175 | 20.8486 | -11.6319 | 21.087 | 0.27581 | -14 | 18 | -1.8143 | 18.2756 | 0.049638
| 10 | -6.2 | 8.5 | 0.78608 | 9.0639 | 0.043364 | -9 | 3.5 | -2.0139 | 4.7068 | 0.21394 | -1 | 12 | 5.9861 | 12.4058 | 0.24126 | -2 | 15 | 4.9861 | 15.3266 | 0.16266 | -9 | 26 | -2.0139 | 26.1898 | 0.038449 | -10 | 11 | -3.0139 | 11.4413 | 0.13171 | -2 | 18.5 | 4.9861 | 18.7658 | 0.13285 | -2 | 7 | 4.9861 | 7.6749 | 0.32483 | -37 | 22 | -30.0139 | 22.2239 | 0.67526 | 81.8333 | 41.6972 | 88.8194 | 41.8158 | 1.062 | 1 | 19.4 | 7.9861 | 19.6536 | 0.20317 | -3.8 | 15 | 3.1861 | 15.3266 | 0.10394 | -6.9861 | 2.4691 | 4.9381 | 4 | -6 | 10.1 | 0.98608 | 10.5789 | 0.046606 | -18.3261 | 20.9118 | -11.34 | 21.1473 | 0.26812 | -10 | 21 | -3.0139 | 21.2345 | 0.070967
| 5 | -2.2 | 8.5 | 0.51879 | 8.6328 | 0.030047 | -5 | 3.5 | -2.2812 | 3.8113 | 0.29927 | 4 | 19 | 6.7188 | 19.0598 | 0.17626 | 16 | 15 | 18.7188 | 15.0757 | 0.62083 | -6 | 40 | -3.2812 | 40.0284 | 0.040986 | -3 | 11 | -0.28121 | 11.103 | 0.012664 | 8 | 27.3 | 10.7188 | 27.3417 | 0.19602 | 1 | 7 | 3.7188 | 7.1607 | 0.25967 | -53 | 30 | -50.2812 | 30.0379 | 0.83696 | 126.3333 | 51.6372 | 129.0521 | 51.6593 | 1.2491 | 8 | 19.4 | 10.7188 | 19.4586 | 0.27543 | 1.1 | 37 | 3.8188 | 37.0307 | 0.051562 | -2.7188 | 2.5931 | 5.1861 | 3 | -3 | 12.2 | -0.28121 | 12.2929 | 0.011438 | -17.2861 | 22.0031 | -14.5673 | 22.0548 | 0.33025 | -3 | 24 | -0.28121 | 24.0474 | 0.0058471
| 2 | 1.6 | 8.5 | 1.4497 | 8.6045 | 0.084244 | -2 | 3.5 | -2.1503 | 3.7466 | 0.28696 | 8 | 19 | 7.8497 | 19.047 | 0.20606 | 47 | 26 | 46.8497 | 26.0343 | 0.89977 | -5 | 49 | -5.1503 | 49.0182 | 0.052534 | -1 | 11 | -1.1503 | 11.0809 | 0.051902 | 10 | 41.5 | 9.8497 | 41.5215 | 0.11861 | 5 | 7 | 4.8497 | 7.1265 | 0.34026 | -144 | 50 | -144.1503 | 50.0179 | 1.441 | [5 empty] | 27 | 28.4 | 26.8497 | 28.4314 | 0.47218 | 5.3 | 37 | 5.1497 | 37.0241 | 0.069546 | 0.15025 | 2.6857 | 5.3713 | 3 | 1 | 13.2 | 0.84975 | 13.2675 | 0.032024 | -11.1614 | 29.6196 | -11.3116 | 29.6498 | 0.19075
| 1 | 0.65 | 8.5 | -0.13201 | 8.5726 | 0.0076994 | -1 | 3.5 | -1.782 | 3.6727 | 0.2426 | 10 | 19 | 9.218 | 19.0326 | 0.24216 | 85 | 35 | 84.218 | 35.0177 | 1.2025 | -11 | 49 | -11.782 | 49.0126 | 0.12019 | 0 | 11 | -0.78201 | 11.0562 | 0.035365 | 13 | 51.2 | 12.218 | 51.2121 | 0.11929 | 6 | 8 | 5.218 | 8.0771 | 0.32301 | -287 | 100 | -287.782 | 100.0062 | 1.4388 | [10 empty] | 3.2 | 37 | 2.418 | 37.0167 | 0.032661 | 0.78201 | 2.7859 | 5.5717 | 3 | 2 | 15.3 | 1.218 | 15.3404 | 0.039699
15 VA | I1 (%IN) | CMI |  |  |  |  | PTB |  |  |  |  | METAS |  |  |  |  | NPL |  |  |  |  | LCOE |  |  |  |  | MIKES |  |  |  |  | BEV |  |  |  |  | VSL |  |  |  |  | SP |  |  |  |  | DMDM |  |  |  |  | GUM |  |  |  |  | BIM |  |  |  |  | weighted mean er (ppm) | u(er) (ppm) | U(er) (ppm) | u(estd) (ppm) | UME |  |  |  |  | LNE |  |  |  |  | INRIM
kI (kA/A) |  | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) |  |  |  |  | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-) | eL (ppm) | u(eL) (ppm) | De (ppm) | u(Destd) (ppm) | E(e) (-)
4/5 | 120 | -131.45 | 8.5 | -3.313 | 9.4984 | 0.1744 | -130 | 4 | -1.863 | 5.8283 | 0.15982 | -121 | 11 | 7.137 | 11.7885 | 0.30271 | -130 | 15 | -1.863 | 15.5875 | 0.059759 | -129 | 21 | -0.86297 | 21.4236 | 0.020141 | -138 | 12 | -9.863 | 12.7267 | 0.38749 | -126 | 18.5 | 2.137 | 18.9794 | 0.056298 | -135 | 9 | -6.863 | 9.9483 | 0.34493 | [5 empty] | -120.8333 | 36.0953 | 7.3037 | 36.3433 | 0.10048 | -120 | 19.4 | 8.137 | 19.8577 | 0.20488 | -123 | 11 | 5.137 | 11.7885 | 0.21788 | -128.137 | 2.6515 | 5.3031 | 5 | -123 | 10.3 | 5.137 | 11.1382 | 0.2306 | -85.8843 | 17.8292 | 42.2528 | 18.3262 | 1.1528 | -126 | 15 | 2.137 | 15.5875 | 0.068549
| 100 | -128.3 | 8.5 | -1.5651 | 10.0776 | 0.077654 | -128 | 4 | -1.2651 | 6.7311 | 0.093977 | -120 | 11 | 6.7349 | 12.26 | 0.27467 | -133 | 15 | -6.2651 | 15.947 | 0.19644 | -126 | 21 | 0.73487 | 21.6866 | 0.016943 | -135 | 11 | -8.2651 | 12.26 | 0.33708 | -124 | 18.5 | 2.7349 | 19.2758 | 0.07094 | -134 | 8 | -7.2651 | 9.6596 | 0.37606 | -136 | 20 | -9.2651 | 20.7197 | 0.22358 | -118.5 | 36.2096 | 8.2349 | 36.6121 | 0.11246 | -118 | 19.4 | 8.7349 | 20.1412 | 0.21684 | -122 | 11 | 4.7349 | 12.26 | 0.1931 | -126.7349 | 2.587 | 5.1739 | 6 | -121 | 10.3 | 5.7349 | 11.636 | 0.24643 | -84.1859 | 17.5681 | 42.549 | 18.3833 | 1.1573 | -125 | 15 | 1.7349 | 15.947 | 0.054395
| 50 | -117 | 8.5 | -1.0523 | 10.0779 | 0.052206 | -117 | 4 | -1.0523 | 6.7316 | 0.078158 | -111 | 11 | 4.9477 | 12.2603 | 0.20178 | -121 | 15 | -5.0523 | 15.9472 | 0.15841 | -113 | 20.5 | 2.9477 | 21.2029 | 0.069513 | -123 | 11 | -7.0523 | 12.2603 | 0.28761 | -114 | 18.5 | 1.9477 | 19.276 | 0.050522 | -121 | 8 | -5.0523 | 9.6599 | 0.26151 | -121 | 20 | -5.0523 | 20.7199 | 0.12192 | -106 | 36.1042 | 9.9477 | 36.5079 | 0.13624 | -109 | 19.4 | 6.9477 | 20.1413 | 0.17247 | -114 | 11 | 1.9477 | 12.2603 | 0.079433 | -115.9477 | 2.5857 | 5.1715 | 6 | -113 | 10.3 | 2.9477 | 11.6363 | 0.12666 | -76.3392 | 17.5022 | 39.6086 | 18.3205 | 1.081 | -116 | 15 | -0.052257 | 15.9472 | 0.0016384
| 20 | -100.85 | 8.5 | -0.71622 | 9.5271 | 0.037588 | -102 | 4 | -1.8662 | 5.8751 | 0.15883 | -97 | 11 | 3.1338 | 11.8117 | 0.13266 | -99 | 15 | 1.1338 | 15.605 | 0.036327 | -97 | 20.5 | 3.1338 | 20.9467 | 0.074803 | -106 | 11 | -5.8662 | 11.8117 | 0.24832 | -94 | 18.5 | 6.1338 | 18.9938 | 0.16147 | -102 | 7 | -1.8662 | 8.2168 | 0.11356 | -104 | 20 | -3.8662 | 20.4577 | 0.094493 | -88.8333 | 36.048 | 11.3004 | 36.3039 | 0.15564 | -87 | 19.4 | 13.1338 | 19.8715 | 0.33047 | -98 | 11 | 2.1338 | 11.8117 | 0.090325 | -100.1338 | 2.5463 | 5.0927 | 5 | -100 | 10.3 | 0.13378 | 11.1627 | 0.0059922 | -68.6825 | 17.8292 | 31.4513 | 18.3412 | 0.8574 | -99 | 15 | 1.1338 | 15.605 | 0.036327
| 10 | -92.275 | 8.5 | 0.26268 | 9.5233 | 0.013792 | -94 | 4 | -1.4623 | 5.8688 | 0.12458 | -89 | 12 | 3.5377 | 12.7453 | 0.13878 | -80 | 15 | 12.5377 | 15.6026 | 0.40178 | -94 | 37 | -1.4623 | 37.2484 | 0.019629 | -101 | 10 | -8.4623 | 10.8831 | 0.38878 | -90 | 18.5 | 2.5377 | 18.9919 | 0.06681 | -94 | 7 | -1.4623 | 8.2123 | 0.089032 | -114 | 22 | -21.4623 | 22.4152 | 0.47874 | -73.1667 | 36.076 | 19.371 | 36.3307 | 0.26659 | -74 | 19.4 | 18.5377 | 19.8696 | 0.46648 | -90 | 11 | 2.5377 | 11.8086 | 0.10745 | -92.5377 | 2.5608 | 5.1216 | 5 | -94 | 10.3 | -1.4623 | 11.1594 | 0.06552 | -65.9774 | 18.0226 | 26.5603 | 18.5272 | 0.71679 | -91 | 15 | 1.5377 | 15.6026 | 0.049276
| 5 | -85.7 | 8.5 | -0.0050327 | 9.4845 | 0.00026531 | -88 | 4 | -2.305 | 5.8057 | 0.19851 | -84 | 18 | 1.695 | 18.4853 | 0.045846 | -61 | 15 | 24.695 | 15.579 | 0.79257 | -91 | 37 | -5.305 | 37.2385 | 0.07123 | -95 | 11 | -9.305 | 11.7774 | 0.39504 | -77 | 27.3 | 8.695 | 27.6224 | 0.15739 | -86 | 7 | -0.30503 | 8.1674 | 0.018674 | -132 | 30 | -46.305 | 30.2937 | 0.76427 | -38.5 | 44.7914 | 47.195 | 44.9886 | 0.52452 | -55 | 19.4 | 30.695 | 19.8511 | 0.77313 | -84 | 15 | 1.695 | 15.579 | 0.054399 | -85.695 | 2.7007 | 5.4015 | 5 | -89 | 12.3 | -3.305 | 12.9998 | 0.12712 | -61.2186 | 19.1823 | 24.4763 | 19.6383 | 0.62318 | -86 | 15 | -0.30503 | 15.579 | 0.0097899
| 2 | -81.2 | 8.5 | -0.75663 | 10.0236 | 0.037742 | -82 | 4 | -1.5566 | 6.6501 | 0.11704 | -77 | 18 | 3.4434 | 18.7676 | 0.091737 | -29 | 26 | 51.4434 | 26.5372 | 0.96927 | -92 | 38 | -11.5566 | 38.3696 | 0.1506 | -89 | 11 | -8.5566 | 12.2157 | 0.35023 | -73 | 41.5 | 7.4434 | 41.8387 | 0.088953 | -79 | 7 | 1.4434 | 8.7877 | 0.082125 | -213 | 50 | -132.5566 | 50.2814 | 1.3181 | [5 empty] | -9 | 28.4 | 71.4434 | 28.8926 | 1.2364 | -80 | 15 | 0.44337 | 15.913 | 0.013931 | -80.4434 | 2.7887 | 5.5773 | 6 | -83 | 13.3 | -2.5566 | 14.3218 | 0.089257 | -65.1343 | 18.5228 | 15.3091 | 19.2696 | 0.39723 | -81 | 18 | -0.55663 | 18.7676 | 0.014829
| 1 | -76.4 | 8.5 | -0.29217 | 11.3131 | 0.012913 | -77 | 4 | -0.89217 | 8.4698 | 0.052668 | -71 | 18 | 5.1078 | 19.4868 | 0.13106 | 8 | 35 | 84.1078 | 35.7874 | 1.1751 | -100 | 38 | -23.8922 | 38.7264 | 0.30847 | -76 | 11 | 0.10783 | 13.2942 | 0.0040556 | -66 | 51.2 | 10.1078 | 51.7414 | 0.097676 | -74 | 7 | 2.1078 | 10.2341 | 0.10298 | -342 | 100 | -265.8922 | 100.2783 | 1.3258 | [10 empty] | -80 | 15 | -3.8922 | 16.7552 | 0.11615 | -76.1078 | 2.8746 | 5.7491 | 8 | -79 | 15.4 | -2.8922 | 17.1142 | 0.084496 | [5 empty] | -76 | 21 | 0.10783 | 22.2876 | 0.0024191
5/5 | 120 | -99.175 | 9 | -2.0623 | 9.9375 | 0.10376 | -99 | 4 | -1.8873 | 5.8098 | 0.16242 | -90 | 11 | 7.1127 | 11.7794 | 0.30191 | -94 | 15 | 3.1127 | 15.5806 | 0.099891 | -98 | 21 | -0.88728 | 21.4185 | 0.020713 | -105 | 14 | -7.8873 | 14.6203 | 0.26974 | -94 | 18.5 | 3.1127 | 18.9738 | 0.082027 | -100 | 7 | -2.8873 | 8.1703 | 0.17669 | [10 empty] | -85 | 19.4 | 12.1127 | 19.8523 | 0.30507 | [5 empty] | -97.1127 | 2.6919 | 5.3837 | 5 | -93 | 10.3 | 4.1127 | 11.1285 | 0.18478 | -75.6384 | 17.615 | 21.4743 | 18.112 | 0.59282 | -92 | 15 | 5.1127 | 15.5806 | 0.16407
| 100 | -97.65 | 8.5 | -2.184 | 9.5238 | 0.11466 | -97 | 4 | -1.534 | 5.8697 | 0.13067 | -89 | 11 | 6.466 | 11.809 | 0.27377 | -98 | 15 | -2.534 | 15.603 | 0.081202 | -95 | 21 | 0.46602 | 21.4349 | 0.010871 | -102 | 12 | -6.534 | 12.7457 | 0.25632 | -92 | 18.5 | 3.466 | 18.9922 | 0.091249 | -99 | 7 | -3.534 | 8.213 | 0.21515 | -98 | 20 | -2.534 | 20.4561 | 0.061937 | -116.3333 | 41.2495 | -20.8673 | 41.4725 | 0.25158 | -87 | 19.4 | 8.466 | 19.8699 | 0.21304 | -88 | 11 | 7.466 | 11.809 | 0.31611 | -95.466 | 2.5586 | 5.1172 | 5 | -91 | 10.3 | 4.466 | 11.1599 | 0.20009 | -76.3814 | 17.4472 | 19.0846 | 17.9682 | 0.53106 | -91 | 15 | 4.466 | 15.603 | 0.14311
| 50 | -87.325 | 8.5 | -0.58329 | 9.5284 | 0.030608 | -88 | 4 | -1.2583 | 5.8771 | 0.10705 | -82 | 11 | 4.7417 | 11.8127 | 0.2007 | -93 | 15 | -6.2583 | 15.6058 | 0.20051 | -82 | 20.5 | 4.7417 | 20.9473 | 0.11318 | -92 | 12 | -5.2583 | 12.7491 | 0.20622 | -83 | 18.5 | 3.7417 | 18.9945 | 0.098495 | -89 | 7 | -2.2583 | 8.2183 | 0.13739 | -92 | 20 | -5.2583 | 20.4582 | 0.12851 | -106.4 | 41.2449 | -19.6583 | 41.469 | 0.23702 | -78 | 19.4 | 8.7417 | 19.8721 | 0.21995 | -84 | 10 | 2.7417 | 10.8876 | 0.12591 | -86.7417 | 2.5417 | 5.0833 | 5 | -83 | 10.2 | 3.7417 | 11.0716 | 0.16898 | -68.0194 | 17.4281 | 18.7223 | 17.9521 | 0.52145 | -84 | 15 | 2.7417 | 15.6058 | 0.087843
| 20 | -73.0875 | 8.5 | -0.0087133 | 9.5311 | 0.0004571 | -74 | 4 | -0.92121 | 5.8815 | 0.078314 | -70 | 11 | 3.0788 | 11.8149 | 0.13029 | -75 | 15 | -1.9212 | 15.6074 | 0.061548 | -71 | 20.5 | 2.0788 | 20.9486 | 0.049616 | -78 | 11 | -4.9212 | 11.8149 | 0.20826 | -65 | 18.5 | 8.0788 | 18.9958 | 0.21265 | -73 | 7 | 0.078787 | 8.2214 | 0.0047915 | -84 | 20 | -10.9212 | 20.4595 | 0.2669 | -86.6667 | 41.2165 | -13.5879 | 41.4414 | 0.16394 | -59 | 19.4 | 14.0788 | 19.8734 | 0.35421 | -73 | 10 | 0.078787 | 10.89 | 0.0036174 | -73.0788 | 2.5314 | 5.0628 | 5 | -73 | 10.2 | 0.078787 | 11.0739 | 0.0035573 | -59.814 | 17.593 | 13.2647 | 18.1137 | 0.36615 | -72 | 15 | 1.0788 | 15.6074 | 0.03456
| 10 | -65.9 | 8.5 | 0.57227 | 9.5238 | 0.030044 | -68 | 4 | -1.5277 | 5.8696 | 0.13014 | -64 | 12 | 2.4723 | 12.7457 | 0.096985 | -57 | 15 | 9.4723 | 15.6029 | 0.30354 | -66 | 37 | 0.47227 | 37.2485 | 0.0063394 | -73 | 11 | -6.5277 | 11.809 | 0.27639 | -61 | 18.5 | 5.4723 | 18.9922 | 0.14407 | -66 | 7 | 0.47227 | 8.2129 | 0.028752 | -92 | 22 | -25.5277 | 22.4154 | 0.56942 | -62.3333 | 41.2032 | 4.1389 | 41.4265 | 0.049955 | -45 | 19.4 | 21.4723 | 19.8699 | 0.54032 | -66 | 10 | 0.47227 | 10.8836 | 0.021696 | -66.4723 | 2.5589 | 5.1178 | 5 | -67 | 10.2 | -0.52773 | 11.0676 | 0.023841 | -53.328 | 17.417 | 13.1443 | 17.9389 | 0.36636 | -65 | 15 | 1.4723 | 15.6029 | 0.047179
| 5 | -60.7375 | 8.5 | -0.25476 | 10.0455 | 0.012681 | -63 | 4 | -2.5173 | 6.6829 | 0.18834 | -58 | 18 | 2.4827 | 18.7793 | 0.066103 | -39 | 15 | 21.4827 | 15.9268 | 0.67442 | -68 | 37 | -7.5173 | 37.3853 | 0.10054 | -68 | 11 | -7.5173 | 12.2336 | 0.30724 | -50 | 27.3 | 10.4827 | 27.82 | 0.1884 | -58 | 7 | 2.4827 | 8.8126 | 0.14086 | -118 | 30 | -57.5173 | 30.4739 | 0.94371 | -11 | 51.191 | 49.4827 | 51.4702 | 0.48069 | -26 | 19.4 | 34.4827 | 20.1251 | 0.85671 | -60 | 15 | 0.48274 | 15.9268 | 0.015155 | -60.4827 | 2.709 | 5.418 | 6 | -62 | 12.3 | -1.5173 | 13.4146 | 0.056553 | -50.6155 | 22.3582 | 9.8672 | 22.9903 | 0.2146 | -60 | 15 | 0.48274 | 15.9268 | 0.015155
| 2 | -56.2563 | 8.5 | -0.26686 | 10.0228 | 0.013313 | -57 | 4 | -1.0106 | 6.6488 | 0.076 | -54 | 18 | 1.9894 | 18.7672 | 0.053002 | -7 | 26 | 48.9894 | 26.5369 | 0.92304 | -69 | 38 | -13.0106 | 38.3693 | 0.16954 | -64 | 11 | -8.0106 | 12.215 | 0.3279 | -46 | 41.5 | 9.9894 | 41.8385 | 0.11938 | -54 | 7 | 1.9894 | 8.7867 | 0.1132 | -220 | 50 | -164.0106 | 50.2813 | 1.6309 | [5 empty] | -2 | 28.4 | 53.9894 | 28.8923 | 0.93432 | -56 | 15 | -0.010612 | 15.9125 | 0.00033346 | -55.9894 | 2.7917 | 5.5833 | 6 | -57 | 13.3 | -1.0106 | 14.3212 | 0.035284 | -47.8331 | 19.4679 | 8.1562 | 20.1794 | 0.20209 | -56 | 18 | -0.010612 | 18.7672 | 0.00028274
| 1 | -52.2313 | 8.5 | 0.29745 | 9.9698 | 0.014918 | -53 | 4 | -0.4713 | 6.5686 | 0.035875 | -51 | 18 | 1.5287 | 18.7389 | 0.04079 | 30 | 35 | 82.5287 | 35.3857 | 1.1661 | -79 | 38 | -26.4713 | 38.3555 | 0.34508 | -56 | 11 | -3.4713 | 12.1716 | 0.1426 | -41 | 51.2 | 11.5287 | 51.4644 | 0.11201 | -49 | 9 | 3.5287 | 10.3994 | 0.16966 | -365 | 100 | -312.4713 | 100.1356 | 1.5602 | [10 empty] | -57 | 15 | -4.4713 | 15.8791 | 0.14079 | -52.5287 | 2.9754 | 5.9508 | 6 | -52 | 15.4 | 0.5287 | 16.2575 | 0.01626 | [5 empty] | -52 | 21 | 0.5287 | 21.6367 | 0.012218
6/5 | 120 | -74.9625 | 8.5 | 1.5447 | 9.5322 | 0.081027 | -78 | 3.5 | -1.4928 | 5.5554 | 0.13435 | -78 | 11 | -1.4928 | 11.8158 | 0.063169 | -70 | 15 | 6.5072 | 15.6081 | 0.20846 | -81 | 26 | -4.4928 | 26.3555 | 0.085234 | -83 | 16 | -6.4928 | 16.5714 | 0.1959 | -73 | 18.5 | 3.5072 | 18.9964 | 0.092313 | -79 | 7 | -2.4928 | 8.2227 | 0.15158 | [10 empty] | -40 | 19.4 | 36.5072 | 19.8739 | 0.91847 | [5 empty] | -76.5072 | 2.5274 | 5.0547 | 5 | -73 | 10.2 | 3.5072 | 11.0749 | 0.15834 | -63.9642 | 18.4376 | 12.543 | 18.9356 | 0.3312 | -79 | 15 | -2.4928 | 15.6081 | 0.079855
| 100 | -73.0625 | 8.5 | 1.7237 | 9.536 | 0.090377 | -76 | 3.5 | -1.2138 | 5.562 | 0.10912 | -76 | 11 | -1.2138 | 11.8189 | 0.051351 | -72 | 15 | 2.7862 | 15.6104 | 0.089241 | -75 | 26 | -0.21382 | 26.3569 | 0.0040562 | -79 | 13 | -4.2138 | 13.6998 | 0.15379 | -71 | 18.5 | 3.7862 | 18.9983 | 0.099645 | -77 | 7 | -2.2138 | 8.2271 | 0.13454 | [10 empty] | -41 | 19.4 | 33.7862 | 19.8757 | 0.84993 | [5 empty] | -74.7862 | 2.5129 | 5.0258 | 5 | -72 | 10.2 | 2.7862 | 11.0782 | 0.12575 | -63.8525 | 19.0974 | 10.9336 | 19.5805 | 0.2792 | -78 | 15 | -3.2138 | 15.6104 | 0.10294
| 50 | -64.175 | 8.5 | 2.1506 | 9.5528 | 0.11256 | -67 | 3.5 | -0.67442 | 5.5907 | 0.060317 | -65 | 11 | 1.3256 | 11.8324 | 0.056015 | -74 | 15 | -7.6744 | 15.6207 | 0.24565 | -68 | 27.5 | -1.6744 | 27.8434 | 0.030068 | -69 | 12 | -2.6744 | 12.7674 | 0.10474 | -63 | 18.5 | 3.3256 | 19.0067 | 0.087484 | -68 | 7 | -1.6744 | 8.2465 | 0.10152 | -71 | 20 | -4.6744 | 20.4696 | 0.11418 | -24.5 | 43.1592 | 41.8256 | 43.3788 | 0.4821 | -33 | 19.4 | 33.3256 | 19.8838 | 0.83801 | -69 | 15 | -2.6744 | 15.6207 | 0.085605 | -66.3256 | 2.4484 | 4.8968 | 5 | -65 | 10.2 | 1.3256 | 11.0926 | 0.059751 | -59.7372 | 18.7535 | 6.5884 | 19.2536 | 0.1711 | -72 | 15 | -5.6744 | 15.6207 | 0.18163
| 20 | -51.3875 | 8.5 | 2.8561 | 9.5557 | 0.14944 | -56 | 3.5 | -1.7564 | 5.5957 | 0.15694 | -48 | 11 | 6.2436 | 11.8348 | 0.26378 | -57 | 15 | -2.7564 | 15.6225 | 0.08822 | -58 | 26 | -3.7564 | 26.364 | 0.071242 | -59 | 11 | -4.7564 | 11.8348 | 0.20095 | -45 | 18.5 | 9.2436 | 19.0082 | 0.24315 | -53 | 7 | 1.2436 | 8.25 | 0.075367 | -67 | 20 | -12.7564 | 20.471 | 0.31157 | -17.6667 | 42.091 | 36.5769 | 42.3169 | 0.43218 | -19 | 19.4 | 35.2436 | 19.8852 | 0.88617 | -58 | 15 | -3.7564 | 15.6225 | 0.12023 | -54.2436 | 2.4367 | 4.8735 | 5 | -53 | 10.2 | 1.2436 | 11.0952 | 0.056041 | -49.5472 | 18.4466 | 4.6963 | 18.9562 | 0.12387 | -62 | 15 | -7.7564 | 15.6225 | 0.24825
| 10 | -44.3 | 8.5 | 2.9882 | 10.1144 | 0.14772 | -49 | 3.5 | -1.7118 | 6.5039 | 0.1316 | -41 | 12 | 6.2882 | 13.1928 | 0.23832 | -40 | 15 | 7.2882 | 15.9703 | 0.22818 | -49 | 26 | -1.7118 | 26.5716 | 0.032212 | -53 | 10 | -5.7118 | 11.404 | 0.25043 | -42 | 18.5 | 5.2882 | 19.2951 | 0.13703 | -46 | 7 | 1.2882 | 8.8911 | 0.072441 | -74 | 22 | -26.7118 | 22.6727 | 0.58908 | 22.8333 | 41.9172 | 70.1215 | 42.2741 | 0.82937 | -5 | 19.4 | 42.2882 | 20.1596 | 1.0488 | -49 | 15 | -1.7118 | 15.9703 | 0.053594 | -47.2882 | 2.439 | 4.878 | 6 | -49 | 10.2 | -1.7118 | 11.5798 | 0.073915 | -45.0363 | 19.4775 | 2.2519 | 20.2343 | 0.055646 | -56 | 15 | -8.7118 | 15.9703 | 0.27275
| 5 | -40 | 8.5 | 2.5702 | 10.0806 | 0.12748 | -45 | 3.5 | -2.4298 | 6.4513 | 0.18832 | -35 | 19 | 7.5702 | 19.7578 | 0.19157 | -22 | 15 | 20.5702 | 15.949 | 0.64487 | -51 | 40 | -8.4298 | 40.3654 | 0.10442 | -48 | 11 | -5.4298 | 12.2625 | 0.2214 | -41 | 27.3 | 1.5702 | 27.8327 | 0.028207 | -40 | 7 | 2.5702 | 8.8526 | 0.14516 | -99 | 30 | -56.4298 | 30.4856 | 0.92552 | 83.8333 | 51.6379 | 126.4035 | 51.9215 | 1.2173 | 0 | 19.4 | 42.5702 | 20.1427 | 1.0567 | -43 | 37 | -0.42984 | 37.3948 | 0.0057473 | -42.5702 | 2.575 | 5.15 | 6 | -45 | 12.3 | -2.4298 | 13.441 | 0.090389 | -41.2881 | 25.3334 | 1.282 | 25.9066 | 0.024744 | -51 | 15 | -8.4298 | 15.949 | 0.26428
| 2 | -32.45 | 8.5 | 5.7887 | 10.6903 | 0.27075 | -40 | 3.5 | -1.7613 | 7.3676 | 0.11953 | -31 | 19 | 7.2387 | 20.0757 | 0.18029 | 10 | 26 | 48.2387 | 26.7961 | 0.90011 | -55 | 49 | -16.7613 | 49.427 | 0.16956 | -47 | 11 | -8.7613 | 12.7684 | 0.34308 | -27 | 41.5 | 11.2387 | 42.0034 | 0.13378 | -36 | 7 | 2.2387 | 9.5411 | 0.11732 | -198 | 50 | -159.7613 | 50.4186 | 1.5843 | [5 empty] | 26 | 28.4 | 64.2387 | 29.1306 | 1.1026 | -36 | 37 | 2.2387 | 37.5637 | 0.029799 | -38.2387 | 2.6397 | 5.2794 | 7 | -40 | 13.2 | -1.7613 | 14.7062 | 0.059882 | -32.7301 | 18.4938 | 5.5087 | 19.5973 | 0.14055 | -46 | 18 | -7.7613 | 19.132 | 0.20284
| 1 | -26.6625 | 8.5 | 7.2273 | 10.672 | 0.33861 | -36 | 3.5 | -2.1102 | 7.3411 | 0.14372 | -26 | 19 | 7.8898 | 20.0659 | 0.1966 | 47 | 35 | 80.8898 | 35.5899 | 1.1364 | -60 | 49 | -26.1102 | 49.4231 | 0.26415 | -35 | 11 | -1.1102 | 12.7531 | 0.043525 | -22 | 51.2 | 11.8898 | 51.6051 | 0.1152 | -32 | 7 | 1.8898 | 9.5206 | 0.09925 | -348 | 100 | -314.1102 | 100.208 | 1.5673 | [10 empty] | -25 | 37 | 8.8898 | 37.5585 | 0.11835 | -33.8898 | 2.7126 | 5.4252 | 7 | -36 | 15.3 | -2.1102 | 16.6052 | 0.063539 | [5 empty] | -36 | 21 | -2.1102 | 21.9691 | 0.048026
8/5 | 120 | -59.15 | 8.5 | -4.6468 | 10.6502 | 0.21816 | -55 | 3.5 | -0.49679 | 7.3093 | 0.033983 | -52 | 11 | 2.5032 | 12.7348 | 0.098282 | -57 | 15 | -2.4968 | 16.3149 | 0.076519 | [10 empty] | -50 | 18.5 | 4.5032 | 19.5813 | 0.11499 | [15 empty] | -41 | 19.4 | 13.5032 | 20.4337 | 0.33042 | [5 empty] | -54.5032 | 2.7971 | 5.5942 | 7 | -49 | 10.2 | 5.5032 | 12.0506 | 0.22834 | [5 empty] | -56 | 15 | -1.4968 | 16.3149 | 0.045872
| 100 | -58.8875 | 8.5 | -6.4565 | 9.5295 | 0.33877 | -53 | 3.5 | -0.56905 | 5.5508 | 0.051258 | -50 | 11 | 2.431 | 11.8136 | 0.10289 | -50 | 15 | 2.431 | 15.6065 | 0.077883 | -57 | 26 | -4.569 | 26.3545 | 0.086684 | -53 | 17 | -0.56905 | 17.5374 | 0.016224 | -48 | 18.5 | 4.431 | 18.995 | 0.11663 | -51 | 7 | 1.431 | 8.2196 | 0.087046 | [10 empty] | -42 | 19.4 | 10.431 | 19.8726 | 0.26245 | [5 empty] | -52.431 | 2.5374 | 5.0749 | 5 | -49 | 10.2 | 3.431 | 11.0726 | 0.15493 | -46.2211 | 21.5784 | 6.2099 | 22.0043 | 0.14111 | -53 | 15 | -0.56905 | 15.6065 | 0.018231
| 50 | -47.625 | 8.5 | -3.3764 | 9.5423 | 0.17692 | -44 | 3.5 | 0.24861 | 5.5727 | 0.022306 | -42 | 11 | 2.2486 | 11.8239 | 0.095087 | -43 | 15 | 1.2486 | 15.6143 | 0.039983 | -42 | 28 | 2.2486 | 28.3338 | 0.039681 | -45 | 12 | -0.75139 | 12.7595 | 0.029444 | -40 | 18.5 | 4.2486 | 19.0015 | 0.1118 | -42 | 8 | 2.2486 | 9.0997 | 0.12355 | -50 | 20 | -5.7514 | 20.4647 | 0.14052 | -54.5 | 43.1361 | -10.2514 | 43.3535 | 0.11823 | -35 | 19.4 | 9.2486 | 19.8788 | 0.23263 | -55 | 15 | -10.7514 | 15.6143 | 0.34428 | -44.2486 | 2.4889 | 4.9778 | 5 | -42 | 10.2 | 2.2486 | 11.0836 | 0.10144 | -44.3761 | 20.9245 | -0.12747 | 21.3691 | 0.0029825 | -49 | 15 | -4.7514 | 15.6143 | 0.15215
| 20 | -35.0625 | 8.5 | -2.41 | 9.5544 | 0.12612 | -33 | 3.5 | -0.34755 | 5.5934 | 0.031068 | -27 | 11 | 5.6525 | 11.8337 | 0.23883 | -33 | 15 | -0.34755 | 15.6217 | 0.011124 | -40 | 26 | -7.3475 | 26.3635 | 0.13935 | -37 | 11 | -4.3475 | 11.8337 | 0.18369 | -23 | 18.5 | 9.6525 | 19.0075 | 0.25391 | -28 | 7 | 4.6525 | 8.2484 | 0.28202 | -48 | 20 | -15.3475 | 20.4704 | 0.37487 | 20 | 43.237 | 52.6525 | 43.4565 | 0.60581 | -20 | 19.4 | 12.6525 | 19.8846 | 0.31815 | -46 | 15 | -13.3475 | 15.6217 | 0.42721 | -32.6525 | 2.4421 | 4.8842 | 5 | -30 | 10.2 | 2.6525 | 11.094 | 0.11954 | -36.1736 | 21.1999 | -3.5212 | 21.6442 | 0.081342 | -42 | 15 | -9.3475 | 15.6217 | 0.29919
| 10 | -27.075 | 8.5 | -1.0381 | 10.1109 | 0.051336 | -27 | 3.5 | -0.96311 | 6.4986 | 0.074102 | -18 | 12 | 8.0369 | 13.1902 | 0.30465 | -16 | 15 | 10.0369 | 15.9681 | 0.31428 | -35 | 26 | -8.9631 | 26.5703 | 0.16867 | -30 | 11 | -3.9631 | 12.2874 | 0.16127 | -20 | 18.5 | 6.0369 | 19.2933 | 0.15645 | -22 | 7 | 4.0369 | 8.8871 | 0.22712 | -57 | 22 | -30.9631 | 22.6712 | 0.68287 | 58.1667 | 42.9838 | 84.2036 | 43.3312 | 0.97163 | -9 | 19.4 | 17.0369 | 20.1579 | 0.42259 | -37 | 15 | -10.9631 | 15.9681 | 0.34328 | -26.0369 | 2.4533 | 4.9066 | 6 | -25 | 10.1 | 1.0369 | 11.4887 | 0.045126 | -30.4331 | 21.0571 | -4.3962 | 21.7573 | 0.10103 | -35 | 15 | -8.9631 | 15.9681 | 0.28066
| 5 | -22.3 | 8.5 | -0.63233 | 10.0819 | 0.03136 | -23 | 3.5 | -1.3323 | 6.4533 | 0.10323 | -13 | 19 | 8.6677 | 19.7584 | 0.21934 | 2 | 15 | 23.6677 | 15.9498 | 0.74194 | -31 | 40 | -9.3323 | 40.3658 | 0.1156 | -28 | 11 | -6.3323 | 12.2636 | 0.25818 | -19 | 27.3 | 2.6677 | 27.8332 | 0.047923 | -18 | 7 | 3.6677 | 8.8541 | 0.20712 | -75 | 30 | -53.3323 | 30.486 | 0.8747 | 105.8333 | 53.4093 | 127.501 | 53.6837 | 1.1875 | -5 | 19.4 | 16.6677 | 20.1434 | 0.41373 | -31 | 37 | -9.3323 | 37.3951 | 0.12478 | -21.6677 | 2.57 | 5.1401 | 6 | -21 | 12.2 | 0.66767 | 13.3505 | 0.025005 | -28.1878 | 22.0396 | -6.5202 | 22.6967 | 0.14364 | -32 | 15 | -10.3323 | 15.9498 | 0.3239
| 2 | -17.55 | 8.5 | -0.46303 | 10.0623 | 0.023008 | -18 | 3.5 | -0.91303 | 6.4227 | 0.071079 | -8 | 19 | 9.087 | 19.7484 | 0.23007 | 34 | 26 | 51.087 | 26.5518 | 0.96202 | -35 | 49 | -17.913 | 49.295 | 0.18169 | -25 | 11 | -7.913 | 12.2475 | 0.32305 | -6 | 41.5 | 11.087 | 41.8479 | 0.13247 | -11 | 7 | 6.087 | 8.8318 | 0.34461 | -168 | 50 | -150.913 | 50.2892 | 1.5005 | [5 empty] | 19 | 28.4 | 36.087 | 28.9061 | 0.62421 | -23 | 37 | -5.913 | 37.3898 | 0.079073 | -17.087 | 2.6456 | 5.2912 | 6 | -17 | 13.2 | 0.086968 | 14.2563 | 0.0030502 | -26.4582 | 20.9293 | -9.3712 | 21.611 | 0.21682 | -29 | 18 | -11.913 | 18.7883 | 0.31703
| 1 | -3.3125 | 8.5 | 9.5856 | 10.672 | 0.4491 | -15 | 3.5 | -2.1019 | 7.3411 | 0.14316 | -5 | 19 | 7.8981 | 20.0659 | 0.1968 | 70 | 35 | 82.8981 | 35.5899 | 1.1646 | -32 | 49 | -19.1019 | 49.4231 | 0.19325 | -20 | 11 | -7.1019 | 12.7531 | 0.27844 | -1 | 51.2 | 11.8981 | 51.6051 | 0.11528 | -11 | 7 | 1.8981 | 9.5206 | 0.099685 | -322 | 100 | -309.1019 | 100.208 | 1.5423 | [10 empty] | -7 | 37 | 5.8981 | 37.5585 | 0.078519 | -12.8981 | 2.7126 | 5.4252 | 7 | -13 | 15.3 | -0.10188 | 16.6052 | 0.0030677 | [5 empty] | -13 | 21 | -0.10188 | 21.9691 | 0.0023187
10/5 | 120 | [5 empty] | -40 | 3.5 | -0.80996 | 5.2428 | 0.077244 | -38 | 11 | 1.19 | 11.6721 | 0.050978 | [15 empty] | -37 | 18.5 | 2.19 | 18.9073 | 0.057915 | [25 empty] | -39.19 | 3.1245 | 6.2491 | 5 | -34 | 10.2 | 5.19 | 10.9214 | 0.23761
| 100 | -36.575 | 8.5 | 1.0526 | 9.4521 | 0.05568 | -39 | 3.5 | -1.3724 | 5.4169 | 0.12668 | -37 | 11 | 0.62758 | 11.7513 | 0.026703 | -42 | 15 | -4.3724 | 15.5593 | 0.14051 | [10 empty] | -35 | 18.5 | 2.6276 | 18.9563 | 0.069306 | [15 empty] | -18 | 19.4 | 19.6276 | 19.8356 | 0.49476 | [5 empty] | -37.6276 | 2.8121 | 5.6241 | 5 | -35 | 10.2 | 2.6276 | 11.006 | 0.11937 | [5 empty] | -29 | 18 | 8.6276 | 18.4687 | 0.23357
| 50 | -27.3 | 8.5 | 0.46058 | 9.5322 | 0.024159 | -30 | 3.5 | -2.2394 | 5.5555 | 0.20155 | -29 | 11 | -1.2394 | 11.8158 | 0.052447 | -27 | 15 | 0.76058 | 15.6081 | 0.024365 | -29 | 26.5 | -1.2394 | 26.8489 | 0.023081 | -28 | 13 | -0.23942 | 13.6972 | 0.0087397 | -16 | 18.5 | 11.7606 | 18.9964 | 0.30955 | -25 | 8 | 2.7606 | 9.0892 | 0.15186 | [5 empty] | 5.3333 | 43.0546 | 33.0939 | 43.2702 | 0.38241 | -14 | 19.4 | 13.7606 | 19.8739 | 0.3462 | -20 | 15 | 7.7606 | 15.6081 | 0.24861 | -27.7606 | 2.5272 | 5.0545 | 5 | -27 | 10.2 | 0.76058 | 11.0749 | 0.034338 | -30.0801 | 21.508 | -2.3196 | 21.9364 | 0.05287 | -22 | 18 | 5.7606 | 18.5098 | 0.15561
| 20 | -15.825 | 8.5 | 1.2551 | 9.5522 | 0.065697 | -19 | 3.5 | -1.9199 | 5.5896 | 0.17174 | -14 | 11 | 3.0801 | 11.8319 | 0.13016 | -22 | 15 | -4.9199 | 15.6203 | 0.15748 | -14 | 26 | 3.0801 | 26.3627 | 0.058418 | -20 | 11 | -2.9199 | 11.8319 | 0.12339 | -9 | 18.5 | 8.0801 | 19.0064 | 0.21256 | -14 | 7 | 3.0801 | 8.2458 | 0.18677 | -34 | 20 | -16.9199 | 20.4693 | 0.4133 | 50.3333 | 42.7072 | 67.4134 | 42.929 | 0.78517 | -2 | 19.4 | 15.0801 | 19.8835 | 0.37921 | -14 | 15 | 3.0801 | 15.6203 | 0.098593 | -17.0801 | 2.4508 | 4.9015 | 5 | -14 | 10.1 | 3.0801 | 11.0002 | 0.14 | -27.2342 | 21.3233 | -10.1541 | 21.7641 | 0.23328 | -13 | 18 | 4.0801 | 18.5201 | 0.11015
| 10 | -9.75 | 8.5 | 1.4347 | 9.5473 | 0.075137 | -13 | 3.5 | -1.8153 | 5.5814 | 0.16262 | -4 | 12 | 7.1847 | 12.7633 | 0.28146 | -6 | 15 | 5.1847 | 15.6173 | 0.16599 | -14 | 26 | -2.8153 | 26.361 | 0.053399 | -16 | 11 | -4.8153 | 11.828 | 0.20355 | -7 | 18.5 | 4.1847 | 19.004 | 0.1101 | -8 | 7 | 3.1847 | 8.2402 | 0.19324 | -42 | 22 | -30.8153 | 22.4255 | 0.68706 | 78.8333 | 42.0047 | 90.018 | 42.2291 | 1.0658 | 1 | 19.4 | 12.1847 | 19.8812 | 0.30644 | -8 | 15 | 3.1847 | 15.6173 | 0.10196 | -11.1847 | 2.4695 | 4.939 | 5 | -11 | 10.1 | 0.18471 | 10.996 | 0.0083992 | -23.5345 | 21.159 | -12.3498 | 21.601 | 0.28586 | -5 | 21 | 6.1847 | 21.4453 | 0.1442
| 5 | -5.25 | 8.5 | 1.4599 | 9.5144 | 0.07672 | -9 | 3.5 | -2.2901 | 5.5248 | 0.20726 | 0 | 19 | 6.7099 | 19.4749 | 0.17227 | 13 | 15 | 19.7099 | 15.5972 | 0.63184 | -7 | 40 | -0.29012 | 40.2278 | 0.0036059 | -12 | 11 | -5.2901 | 11.8014 | 0.22413 | 4 | 27.3 | 10.7099 | 27.6326 | 0.19379 | -3 | 7 | 3.7099 | 8.202 | 0.22616 | -58 | 30 | -51.2901 | 30.303 | 0.84629 | 128.5 | 52.5689 | 135.2099 | 52.7424 | 1.2818 | 8 | 19.4 | 14.7099 | 19.8654 | 0.37024 | -3 | 37 | 3.7099 | 37.2461 | 0.049802 | -6.7099 | 2.5937 | 5.1873 | 5 | -8 | 12.2 | -1.2901 | 12.9272 | 0.049899 | -16.9528 | 22.2971 | -10.2429 | 22.7032 | 0.22558 | -1 | 24 | 5.7099 | 24.3777 | 0.11711
| 2 | -0.9 | 8.5 | 2.4933 | 10.6811 | 0.11671 | -5 | 3.5 | -1.6067 | 7.3543 | 0.10924 | 4 | 19 | 7.3933 | 20.0708 | 0.18418 | 45 | 26 | 48.3933 | 26.7924 | 0.90311 | -7 | 49 | -3.6067 | 49.425 | 0.036487 | -10 | 11 | -6.6067 | 12.7607 | 0.25887 | 5 | 41.5 | 8.3933 | 42.001 | 0.099917 | 1 | 7 | 4.3933 | 9.5308 | 0.23048 | -149 | 50 | -145.6067 | 50.4166 | 1.444 | [5 empty] | 27 | 28.4 | 30.3933 | 29.1272 | 0.52173 | 3 | 37 | 6.3933 | 37.5611 | 0.085105 | -3.3933 | 2.6767 | 5.3534 | 7 | -3 | 13.2 | 0.39327 | 14.6995 | 0.013377 | -13.1614 | 21.9843 | -9.7681 | 22.916 | 0.21313
| 1 | -1.3 | 8.5 | -0.59452 | 11.3475 | 0.026196 | -2 | 3.5 | -1.2945 | 8.2926 | 0.078053 | 8 | 19 | 8.7055 | 20.4332 | 0.21302 | 82 | 35 | 82.7055 | 35.7983 | 1.1552 | -12 | 49 | -11.2945 | 49.5733 | 0.11392 | -4 | 11 | -3.2945 | 13.3235 | 0.12364 | 10 | 51.2 | 10.7055 | 51.749 | 0.10344 | 3 | 7 | 3.7055 | 10.2721 | 0.18037 | -291 | 100 | -290.2945 | 100.2822 | 1.4474 | [10 empty] | 0 | 37 | 0.70548 | 37.756 | 0.0093426 | -0.70548 | 2.7355 | 5.4711 | 8 | 0 | 15.3 | 0.70548 | 17.0472 | 0.020692
Annex 4 TAB A4-1. Calculation from results of all labs.
Meaning of individual symbols:
| I1 | Primary current (%)IN
| kI | Transformation ratio I (kA/A)
| B | Burden (VA)
| I1N | Rated primary current (A)
| eL | Ratio error of each participating laboratory (ppm)
| u(eL) | Standard uncertainty of ratio error of each participating laboratory (ppm)
| er | CRV for ratio error calculated from results of all laboratories see eq. (2) (ppm)
| u(er) | CRV uncertainty calculated from results of all laboratories see eq. (3) (ppm)
| De | Difference between results of individual laboratory eL and er see eq. (5) (ppm)
| u(Destd) | The uncertainty of the difference De according to (8) with transfer standard uncertaity (ppm)
| u(estd) | Pilot laboratory evaluated uncertainty component of ratio error of transfer standard (ppm)
| E(e) | confidence coefficient of ratio error calculated from results of all laboratories see eq. (9)
red and underlined results are for E(e) > 1.0
5 VA | I1 (%IN) | CMI | CMI | CMI | CMI | CMI | PTB | PTB | PTB | PTB | PTB | METAS | METAS | METAS | METAS | METAS | NPL | NPL | NPL | NPL | NPL | LCOE | LCOE | LCOE | LCOE | LCOE | MIKES | MIKES | MIKES | MIKES | MIKES | BEV | BEV | BEV | BEV | BEV | VSL | VSL | VSL | VSL | VSL | SP | SP | SP | SP | SP | DMDM | DMDM | DMDM | DMDM | DMDM | GUM | GUM | GUM | GUM | GUM | BIM | BIM | BIM | BIM | BIM | weighted mean erC (ppm) | u(erC) (ppm) | U(erC) (ppm) | u(estd) (ppm) | UME | UME | UME | UME | UME | LNE | LNE | LNE | LNE | LNE | INRIM | INRIM | INRIM | INRIM | INRIM
kI (kA/A) |  | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) |  |  |  |  | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-)
4/5 | 120 | -92.525 | 8.5 | -0.35875 | 9.516 | 0.01885 | -92 | 4 | 0.16625 | 5.857 | 0.014192 | -86 | 11 | 6.1662 | 11.8027 | 0.26122 | -94 | 15 | -1.8338 | 15.5982 | 0.058781 | -91 | 21 | 1.1662 | 21.4314 | 0.027209 | -105 | 13 | -12.8338 | 13.6859 | 0.46887 | -88 | 18.5 | 4.1662 | 18.9883 | 0.10971 | -98 | 7 | -5.8338 | 8.2039 | 0.35555 | [5 empty] | -84.8333 | 36.0258 | 7.3329 | 36.279 | 0.10106 | -87 | 19.4 | 5.1662 | 19.8662 | 0.13003 | -84 | 11 | 8.1662 | 11.8027 | 0.34595 | -92.1662 | 2.5876 | 5.1753 | 5 | -89 | 10.3 | 3.1662 | 11.1532 | 0.14194 | -91.8063 | 17.6245 | 0.35992 | 18.1363 | 0.0099226 | -89 | 15 | 3.1662 | 15.5982 | 0.10149
| 100 | -90.1833 | 8.5 | 0.50235 | 9.5281 | 0.026361 | -91 | 4 | -0.31432 | 5.8767 | 0.026743 | -84 | 11 | 6.6857 | 11.8125 | 0.28299 | -98 | 15 | -7.3143 | 15.6056 | 0.23435 | -89 | 21 | 1.6857 | 21.4368 | 0.039318 | -101 | 12 | -10.3143 | 12.7489 | 0.40452 | -85 | 18.5 | 5.6857 | 18.9943 | 0.14967 | -96 | 7 | -5.3143 | 8.218 | 0.32333 | -92 | 20 | -1.3143 | 20.4581 | 0.032122 | -82.5 | 36.0295 | 8.1857 | 36.2858 | 0.11279 | -85 | 19.4 | 5.6857 | 19.872 | 0.14306 | -84 | 10 | 6.6857 | 10.8874 | 0.30704 | -90.6857 | 2.5426 | 5.0851 | 5 | -86 | 10.3 | 4.6857 | 11.1636 | 0.20986 | -89.4537 | 17.4147 | 1.232 | 17.939 | 0.034338 | -87 | 15 | 3.6857 | 15.6056 | 0.11809
| 50 | -80.5 | 8.5 | 1.4482 | 9.036 | 0.080135 | -83 | 4 | -1.0518 | 5.0398 | 0.10435 | -76 | 11 | 5.9482 | 11.4192 | 0.26045 | -90 | 15 | -8.0518 | 15.3101 | 0.26296 | -77 | 20.5 | 4.9482 | 20.728 | 0.11936 | -89 | 11 | -7.0518 | 11.4192 | 0.30877 | -78 | 18.5 | 3.9482 | 18.7523 | 0.10527 | -85 | 8 | -3.0518 | 8.5673 | 0.17811 | -86 | 20 | -4.0518 | 20.2336 | 0.10013 | -75.1667 | 36.0196 | 6.7815 | 36.1499 | 0.093797 | -74 | 19.4 | 7.9482 | 19.6408 | 0.20234 | -78 | 10 | 3.9482 | 10.4594 | 0.18874 | -81.9482 | 2.5692 | 5.1384 | 4 | -78 | 10.2 | 3.9482 | 10.6508 | 0.18535 | -79.4339 | 17.484 | 2.5143 | 17.7508 | 0.070823 | -79 | 15 | 2.9482 | 15.3101 | 0.096283
| 20 | -70.95 | 8.5 | 0.78441 | 9.0472 | 0.043351 | -73 | 4 | -1.2656 | 5.0597 | 0.12506 | -67 | 11 | 4.7344 | 11.4281 | 0.20714 | -72 | 15 | -0.26559 | 15.3167 | 0.0086699 | -74 | 20.5 | -2.2656 | 20.7328 | 0.054638 | -76 | 11 | -4.2656 | 11.4281 | 0.18663 | -63 | 18.5 | 8.7344 | 18.7577 | 0.23282 | -73 | 7 | -1.2656 | 7.6551 | 0.082663 | -85 | 20 | -13.2656 | 20.2386 | 0.32773 | -62.8333 | 36.0184 | 8.9011 | 36.1514 | 0.12311 | -59 | 19.4 | 12.7344 | 19.6459 | 0.3241 | -68 | 10 | 3.7344 | 10.4691 | 0.17835 | -71.7344 | 2.5296 | 5.0592 | 4 | -70 | 10.2 | 1.7344 | 10.6603 | 0.081349 | -72.4854 | 17.4951 | -0.75102 | 17.7674 | 0.021135 | -69 | 15 | 2.7344 | 15.3167 | 0.089263
| 10 | -65.1167 | 8.5 | 1.8579 | 9.039 | 0.10277 | -68 | 4 | -1.0255 | 5.0452 | 0.10163 | -63 | 12 | 3.9745 | 12.3877 | 0.16042 | -57 | 15 | 9.9745 | 15.3119 | 0.32571 | -70 | 37 | -3.0255 | 37.1275 | 0.040744 | -77 | 11 | -10.0255 | 11.4217 | 0.43888 | -62 | 18.5 | 4.9745 | 18.7538 | 0.13263 | -68 | 7 | -1.0255 | 7.6455 | 0.067062 | -92 | 22 | -25.0255 | 22.2138 | 0.56329 | -44 | 36.0153 | 22.9745 | 36.1463 | 0.3178 | -48 | 19.4 | 18.9745 | 19.6422 | 0.48301 | -63 | 10 | 3.9745 | 10.462 | 0.18995 | -66.9745 | 2.5585 | 5.1169 | 4 | -67 | 10.2 | -0.025454 | 10.6534 | 0.0011947 | -69.1213 | 17.7714 | -2.1468 | 18.0354 | 0.059516 | -65 | 15 | 1.9745 | 15.3119 | 0.064477
| 5 | -61.475 | 8.5 | 0.92872 | 8.9988 | 0.051603 | -64 | 4 | -1.5963 | 4.9727 | 0.1605 | -59 | 18 | 3.4037 | 18.2408 | 0.093299 | -39 | 15 | 23.4037 | 15.2882 | 0.76542 | -68 | 37 | -5.5963 | 37.1178 | 0.075385 | -68 | 11 | -5.5963 | 11.3898 | 0.24567 | -52 | 27.3 | 10.4037 | 27.4594 | 0.18944 | -64 | 7 | -1.5963 | 7.5979 | 0.10505 | -111 | 30 | -48.5963 | 30.1451 | 0.80604 | -16 | 44.7547 | 46.4037 | 44.8521 | 0.5173 | -32 | 19.4 | 30.4037 | 19.6237 | 0.77467 | -60 | 15 | 2.4037 | 15.2882 | 0.078614 | -62.4037 | 2.6967 | 5.3934 | 4 | -64 | 12.3 | -1.5963 | 12.6498 | 0.063095 | -65.7714 | 17.8798 | -3.3676 | 18.1223 | 0.092914 | -62 | 15 | 0.40372 | 15.2882 | 0.013204
| 2 | -58.1 | 8.5 | 0.83168 | 9.4494 | 0.044007 | -60 | 4 | -1.0683 | 5.7482 | 0.092926 | -55 | 18 | 3.9317 | 18.4673 | 0.10645 | -10 | 26 | 48.9317 | 26.3257 | 0.92935 | -71 | 38 | -12.0683 | 38.2236 | 0.15786 | -64 | 12 | -5.0683 | 12.6902 | 0.19969 | -50 | 41.5 | 8.9317 | 41.7048 | 0.10708 | -58 | 7 | 0.93168 | 8.1266 | 0.057323 | -193 | 50 | -134.0683 | 50.3285 | 1.3319 | [5 empty] | 5 | 28.4 | 63.9317 | 28.9744 | 1.1032 | -58 | 15 | 0.93168 | 15.5577 | 0.029943 | -58.9317 | 2.821 | 5.642 | 5 | -60 | 13.3 | -1.0683 | 13.9259 | 0.038357 | -61.5803 | 18.4313 | -2.6486 | 18.888 | 0.070114 | -59 | 18 | -0.068317 | 18.4673 | 0.0018497
| 1 | -56.2833 | 8.5 | -0.12577 | 9.4299 | 0.0066688 | -57 | 4 | -0.84244 | 5.7161 | 0.07369 | -51 | 18 | 5.1576 | 18.4574 | 0.13972 | 26 | 35 | 82.1576 | 35.4729 | 1.158 | -81 | 38 | -24.8424 | 38.2188 | 0.325 | -53 | 11 | 3.1576 | 11.7334 | 0.13455 | -46 | 51.2 | 10.1576 | 51.3626 | 0.098881 | -54 | 7 | 2.1576 | 8.1039 | 0.13312 | -329 | 100 | -272.8424 | 100.1665 | 1.3619 | [10 empty] | -60 | 15 | -3.8424 | 15.5459 | 0.12358 | -56.1576 | 2.8855 | 5.771 | 5 | -56 | 15.4 | 0.15756 | 15.9322 | 0.0049447 | [5 empty] | -57 | 21 | -0.84244 | 21.3933 | 0.019689
10/5 | 120 | [5 empty] | -34 | 3.5 | -0.85164 | 4.2997 | 0.099035 | -33 | 11 | 0.14836 | 11.28 | 0.0065764 | [15 empty] | -30 | 18.5 | 3.1484 | 18.6678 | 0.084326 | [25 empty] | -33.1484 | 3.1245 | 6.2491 | 4 | -27 | 10.2 | 6.1484 | 10.5013 | 0.29274
| 100 | -30.75 | 8.5 | 0.90801 | 8.9634 | 0.050651 | -33 | 3.5 | -1.342 | 4.5102 | 0.14877 | -32 | 11 | -0.34199 | 11.3619 | 0.01505 | -33 | 15 | -1.342 | 15.2674 | 0.04395 | [10 empty] | -28 | 18.5 | 3.658 | 18.7174 | 0.097717 | [15 empty] | -18 | 19.4 | 13.658 | 19.6075 | 0.34829 | [5 empty] | -31.658 | 2.8121 | 5.6241 | 4 | -28 | 10.2 | 3.658 | 10.5893 | 0.17272 | [5 empty] | -24 | 18 | 7.658 | 18.2234 | 0.21011
| 50 | -21.3 | 8.5 | 0.62389 | 9.0589 | 0.034435 | -24 | 3.5 | -2.0761 | 4.6972 | 0.22099 | -24 | 11 | -2.0761 | 11.4374 | 0.09076 | -22 | 15 | -0.076107 | 15.3236 | 0.0024833 | -24 | 26.5 | -2.0761 | 26.6845 | 0.038901 | -22 | 13 | -0.076107 | 13.3721 | 0.0028457 | -9 | 18.5 | 12.9239 | 18.7634 | 0.34439 | -19 | 7 | 2.9239 | 7.669 | 0.19063 | [5 empty] | 12.3333 | 43.2655 | 34.2572 | 43.3787 | 0.39486 | -14 | 19.4 | 7.9239 | 19.6513 | 0.20161 | -14 | 15 | 7.9239 | 15.3236 | 0.25855 | -21.9239 | 2.4873 | 4.9745 | 4 | -21 | 10.1 | 0.92389 | 10.5747 | 0.043684 | -28.3301 | 21.0096 | -6.4062 | 21.2419 | 0.15079 | -17 | 18 | 4.9239 | 18.2706 | 0.13475
| 20 | -11.1 | 8.5 | 1.0857 | 9.0689 | 0.059857 | -14 | 3.5 | -1.8143 | 4.7165 | 0.19234 | -9 | 11 | 3.1857 | 11.4453 | 0.13917 | -18 | 15 | -5.8143 | 15.3296 | 0.18964 | -7 | 26.5 | 5.1857 | 26.6879 | 0.097154 | -12 | 11 | 0.18567 | 11.4453 | 0.0081113 | -3 | 18.5 | 9.1857 | 18.7682 | 0.24471 | -9 | 7 | 3.1857 | 7.6809 | 0.20738 | -27 | 20 | -14.8143 | 20.2484 | 0.36582 | 54.6667 | 42.3705 | 66.8523 | 42.4883 | 0.78671 | -2 | 19.4 | 10.1857 | 19.6559 | 0.2591 | -8.7 | 15 | 3.4857 | 15.3296 | 0.11369 | -12.1857 | 2.4504 | 4.9007 | 4 | -10 | 10.1 | 2.1857 | 10.5833 | 0.10326 | -23.8175 | 20.8486 | -11.6319 | 21.087 | 0.27581 | -14 | 18 | -1.8143 | 18.2756 | 0.049638
| 10 | -6.2 | 8.5 | 1.0986 | 9.0627 | 0.060611 | -9 | 3.5 | -1.7014 | 4.7045 | 0.18083 | -1 | 12 | 6.2986 | 12.4049 | 0.25388 | -2 | 15 | 5.2986 | 15.3259 | 0.17286 | -9 | 26 | -1.7014 | 26.1894 | 0.032483 | -10 | 11 | -2.7014 | 11.4404 | 0.11806 | -2 | 18.5 | 5.2986 | 18.7652 | 0.14118 | -2 | 7 | 5.2986 | 7.6735 | 0.34525 | -37 | 22 | -29.7014 | 22.2235 | 0.66824 | 81.8333 | 41.6972 | 89.1319 | 41.9616 | 1.0621 | 1 | 19.4 | 8.2986 | 19.653 | 0.21113 | -3.8 | 15 | 3.4986 | 15.3259 | 0.11414 | -7.2986 | 2.4734 | 4.9468 | 4 | -6 | 10.1 | 1.2986 | 10.5779 | 0.061383 | -18.3261 | 20.9118 | -11.0275 | 21.1468 | 0.26074 | -10 | 21 | -2.7014 | 21.234 | 0.06361
| 5 | -2.2 | 8.5 | 0.84505 | 8.6319 | 0.048949 | -5 | 3.5 | -1.955 | 3.8091 | 0.25662 | 4 | 19 | 7.045 | 19.0594 | 0.18482 | 16 | 15 | 19.045 | 15.0751 | 0.63167 | -6 | 40 | -2.955 | 40.0282 | 0.036911 | -3 | 11 | 0.045045 | 11.1022 | 0.0020287 | 8 | 27.3 | 11.045 | 27.3413 | 0.20198 | 1 | 7 | 4.045 | 7.1595 | 0.28249 | -53 | 30 | -49.955 | 30.0376 | 0.83154 | 126.3333 | 51.6372 | 129.3784 | 51.7894 | 1.2491 | 8 | 19.4 | 11.045 | 19.4581 | 0.28382 | 1.1 | 37 | 4.145 | 37.0305 | 0.055968 | -3.045 | 2.5963 | 5.1927 | 3 | -3 | 12.2 | 0.045045 | 12.2922 | 0.0018323 | -17.2861 | 22.0031 | -14.241 | 22.0544 | 0.32286 | -3 | 24 | 0.045045 | 24.047 | 0.00093661
| 2 | 1.6 | 8.5 | 1.0326 | 8.6033 | 0.060015 | -2 | 3.5 | -2.5674 | 3.7438 | 0.34288 | 8 | 19 | 7.4326 | 19.0464 | 0.19512 | 47 | 26 | 46.4326 | 26.0339 | 0.89177 | -5 | 49 | -5.5674 | 49.018 | 0.056789 | -1 | 11 | -1.5674 | 11.08 | 0.070729 | 10 | 41.5 | 9.4326 | 41.5213 | 0.11359 | 5 | 7 | 4.4326 | 7.125 | 0.31106 | -144 | 50 | -144.5674 | 50.1621 | 1.441 | [5 empty] | 27 | 28.4 | 26.4326 | 28.4311 | 0.46485 | 5.3 | 37 | 4.7326 | 37.0239 | 0.063914 | 0.56735 | 2.6896 | 5.3791 | 3 | 1 | 13.2 | 0.43265 | 13.2667 | 0.016306 | -11.1614 | 29.6196 | -11.7287 | 29.6494 | 0.19779
| 1 | 0.65 | 8.5 | 0.18043 | 8.5693 | 0.010528 | -1 | 3.5 | -1.4696 | 3.6652 | 0.20048 | 10 | 19 | 9.5304 | 19.0311 | 0.25039 | 85 | 35 | 84.5304 | 35.2394 | 1.1994 | -11 | 49 | -11.4696 | 49.0121 | 0.11701 | 0 | 11 | -0.46957 | 11.0537 | 0.02124 | 13 | 51.2 | 12.5304 | 51.2116 | 0.12234 | 6 | 8 | 5.5304 | 8.0736 | 0.3425 | -287 | 100 | -287.4696 | 100.084 | 1.4361 | [10 empty] | 3.2 | 37 | 2.7304 | 37.016 | 0.036882 | 0.46957 | 2.7958 | 5.5916 | 3 | 2 | 15.3 | 1.5304 | 15.3386 | 0.049888
15 VA | I1 (%IN) | CMI | CMI | CMI | CMI | CMI | PTB | PTB | PTB | PTB | PTB | METAS | METAS | METAS | METAS | METAS | NPL | NPL | NPL | NPL | NPL | LCOE | LCOE | LCOE | LCOE | LCOE | MIKES | MIKES | MIKES | MIKES | MIKES | BEV | BEV | BEV | BEV | BEV | VSL | VSL | VSL | VSL | VSL | SP | SP | SP | SP | SP | DMDM | DMDM | DMDM | DMDM | DMDM | GUM | GUM | GUM | GUM | GUM | BIM | BIM | BIM | BIM | BIM | weighted mean erC (ppm) | u(erC) (ppm) | U(erC) (ppm) | u(estd) (ppm) | UME | UME | UME | UME | UME | LNE | LNE | LNE | LNE | LNE | INRIM | INRIM | INRIM | INRIM | INRIM
kI (kA/A) |  | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) |  |  |  |  | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-) | eL (ppm) | u(eL) (ppm) | DeC (ppm) | u(DestdC) (ppm) | EC(e) (-)
4/5 | 120 | -131.45 | 8.5 | -2.3573 | 9.49 | 0.1242 | -130 | 4 | -0.90733 | 5.8147 | 0.07802 | -121 | 11 | 8.0927 | 11.7818 | 0.34344 | -130 | 15 | -0.90733 | 15.5824 | 0.029114 | -129 | 21 | 0.092674 | 21.4199 | 0.0021633 | -138 | 12 | -8.9073 | 12.7205 | 0.35012 | -126 | 18.5 | 3.0927 | 18.9753 | 0.081492 | -135 | 9 | -5.9073 | 9.9403 | 0.29714 | [5 empty] | -120.8333 | 36.0953 | 8.2593 | 36.3412 | 0.11364 | -120 | 19.4 | 9.0927 | 19.8537 | 0.22899 | -123 | 11 | 6.0927 | 11.7818 | 0.25856 | -129.0927 | 2.6814 | 5.3627 | 5 | -123 | 10.3 | 6.0927 | 11.1311 | 0.27368 | -85.8843 | 17.8292 | 43.2084 | 18.7102 | 1.1547 | -126 | 15 | 3.0927 | 15.5824 | 0.099236
| 100 | -128.3 | 8.5 | -0.62206 | 10.0702 | 0.030886 | -128 | 4 | -0.32206 | 6.7201 | 0.023963 | -120 | 11 | 7.6779 | 12.2539 | 0.31328 | -133 | 15 | -5.3221 | 15.9424 | 0.16692 | -126 | 21 | 1.6779 | 21.6832 | 0.038692 | -135 | 11 | -7.3221 | 12.2539 | 0.29876 | -124 | 18.5 | 3.6779 | 19.272 | 0.095422 | -134 | 8 | -6.3221 | 9.6519 | 0.3275 | -136 | 20 | -8.3221 | 20.7162 | 0.20086 | -118.5 | 36.2096 | 9.1779 | 36.61 | 0.12535 | -118 | 19.4 | 9.6779 | 20.1375 | 0.2403 | -122 | 11 | 5.6779 | 12.2539 | 0.23168 | -127.6779 | 2.6155 | 5.231 | 6 | -121 | 10.3 | 6.6779 | 11.6297 | 0.28711 | -84.1859 | 17.5681 | 43.492 | 18.7477 | 1.1599 | -125 | 15 | 2.6779 | 15.9424 | 0.083988
| 50 | -117 | 8.5 | -0.16844 | 10.0705 | 0.008363 | -117 | 4 | -0.16844 | 6.7205 | 0.012532 | -111 | 11 | 5.8316 | 12.2542 | 0.23794 | -121 | 15 | -4.1684 | 15.9425 | 0.13073 | -113 | 20.5 | 3.8316 | 21.1994 | 0.09037 | -123 | 11 | -6.1684 | 12.2542 | 0.25169 | -114 | 18.5 | 2.8316 | 19.2721 | 0.073463 | -121 | 8 | -4.1684 | 9.6522 | 0.21593 | -121 | 20 | -4.1684 | 20.7163 | 0.10061 | -106 | 36.1042 | 10.8316 | 36.5059 | 0.14835 | -109 | 19.4 | 7.8316 | 20.1376 | 0.19445 | -114 | 11 | 2.8316 | 12.2542 | 0.11553 | -116.8316 | 2.6144 | 5.2289 | 6 | -113 | 10.3 | 3.8316 | 11.6299 | 0.16473 | -76.3392 | 17.5022 | 40.4924 | 18.6858 | 1.0835 | -116 | 15 | 0.83156 | 15.9425 | 0.02608
| 20 | -100.85 | 8.5 | -0.71622 | 9.5271 | 0.037588 | -102 | 4 | -1.8662 | 5.8751 | 0.15883 | -97 | 11 | 3.1338 | 11.8117 | 0.13266 | -99 | 15 | 1.1338 | 15.605 | 0.036327 | -97 | 20.5 | 3.1338 | 20.9467 | 0.074803 | -106 | 11 | -5.8662 | 11.8117 | 0.24832 | -94 | 18.5 | 6.1338 | 18.9938 | 0.16147 | -102 | 7 | -1.8662 | 8.2168 | 0.11356 | -104 | 20 | -3.8662 | 20.4577 | 0.094493 | -88.8333 | 36.048 | 11.3004 | 36.3039 | 0.15564 | -87 | 19.4 | 13.1338 | 19.8715 | 0.33047 | -98 | 11 | 2.1338 | 11.8117 | 0.090325 | -100.1338 | 2.5463 | 5.0927 | 5 | -100 | 10.3 | 0.13378 | 11.1627 | 0.0059922 | -68.6825 | 17.8292 | 31.4513 | 18.3412 | 0.8574 | -99 | 15 | 1.1338 | 15.605 | 0.036327
| 10 | -92.275 | 8.5 | 0.26268 | 9.5233 | 0.013792 | -94 | 4 | -1.4623 | 5.8688 | 0.12458 | -89 | 12 | 3.5377 | 12.7453 | 0.13878 | -80 | 15 | 12.5377 | 15.6026 | 0.40178 | -94 | 37 | -1.4623 | 37.2484 | 0.019629 | -101 | 10 | -8.4623 | 10.8831 | 0.38878 | -90 | 18.5 | 2.5377 | 18.9919 | 0.06681 | -94 | 7 | -1.4623 | 8.2123 | 0.089032 | -114 | 22 | -21.4623 | 22.4152 | 0.47874 | -73.1667 | 36.076 | 19.371 | 36.3307 | 0.26659 | -74 | 19.4 | 18.5377 | 19.8696 | 0.46648 | -90 | 11 | 2.5377 | 11.8086 | 0.10745 | -92.5377 | 2.5608 | 5.1216 | 5 | -94 | 10.3 | -1.4623 | 11.1594 | 0.06552 | -65.9774 | 18.0226 | 26.5603 | 18.5272 | 0.71679 | -91 | 15 | 1.5377 | 15.6026 | 0.049276
| 5 | -85.7 | 8.5 | -0.0050327 | 9.4845 | 0.00026531 | -88 | 4 | -2.305 | 5.8057 | 0.19851 | -84 | 18 | 1.695 | 18.4853 | 0.045846 | -61 | 15 | 24.695 | 15.579 | 0.79257 | -91 | 37 | -5.305 | 37.2385 | 0.07123 | -95 | 11 | -9.305 | 11.7774 | 0.39504 | -77 | 27.3 | 8.695 | 27.6224 | 0.15739 | -86 | 7 | -0.30503 | 8.1674 | 0.018674 | -132 | 30 | -46.305 | 30.2937 | 0.76427 | -38.5 | 44.7914 | 47.195 | 44.9886 | 0.52452 | -55 | 19.4 | 30.695 | 19.8511 | 0.77313 | -84 | 15 | 1.695 | 15.579 | 0.054399 | -85.695 | 2.7007 | 5.4015 | 5 | -89 | 12.3 | -3.305 | 12.9998 | 0.12712 | -61.2186 | 19.1823 | 24.4763 | 19.6383 | 0.62318 | -86 | 15 | -0.30503 | 15.579 | 0.0097899
| 2 | -81.2 | 8.5 | -0.47656 | 10.0186 | 0.023784 | -82 | 4 | -1.2766 | 6.6425 | 0.09609 | -77 | 18 | 3.7234 | 18.7649 | 0.099213 | -29 | 26 | 51.7234 | 26.5353 | 0.97462 | -92 | 38 | -11.2766 | 38.3683 | 0.14695 | -89 | 11 | -8.2766 | 12.2116 | 0.33888 | -73 | 41.5 | 7.7234 | 41.8375 | 0.092303 | -79 | 7 | 1.7234 | 8.782 | 0.098124 | -213 | 50 | -132.2766 | 50.4369 | 1.3113 | [5 empty] | -9 | 28.4 | 71.7234 | 29.1623 | 1.2297 | -80 | 15 | 0.72344 | 15.9098 | 0.022736 | -80.7234 | 2.8066 | 5.6132 | 6 | -83 | 13.3 | -2.2766 | 14.3183 | 0.079498 | -65.1343 | 18.5228 | 15.5891 | 19.267 | 0.40455 | -81 | 18 | -0.27656 | 18.7649 | 0.007369
| 1 | -76.4 | 8.5 | 0.058119 | 11.3103 | 0.0025693 | -77 | 4 | -0.54188 | 8.466 | 0.032003 | -71 | 18 | 5.4581 | 19.4852 | 0.14006 | 8 | 35 | 84.4581 | 36.0184 | 1.1724 | -100 | 38 | -23.5419 | 38.7256 | 0.30396 | -76 | 11 | 0.45812 | 13.2919 | 0.017233 | -66 | 51.2 | 10.4581 | 51.7408 | 0.10106 | -74 | 7 | 2.4581 | 10.231 | 0.12013 | -342 | 100 | -265.5419 | 100.361 | 1.3229 | [10 empty] | -80 | 15 | -3.5419 | 16.7533 | 0.10571 | -76.4581 | 2.8855 | 5.771 | 8 | -79 | 15.4 | -2.5419 | 17.1124 | 0.07427 | [5 empty] | -76 | 21 | 0.45812 | 22.2862 | 0.010278
5/5 | 120 | -99.175 | 9 | -2.0623 | 9.9375 | 0.10376 | -99 | 4 | -1.8873 | 5.8098 | 0.16242 | -90 | 11 | 7.1127 | 11.7794 | 0.30191 | -94 | 15 | 3.1127 | 15.5806 | 0.099891 | -98 | 21 | -0.88728 | 21.4185 | 0.020713 | -105 | 14 | -7.8873 | 14.6203 | 0.26974 | -94 | 18.5 | 3.1127 | 18.9738 | 0.082027 | -100 | 7 | -2.8873 | 8.1703 | 0.17669 | [10 empty] | -85 | 19.4 | 12.1127 | 19.8523 | 0.30507 | [5 empty] | -97.1127 | 2.6919 | 5.3837 | 5 | -93 | 10.3 | 4.1127 | 11.1285 | 0.18478 | -75.6384 | 17.615 | 21.4743 | 18.112 | 0.59282 | -92 | 15 | 5.1127 | 15.5806 | 0.16407
| 100 | -97.65 | 8.5 | -2.184 | 9.5238 | 0.11466 | -97 | 4 | -1.534 | 5.8697 | 0.13067 | -89 | 11 | 6.466 | 11.809 | 0.27377 | -98 | 15 | -2.534 | 15.603 | 0.081202 | -95 | 21 | 0.46602 | 21.4349 | 0.010871 | -102 | 12 | -6.534 | 12.7457 | 0.25632 | -92 | 18.5 | 3.466 | 18.9922 | 0.091249 | -99 | 7 | -3.534 | 8.213 | 0.21515 | -98 | 20 | -2.534 | 20.4561 | 0.061937 | -116.3333 | 41.2495 | -20.8673 | 41.4725 | 0.25158 | -87 | 19.4 | 8.466 | 19.8699 | 0.21304 | -88 | 11 | 7.466 | 11.809 | 0.31611 | -95.466 | 2.5586 | 5.1172 | 5 | -91 | 10.3 | 4.466 | 11.1599 | 0.20009 | -76.3814 | 17.4472 | 19.0846 | 17.9682 | 0.53106 | -91 | 15 | 4.466 | 15.603 | 0.14311
| 50 | -87.325 | 8.5 | -0.58329 | 9.5284 | 0.030608 | -88 | 4 | -1.2583 | 5.8771 | 0.10705 | -82 | 11 | 4.7417 | 11.8127 | 0.2007 | -93 | 15 | -6.2583 | 15.6058 | 0.20051 | -82 | 20.5 | 4.7417 | 20.9473 | 0.11318 | -92 | 12 | -5.2583 | 12.7491 | 0.20622 | -83 | 18.5 | 3.7417 | 18.9945 | 0.098495 | -89 | 7 | -2.2583 | 8.2183 | 0.13739 | -92 | 20 | -5.2583 | 20.4582 | 0.12851 | -106.4 | 41.2449 | -19.6583 | 41.469 | 0.23702 | -78 | 19.4 | 8.7417 | 19.8721 | 0.21995 | -84 | 10 | 2.7417 | 10.8876 | 0.12591 | -86.7417 | 2.5417 | 5.0833 | 5 | -83 | 10.2 | 3.7417 | 11.0716 | 0.16898 | -68.0194 | 17.4281 | 18.7223 | 17.9521 | 0.52145 | -84 | 15 | 2.7417 | 15.6058 | 0.087843
| 20 | -73.0875 | 8.5 | -0.0087133 | 9.5311 | 0.0004571 | -74 | 4 | -0.92121 | 5.8815 | 0.078314 | -70 | 11 | 3.0788 | 11.8149 | 0.13029 | -75 | 15 | -1.9212 | 15.6074 | 0.061548 | -71 | 20.5 | 2.0788 | 20.9486 | 0.049616 | -78 | 11 | -4.9212 | 11.8149 | 0.20826 | -65 | 18.5 | 8.0788 | 18.9958 | 0.21265 | -73 | 7 | 0.078787 | 8.2214 | 0.0047915 | -84 | 20 | -10.9212 | 20.4595 | 0.2669 | -86.6667 | 41.2165 | -13.5879 | 41.4414 | 0.16394 | -59 | 19.4 | 14.0788 | 19.8734 | 0.35421 | -73 | 10 | 0.078787 | 10.89 | 0.0036174 | -73.0788 | 2.5314 | 5.0628 | 5 | -73 | 10.2 | 0.078787 | 11.0739 | 0.0035573 | -59.814 | 17.593 | 13.2647 | 18.1137 | 0.36615 | -72 | 15 | 1.0788 | 15.6074 | 0.03456
| 10 | -65.9 | 8.5 | 0.57227 | 9.5238 | 0.030044 | -68 | 4 | -1.5277 | 5.8696 | 0.13014 | -64 | 12 | 2.4723 | 12.7457 | 0.096985 | -57 | 15 | 9.4723 | 15.6029 | 0.30354 | -66 | 37 | 0.47227 | 37.2485 | 0.0063394 | -73 | 11 | -6.5277 | 11.809 | 0.27639 | -61 | 18.5 | 5.4723 | 18.9922 | 0.14407 | -66 | 7 | 0.47227 | 8.2129 | 0.028752 | -92 | 22 | -25.5277 | 22.4154 | 0.56942 | -62.3333 | 41.2032 | 4.1389 | 41.4265 | 0.049955 | -45 | 19.4 | 21.4723 | 19.8699 | 0.54032 | -66 | 10 | 0.47227 | 10.8836 | 0.021696 | -66.4723 | 2.5589 | 5.1178 | 5 | -67 | 10.2 | -0.52773 | 11.0676 | 0.023841 | -53.328 | 17.417 | 13.1443 | 17.9389 | 0.36636 | -65 | 15 | 1.4723 | 15.6029 | 0.047179
| 5 | -60.7375 | 8.5 | -0.25476 | 10.0455 | 0.012681 | -63 | 4 | -2.5173 | 6.6829 | 0.18834 | -58 | 18 | 2.4827 | 18.7793 | 0.066103 | -39 | 15 | 21.4827 | 15.9268 | 0.67442 | -68 | 37 | -7.5173 | 37.3853 | 0.10054 | -68 | 11 | -7.5173 | 12.2336 | 0.30724 | -50 | 27.3 | 10.4827 | 27.82 | 0.1884 | -58 | 7 | 2.4827 | 8.8126 | 0.14086 | -118 | 30 | -57.5173 | 30.4739 | 0.94371 | -11 | 51.191 | 49.4827 | 51.4702 | 0.48069 | -26 | 19.4 | 34.4827 | 20.1251 | 0.85671 | -60 | 15 | 0.48274 | 15.9268 | 0.015155 | -60.4827 | 2.709 | 5.418 | 6 | -62 | 12.3 | -1.5173 | 13.4146 | 0.056553 | -50.6155 | 22.3582 | 9.8672 | 22.9903 | 0.2146 | -60 | 15 | 0.48274 | 15.9268 | 0.015155
| 2 | -56.2563 | 8.5 | -0.77974 | 10.0216 | 0.038903 | -57 | 4 | -1.5235 | 6.647 | 0.1146 | -54 | 18 | 1.4765 | 18.7665 | 0.039339 | -7 | 26 | 48.4765 | 26.5364 | 0.9134 | -69 | 38 | -13.5235 | 38.369 | 0.17623 | -64 | 11 | -8.5235 | 12.214 | 0.34892 | -46 | 41.5 | 9.4765 | 41.8382 | 0.11325 | -54 | 7 | 1.4765 | 8.7853 | 0.084033 | -220 | 50 | -164.5235 | 50.4363 | 1.631 | [5 empty] | -2 | 28.4 | 53.4765 | 28.8919 | 0.92546 | -56 | 15 | -0.52349 | 15.9117 | 0.01645 | -55.4765 | 2.796 | 5.592 | 6 | -57 | 13.3 | -1.5235 | 14.3203 | 0.053193 | -47.8331 | 19.4679 | 7.6434 | 20.1788 | 0.18939 | -56 | 18 | -0.52349 | 18.7665 | 0.013947
| 1 | -52.2313 | 8.5 | 0.61987 | 9.9662 | 0.031099 | -53 | 4 | -0.14888 | 6.5631 | 0.011342 | -51 | 18 | 1.8511 | 18.737 | 0.049397 | 30 | 35 | 82.8511 | 35.636 | 1.1625 | -79 | 38 | -26.1489 | 38.3546 | 0.34088 | -56 | 11 | -3.1489 | 12.1686 | 0.12939 | -41 | 51.2 | 11.8511 | 51.4637 | 0.11514 | -49 | 9 | 3.8511 | 10.3959 | 0.18522 | -365 | 100 | -312.1489 | 100.2244 | 1.5573 | [10 empty] | -57 | 15 | -4.1489 | 15.8769 | 0.13066 | -52.8511 | 2.9875 | 5.9751 | 6 | -52 | 15.4 | 0.85112 | 16.2553 | 0.02618 | [5 empty] | -52 | 21 | 0.85112 | 21.635 | 0.01967
6/5 | 120 | -74.9625 | 8.5 | 1.5447 | 9.5322 | 0.081027 | -78 | 3.5 | -1.4928 | 5.5554 | 0.13435 | -78 | 11 | -1.4928 | 11.8158 | 0.063169 | -70 | 15 | 6.5072 | 15.6081 | 0.20846 | -81 | 26 | -4.4928 | 26.3555 | 0.085234 | -83 | 16 | -6.4928 | 16.5714 | 0.1959 | -73 | 18.5 | 3.5072 | 18.9964 | 0.092313 | -79 | 7 | -2.4928 | 8.2227 | 0.15158 | [10 empty] | -40 | 19.4 | 36.5072 | 19.8739 | 0.91847 | [5 empty] | -76.5072 | 2.5274 | 5.0547 | 5 | -73 | 10.2 | 3.5072 | 11.0749 | 0.15834 | -63.9642 | 18.4376 | 12.543 | 18.9356 | 0.3312 | -79 | 15 | -2.4928 | 15.6081 | 0.079855
| 100 | -73.0625 | 8.5 | 1.7237 | 9.536 | 0.090377 | -76 | 3.5 | -1.2138 | 5.562 | 0.10912 | -76 | 11 | -1.2138 | 11.8189 | 0.051351 | -72 | 15 | 2.7862 | 15.6104 | 0.089241 | -75 | 26 | -0.21382 | 26.3569 | 0.0040562 | -79 | 13 | -4.2138 | 13.6998 | 0.15379 | -71 | 18.5 | 3.7862 | 18.9983 | 0.099645 | -77 | 7 | -2.2138 | 8.2271 | 0.13454 | [10 empty] | -41 | 19.4 | 33.7862 | 19.8757 | 0.84993 | [5 empty] | -74.7862 | 2.5129 | 5.0258 | 5 | -72 | 10.2 | 2.7862 | 11.0782 | 0.12575 | -63.8525 | 19.0974 | 10.9336 | 19.5805 | 0.2792 | -78 | 15 | -3.2138 | 15.6104 | 0.10294
| 50 | -64.175 | 8.5 | 2.1506 | 9.5528 | 0.11256 | -67 | 3.5 | -0.67442 | 5.5907 | 0.060317 | -65 | 11 | 1.3256 | 11.8324 | 0.056015 | -74 | 15 | -7.6744 | 15.6207 | 0.24565 | -68 | 27.5 | -1.6744 | 27.8434 | 0.030068 | -69 | 12 | -2.6744 | 12.7674 | 0.10474 | -63 | 18.5 | 3.3256 | 19.0067 | 0.087484 | -68 | 7 | -1.6744 | 8.2465 | 0.10152 | -71 | 20 | -4.6744 | 20.4696 | 0.11418 | -24.5 | 43.1592 | 41.8256 | 43.3788 | 0.4821 | -33 | 19.4 | 33.3256 | 19.8838 | 0.83801 | -69 | 15 | -2.6744 | 15.6207 | 0.085605 | -66.3256 | 2.4484 | 4.8968 | 5 | -65 | 10.2 | 1.3256 | 11.0926 | 0.059751 | -59.7372 | 18.7535 | 6.5884 | 19.2536 | 0.1711 | -72 | 15 | -5.6744 | 15.6207 | 0.18163
| 20 | -51.3875 | 8.5 | 2.8561 | 9.5557 | 0.14944 | -56 | 3.5 | -1.7564 | 5.5957 | 0.15694 | -48 | 11 | 6.2436 | 11.8348 | 0.26378 | -57 | 15 | -2.7564 | 15.6225 | 0.08822 | -58 | 26 | -3.7564 | 26.364 | 0.071242 | -59 | 11 | -4.7564 | 11.8348 | 0.20095 | -45 | 18.5 | 9.2436 | 19.0082 | 0.24315 | -53 | 7 | 1.2436 | 8.25 | 0.075367 | -67 | 20 | -12.7564 | 20.471 | 0.31157 | -17.6667 | 42.091 | 36.5769 | 42.3169 | 0.43218 | -19 | 19.4 | 35.2436 | 19.8852 | 0.88617 | -58 | 15 | -3.7564 | 15.6225 | 0.12023 | -54.2436 | 2.4367 | 4.8735 | 5 | -53 | 10.2 | 1.2436 | 11.0952 | 0.056041 | -49.5472 | 18.4466 | 4.6963 | 18.9562 | 0.12387 | -62 | 15 | -7.7564 | 15.6225 | 0.24825
| 10 | -44.3 | 8.5 | 3.6673 | 10.1097 | 0.18138 | -49 | 3.5 | -1.0327 | 6.4966 | 0.079479 | -41 | 12 | 6.9673 | 13.1892 | 0.26413 | -40 | 15 | 7.9673 | 15.9673 | 0.24949 | -49 | 26 | -1.0327 | 26.5698 | 0.019434 | -53 | 10 | -5.0327 | 11.3998 | 0.22074 | -42 | 18.5 | 5.9673 | 19.2926 | 0.15465 | -46 | 7 | 1.9673 | 8.8857 | 0.1107 | -74 | 22 | -26.0327 | 22.6706 | 0.57415 | 22.8333 | 41.9172 | 70.8006 | 42.273 | 0.83742 | -5 | 19.4 | 42.9673 | 20.4549 | 1.0503 | -49 | 15 | -1.0327 | 15.9673 | 0.032338 | -47.9673 | 2.4585 | 4.9171 | 6 | -49 | 10.2 | -1.0327 | 11.5756 | 0.044606 | -45.0363 | 19.4775 | 2.931 | 20.2319 | 0.072436 | -56 | 15 | -8.0327 | 15.9673 | 0.25154
| 5 | -40 | 8.5 | 3.6563 | 10.0739 | 0.18148 | -45 | 3.5 | -1.3437 | 6.4407 | 0.10431 | -35 | 19 | 8.6563 | 19.7543 | 0.2191 | -22 | 15 | 21.6563 | 15.9447 | 0.67911 | -51 | 40 | -7.3437 | 40.3638 | 0.090969 | -48 | 11 | -4.3437 | 12.257 | 0.17719 | -41 | 27.3 | 2.6563 | 27.8303 | 0.047724 | -40 | 7 | 3.6563 | 8.845 | 0.20669 | -99 | 30 | -55.3437 | 30.4833 | 0.90777 | 83.8333 | 51.6379 | 127.4897 | 52.0504 | 1.2247 | 0 | 19.4 | 43.6563 | 20.4726 | 1.0662 | -43 | 37 | 0.65633 | 37.393 | 0.0087761 | -43.6563 | 2.6013 | 5.2026 | 6 | -45 | 12.3 | -1.3437 | 13.4359 | 0.050003 | -41.2881 | 25.3334 | 2.3682 | 25.904 | 0.045711 | -51 | 15 | -7.3437 | 15.9447 | 0.23029
| 2 | -32.45 | 8.5 | 5.8997 | 10.6865 | 0.27603 | -40 | 3.5 | -1.6503 | 7.3622 | 0.11208 | -31 | 19 | 7.3497 | 20.0736 | 0.18307 | 10 | 26 | 48.3497 | 26.7946 | 0.90223 | -55 | 49 | -16.6503 | 49.4262 | 0.16844 | -47 | 11 | -8.6503 | 12.7652 | 0.33882 | -27 | 41.5 | 11.3497 | 42.0024 | 0.13511 | -36 | 7 | 2.3497 | 9.5368 | 0.12319 | -198 | 50 | -159.6503 | 50.5574 | 1.5789 | [5 empty] | 26 | 28.4 | 64.3497 | 29.3702 | 1.0955 | -36 | 37 | 2.3497 | 37.5626 | 0.031277 | -38.3497 | 2.6549 | 5.3098 | 7 | -40 | 13.2 | -1.6503 | 14.7034 | 0.05612 | -32.7301 | 18.4938 | 5.6196 | 19.5952 | 0.14339 | -46 | 18 | -7.6503 | 19.1299 | 0.19996
| 1 | -26.6625 | 8.5 | 7.4838 | 10.6697 | 0.35071 | -36 | 3.5 | -1.8537 | 7.3377 | 0.12631 | -26 | 19 | 8.1463 | 20.0647 | 0.203 | 47 | 35 | 81.1463 | 35.7968 | 1.1334 | -60 | 49 | -25.8537 | 49.4226 | 0.26156 | -35 | 11 | -0.85367 | 12.7511 | 0.033474 | -22 | 51.2 | 12.1463 | 51.6046 | 0.11769 | -32 | 7 | 2.1463 | 9.518 | 0.11275 | -348 | 100 | -313.8537 | 100.2816 | 1.5649 | [10 empty] | -25 | 37 | 9.1463 | 37.5578 | 0.12176 | -34.1463 | 2.7218 | 5.4436 | 7 | -36 | 15.3 | -1.8537 | 16.6037 | 0.055821 | [5 empty] | -36 | 21 | -1.8537 | 21.968 | 0.04219
8/5 | 120 | -59.15 | 8.5 | -4.6468 | 10.6502 | 0.21816 | -55 | 3.5 | -0.49679 | 7.3093 | 0.033983 | -52 | 11 | 2.5032 | 12.7348 | 0.098282 | -57 | 15 | -2.4968 | 16.3149 | 0.076519 | [10 empty] | -50 | 18.5 | 4.5032 | 19.5813 | 0.11499 | [15 empty] | -41 | 19.4 | 13.5032 | 20.4337 | 0.33042 | [5 empty] | -54.5032 | 2.7971 | 5.5942 | 7 | -49 | 10.2 | 5.5032 | 12.0506 | 0.22834 | [5 empty] | -56 | 15 | -1.4968 | 16.3149 | 0.045872
| 100 | -58.8875 | 8.5 | -6.4565 | 9.5295 | 0.33877 | -53 | 3.5 | -0.56905 | 5.5508 | 0.051258 | -50 | 11 | 2.431 | 11.8136 | 0.10289 | -50 | 15 | 2.431 | 15.6065 | 0.077883 | -57 | 26 | -4.569 | 26.3545 | 0.086684 | -53 | 17 | -0.56905 | 17.5374 | 0.016224 | -48 | 18.5 | 4.431 | 18.995 | 0.11663 | -51 | 7 | 1.431 | 8.2196 | 0.087046 | [10 empty] | -42 | 19.4 | 10.431 | 19.8726 | 0.26245 | [5 empty] | -52.431 | 2.5374 | 5.0749 | 5 | -49 | 10.2 | 3.431 | 11.0726 | 0.15493 | -46.2211 | 21.5784 | 6.2099 | 22.0043 | 0.14111 | -53 | 15 | -0.56905 | 15.6065 | 0.018231
| 50 | -47.625 | 8.5 | -3.3764 | 9.5423 | 0.17692 | -44 | 3.5 | 0.24861 | 5.5727 | 0.022306 | -42 | 11 | 2.2486 | 11.8239 | 0.095087 | -43 | 15 | 1.2486 | 15.6143 | 0.039983 | -42 | 28 | 2.2486 | 28.3338 | 0.039681 | -45 | 12 | -0.75139 | 12.7595 | 0.029444 | -40 | 18.5 | 4.2486 | 19.0015 | 0.1118 | -42 | 8 | 2.2486 | 9.0997 | 0.12355 | -50 | 20 | -5.7514 | 20.4647 | 0.14052 | -54.5 | 43.1361 | -10.2514 | 43.3535 | 0.11823 | -35 | 19.4 | 9.2486 | 19.8788 | 0.23263 | -55 | 15 | -10.7514 | 15.6143 | 0.34428 | -44.2486 | 2.4889 | 4.9778 | 5 | -42 | 10.2 | 2.2486 | 11.0836 | 0.10144 | -44.3761 | 20.9245 | -0.12747 | 21.3691 | 0.0029825 | -49 | 15 | -4.7514 | 15.6143 | 0.15215
| 20 | -35.0625 | 8.5 | -2.41 | 9.5544 | 0.12612 | -33 | 3.5 | -0.34755 | 5.5934 | 0.031068 | -27 | 11 | 5.6525 | 11.8337 | 0.23883 | -33 | 15 | -0.34755 | 15.6217 | 0.011124 | -40 | 26 | -7.3475 | 26.3635 | 0.13935 | -37 | 11 | -4.3475 | 11.8337 | 0.18369 | -23 | 18.5 | 9.6525 | 19.0075 | 0.25391 | -28 | 7 | 4.6525 | 8.2484 | 0.28202 | -48 | 20 | -15.3475 | 20.4704 | 0.37487 | 20 | 43.237 | 52.6525 | 43.4565 | 0.60581 | -20 | 19.4 | 12.6525 | 19.8846 | 0.31815 | -46 | 15 | -13.3475 | 15.6217 | 0.42721 | -32.6525 | 2.4421 | 4.8842 | 5 | -30 | 10.2 | 2.6525 | 11.094 | 0.11954 | -36.1736 | 21.1999 | -3.5212 | 21.6442 | 0.081342 | -42 | 15 | -9.3475 | 15.6217 | 0.29919
| 10 | -27.075 | 8.5 | -1.0381 | 10.1109 | 0.051336 | -27 | 3.5 | -0.96311 | 6.4986 | 0.074102 | -18 | 12 | 8.0369 | 13.1902 | 0.30465 | -16 | 15 | 10.0369 | 15.9681 | 0.31428 | -35 | 26 | -8.9631 | 26.5703 | 0.16867 | -30 | 11 | -3.9631 | 12.2874 | 0.16127 | -20 | 18.5 | 6.0369 | 19.2933 | 0.15645 | -22 | 7 | 4.0369 | 8.8871 | 0.22712 | -57 | 22 | -30.9631 | 22.6712 | 0.68287 | 58.1667 | 42.9838 | 84.2036 | 43.3312 | 0.97163 | -9 | 19.4 | 17.0369 | 20.1579 | 0.42259 | -37 | 15 | -10.9631 | 15.9681 | 0.34328 | -26.0369 | 2.4533 | 4.9066 | 6 | -25 | 10.1 | 1.0369 | 11.4887 | 0.045126 | -30.4331 | 21.0571 | -4.3962 | 21.7573 | 0.10103 | -35 | 15 | -8.9631 | 15.9681 | 0.28066
| 5 | -22.3 | 8.5 | -0.33642 | 10.0811 | 0.016685 | -23 | 3.5 | -1.0364 | 6.4521 | 0.080316 | -13 | 19 | 8.9636 | 19.758 | 0.22683 | 2 | 15 | 23.9636 | 15.9493 | 0.75124 | -31 | 40 | -9.0364 | 40.3656 | 0.11193 | -28 | 11 | -6.0364 | 12.2629 | 0.24612 | -19 | 27.3 | 2.9636 | 27.8329 | 0.053239 | -18 | 7 | 3.9636 | 8.8532 | 0.22385 | -75 | 30 | -53.0364 | 30.4857 | 0.86986 | 105.8333 | 53.4093 | 127.7969 | 53.8068 | 1.1876 | -5 | 19.4 | 16.9636 | 20.143 | 0.42108 | -31 | 37 | -9.0364 | 37.3949 | 0.12082 | -21.9636 | 2.573 | 5.146 | 6 | -21 | 12.2 | 0.96358 | 13.3499 | 0.03609 | -28.1878 | 22.0396 | -6.2242 | 22.6964 | 0.13712 | -32 | 15 | -10.0364 | 15.9493 | 0.31464
| 2 | -17.55 | 8.5 | -0.88673 | 10.0614 | 0.044066 | -18 | 3.5 | -1.3367 | 6.4211 | 0.10409 | -8 | 19 | 8.6633 | 19.7479 | 0.21935 | 34 | 26 | 50.6633 | 26.5515 | 0.95406 | -35 | 49 | -18.3367 | 49.2948 | 0.18599 | -25 | 11 | -8.3367 | 12.2467 | 0.34037 | -6 | 41.5 | 10.6633 | 41.8477 | 0.12741 | -11 | 7 | 5.6633 | 8.8307 | 0.32066 | -168 | 50 | -151.3367 | 50.4284 | 1.5005 | [5 empty] | 19 | 28.4 | 35.6633 | 28.9057 | 0.61689 | -23 | 37 | -6.3367 | 37.3896 | 0.084739 | -16.6633 | 2.6493 | 5.2987 | 6 | -17 | 13.2 | -0.33673 | 14.2556 | 0.011811 | -26.4582 | 20.9293 | -9.7949 | 21.6106 | 0.22662 | -29 | 18 | -12.3367 | 18.7878 | 0.32832
| 1 | -3.3125 | 8.5 | 9.858 | 10.6697 | 0.46196 | -15 | 3.5 | -1.8295 | 7.3377 | 0.12467 | -5 | 19 | 8.1705 | 20.0647 | 0.2036 | 70 | 35 | 83.1705 | 35.7968 | 1.1617 | -32 | 49 | -18.8295 | 49.4226 | 0.1905 | -20 | 11 | -6.8295 | 12.7511 | 0.2678 | -1 | 51.2 | 12.1705 | 51.6046 | 0.11792 | -11 | 7 | 2.1705 | 9.518 | 0.11402 | -322 | 100 | -308.8295 | 100.2816 | 1.5398 | [10 empty] | -7 | 37 | 6.1705 | 37.5578 | 0.082146 | -13.1705 | 2.7218 | 5.4436 | 7 | -13 | 15.3 | 0.17046 | 16.6037 | 0.0051332 | [5 empty] | -13 | 21 | 0.17046 | 21.968 | 0.0038798
10/5 | 120 | [5 empty] | -40 | 3.5 | -0.80996 | 5.2428 | 0.077244 | -38 | 11 | 1.19 | 11.6721 | 0.050978 | [15 empty] | -37 | 18.5 | 2.19 | 18.9073 | 0.057915 | [25 empty] | -39.19 | 3.1245 | 6.2491 | 5 | -34 | 10.2 | 5.19 | 10.9214 | 0.23761
| 100 | -36.575 | 8.5 | 1.0526 | 9.4521 | 0.05568 | -39 | 3.5 | -1.3724 | 5.4169 | 0.12668 | -37 | 11 | 0.62758 | 11.7513 | 0.026703 | -42 | 15 | -4.3724 | 15.5593 | 0.14051 | [10 empty] | -35 | 18.5 | 2.6276 | 18.9563 | 0.069306 | [15 empty] | -18 | 19.4 | 19.6276 | 19.8356 | 0.49476 | [5 empty] | -37.6276 | 2.8121 | 5.6241 | 5 | -35 | 10.2 | 2.6276 | 11.006 | 0.11937 | [5 empty] | -29 | 18 | 8.6276 | 18.4687 | 0.23357
| 50 | -27.3 | 8.5 | 0.46058 | 9.5322 | 0.024159 | -30 | 3.5 | -2.2394 | 5.5555 | 0.20155 | -29 | 11 | -1.2394 | 11.8158 | 0.052447 | -27 | 15 | 0.76058 | 15.6081 | 0.024365 | -29 | 26.5 | -1.2394 | 26.8489 | 0.023081 | -28 | 13 | -0.23942 | 13.6972 | 0.0087397 | -16 | 18.5 | 11.7606 | 18.9964 | 0.30955 | -25 | 8 | 2.7606 | 9.0892 | 0.15186 | [5 empty] | 5.3333 | 43.0546 | 33.0939 | 43.2702 | 0.38241 | -14 | 19.4 | 13.7606 | 19.8739 | 0.3462 | -20 | 15 | 7.7606 | 15.6081 | 0.24861 | -27.7606 | 2.5272 | 5.0545 | 5 | -27 | 10.2 | 0.76058 | 11.0749 | 0.034338 | -30.0801 | 21.508 | -2.3196 | 21.9364 | 0.05287 | -22 | 18 | 5.7606 | 18.5098 | 0.15561
| 20 | -15.825 | 8.5 | 1.2551 | 9.5522 | 0.065697 | -19 | 3.5 | -1.9199 | 5.5896 | 0.17174 | -14 | 11 | 3.0801 | 11.8319 | 0.13016 | -22 | 15 | -4.9199 | 15.6203 | 0.15748 | -14 | 26 | 3.0801 | 26.3627 | 0.058418 | -20 | 11 | -2.9199 | 11.8319 | 0.12339 | -9 | 18.5 | 8.0801 | 19.0064 | 0.21256 | -14 | 7 | 3.0801 | 8.2458 | 0.18677 | -34 | 20 | -16.9199 | 20.4693 | 0.4133 | 50.3333 | 42.7072 | 67.4134 | 42.929 | 0.78517 | -2 | 19.4 | 15.0801 | 19.8835 | 0.37921 | -14 | 15 | 3.0801 | 15.6203 | 0.098593 | -17.0801 | 2.4508 | 4.9015 | 5 | -14 | 10.1 | 3.0801 | 11.0002 | 0.14 | -27.2342 | 21.3233 | -10.1541 | 21.7641 | 0.23328 | -13 | 18 | 4.0801 | 18.5201 | 0.11015
| 10 | -9.75 | 8.5 | 1.7469 | 9.5462 | 0.091499 | -13 | 3.5 | -1.5031 | 5.5795 | 0.1347 | -4 | 12 | 7.4969 | 12.7625 | 0.29371 | -6 | 15 | 5.4969 | 15.6167 | 0.176 | -14 | 26 | -2.5031 | 26.3606 | 0.047477 | -16 | 11 | -4.5031 | 11.8271 | 0.19037 | -7 | 18.5 | 4.4969 | 19.0034 | 0.11832 | -8 | 7 | 3.4969 | 8.239 | 0.21222 | -42 | 22 | -30.5031 | 22.425 | 0.68011 | 78.8333 | 42.0047 | 90.3303 | 42.3735 | 1.0659 | 1 | 19.4 | 12.4969 | 19.8807 | 0.3143 | -8 | 15 | 3.4969 | 15.6167 | 0.11196 | -11.4969 | 2.4738 | 4.9476 | 5 | -11 | 10.1 | 0.49693 | 10.995 | 0.022598 | -23.5345 | 21.159 | -12.0375 | 21.6005 | 0.27864 | -5 | 21 | 6.4969 | 21.4448 | 0.15148
| 5 | -5.25 | 8.5 | 1.7898 | 9.5135 | 0.094068 | -9 | 3.5 | -1.9602 | 5.5233 | 0.17745 | 0 | 19 | 7.0398 | 19.4745 | 0.18074 | 13 | 15 | 20.0398 | 15.5967 | 0.64244 | -7 | 40 | 0.039822 | 40.2276 | 0.00049495 | -12 | 11 | -4.9602 | 11.8007 | 0.21016 | 4 | 27.3 | 11.0398 | 27.6323 | 0.19976 | -3 | 7 | 4.0398 | 8.201 | 0.2463 | -58 | 30 | -50.9602 | 30.3027 | 0.84085 | 128.5 | 52.5689 | 135.5398 | 52.87 | 1.2818 | 8 | 19.4 | 15.0398 | 19.865 | 0.37855 | -3 | 37 | 4.0398 | 37.2459 | 0.054232 | -7.0398 | 2.5968 | 5.1936 | 5 | -8 | 12.2 | -0.96018 | 12.9266 | 0.03714 | -16.9528 | 22.2971 | -9.9129 | 22.7028 | 0.21832 | -1 | 24 | 6.0398 | 24.3774 | 0.12388
| 2 | -0.9 | 8.5 | 2.0748 | 10.6801 | 0.097133 | -5 | 3.5 | -2.0252 | 7.3529 | 0.13772 | 4 | 19 | 6.9748 | 20.0702 | 0.17376 | 45 | 26 | 47.9748 | 26.7921 | 0.89532 | -7 | 49 | -4.0252 | 49.4248 | 0.040721 | -10 | 11 | -7.0252 | 12.7599 | 0.27529 | 5 | 41.5 | 7.9748 | 42.0008 | 0.094936 | 1 | 7 | 3.9748 | 9.5297 | 0.20855 | -149 | 50 | -146.0252 | 50.5587 | 1.4441 | [5 empty] | 27 | 28.4 | 29.9748 | 29.1269 | 0.51456 | 3 | 37 | 5.9748 | 37.5608 | 0.079535 | -2.9748 | 2.6806 | 5.3611 | 7 | -3 | 13.2 | -0.025226 | 14.6988 | 0.00085808 | -13.1614 | 21.9843 | -10.1866 | 22.9156 | 0.22226
| 1 | -1.3 | 8.5 | -0.30454 | 11.3453 | 0.013421 | -2 | 3.5 | -1.0045 | 8.2895 | 0.060591 | 8 | 19 | 8.9955 | 20.432 | 0.22013 | 82 | 35 | 82.9955 | 36.0074 | 1.1525 | -12 | 49 | -11.0045 | 49.5728 | 0.11099 | -4 | 11 | -3.0045 | 13.3216 | 0.11277 | 10 | 51.2 | 10.9955 | 51.7485 | 0.10624 | 3 | 7 | 3.9955 | 10.2696 | 0.19453 | -291 | 100 | -290.0045 | 100.357 | 1.4449 | [10 empty] | 0 | 37 | 0.99546 | 37.7553 | 0.013183 | -0.99546 | 2.745 | 5.4899 | 8 | 0 | 15.3 | 0.99546 | 17.0457 | 0.0292
Annex 4. TAB A4-2 Calculation without outliers.
Meaning of individual symbols:
| I1 | Primary current (%)IN
| kI | Transformation ratio I (kA/A)
| B | Burden (VA)
| I1N | Rated primary current (A)
| eL | Ratio error of each participating laboratory (ppm)
| u(eL) | Standard uncertainty of ratio error of each participating laboratory (ppm)
| erC | CRVC for ratio error according to (2) calculated without outliers (ppm)
| u(erC) | CRVC uncertainty according to (3) calculated without outliers (ppm)
| DeC | Difference between results of individual laboratory eL and erC without outliers, see eq. (10) (ppm)
| u(DestdC) | The uncertainty of the difference DeC without outliers according to (11) with transfer standard uncertaity (ppm)
| EC(e) | Corrected confidence coefficient without outliers, accoring to (12)
5 VA | I1 (%IN) |  |  | CMI | CMI |  |  |  | PTB | PTB |  |  |  | METAS | METAS |  |  |  | NPL | NPL |  |  |  | LCOE | LCOE |  |  |  | MIKES | MIKES |  |  |  | BEV | BEV |  |  |  | VSL | VSL |  |  |  | SP | SP |  |  |  | DMDM | DMDM |  |  |  | GUM | GUM |  |  |  | BIM | BIM |  |  |  |  | values for graphs |  |  | UME | UME |  |  |  | LNE | LNE |  |  |  | INRIM | INRIM
kI (kA/A) |  |  |  | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) |  | -U(erC) (ppm) | U(erC) (ppm) |  | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm) | DeC (ppm) | U(DestdC) (ppm) | novalue | DeC (ppm) | U(DestdC) (ppm)
4/5 | 120 |  |  | -0.35875 | 19.032 | #N/A | -0.35875 | 19.032 | 0.16625 | 11.7139 | #N/A | 0.16625 | 11.7139 | 6.1662 | 23.6054 | #N/A | 6.1662 | 23.6054 | -1.8338 | 31.1964 | #N/A | -1.8338 | 31.1964 | 1.1662 | 42.8628 | #N/A | 1.1662 | 42.8628 | -12.8338 | 27.3718 | #N/A | -12.8338 | 27.3718 | 4.1662 | 37.9765 | #N/A | 4.1662 | 37.9765 | -5.8338 | 16.4078 | #N/A | -5.8338 | 16.4078 |  |  | 0 | #N/A | #N/A | 7.3329 | 72.5579 | #N/A | 7.3329 | 72.5579 | 5.1662 | 39.7323 | #N/A | 5.1662 | 39.7323 | 8.1662 | 23.6054 | #N/A | 8.1662 | 23.6054 |  | -5.1753 | 5.1753 |  | 3.1662 | 22.3064 | #N/A | 3.1662 | 22.3064 | 0.35992 | 36.2727 | #N/A | 0.35992 | 36.2727 | 3.1662 | 31.1964 | #N/A | 3.1662 | 31.1964
| 100 |  |  | 0.50235 | 19.0563 | #N/A | 0.50235 | 19.0563 | -0.31432 | 11.7534 | #N/A | -0.31432 | 11.7534 | 6.6857 | 23.625 | #N/A | 6.6857 | 23.625 | -7.3143 | 31.2112 | #N/A | -7.3143 | 31.2112 | 1.6857 | 42.8736 | #N/A | 1.6857 | 42.8736 | -10.3143 | 25.4979 | #N/A | -10.3143 | 25.4979 | 5.6857 | 37.9887 | #N/A | 5.6857 | 37.9887 | -5.3143 | 16.436 | #N/A | -5.3143 | 16.436 | -1.3143 | 40.9163 | #N/A | -1.3143 | 40.9163 | 8.1857 | 72.5716 | #N/A | 8.1857 | 72.5716 | 5.6857 | 39.7439 | #N/A | 5.6857 | 39.7439 | 6.6857 | 21.7748 | #N/A | 6.6857 | 21.7748 |  | -5.0851 | 5.0851 |  | 4.6857 | 22.3271 | #N/A | 4.6857 | 22.3271 | 1.232 | 35.878 | #N/A | 1.232 | 35.878 | 3.6857 | 31.2112 | #N/A | 3.6857 | 31.2112
| 50 |  |  | 1.4482 | 18.072 | #N/A | 1.4482 | 18.072 | -1.0518 | 10.0795 | #N/A | -1.0518 | 10.0795 | 5.9482 | 22.8385 | #N/A | 5.9482 | 22.8385 | -8.0518 | 30.6202 | #N/A | -8.0518 | 30.6202 | 4.9482 | 41.456 | #N/A | 4.9482 | 41.456 | -7.0518 | 22.8385 | #N/A | -7.0518 | 22.8385 | 3.9482 | 37.5046 | #N/A | 3.9482 | 37.5046 | -3.0518 | 17.1347 | #N/A | -3.0518 | 17.1347 | -4.0518 | 40.4672 | #N/A | -4.0518 | 40.4672 | 6.7815 | 72.2997 | #N/A | 6.7815 | 72.2997 | 7.9482 | 39.2815 | #N/A | 7.9482 | 39.2815 | 3.9482 | 20.9188 | #N/A | 3.9482 | 20.9188 |  | -5.1384 | 5.1384 |  | 3.9482 | 21.3016 | #N/A | 3.9482 | 21.3016 | 2.5143 | 35.5015 | #N/A | 2.5143 | 35.5015 | 2.9482 | 30.6202 | #N/A | 2.9482 | 30.6202
| 20 |  |  | 0.78441 | 18.0943 | #N/A | 0.78441 | 18.0943 | -1.2656 | 10.1195 | #N/A | -1.2656 | 10.1195 | 4.7344 | 22.8562 | #N/A | 4.7344 | 22.8562 | -0.26559 | 30.6334 | #N/A | -0.26559 | 30.6334 | -2.2656 | 41.4657 | #N/A | -2.2656 | 41.4657 | -4.2656 | 22.8562 | #N/A | -4.2656 | 22.8562 | 8.7344 | 37.5154 | #N/A | 8.7344 | 37.5154 | -1.2656 | 15.3103 | #N/A | -1.2656 | 15.3103 | -13.2656 | 40.4772 | #N/A | -13.2656 | 40.4772 | 8.9011 | 72.3028 | #N/A | 8.9011 | 72.3028 | 12.7344 | 39.2918 | #N/A | 12.7344 | 39.2918 | 3.7344 | 20.9381 | #N/A | 3.7344 | 20.9381 |  | -5.0592 | 5.0592 |  | 1.7344 | 21.3205 | #N/A | 1.7344 | 21.3205 | -0.75102 | 35.5347 | #N/A | -0.75102 | 35.5347 | 2.7344 | 30.6334 | #N/A | 2.7344 | 30.6334
| 10 |  |  | 1.8579 | 18.0781 | #N/A | 1.8579 | 18.0781 | -1.0255 | 10.0904 | #N/A | -1.0255 | 10.0904 | 3.9745 | 24.7753 | #N/A | 3.9745 | 24.7753 | 9.9745 | 30.6238 | #N/A | 9.9745 | 30.6238 | -3.0255 | 74.2551 | #N/A | -3.0255 | 74.2551 | -10.0255 | 22.8433 | #N/A | -10.0255 | 22.8433 | 4.9745 | 37.5076 | #N/A | 4.9745 | 37.5076 | -1.0255 | 15.2911 | #N/A | -1.0255 | 15.2911 | -25.0255 | 44.4277 | #N/A | -25.0255 | 44.4277 | 22.9745 | 72.2927 | #N/A | 22.9745 | 72.2927 | 18.9745 | 39.2843 | #N/A | 18.9745 | 39.2843 | 3.9745 | 20.9241 | #N/A | 3.9745 | 20.9241 |  | -5.1169 | 5.1169 |  | -0.025454 | 21.3067 | #N/A | -0.025454 | 21.3067 | -2.1468 | 36.0708 | #N/A | -2.1468 | 36.0708 | 1.9745 | 30.6238 | #N/A | 1.9745 | 30.6238
| 5 |  |  | 0.92872 | 17.9975 | #N/A | 0.92872 | 17.9975 | -1.5963 | 9.9454 | #N/A | -1.5963 | 9.9454 | 3.4037 | 36.4817 | #N/A | 3.4037 | 36.4817 | 23.4037 | 30.5763 | #N/A | 23.4037 | 30.5763 | -5.5963 | 74.2355 | #N/A | -5.5963 | 74.2355 | -5.5963 | 22.7796 | #N/A | -5.5963 | 22.7796 | 10.4037 | 54.9188 | #N/A | 10.4037 | 54.9188 | -1.5963 | 15.1958 | #N/A | -1.5963 | 15.1958 | -48.5963 | 60.2902 | #N/A | -48.5963 | 60.2902 | 46.4037 | 89.7042 | #N/A | 46.4037 | 89.7042 | 30.4037 | 39.2473 | #N/A | 30.4037 | 39.2473 | 2.4037 | 30.5763 | #N/A | 2.4037 | 30.5763 |  | -5.3934 | 5.3934 |  | -1.5963 | 25.2996 | #N/A | -1.5963 | 25.2996 | -3.3676 | 36.2445 | #N/A | -3.3676 | 36.2445 | 0.40372 | 30.5763 | #N/A | 0.40372 | 30.5763
| 2 |  |  | 0.83168 | 18.8989 | #N/A | 0.83168 | 18.8989 | -1.0683 | 11.4964 | #N/A | -1.0683 | 11.4964 | 3.9317 | 36.9346 | #N/A | 3.9317 | 36.9346 | 48.9317 | 52.6514 | #N/A | 48.9317 | 52.6514 | -12.0683 | 76.4472 | #N/A | -12.0683 | 76.4472 | -5.0683 | 25.3805 | #N/A | -5.0683 | 25.3805 | 8.9317 | 83.4096 | #N/A | 8.9317 | 83.4096 | 0.93168 | 16.2532 | #N/A | 0.93168 | 16.2532 | -134.0683 | 100.657 | #N/A | -134.0683 | 100.657 |  |  | 0 | #N/A | #N/A | 63.9317 | 57.9489 | #N/A | 63.9317 | 57.9489 | 0.93168 | 31.1154 | #N/A | 0.93168 | 31.1154 |  | -5.642 | 5.642 |  | -1.0683 | 27.8519 | #N/A | -1.0683 | 27.8519 | -2.6486 | 37.7759 | #N/A | -2.6486 | 37.7759 | -0.068317 | 36.9346 | #N/A | -0.068317 | 36.9346
| 1 |  |  | -0.12577 | 18.8599 | #N/A | -0.12577 | 18.8599 | -0.84244 | 11.4322 | #N/A | -0.84244 | 11.4322 | 5.1576 | 36.9147 | #N/A | 5.1576 | 36.9147 | 82.1576 | 70.9458 | #N/A | 82.1576 | 70.9458 | -24.8424 | 76.4375 | #N/A | -24.8424 | 76.4375 | 3.1576 | 23.4669 | #N/A | 3.1576 | 23.4669 | 10.1576 | 102.7251 | #N/A | 10.1576 | 102.7251 | 2.1576 | 16.2079 | #N/A | 2.1576 | 16.2079 | -272.8424 | 200.333 | #N/A | -272.8424 | 200.333 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | -3.8424 | 31.0917 | #N/A | -3.8424 | 31.0917 |  | -5.771 | 5.771 |  | 0.15756 | 31.8643 | #N/A | 0.15756 | 31.8643 |  |  | 0 | #N/A | #N/A | -0.84244 | 42.7866 | #N/A | -0.84244 | 42.7866
10/5 | 120 |  |  |  |  | 0 | #N/A | #N/A | -0.85164 | 8.5994 | #N/A | -0.85164 | 8.5994 | 0.14836 | 22.5599 | #N/A | 0.14836 | 22.5599 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 3.1484 | 37.3356 | #N/A | 3.1484 | 37.3356 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  | -6.2491 | 6.2491 |  | 6.1484 | 21.0026 | #N/A | 6.1484 | 21.0026 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A
| 100 |  |  | 0.90801 | 17.9268 | #N/A | 0.90801 | 17.9268 | -1.342 | 9.0205 | #N/A | -1.342 | 9.0205 | -0.34199 | 22.7238 | #N/A | -0.34199 | 22.7238 | -1.342 | 30.5347 | #N/A | -1.342 | 30.5347 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 3.658 | 37.4349 | #N/A | 3.658 | 37.4349 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 13.658 | 39.2149 | #N/A | 13.658 | 39.2149 |  |  | 0 | #N/A | #N/A |  | -5.6241 | 5.6241 |  | 3.658 | 21.1785 | #N/A | 3.658 | 21.1785 |  |  | 0 | #N/A | #N/A | 7.658 | 36.4468 | #N/A | 7.658 | 36.4468
| 50 |  |  | 0.62389 | 18.1178 | #N/A | 0.62389 | 18.1178 | -2.0761 | 9.3944 | #N/A | -2.0761 | 9.3944 | -2.0761 | 22.8748 | #N/A | -2.0761 | 22.8748 | -0.076107 | 30.6473 | #N/A | -0.076107 | 30.6473 | -2.0761 | 53.369 | #N/A | -2.0761 | 53.369 | -0.076107 | 26.7442 | #N/A | -0.076107 | 26.7442 | 12.9239 | 37.5267 | #N/A | 12.9239 | 37.5267 | 2.9239 | 15.338 | #N/A | 2.9239 | 15.338 |  |  | 0 | #N/A | #N/A | 34.2572 | 86.7574 | #N/A | 34.2572 | 86.7574 | 7.9239 | 39.3026 | #N/A | 7.9239 | 39.3026 | 7.9239 | 30.6473 | #N/A | 7.9239 | 30.6473 |  | -4.9745 | 4.9745 |  | 0.92389 | 21.1493 | #N/A | 0.92389 | 21.1493 | -6.4062 | 42.4838 | #N/A | -6.4062 | 42.4838 | 4.9239 | 36.5411 | #N/A | 4.9239 | 36.5411
| 20 |  |  | 1.0857 | 18.1379 | #N/A | 1.0857 | 18.1379 | -1.8143 | 9.4331 | #N/A | -1.8143 | 9.4331 | 3.1857 | 22.8907 | #N/A | 3.1857 | 22.8907 | -5.8143 | 30.6591 | #N/A | -5.8143 | 30.6591 | 5.1857 | 53.3759 | #N/A | 5.1857 | 53.3759 | 0.18567 | 22.8907 | #N/A | 0.18567 | 22.8907 | 9.1857 | 37.5364 | #N/A | 9.1857 | 37.5364 | 3.1857 | 15.3617 | #N/A | 3.1857 | 15.3617 | -14.8143 | 40.4967 | #N/A | -14.8143 | 40.4967 | 66.8523 | 84.9766 | #N/A | 66.8523 | 84.9766 | 10.1857 | 39.3119 | #N/A | 10.1857 | 39.3119 | 3.4857 | 30.6591 | #N/A | 3.4857 | 30.6591 |  | -4.9007 | 4.9007 |  | 2.1857 | 21.1666 | #N/A | 2.1857 | 21.1666 | -11.6319 | 42.1739 | #N/A | -11.6319 | 42.1739 | -1.8143 | 36.5511 | #N/A | -1.8143 | 36.5511
| 10 |  |  | 1.0986 | 18.1254 | #N/A | 1.0986 | 18.1254 | -1.7014 | 9.409 | #N/A | -1.7014 | 9.409 | 6.2986 | 24.8099 | #N/A | 6.2986 | 24.8099 | 5.2986 | 30.6517 | #N/A | 5.2986 | 30.6517 | -1.7014 | 52.3787 | #N/A | -1.7014 | 52.3787 | -2.7014 | 22.8808 | #N/A | -2.7014 | 22.8808 | 5.2986 | 37.5304 | #N/A | 5.2986 | 37.5304 | 5.2986 | 15.347 | #N/A | 5.2986 | 15.347 | -29.7014 | 44.4469 | #N/A | -29.7014 | 44.4469 | 89.1319 | 83.9232 | #N/A | 89.1319 | 83.9232 | 8.2986 | 39.3061 | #N/A | 8.2986 | 39.3061 | 3.4986 | 30.6517 | #N/A | 3.4986 | 30.6517 |  | -4.9468 | 4.9468 |  | 1.2986 | 21.1558 | #N/A | 1.2986 | 21.1558 | -11.0275 | 42.2936 | #N/A | -11.0275 | 42.2936 | -2.7014 | 42.468 | #N/A | -2.7014 | 42.468
| 5 |  |  | 0.84505 | 17.2637 | #N/A | 0.84505 | 17.2637 | -1.955 | 7.6181 | #N/A | -1.955 | 7.6181 | 7.045 | 38.1187 | #N/A | 7.045 | 38.1187 | 19.045 | 30.1502 | #N/A | 19.045 | 30.1502 | -2.955 | 80.0565 | #N/A | -2.955 | 80.0565 | 0.045045 | 22.2044 | #N/A | 0.045045 | 22.2044 | 11.045 | 54.6827 | #N/A | 11.045 | 54.6827 | 4.045 | 14.3191 | #N/A | 4.045 | 14.3191 | -49.955 | 60.0753 | #N/A | -49.955 | 60.0753 | 129.3784 | 103.5789 | #N/A | 129.3784 | 103.5789 | 11.045 | 38.9163 | #N/A | 11.045 | 38.9163 | 4.145 | 74.061 | #N/A | 4.145 | 74.061 |  | -5.1927 | 5.1927 |  | 0.045045 | 24.5845 | #N/A | 0.045045 | 24.5845 | -14.241 | 44.1088 | #N/A | -14.241 | 44.1088 | 0.045045 | 48.094 | #N/A | 0.045045 | 48.094
| 2 |  |  | 1.0326 | 17.2065 | #N/A | 1.0326 | 17.2065 | -2.5674 | 7.4877 | #N/A | -2.5674 | 7.4877 | 7.4326 | 38.0929 | #N/A | 7.4326 | 38.0929 | 46.4326 | 52.0679 | #N/A | 46.4326 | 52.0679 | -5.5674 | 98.036 | #N/A | -5.5674 | 98.036 | -1.5674 | 22.16 | #N/A | -1.5674 | 22.16 | 9.4326 | 83.0426 | #N/A | 9.4326 | 83.0426 | 4.4326 | 14.2501 | #N/A | 4.4326 | 14.2501 | -144.5674 | 100.3241 | #N/A | -144.5674 | 100.3241 |  |  | 0 | #N/A | #N/A | 26.4326 | 56.8622 | #N/A | 26.4326 | 56.8622 | 4.7326 | 74.0477 | #N/A | 4.7326 | 74.0477 |  | -5.3791 | 5.3791 |  | 0.43265 | 26.5335 | #N/A | 0.43265 | 26.5335 | -11.7287 | 59.2988 | #N/A | -11.7287 | 59.2988 |  |  | 0 | #N/A | #N/A
| 1 |  |  | 0.18043 | 17.1387 | #N/A | 0.18043 | 17.1387 | -1.4696 | 7.3303 | #N/A | -1.4696 | 7.3303 | 9.5304 | 38.0622 | #N/A | 9.5304 | 38.0622 | 84.5304 | 70.4788 | #N/A | 84.5304 | 70.4788 | -11.4696 | 98.0241 | #N/A | -11.4696 | 98.0241 | -0.46957 | 22.1073 | #N/A | -0.46957 | 22.1073 | 12.5304 | 102.4231 | #N/A | 12.5304 | 102.4231 | 5.5304 | 16.1473 | #N/A | 5.5304 | 16.1473 | -287.4696 | 200.1681 | #N/A | -287.4696 | 200.1681 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 2.7304 | 74.032 | #N/A | 2.7304 | 74.032 |  | -5.5916 | 5.5916 |  | 1.5304 | 30.6773 | #N/A | 1.5304 | 30.6773 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A
15 VA | I1 (%IN) |  |  | CMI | CMI |  |  |  | PTB | PTB |  |  |  | METAS | METAS |  |  |  | NPL | NPL |  |  |  | LCOE | LCOE |  |  |  | MIKES | MIKES |  |  |  | BEV | BEV |  |  |  | VSL | VSL |  |  |  | SP | SP |  |  |  | DMDM | DMDM |  |  |  | GUM | GUM |  |  |  | BIM | BIM |  |  |  |  | -U(erC) (ppm) | U(erC) (ppm) |  | UME | UME |  |  |  | LNE | LNE |  |  |  | INRIM | INRIM
kI (kA/A) |  |  |  | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm) | [7 empty] | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm) |  |  |  | DeC (ppm) | U(DestdC) (ppm)
4/5 | 120 |  |  | -2.3573 | 18.98 | #N/A | -2.3573 | 18.98 | -0.90733 | 11.6293 | #N/A | -0.90733 | 11.6293 | 8.0927 | 23.5636 | #N/A | 8.0927 | 23.5636 | -0.90733 | 31.1647 | #N/A | -0.90733 | 31.1647 | 0.092674 | 42.8397 | #N/A | 0.092674 | 42.8397 | -8.9073 | 25.4409 | #N/A | -8.9073 | 25.4409 | 3.0927 | 37.9505 | #N/A | 3.0927 | 37.9505 | -5.9073 | 19.8807 | #N/A | -5.9073 | 19.8807 |  |  | 0 | #N/A | #N/A | 8.2593 | 72.6823 | #N/A | 8.2593 | 72.6823 | 9.0927 | 39.7074 | #N/A | 9.0927 | 39.7074 | 6.0927 | 23.5636 | #N/A | 6.0927 | 23.5636 |  | -5.3627 | 5.3627 |  | 6.0927 | 22.2621 | #N/A | 6.0927 | 22.2621 | 43.2084 | 37.4204 | #N/A | 43.2084 | 37.4204 | 3.0927 | 31.1647 | #N/A | 3.0927 | 31.1647
| 100 |  |  | -0.62206 | 20.1404 | #N/A | -0.62206 | 20.1404 | -0.32206 | 13.4401 | #N/A | -0.32206 | 13.4401 | 7.6779 | 24.5079 | #N/A | 7.6779 | 24.5079 | -5.3221 | 31.8847 | #N/A | -5.3221 | 31.8847 | 1.6779 | 43.3663 | #N/A | 1.6779 | 43.3663 | -7.3221 | 24.5079 | #N/A | -7.3221 | 24.5079 | 3.6779 | 38.544 | #N/A | 3.6779 | 38.544 | -6.3221 | 19.3038 | #N/A | -6.3221 | 19.3038 | -8.3221 | 41.4323 | #N/A | -8.3221 | 41.4323 | 9.1779 | 73.2201 | #N/A | 9.1779 | 73.2201 | 9.6779 | 40.275 | #N/A | 9.6779 | 40.275 | 5.6779 | 24.5079 | #N/A | 5.6779 | 24.5079 |  | -5.231 | 5.231 |  | 6.6779 | 23.2593 | #N/A | 6.6779 | 23.2593 | 43.492 | 37.4955 | #N/A | 43.492 | 37.4955 | 2.6779 | 31.8847 | #N/A | 2.6779 | 31.8847
| 50 |  |  | -0.16844 | 20.141 | #N/A | -0.16844 | 20.141 | -0.16844 | 13.4409 | #N/A | -0.16844 | 13.4409 | 5.8316 | 24.5083 | #N/A | 5.8316 | 24.5083 | -4.1684 | 31.8851 | #N/A | -4.1684 | 31.8851 | 3.8316 | 42.3988 | #N/A | 3.8316 | 42.3988 | -6.1684 | 24.5083 | #N/A | -6.1684 | 24.5083 | 2.8316 | 38.5442 | #N/A | 2.8316 | 38.5442 | -4.1684 | 19.3044 | #N/A | -4.1684 | 19.3044 | -4.1684 | 41.4326 | #N/A | -4.1684 | 41.4326 | 10.8316 | 73.0118 | #N/A | 10.8316 | 73.0118 | 7.8316 | 40.2753 | #N/A | 7.8316 | 40.2753 | 2.8316 | 24.5083 | #N/A | 2.8316 | 24.5083 |  | -5.2289 | 5.2289 |  | 3.8316 | 23.2598 | #N/A | 3.8316 | 23.2598 | 40.4924 | 37.3717 | #N/A | 40.4924 | 37.3717 | 0.83156 | 31.8851 | #N/A | 0.83156 | 31.8851
| 20 |  |  | -0.71622 | 19.0543 | #N/A | -0.71622 | 19.0543 | -1.8662 | 11.7501 | #N/A | -1.8662 | 11.7501 | 3.1338 | 23.6234 | #N/A | 3.1338 | 23.6234 | 1.1338 | 31.21 | #N/A | 1.1338 | 31.21 | 3.1338 | 41.8935 | #N/A | 3.1338 | 41.8935 | -5.8662 | 23.6234 | #N/A | -5.8662 | 23.6234 | 6.1338 | 37.9877 | #N/A | 6.1338 | 37.9877 | -1.8662 | 16.4337 | #N/A | -1.8662 | 16.4337 | -3.8662 | 40.9153 | #N/A | -3.8662 | 40.9153 | 11.3004 | 72.6079 | #N/A | 11.3004 | 72.6079 | 13.1338 | 39.743 | #N/A | 13.1338 | 39.743 | 2.1338 | 23.6234 | #N/A | 2.1338 | 23.6234 |  | -5.0927 | 5.0927 |  | 0.13378 | 22.3254 | #N/A | 0.13378 | 22.3254 | 31.4513 | 36.6823 | #N/A | 31.4513 | 36.6823 | 1.1338 | 31.21 | #N/A | 1.1338 | 31.21
| 10 |  |  | 0.26268 | 19.0465 | #N/A | 0.26268 | 19.0465 | -1.4623 | 11.7375 | #N/A | -1.4623 | 11.7375 | 3.5377 | 25.4906 | #N/A | 3.5377 | 25.4906 | 12.5377 | 31.2053 | #N/A | 12.5377 | 31.2053 | -1.4623 | 74.4968 | #N/A | -1.4623 | 74.4968 | -8.4623 | 21.7663 | #N/A | -8.4623 | 21.7663 | 2.5377 | 37.9838 | #N/A | 2.5377 | 37.9838 | -1.4623 | 16.4247 | #N/A | -1.4623 | 16.4247 | -21.4623 | 44.8305 | #N/A | -21.4623 | 44.8305 | 19.371 | 72.6615 | #N/A | 19.371 | 72.6615 | 18.5377 | 39.7393 | #N/A | 18.5377 | 39.7393 | 2.5377 | 23.6171 | #N/A | 2.5377 | 23.6171 |  | -5.1216 | 5.1216 |  | -1.4623 | 22.3188 | #N/A | -1.4623 | 22.3188 | 26.5603 | 37.0544 | #N/A | 26.5603 | 37.0544 | 1.5377 | 31.2053 | #N/A | 1.5377 | 31.2053
| 5 |  |  | -0.0050327 | 18.969 | #N/A | -0.0050327 | 18.969 | -2.305 | 11.6114 | #N/A | -2.305 | 11.6114 | 1.695 | 36.9706 | #N/A | 1.695 | 36.9706 | 24.695 | 31.1581 | #N/A | 24.695 | 31.1581 | -5.305 | 74.477 | #N/A | -5.305 | 74.477 | -9.305 | 23.5547 | #N/A | -9.305 | 23.5547 | 8.695 | 55.2448 | #N/A | 8.695 | 55.2448 | -0.30503 | 16.3348 | #N/A | -0.30503 | 16.3348 | -46.305 | 60.5873 | #N/A | -46.305 | 60.5873 | 47.195 | 89.9772 | #N/A | 47.195 | 89.9772 | 30.695 | 39.7022 | #N/A | 30.695 | 39.7022 | 1.695 | 31.1581 | #N/A | 1.695 | 31.1581 |  | -5.4015 | 5.4015 |  | -3.305 | 25.9997 | #N/A | -3.305 | 25.9997 | 24.4763 | 39.2767 | #N/A | 24.4763 | 39.2767 | -0.30503 | 31.1581 | #N/A | -0.30503 | 31.1581
| 2 |  |  | -0.47656 | 20.0373 | #N/A | -0.47656 | 20.0373 | -1.2766 | 13.285 | #N/A | -1.2766 | 13.285 | 3.7234 | 37.5299 | #N/A | 3.7234 | 37.5299 | 51.7234 | 53.0706 | #N/A | 51.7234 | 53.0706 | -11.2766 | 76.7365 | #N/A | -11.2766 | 76.7365 | -8.2766 | 24.4232 | #N/A | -8.2766 | 24.4232 | 7.7234 | 83.6749 | #N/A | 7.7234 | 83.6749 | 1.7234 | 17.5639 | #N/A | 1.7234 | 17.5639 | -132.2766 | 100.8737 | #N/A | -132.2766 | 100.8737 |  |  | 0 | #N/A | #N/A | 71.7234 | 58.3245 | #N/A | 71.7234 | 58.3245 | 0.72344 | 31.8197 | #N/A | 0.72344 | 31.8197 |  | -5.6132 | 5.6132 |  | -2.2766 | 28.6365 | #N/A | -2.2766 | 28.6365 | 15.5891 | 38.5341 | #N/A | 15.5891 | 38.5341 | -0.27656 | 37.5299 | #N/A | -0.27656 | 37.5299
| 1 |  |  | 0.058119 | 22.6207 | #N/A | 0.058119 | 22.6207 | -0.54188 | 16.9321 | #N/A | -0.54188 | 16.9321 | 5.4581 | 38.9704 | #N/A | 5.4581 | 38.9704 | 84.4581 | 72.0368 | #N/A | 84.4581 | 72.0368 | -23.5419 | 77.4512 | #N/A | -23.5419 | 77.4512 | 0.45812 | 26.5837 | #N/A | 0.45812 | 26.5837 | 10.4581 | 103.4817 | #N/A | 10.4581 | 103.4817 | 2.4581 | 20.462 | #N/A | 2.4581 | 20.462 | -265.5419 | 200.722 | #N/A | -265.5419 | 200.722 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | -3.5419 | 33.5066 | #N/A | -3.5419 | 33.5066 |  | -5.771 | 5.771 |  | -2.5419 | 34.2248 | #N/A | -2.5419 | 34.2248 |  |  | 0 | #N/A | #N/A | 0.45812 | 44.5724 | #N/A | 0.45812 | 44.5724
5/5 | 120 |  |  | -2.0623 | 19.875 | #N/A | -2.0623 | 19.875 | -1.8873 | 11.6196 | #N/A | -1.8873 | 11.6196 | 7.1127 | 23.5588 | #N/A | 7.1127 | 23.5588 | 3.1127 | 31.1611 | #N/A | 3.1127 | 31.1611 | -0.88728 | 42.8371 | #N/A | -0.88728 | 42.8371 | -7.8873 | 29.2406 | #N/A | -7.8873 | 29.2406 | 3.1127 | 37.9475 | #N/A | 3.1127 | 37.9475 | -2.8873 | 16.3406 | #N/A | -2.8873 | 16.3406 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 12.1127 | 39.7046 | #N/A | 12.1127 | 39.7046 |  |  | 0 | #N/A | #N/A |  | -5.3837 | 5.3837 |  | 4.1127 | 22.257 | #N/A | 4.1127 | 22.257 | 21.4743 | 36.224 | #N/A | 21.4743 | 36.224 | 5.1127 | 31.1611 | #N/A | 5.1127 | 31.1611
| 100 |  |  | -2.184 | 19.0477 | #N/A | -2.184 | 19.0477 | -1.534 | 11.7394 | #N/A | -1.534 | 11.7394 | 6.466 | 23.6181 | #N/A | 6.466 | 23.6181 | -2.534 | 31.206 | #N/A | -2.534 | 31.206 | 0.46602 | 42.8697 | #N/A | 0.46602 | 42.8697 | -6.534 | 25.4915 | #N/A | -6.534 | 25.4915 | 3.466 | 37.9844 | #N/A | 3.466 | 37.9844 | -3.534 | 16.426 | #N/A | -3.534 | 16.426 | -2.534 | 40.9123 | #N/A | -2.534 | 40.9123 | -20.8673 | 82.9451 | #N/A | -20.8673 | 82.9451 | 8.466 | 39.7398 | #N/A | 8.466 | 39.7398 | 7.466 | 23.6181 | #N/A | 7.466 | 23.6181 |  | -5.1172 | 5.1172 |  | 4.466 | 22.3198 | #N/A | 4.466 | 22.3198 | 19.0846 | 35.9365 | #N/A | 19.0846 | 35.9365 | 4.466 | 31.206 | #N/A | 4.466 | 31.206
| 50 |  |  | -0.58329 | 19.0567 | #N/A | -0.58329 | 19.0567 | -1.2583 | 11.7541 | #N/A | -1.2583 | 11.7541 | 4.7417 | 23.6254 | #N/A | 4.7417 | 23.6254 | -6.2583 | 31.2115 | #N/A | -6.2583 | 31.2115 | 4.7417 | 41.8946 | #N/A | 4.7417 | 41.8946 | -5.2583 | 25.4982 | #N/A | -5.2583 | 25.4982 | 3.7417 | 37.9889 | #N/A | 3.7417 | 37.9889 | -2.2583 | 16.4365 | #N/A | -2.2583 | 16.4365 | -5.2583 | 40.9165 | #N/A | -5.2583 | 40.9165 | -19.6583 | 82.9381 | #N/A | -19.6583 | 82.9381 | 8.7417 | 39.7442 | #N/A | 8.7417 | 39.7442 | 2.7417 | 21.7752 | #N/A | 2.7417 | 21.7752 |  | -5.0833 | 5.0833 |  | 3.7417 | 22.1432 | #N/A | 3.7417 | 22.1432 | 18.7223 | 35.9043 | #N/A | 18.7223 | 35.9043 | 2.7417 | 31.2115 | #N/A | 2.7417 | 31.2115
| 20 |  |  | -0.0087133 | 19.0622 | #N/A | -0.0087133 | 19.0622 | -0.92121 | 11.763 | #N/A | -0.92121 | 11.763 | 3.0788 | 23.6298 | #N/A | 3.0788 | 23.6298 | -1.9212 | 31.2149 | #N/A | -1.9212 | 31.2149 | 2.0788 | 41.8971 | #N/A | 2.0788 | 41.8971 | -4.9212 | 23.6298 | #N/A | -4.9212 | 23.6298 | 8.0788 | 37.9917 | #N/A | 8.0788 | 37.9917 | 0.078787 | 16.4429 | #N/A | 0.078787 | 16.4429 | -10.9212 | 40.919 | #N/A | -10.9212 | 40.919 | -13.5879 | 82.8828 | #N/A | -13.5879 | 82.8828 | 14.0788 | 39.7468 | #N/A | 14.0788 | 39.7468 | 0.078787 | 21.78 | #N/A | 0.078787 | 21.78 |  | -5.0628 | 5.0628 |  | 0.078787 | 22.1479 | #N/A | 0.078787 | 22.1479 | 13.2647 | 36.2275 | #N/A | 13.2647 | 36.2275 | 1.0788 | 31.2149 | #N/A | 1.0788 | 31.2149
| 10 |  |  | 0.57227 | 19.0475 | #N/A | 0.57227 | 19.0475 | -1.5277 | 11.7392 | #N/A | -1.5277 | 11.7392 | 2.4723 | 25.4913 | #N/A | 2.4723 | 25.4913 | 9.4723 | 31.2059 | #N/A | 9.4723 | 31.2059 | 0.47227 | 74.497 | #N/A | 0.47227 | 74.497 | -6.5277 | 23.618 | #N/A | -6.5277 | 23.618 | 5.4723 | 37.9843 | #N/A | 5.4723 | 37.9843 | 0.47227 | 16.4258 | #N/A | 0.47227 | 16.4258 | -25.5277 | 44.8309 | #N/A | -25.5277 | 44.8309 | 4.1389 | 82.8531 | #N/A | 4.1389 | 82.8531 | 21.4723 | 39.7398 | #N/A | 21.4723 | 39.7398 | 0.47227 | 21.7671 | #N/A | 0.47227 | 21.7671 |  | -5.1178 | 5.1178 |  | -0.52773 | 22.1352 | #N/A | -0.52773 | 22.1352 | 13.1443 | 35.8778 | #N/A | 13.1443 | 35.8778 | 1.4723 | 31.2059 | #N/A | 1.4723 | 31.2059
| 5 |  |  | -0.25476 | 20.0909 | #N/A | -0.25476 | 20.0909 | -2.5173 | 13.3658 | #N/A | -2.5173 | 13.3658 | 2.4827 | 37.5586 | #N/A | 2.4827 | 37.5586 | 21.4827 | 31.8535 | #N/A | 21.4827 | 31.8535 | -7.5173 | 74.7706 | #N/A | -7.5173 | 74.7706 | -7.5173 | 24.4672 | #N/A | -7.5173 | 24.4672 | 10.4827 | 55.64 | #N/A | 10.4827 | 55.64 | 2.4827 | 17.6251 | #N/A | 2.4827 | 17.6251 | -57.5173 | 60.9479 | #N/A | -57.5173 | 60.9479 | 49.4827 | 102.9403 | #N/A | 49.4827 | 102.9403 | 34.4827 | 40.2503 | #N/A | 34.4827 | 40.2503 | 0.48274 | 31.8535 | #N/A | 0.48274 | 31.8535 |  | -5.418 | 5.418 |  | -1.5173 | 26.8292 | #N/A | -1.5173 | 26.8292 | 9.8672 | 45.9805 | #N/A | 9.8672 | 45.9805 | 0.48274 | 31.8535 | #N/A | 0.48274 | 31.8535
| 2 |  |  | -0.77974 | 20.0432 | #N/A | -0.77974 | 20.0432 | -1.5235 | 13.294 | #N/A | -1.5235 | 13.294 | 1.4765 | 37.533 | #N/A | 1.4765 | 37.533 | 48.4765 | 53.0729 | #N/A | 48.4765 | 53.0729 | -13.5235 | 76.7381 | #N/A | -13.5235 | 76.7381 | -8.5235 | 24.428 | #N/A | -8.5235 | 24.428 | 9.4765 | 83.6763 | #N/A | 9.4765 | 83.6763 | 1.4765 | 17.5707 | #N/A | 1.4765 | 17.5707 | -164.5235 | 100.8725 | #N/A | -164.5235 | 100.8725 |  |  | 0 | #N/A | #N/A | 53.4765 | 57.7838 | #N/A | 53.4765 | 57.7838 | -0.52349 | 31.8234 | #N/A | -0.52349 | 31.8234 |  | -5.592 | 5.592 |  | -1.5235 | 28.6407 | #N/A | -1.5235 | 28.6407 | 7.6434 | 40.3575 | #N/A | 7.6434 | 40.3575 | -0.52349 | 37.533 | #N/A | -0.52349 | 37.533
| 1 |  |  | 0.61987 | 19.9323 | #N/A | 0.61987 | 19.9323 | -0.14888 | 13.1263 | #N/A | -0.14888 | 13.1263 | 1.8511 | 37.474 | #N/A | 1.8511 | 37.474 | 82.8511 | 71.272 | #N/A | 82.8511 | 71.272 | -26.1489 | 76.7092 | #N/A | -26.1489 | 76.7092 | -3.1489 | 24.3372 | #N/A | -3.1489 | 24.3372 | 11.8511 | 102.9274 | #N/A | 11.8511 | 102.9274 | 3.8511 | 20.7918 | #N/A | 3.8511 | 20.7918 | -312.1489 | 200.4488 | #N/A | -312.1489 | 200.4488 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | -4.1489 | 31.7537 | #N/A | -4.1489 | 31.7537 |  | -5.9751 | 5.9751 |  | 0.85112 | 32.5106 | #N/A | 0.85112 | 32.5106 |  |  | 0 | #N/A | #N/A | 0.85112 | 43.2701 | #N/A | 0.85112 | 43.2701
6/5 | 120 |  |  | 1.5447 | 19.0644 | #N/A | 1.5447 | 19.0644 | -1.4928 | 11.1108 | #N/A | -1.4928 | 11.1108 | -1.4928 | 23.6315 | #N/A | -1.4928 | 23.6315 | 6.5072 | 31.2162 | #N/A | 6.5072 | 31.2162 | -4.4928 | 52.711 | #N/A | -4.4928 | 52.711 | -6.4928 | 33.1429 | #N/A | -6.4928 | 33.1429 | 3.5072 | 37.9928 | #N/A | 3.5072 | 37.9928 | -2.4928 | 16.4453 | #N/A | -2.4928 | 16.4453 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 36.5072 | 39.7478 | #N/A | 36.5072 | 39.7478 |  |  | 0 | #N/A | #N/A |  | -5.0547 | 5.0547 |  | 3.5072 | 22.1497 | #N/A | 3.5072 | 22.1497 | 12.543 | 37.8713 | #N/A | 12.543 | 37.8713 | -2.4928 | 31.2162 | #N/A | -2.4928 | 31.2162
| 100 |  |  | 1.7237 | 19.072 | #N/A | 1.7237 | 19.072 | -1.2138 | 11.1239 | #N/A | -1.2138 | 11.1239 | -1.2138 | 23.6377 | #N/A | -1.2138 | 23.6377 | 2.7862 | 31.2209 | #N/A | 2.7862 | 31.2209 | -0.21382 | 52.7138 | #N/A | -0.21382 | 52.7138 | -4.2138 | 27.3997 | #N/A | -4.2138 | 27.3997 | 3.7862 | 37.9966 | #N/A | 3.7862 | 37.9966 | -2.2138 | 16.4542 | #N/A | -2.2138 | 16.4542 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 33.7862 | 39.7515 | #N/A | 33.7862 | 39.7515 |  |  | 0 | #N/A | #N/A |  | -5.0258 | 5.0258 |  | 2.7862 | 22.1563 | #N/A | 2.7862 | 22.1563 | 10.9336 | 39.1609 | #N/A | 10.9336 | 39.1609 | -3.2138 | 31.2209 | #N/A | -3.2138 | 31.2209
| 50 |  |  | 2.1506 | 19.1055 | #N/A | 2.1506 | 19.1055 | -0.67442 | 11.1813 | #N/A | -0.67442 | 11.1813 | 1.3256 | 23.6648 | #N/A | 1.3256 | 23.6648 | -7.6744 | 31.2413 | #N/A | -7.6744 | 31.2413 | -1.6744 | 55.6868 | #N/A | -1.6744 | 55.6868 | -2.6744 | 25.5347 | #N/A | -2.6744 | 25.5347 | 3.3256 | 38.0134 | #N/A | 3.3256 | 38.0134 | -1.6744 | 16.4931 | #N/A | -1.6744 | 16.4931 | -4.6744 | 40.9392 | #N/A | -4.6744 | 40.9392 | 41.8256 | 86.7577 | #N/A | 41.8256 | 86.7577 | 33.3256 | 39.7676 | #N/A | 33.3256 | 39.7676 | -2.6744 | 31.2413 | #N/A | -2.6744 | 31.2413 |  | -4.8968 | 4.8968 |  | 1.3256 | 22.1852 | #N/A | 1.3256 | 22.1852 | 6.5884 | 38.5072 | #N/A | 6.5884 | 38.5072 | -5.6744 | 31.2413 | #N/A | -5.6744 | 31.2413
| 20 |  |  | 2.8561 | 19.1115 | #N/A | 2.8561 | 19.1115 | -1.7564 | 11.1915 | #N/A | -1.7564 | 11.1915 | 6.2436 | 23.6696 | #N/A | 6.2436 | 23.6696 | -2.7564 | 31.245 | #N/A | -2.7564 | 31.245 | -3.7564 | 52.7281 | #N/A | -3.7564 | 52.7281 | -4.7564 | 23.6696 | #N/A | -4.7564 | 23.6696 | 9.2436 | 38.0164 | #N/A | 9.2436 | 38.0164 | 1.2436 | 16.5 | #N/A | 1.2436 | 16.5 | -12.7564 | 40.942 | #N/A | -12.7564 | 40.942 | 36.5769 | 84.6338 | #N/A | 36.5769 | 84.6338 | 35.2436 | 39.7705 | #N/A | 35.2436 | 39.7705 | -3.7564 | 31.245 | #N/A | -3.7564 | 31.245 |  | -4.8735 | 4.8735 |  | 1.2436 | 22.1903 | #N/A | 1.2436 | 22.1903 | 4.6963 | 37.9125 | #N/A | 4.6963 | 37.9125 | -7.7564 | 31.245 | #N/A | -7.7564 | 31.245
| 10 |  |  | 3.6673 | 20.2194 | #N/A | 3.6673 | 20.2194 | -1.0327 | 12.9932 | #N/A | -1.0327 | 12.9932 | 6.9673 | 26.3784 | #N/A | 6.9673 | 26.3784 | 7.9673 | 31.9347 | #N/A | 7.9673 | 31.9347 | -1.0327 | 53.1397 | #N/A | -1.0327 | 53.1397 | -5.0327 | 22.7996 | #N/A | -5.0327 | 22.7996 | 5.9673 | 38.5853 | #N/A | 5.9673 | 38.5853 | 1.9673 | 17.7714 | #N/A | 1.9673 | 17.7714 | -26.0327 | 45.3412 | #N/A | -26.0327 | 45.3412 | 70.8006 | 84.546 | #N/A | 70.8006 | 84.546 | 42.9673 | 40.9099 | #N/A | 42.9673 | 40.9099 | -1.0327 | 31.9347 | #N/A | -1.0327 | 31.9347 |  | -4.9171 | 4.9171 |  | -1.0327 | 23.1513 | #N/A | -1.0327 | 23.1513 | 2.931 | 40.4638 | #N/A | 2.931 | 40.4638 | -8.0327 | 31.9347 | #N/A | -8.0327 | 31.9347
| 5 |  |  | 3.6563 | 20.1478 | #N/A | 3.6563 | 20.1478 | -1.3437 | 12.8815 | #N/A | -1.3437 | 12.8815 | 8.6563 | 39.5086 | #N/A | 8.6563 | 39.5086 | 21.6563 | 31.8894 | #N/A | 21.6563 | 31.8894 | -7.3437 | 80.7275 | #N/A | -7.3437 | 80.7275 | -4.3437 | 24.5139 | #N/A | -4.3437 | 24.5139 | 2.6563 | 55.6605 | #N/A | 2.6563 | 55.6605 | 3.6563 | 17.6899 | #N/A | 3.6563 | 17.6899 | -55.3437 | 60.9667 | #N/A | -55.3437 | 60.9667 | 127.4897 | 104.1008 | #N/A | 127.4897 | 104.1008 | 43.6563 | 40.9452 | #N/A | 43.6563 | 40.9452 | 0.65633 | 74.7859 | #N/A | 0.65633 | 74.7859 |  | -5.2026 | 5.2026 |  | -1.3437 | 26.8718 | #N/A | -1.3437 | 26.8718 | 2.3682 | 51.808 | #N/A | 2.3682 | 51.808 | -7.3437 | 31.8894 | #N/A | -7.3437 | 31.8894
| 2 |  |  | 5.8997 | 21.373 | #N/A | 5.8997 | 21.373 | -1.6503 | 14.7243 | #N/A | -1.6503 | 14.7243 | 7.3497 | 40.1473 | #N/A | 7.3497 | 40.1473 | 48.3497 | 53.5892 | #N/A | 48.3497 | 53.5892 | -16.6503 | 98.8524 | #N/A | -16.6503 | 98.8524 | -8.6503 | 25.5305 | #N/A | -8.6503 | 25.5305 | 11.3497 | 84.0048 | #N/A | 11.3497 | 84.0048 | 2.3497 | 19.0737 | #N/A | 2.3497 | 19.0737 | -159.6503 | 101.1148 | #N/A | -159.6503 | 101.1148 |  |  | 0 | #N/A | #N/A | 64.3497 | 58.7404 | #N/A | 64.3497 | 58.7404 | 2.3497 | 75.1253 | #N/A | 2.3497 | 75.1253 |  | -5.3098 | 5.3098 |  | -1.6503 | 29.4069 | #N/A | -1.6503 | 29.4069 | 5.6196 | 39.1904 | #N/A | 5.6196 | 39.1904 | -7.6503 | 38.2597 | #N/A | -7.6503 | 38.2597
| 1 |  |  | 7.4838 | 21.3393 | #N/A | 7.4838 | 21.3393 | -1.8537 | 14.6754 | #N/A | -1.8537 | 14.6754 | 8.1463 | 40.1294 | #N/A | 8.1463 | 40.1294 | 81.1463 | 71.5935 | #N/A | 81.1463 | 71.5935 | -25.8537 | 98.8452 | #N/A | -25.8537 | 98.8452 | -0.85367 | 25.5023 | #N/A | -0.85367 | 25.5023 | 12.1463 | 103.2091 | #N/A | 12.1463 | 103.2091 | 2.1463 | 19.0359 | #N/A | 2.1463 | 19.0359 | -313.8537 | 200.5633 | #N/A | -313.8537 | 200.5633 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 9.1463 | 75.1157 | #N/A | 9.1463 | 75.1157 |  | -5.4436 | 5.4436 |  | -1.8537 | 33.2073 | #N/A | -1.8537 | 33.2073 |  |  | 0 | #N/A | #N/A | -1.8537 | 43.9359 | #N/A | -1.8537 | 43.9359
8/5 | 120 |  |  | -4.6468 | 21.3004 | #N/A | -4.6468 | 21.3004 | -0.49679 | 14.6187 | #N/A | -0.49679 | 14.6187 | 2.5032 | 25.4697 | #N/A | 2.5032 | 25.4697 | -2.4968 | 32.6298 | #N/A | -2.4968 | 32.6298 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 4.5032 | 39.1626 | #N/A | 4.5032 | 39.1626 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 13.5032 | 40.8674 | #N/A | 13.5032 | 40.8674 |  |  | 0 | #N/A | #N/A |  | -5.5942 | 5.5942 |  | 5.5032 | 24.1011 | #N/A | 5.5032 | 24.1011 |  |  | 0 | #N/A | #N/A | -1.4968 | 32.6298 | #N/A | -1.4968 | 32.6298
| 100 |  |  | -6.4565 | 19.059 | #N/A | -6.4565 | 19.059 | -0.56905 | 11.1016 | #N/A | -0.56905 | 11.1016 | 2.431 | 23.6272 | #N/A | 2.431 | 23.6272 | 2.431 | 31.2129 | #N/A | 2.431 | 31.2129 | -4.569 | 52.7091 | #N/A | -4.569 | 52.7091 | -0.56905 | 35.0749 | #N/A | -0.56905 | 35.0749 | 4.431 | 37.9901 | #N/A | 4.431 | 37.9901 | 1.431 | 16.4391 | #N/A | 1.431 | 16.4391 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 10.431 | 39.7453 | #N/A | 10.431 | 39.7453 |  |  | 0 | #N/A | #N/A |  | -5.0749 | 5.0749 |  | 3.431 | 22.1451 | #N/A | 3.431 | 22.1451 | 6.2099 | 44.0085 | #N/A | 6.2099 | 44.0085 | -0.56905 | 31.2129 | #N/A | -0.56905 | 31.2129
| 50 |  |  | -3.3764 | 19.0846 | #N/A | -3.3764 | 19.0846 | 0.24861 | 11.1455 | #N/A | 0.24861 | 11.1455 | 2.2486 | 23.6479 | #N/A | 2.2486 | 23.6479 | 1.2486 | 31.2285 | #N/A | 1.2486 | 31.2285 | 2.2486 | 56.6676 | #N/A | 2.2486 | 56.6676 | -0.75139 | 25.519 | #N/A | -0.75139 | 25.519 | 4.2486 | 38.0029 | #N/A | 4.2486 | 38.0029 | 2.2486 | 18.1995 | #N/A | 2.2486 | 18.1995 | -5.7514 | 40.9295 | #N/A | -5.7514 | 40.9295 | -10.2514 | 86.707 | #N/A | -10.2514 | 86.707 | 9.2486 | 39.7575 | #N/A | 9.2486 | 39.7575 | -10.7514 | 31.2285 | #N/A | -10.7514 | 31.2285 |  | -4.9778 | 4.9778 |  | 2.2486 | 22.1671 | #N/A | 2.2486 | 22.1671 | -0.12747 | 42.7383 | #N/A | -0.12747 | 42.7383 | -4.7514 | 31.2285 | #N/A | -4.7514 | 31.2285
| 20 |  |  | -2.41 | 19.1087 | #N/A | -2.41 | 19.1087 | -0.34755 | 11.1868 | #N/A | -0.34755 | 11.1868 | 5.6525 | 23.6674 | #N/A | 5.6525 | 23.6674 | -0.34755 | 31.2433 | #N/A | -0.34755 | 31.2433 | -7.3475 | 52.7271 | #N/A | -7.3475 | 52.7271 | -4.3475 | 23.6674 | #N/A | -4.3475 | 23.6674 | 9.6525 | 38.0151 | #N/A | 9.6525 | 38.0151 | 4.6525 | 16.4968 | #N/A | 4.6525 | 16.4968 | -15.3475 | 40.9407 | #N/A | -15.3475 | 40.9407 | 52.6525 | 86.9131 | #N/A | 52.6525 | 86.9131 | 12.6525 | 39.7691 | #N/A | 12.6525 | 39.7691 | -13.3475 | 31.2433 | #N/A | -13.3475 | 31.2433 |  | -4.8842 | 4.8842 |  | 2.6525 | 22.1879 | #N/A | 2.6525 | 22.1879 | -3.5212 | 43.2884 | #N/A | -3.5212 | 43.2884 | -9.3475 | 31.2433 | #N/A | -9.3475 | 31.2433
| 10 |  |  | -1.0381 | 20.2219 | #N/A | -1.0381 | 20.2219 | -0.96311 | 12.9971 | #N/A | -0.96311 | 12.9971 | 8.0369 | 26.3804 | #N/A | 8.0369 | 26.3804 | 10.0369 | 31.9363 | #N/A | 10.0369 | 31.9363 | -8.9631 | 53.1406 | #N/A | -8.9631 | 53.1406 | -3.9631 | 24.5749 | #N/A | -3.9631 | 24.5749 | 6.0369 | 38.5866 | #N/A | 6.0369 | 38.5866 | 4.0369 | 17.7743 | #N/A | 4.0369 | 17.7743 | -30.9631 | 45.3423 | #N/A | -30.9631 | 45.3423 | 84.2036 | 86.6624 | #N/A | 84.2036 | 86.6624 | 17.0369 | 40.3158 | #N/A | 17.0369 | 40.3158 | -10.9631 | 31.9363 | #N/A | -10.9631 | 31.9363 |  | -4.9066 | 4.9066 |  | 1.0369 | 22.9775 | #N/A | 1.0369 | 22.9775 | -4.3962 | 43.5147 | #N/A | -4.3962 | 43.5147 | -8.9631 | 31.9363 | #N/A | -8.9631 | 31.9363
| 5 |  |  | -0.33642 | 20.1623 | #N/A | -0.33642 | 20.1623 | -1.0364 | 12.9042 | #N/A | -1.0364 | 12.9042 | 8.9636 | 39.5161 | #N/A | 8.9636 | 39.5161 | 23.9636 | 31.8986 | #N/A | 23.9636 | 31.8986 | -9.0364 | 80.7311 | #N/A | -9.0364 | 80.7311 | -6.0364 | 24.5259 | #N/A | -6.0364 | 24.5259 | 2.9636 | 55.6658 | #N/A | 2.9636 | 55.6658 | 3.9636 | 17.7064 | #N/A | 3.9636 | 17.7064 | -53.0364 | 60.9715 | #N/A | -53.0364 | 60.9715 | 127.7969 | 107.6135 | #N/A | 127.7969 | 107.6135 | 16.9636 | 40.286 | #N/A | 16.9636 | 40.286 | -9.0364 | 74.7898 | #N/A | -9.0364 | 74.7898 |  | -5.146 | 5.146 |  | 0.96358 | 26.6998 | #N/A | 0.96358 | 26.6998 | -6.2242 | 45.3928 | #N/A | -6.2242 | 45.3928 | -10.0364 | 31.8986 | #N/A | -10.0364 | 31.8986
| 2 |  |  | -0.88673 | 20.1227 | #N/A | -0.88673 | 20.1227 | -1.3367 | 12.8423 | #N/A | -1.3367 | 12.8423 | 8.6633 | 39.4959 | #N/A | 8.6633 | 39.4959 | 50.6633 | 53.103 | #N/A | 50.6633 | 53.103 | -18.3367 | 98.5897 | #N/A | -18.3367 | 98.5897 | -8.3367 | 24.4934 | #N/A | -8.3367 | 24.4934 | 10.6633 | 83.6954 | #N/A | 10.6633 | 83.6954 | 5.6633 | 17.6614 | #N/A | 5.6633 | 17.6614 | -151.3367 | 100.8567 | #N/A | -151.3367 | 100.8567 |  |  | 0 | #N/A | #N/A | 35.6633 | 57.8115 | #N/A | 35.6633 | 57.8115 | -6.3367 | 74.7792 | #N/A | -6.3367 | 74.7792 |  | -5.2987 | 5.2987 |  | -0.33673 | 28.5111 | #N/A | -0.33673 | 28.5111 | -9.7949 | 43.2211 | #N/A | -9.7949 | 43.2211 | -12.3367 | 37.5756 | #N/A | -12.3367 | 37.5756
| 1 |  |  | 9.858 | 21.3393 | #N/A | 9.858 | 21.3393 | -1.8295 | 14.6754 | #N/A | -1.8295 | 14.6754 | 8.1705 | 40.1294 | #N/A | 8.1705 | 40.1294 | 83.1705 | 71.5935 | #N/A | 83.1705 | 71.5935 | -18.8295 | 98.8452 | #N/A | -18.8295 | 98.8452 | -6.8295 | 25.5023 | #N/A | -6.8295 | 25.5023 | 12.1705 | 103.2091 | #N/A | 12.1705 | 103.2091 | 2.1705 | 19.0359 | #N/A | 2.1705 | 19.0359 | -308.8295 | 200.5633 | #N/A | -308.8295 | 200.5633 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 6.1705 | 75.1157 | #N/A | 6.1705 | 75.1157 |  | -5.4436 | 5.4436 |  | 0.17046 | 33.2073 | #N/A | 0.17046 | 33.2073 |  |  | 0 | #N/A | #N/A | 0.17046 | 43.9359 | #N/A | 0.17046 | 43.9359
10/5 | 120 |  |  |  |  | 0 | #N/A | #N/A | -0.80996 | 10.4857 | #N/A | -0.80996 | 10.4857 | 1.19 | 23.3441 | #N/A | 1.19 | 23.3441 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 2.19 | 37.8147 | #N/A | 2.19 | 37.8147 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  | -6.2491 | 6.2491 |  | 5.19 | 21.8428 | #N/A | 5.19 | 21.8428 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A
| 100 |  |  | 1.0526 | 18.9042 | #N/A | 1.0526 | 18.9042 | -1.3724 | 10.8337 | #N/A | -1.3724 | 10.8337 | 0.62758 | 23.5025 | #N/A | 0.62758 | 23.5025 | -4.3724 | 31.1186 | #N/A | -4.3724 | 31.1186 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 2.6276 | 37.9127 | #N/A | 2.6276 | 37.9127 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 19.6276 | 39.6713 | #N/A | 19.6276 | 39.6713 |  |  | 0 | #N/A | #N/A |  | -5.6241 | 5.6241 |  | 2.6276 | 22.012 | #N/A | 2.6276 | 22.012 |  |  | 0 | #N/A | #N/A | 8.6276 | 36.9374 | #N/A | 8.6276 | 36.9374
| 50 |  |  | 0.46058 | 19.0644 | #N/A | 0.46058 | 19.0644 | -2.2394 | 11.1109 | #N/A | -2.2394 | 11.1109 | -1.2394 | 23.6316 | #N/A | -1.2394 | 23.6316 | 0.76058 | 31.2162 | #N/A | 0.76058 | 31.2162 | -1.2394 | 53.6978 | #N/A | -1.2394 | 53.6978 | -0.23942 | 27.3944 | #N/A | -0.23942 | 27.3944 | 11.7606 | 37.9928 | #N/A | 11.7606 | 37.9928 | 2.7606 | 18.1784 | #N/A | 2.7606 | 18.1784 |  |  | 0 | #N/A | #N/A | 33.0939 | 86.5405 | #N/A | 33.0939 | 86.5405 | 13.7606 | 39.7479 | #N/A | 13.7606 | 39.7479 | 7.7606 | 31.2162 | #N/A | 7.7606 | 31.2162 |  | -5.0545 | 5.0545 |  | 0.76058 | 22.1498 | #N/A | 0.76058 | 22.1498 | -2.3196 | 43.8728 | #N/A | -2.3196 | 43.8728 | 5.7606 | 37.0196 | #N/A | 5.7606 | 37.0196
| 20 |  |  | 1.2551 | 19.1043 | #N/A | 1.2551 | 19.1043 | -1.9199 | 11.1792 | #N/A | -1.9199 | 11.1792 | 3.0801 | 23.6638 | #N/A | 3.0801 | 23.6638 | -4.9199 | 31.2406 | #N/A | -4.9199 | 31.2406 | 3.0801 | 52.7255 | #N/A | 3.0801 | 52.7255 | -2.9199 | 23.6638 | #N/A | -2.9199 | 23.6638 | 8.0801 | 38.0128 | #N/A | 8.0801 | 38.0128 | 3.0801 | 16.4917 | #N/A | 3.0801 | 16.4917 | -16.9199 | 40.9387 | #N/A | -16.9199 | 40.9387 | 67.4134 | 85.858 | #N/A | 67.4134 | 85.858 | 15.0801 | 39.767 | #N/A | 15.0801 | 39.767 | 3.0801 | 31.2406 | #N/A | 3.0801 | 31.2406 |  | -4.9015 | 4.9015 |  | 3.0801 | 22.0003 | #N/A | 3.0801 | 22.0003 | -10.1541 | 43.5282 | #N/A | -10.1541 | 43.5282 | 4.0801 | 37.0402 | #N/A | 4.0801 | 37.0402
| 10 |  |  | 1.7469 | 19.0924 | #N/A | 1.7469 | 19.0924 | -1.5031 | 11.1589 | #N/A | -1.5031 | 11.1589 | 7.4969 | 25.5249 | #N/A | 7.4969 | 25.5249 | 5.4969 | 31.2333 | #N/A | 5.4969 | 31.2333 | -2.5031 | 52.7212 | #N/A | -2.5031 | 52.7212 | -4.5031 | 23.6542 | #N/A | -4.5031 | 23.6542 | 4.4969 | 38.0069 | #N/A | 4.4969 | 38.0069 | 3.4969 | 16.4779 | #N/A | 3.4969 | 16.4779 | -30.5031 | 44.85 | #N/A | -30.5031 | 44.85 | 90.3303 | 84.7469 | #N/A | 90.3303 | 84.7469 | 12.4969 | 39.7613 | #N/A | 12.4969 | 39.7613 | 3.4969 | 31.2333 | #N/A | 3.4969 | 31.2333 |  | -4.9476 | 4.9476 |  | 0.49693 | 21.99 | #N/A | 0.49693 | 21.99 | -12.0375 | 43.201 | #N/A | -12.0375 | 43.201 | 6.4969 | 42.8896 | #N/A | 6.4969 | 42.8896
| 5 |  |  | 1.7898 | 19.027 | #N/A | 1.7898 | 19.027 | -1.9602 | 11.0465 | #N/A | -1.9602 | 11.0465 | 7.0398 | 38.949 | #N/A | 7.0398 | 38.949 | 20.0398 | 31.1934 | #N/A | 20.0398 | 31.1934 | 0.039822 | 80.4551 | #N/A | 0.039822 | 80.4551 | -4.9602 | 23.6014 | #N/A | -4.9602 | 23.6014 | 11.0398 | 55.2647 | #N/A | 11.0398 | 55.2647 | 4.0398 | 16.402 | #N/A | 4.0398 | 16.402 | -50.9602 | 60.6055 | #N/A | -50.9602 | 60.6055 | 135.5398 | 105.7399 | #N/A | 135.5398 | 105.7399 | 15.0398 | 39.7299 | #N/A | 15.0398 | 39.7299 | 4.0398 | 74.4918 | #N/A | 4.0398 | 74.4918 |  | -5.1936 | 5.1936 |  | -0.96018 | 25.8532 | #N/A | -0.96018 | 25.8532 | -9.9129 | 45.4056 | #N/A | -9.9129 | 45.4056 | 6.0398 | 48.7548 | #N/A | 6.0398 | 48.7548
| 2 |  |  | 2.0748 | 21.3602 | #N/A | 2.0748 | 21.3602 | -2.0252 | 14.7057 | #N/A | -2.0252 | 14.7057 | 6.9748 | 40.1405 | #N/A | 6.9748 | 40.1405 | 47.9748 | 53.5841 | #N/A | 47.9748 | 53.5841 | -4.0252 | 98.8497 | #N/A | -4.0252 | 98.8497 | -7.0252 | 25.5198 | #N/A | -7.0252 | 25.5198 | 7.9748 | 84.0015 | #N/A | 7.9748 | 84.0015 | 3.9748 | 19.0593 | #N/A | 3.9748 | 19.0593 | -146.0252 | 101.1175 | #N/A | -146.0252 | 101.1175 |  |  | 0 | #N/A | #N/A | 29.9748 | 58.2537 | #N/A | 29.9748 | 58.2537 | 5.9748 | 75.1216 | #N/A | 5.9748 | 75.1216 |  | -5.3611 | 5.3611 |  | -0.025226 | 29.3976 | #N/A | -0.025226 | 29.3976 | -10.1866 | 45.8311 | #N/A | -10.1866 | 45.8311 |  |  | 0 | #N/A | #N/A
| 1 |  |  | -0.30454 | 22.6905 | #N/A | -0.30454 | 22.6905 | -1.0045 | 16.5789 | #N/A | -1.0045 | 16.5789 | 8.9955 | 40.8639 | #N/A | 8.9955 | 40.8639 | 82.9955 | 72.0149 | #N/A | 82.9955 | 72.0149 | -11.0045 | 99.1457 | #N/A | -11.0045 | 99.1457 | -3.0045 | 26.6432 | #N/A | -3.0045 | 26.6432 | 10.9955 | 103.497 | #N/A | 10.9955 | 103.497 | 3.9955 | 20.5392 | #N/A | 3.9955 | 20.5392 | -290.0045 | 200.7141 | #N/A | -290.0045 | 200.7141 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 0.99546 | 75.5107 | #N/A | 0.99546 | 75.5107 |  | -5.4899 | 5.4899 |  | 0.99546 | 34.0914 | #N/A | 0.99546 | 34.0914 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A
Annex 4. TAB A4-3 auxiliary calculation for graphs

## del
Estd-Limit(d) |  | 1 | [80 empty] | last rev. 23.1.2018 by JH-CMI
5 VA |  | CMI |  |  |  |  | PTB |  |  |  |  | METAS |  |  |  |  | NPL |  |  |  |  | LCOE |  |  |  |  | MIKES |  |  |  |  | BEV |  |  |  |  | VSL |  |  |  |  | SP |  |  |  |  | DMDM |  |  |  |  | GUM |  |  |  |  | BIM |  |  |  |  | weighted mean dr (mrad) | u(dr) (mrad) | U(dr) (mrad) | u(dstd) (mrad) | UME |  |  |  |  | LNE |  |  |  |  | INRIM
kI (kA/A) | I1 (%IN) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) |  |  |  |  | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-)
4/5 | 120 | 159.1334 | 30 | 7.1284 | 31.2466 | 0.11407 | 144 | 10 | -8.0049 | 13.2796 | 0.3014 | 165.87 | 20.37 | 13.8651 | 22.165 | 0.31277 | 144 | 15 | -8.0049 | 17.3594 | 0.23056 | 224.07 | 91.665 | 72.0651 | 92.0805 | 0.39132 | 165.87 | 26.19 | 13.8651 | 27.6091 | 0.2511 | 145.5 | 20.37 | -6.5049 | 22.165 | 0.14674 | 156 | 20 | 3.9951 | 21.8254 | 0.091524 | [5 empty] | 215.1605 | 46.4616 | 63.1556 | 47.2761 | 0.66794 | 177.51 | 40.74 | 25.5051 | 41.6665 | 0.30606 | 119.892 | 13.386 | -32.1129 | 15.9854 | 1.0044 | 152.0049 | 4.8633 | 9.7265 | 10 | 177.51 | 11.64 | 25.5051 | 14.5547 | 0.87618 | 131.5819 | 18.1683 | -20.423 | 20.1603 | 0.50652 | 229 | 36 | 76.9951 | 37.0452 | 1.0392
| 100 | 168.8431 | 30 | 7.8321 | 30.7363 | 0.12741 | 151 | 10 | -10.0109 | 12.0299 | 0.41608 | 174.6 | 20.37 | 13.5891 | 21.4396 | 0.31692 | 153 | 15 | -8.0109 | 16.4231 | 0.24389 | 235.71 | 91.665 | 74.6991 | 91.9086 | 0.40638 | 177.51 | 26.19 | 16.4991 | 27.0303 | 0.3052 | 174.6 | 20.37 | 13.5891 | 21.4396 | 0.31692 | 166 | 20 | 4.9891 | 21.0884 | 0.11829 | 154.23 | 10.185 | -6.7809 | 12.1842 | 0.27827 | 228.7018 | 46.5304 | 67.6908 | 47.0085 | 0.71999 | 171.69 | 40.74 | 10.6791 | 41.2852 | 0.12933 | 127.458 | 13.386 | -33.5529 | 14.9634 | 1.1212 | 161.0109 | 4.391 | 8.782 | 8 | 192.06 | 11.64 | 31.0491 | 13.4242 | 1.1565 | 139.8533 | 18.3329 | -21.1576 | 19.5144 | 0.5421 | 245 | 36 | 83.9891 | 36.6158 | 1.1469
| 50 | 208.6034 | 30 | 10.4806 | 30.748 | 0.17043 | 190 | 10 | -8.1227 | 12.0598 | 0.33677 | 206.61 | 20.37 | 8.4873 | 21.4564 | 0.19778 | 201 | 15 | 2.8773 | 16.445 | 0.087481 | 215.34 | 23.28 | 17.2173 | 24.2363 | 0.3552 | 215.34 | 26.19 | 17.2173 | 27.0436 | 0.31832 | 203.7 | 20.37 | 5.5773 | 21.4564 | 0.12997 | 204 | 19 | 5.8773 | 20.1603 | 0.14576 | 192.06 | 10.185 | -6.0627 | 12.2137 | 0.24819 | 259.2034 | 46.5609 | 61.0807 | 47.0464 | 0.64915 | 206.61 | 40.74 | 8.4873 | 41.2939 | 0.10277 | 163.251 | 13.386 | -34.8717 | 14.9875 | 1.1634 | 198.1227 | 4.3082 | 8.6164 | 8 | 226.98 | 11.64 | 28.8573 | 13.451 | 1.0727 | 167.1198 | 18.2824 | -31.003 | 19.4855 | 0.79554 | 262 | 36 | 63.8773 | 36.6257 | 0.87203
| 20 | 248.3976 | 30 | 9.3667 | 30.1037 | 0.15557 | 232 | 10 | -7.0309 | 10.3069 | 0.34108 | 241.53 | 20.37 | 2.4991 | 20.5224 | 0.060887 | 252 | 15 | 12.9691 | 15.2063 | 0.42644 | 247.35 | 23.28 | 8.3191 | 23.4135 | 0.17766 | 256.08 | 26.19 | 17.0491 | 26.3087 | 0.32402 | 232.8 | 20.37 | -6.2309 | 20.5224 | 0.15181 | 248 | 19 | 8.9691 | 19.1633 | 0.23402 | 235.71 | 10.185 | -3.3209 | 10.4865 | 0.15834 | 292.4162 | 46.5266 | 53.3853 | 46.5936 | 0.57288 | 244.44 | 40.74 | 5.4091 | 40.8164 | 0.066261 | 213.303 | 13.386 | -25.7279 | 13.6168 | 0.94471 | 239.0309 | 4.3322 | 8.6645 | 5 | 253.17 | 11.64 | 14.1391 | 11.9047 | 0.59385 | 217.9166 | 20.4324 | -21.1144 | 20.5844 | 0.51287 | 291 | 36 | 51.9691 | 36.0864 | 0.72006
| 10 | 269.8637 | 35 | 4.4919 | 35.0522 | 0.064075 | 257 | 10 | -8.3718 | 10.1811 | 0.41115 | 264.81 | 29.1 | -0.56179 | 29.1627 | 0.0096321 | 283 | 15 | 17.6282 | 15.1213 | 0.58289 | 288.09 | 42.195 | 22.7182 | 42.2383 | 0.26893 | 276.45 | 26.19 | 11.0782 | 26.2597 | 0.21094 | 291 | 20.37 | 25.6282 | 20.4595 | 0.62632 | 271 | 19 | 5.6282 | 19.0959 | 0.14737 | 256.08 | 11.64 | -9.2918 | 11.7959 | 0.39386 | 314.7505 | 46.5251 | 49.3787 | 46.5644 | 0.53022 | 256.08 | 40.74 | -9.2918 | 40.7848 | 0.11391 | 240.948 | 13.386 | -24.4238 | 13.5218 | 0.90313 | 265.3718 | 4.6202 | 9.2404 | 5 | 276.45 | 11.64 | 11.0782 | 11.7959 | 0.46958 | 231.5622 | 31.1953 | -33.8096 | 31.2538 | 0.54089 | 313 | 36 | 47.6282 | 36.0507 | 0.66057
| 5 | 289.7875 | 35 | 3.962 | 34.9675 | 0.056653 | 275 | 10 | -10.8255 | 9.8857 | 0.54753 | 282.27 | 40.74 | -3.5555 | 40.7121 | 0.043666 | 309 | 15 | 23.1745 | 14.9241 | 0.77641 | 320.1 | 48.015 | 34.2745 | 47.9913 | 0.35709 | 299.73 | 26.19 | 13.9045 | 26.1466 | 0.2659 | 291 | 32.01 | 5.1745 | 31.9745 | 0.080916 | 294 | 19 | 8.1745 | 18.9401 | 0.2158 | 279.36 | 14.55 | -6.4655 | 14.4717 | 0.22338 | 331.3229 | 55.7707 | 45.4974 | 55.7503 | 0.40805 | 258.99 | 40.74 | -26.8355 | 40.7121 | 0.32958 | 262.191 | 17.751 | -23.6345 | 17.6869 | 0.66814 | 285.8255 | 5.2223 | 10.4445 | 5 | 296.82 | 14.55 | 10.9945 | 14.4717 | 0.37986 | 245.2277 | 34.2977 | -40.5978 | 34.2645 | 0.59242 | 326 | 36 | 40.1745 | 35.9684 | 0.55847
| 2 | 304.7425 | 40 | -1.2892 | 39.8931 | 0.016158 | 295 | 10 | -11.0317 | 9.5637 | 0.57675 | 299.73 | 40.74 | -6.3017 | 40.6351 | 0.07754 | 336 | 20 | 29.9683 | 19.7854 | 0.75733 | 328.83 | 43.65 | 22.7983 | 43.5521 | 0.26174 | 317.19 | 26.19 | 11.1583 | 26.0265 | 0.21436 | 320.1 | 46.56 | 14.0683 | 46.4682 | 0.15138 | 313 | 19 | 6.9683 | 18.774 | 0.18558 | 325.92 | 24.735 | 19.8883 | 24.5618 | 0.40486 | [5 empty] | 215.34 | 75.66 | -90.6917 | 75.6036 | 0.59978 | 278.196 | 17.751 | -27.8357 | 17.5089 | 0.7949 | 306.0317 | 5.7911 | 11.5822 | 5 | 317.19 | 14.55 | 11.1583 | 14.2536 | 0.39142 | 264.2293 | 45.7554 | -41.8024 | 45.662 | 0.45774 | 342 | 38 | 35.9683 | 37.8875 | 0.47467
| 1 | 315.9314 | 40 | -0.7507 | 39.8372 | 0.0094221 | 305 | 10 | -11.6821 | 9.3274 | 0.62623 | 308.46 | 40.74 | -8.2221 | 40.5801 | 0.10131 | 353 | 20 | 36.3179 | 19.6723 | 0.92307 | 337.56 | 43.65 | 20.8779 | 43.5008 | 0.23997 | 320.1 | 26.19 | 3.4179 | 25.9406 | 0.065879 | 320.1 | 46.56 | 3.4179 | 46.4202 | 0.036815 | 321 | 19 | 4.3179 | 18.6548 | 0.11573 | 378.3 | 58.2 | 61.6179 | 58.0882 | 0.53038 | [10 empty] | 286.635 | 17.751 | -30.0471 | 17.381 | 0.86437 | 316.6821 | 6.1644 | 12.3288 | 5 | 334.65 | 17.46 | 17.9679 | 17.0837 | 0.52588 | [5 empty] | 349 | 40 | 32.3179 | 39.8372 | 0.40562
10/5 | 120 | [5 empty] | 67 | 7.5 | -4.9281 | 6.9205 | 0.35605 | 81.48 | 20.37 | 9.5519 | 20.1638 | 0.23686 | [15 empty] | 87.3 | 20.37 | 15.3719 | 20.1638 | 0.38118 | [25 empty] | 71.9281 | 5.7756 | 11.5512 | 5 | 75.66 | 11.64 | 3.7319 | 11.2753 | 0.16549
| 100 | 81.1163 | 30 | 1.8877 | 29.9633 | 0.031499 | 71 | 7.5 | -8.2286 | 7.3517 | 0.55964 | 81.48 | 20.37 | 2.2514 | 20.3159 | 0.05541 | 98 | 15 | 18.7714 | 14.9264 | 0.6288 | [10 empty] | 87.3 | 20.37 | 8.0714 | 20.3159 | 0.19865 | [15 empty] | 107.67 | 40.74 | 28.4414 | 40.713 | 0.34929 | [5 empty] | 79.2286 | 5.2156 | 10.4312 | 5 | 78.57 | 11.64 | -0.6586 | 11.545 | 0.028523 | [5 empty] | 117 | 40 | 37.7714 | 39.9725 | 0.47247
| 50 | 97.5723 | 30 | 7.2457 | 30.0587 | 0.12052 | 89 | 7.5 | -1.3266 | 7.7315 | 0.085795 | 90.21 | 20.37 | -0.11665 | 20.4564 | 0.0028512 | 89 | 15 | -1.3266 | 15.1171 | 0.043879 | 96.03 | 29.1 | 5.7034 | 29.1605 | 0.097792 | 113.49 | 26.19 | 23.1634 | 26.2572 | 0.44109 | 87.3 | 20.37 | -3.0266 | 20.4564 | 0.073978 | 93 | 19 | 2.6734 | 19.0926 | 0.07001 | [5 empty] | 258.859 | 68.1058 | 168.5324 | 68.1317 | 1.2368 | 107.67 | 40.74 | 17.3434 | 40.7832 | 0.21263 | 64.311 | 17.169 | -26.0156 | 17.2714 | 0.75314 | 90.3266 | 4.634 | 9.2681 | 5 | 90.21 | 11.64 | -0.11665 | 11.7905 | 0.0049468 | 95.723 | 35.5998 | 5.3963 | 35.6493 | 0.075686 | 121 | 40 | 30.6734 | 40.044 | 0.383
| 20 | 114.9014 | 30 | 3.3064 | 29.9686 | 0.055165 | 107 | 7.5 | -4.5949 | 7.3734 | 0.31159 | 104.76 | 20.37 | -6.8349 | 20.3237 | 0.16815 | 112 | 15 | 0.40507 | 14.9371 | 0.013559 | 142.59 | 33.465 | 30.9951 | 33.4368 | 0.46349 | 128.04 | 26.19 | 16.4451 | 26.154 | 0.31439 | 116.4 | 20.37 | 4.8051 | 20.3237 | 0.11821 | 110 | 19 | -1.5949 | 18.9504 | 0.042082 | 119.31 | 10.185 | 7.7151 | 10.0921 | 0.38223 | 241.8889 | 68.3626 | 130.294 | 68.3488 | 0.95315 | 116.4 | 40.74 | 4.8051 | 40.7169 | 0.059006 | 84.972 | 17.169 | -26.6229 | 17.1141 | 0.77781 | 111.5949 | 4.2289 | 8.4578 | 4 | 113.49 | 11.64 | 1.8951 | 11.5588 | 0.081975 | 100.2406 | 35.618 | -11.3543 | 35.5915 | 0.15951 | 134 | 40 | 22.4051 | 39.9764 | 0.28023
| 10 | 125.6393 | 35 | 1.6594 | 34.9555 | 0.023736 | 119 | 7.5 | -4.9798 | 7.2896 | 0.34157 | 113.49 | 29.1 | -10.4898 | 29.0465 | 0.18057 | 129 | 15 | 5.0202 | 14.8959 | 0.16851 | 128.04 | 29.1 | 4.0602 | 29.0465 | 0.069891 | 139.68 | 26.19 | 15.7002 | 26.1305 | 0.30042 | 145.5 | 20.37 | 21.5202 | 20.2935 | 0.53022 | 121 | 19 | -2.9798 | 18.9179 | 0.078756 | 133.86 | 11.64 | 9.8802 | 11.5055 | 0.42937 | 254.1594 | 69.247 | 130.1796 | 69.2245 | 0.94027 | 119.31 | 40.74 | -4.6698 | 40.7018 | 0.057366 | 101.268 | 17.169 | -22.7118 | 17.0781 | 0.66494 | 123.9798 | 4.3718 | 8.7436 | 4 | 122.22 | 11.64 | -1.7598 | 11.5055 | 0.076477 | 108.7163 | 36.4661 | -15.2635 | 36.4234 | 0.20953 | 142 | 42 | 18.0202 | 41.9629 | 0.21472
| 5 | 134.5147 | 35 | 0.73411 | 34.8855 | 0.010522 | 128 | 7.5 | -5.7806 | 6.9462 | 0.4161 | 125.13 | 40.74 | -8.6506 | 40.6417 | 0.10643 | 145 | 15 | 11.2194 | 14.7309 | 0.38081 | 145.5 | 39.285 | 11.7194 | 39.183 | 0.14955 | 151.32 | 26.19 | 17.5394 | 26.0368 | 0.33682 | 145.5 | 32.01 | 11.7194 | 31.8848 | 0.18378 | 132 | 19 | -1.7806 | 18.7883 | 0.047387 | 139.68 | 14.55 | 5.8994 | 14.2724 | 0.20667 | 262.0989 | 82.1776 | 128.3182 | 82.1289 | 0.7812 | 119.31 | 40.74 | -14.4706 | 40.6417 | 0.17803 | 112.326 | 37.83 | -21.4546 | 37.7241 | 0.28436 | 133.7806 | 4.899 | 9.7981 | 4 | 130.95 | 14.55 | -2.8306 | 14.2724 | 0.099165 | 115.6842 | 36.2812 | -18.0964 | 36.1707 | 0.25015 | 176 | 44 | 42.2194 | 43.909 | 0.48076
| 2 | 143.2011 | 40 | 0.69229 | 39.8366 | 0.0086891 | 137 | 7.5 | -5.5088 | 6.5732 | 0.41903 | 133.86 | 40.74 | -8.6488 | 40.5796 | 0.10657 | 163 | 20 | 20.4912 | 19.6712 | 0.52084 | 148.41 | 48.015 | 5.9012 | 47.879 | 0.061626 | 160.05 | 26.19 | 17.5412 | 25.9398 | 0.33811 | 145.5 | 46.56 | 2.9912 | 46.4197 | 0.032219 | 140 | 19 | -2.5088 | 18.6536 | 0.067247 | 180.42 | 24.735 | 37.9112 | 24.4699 | 0.77465 | [5 empty] | 93.12 | 75.66 | -49.3888 | 75.5738 | 0.32676 | 120.183 | 37.83 | -22.3258 | 37.6572 | 0.29643 | 142.5088 | 5.3891 | 10.7783 | 4 | 142.59 | 14.55 | 0.081189 | 14.0947 | 0.0028801 | 126.4312 | 41.8072 | -16.0776 | 41.6509 | 0.193
| 1 | 147.6534 | 40 | 0.058197 | 39.788 | 0.00073134 | 142 | 7.5 | -5.5952 | 6.2715 | 0.44608 | 139.68 | 40.74 | -7.9152 | 40.5318 | 0.097642 | 174 | 20 | 26.4048 | 19.5725 | 0.67454 | 154.23 | 48.015 | 6.6348 | 47.8385 | 0.069346 | 174.6 | 34.92 | 27.0048 | 34.6769 | 0.38938 | 145.5 | 46.56 | -2.0952 | 46.378 | 0.022588 | 145 | 19 | -2.5952 | 18.5494 | 0.069954 | 238.62 | 49.47 | 91.0248 | 49.2987 | 0.9232 | [10 empty] | 121.056 | 37.83 | -26.5392 | 37.6057 | 0.35286 | 147.5952 | 5.7375 | 11.4749 | 4 | 148.41 | 17.46 | 0.8148 | 16.9686 | 0.024009
15 VA |  | CMI | CMI | CMI | CMI | CMI | PTB | PTB | PTB | PTB | PTB | METAS | METAS | METAS | METAS | METAS | NPL |  |  |  |  | LCOE |  |  |  |  | MIKES |  |  |  |  | BEV |  |  |  |  | VSL |  |  |  |  | SP |  |  |  |  | DMDM |  |  |  |  | GUM |  |  |  |  | BIM |  |  |  |  | weighted mean dr (mrad) | u(dr) (mrad) | U(dr) (mrad) | u(dstd) (mrad) | UME |  |  |  |  | LNE |  |  |  |  | INRIM
kI (kA/A) | I1 (%IN) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) |  |  |  |  | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-) | dL (mrad) | u(dL) (mrad) | Dd (mrad) | u(Ddstd) (mrad) | E(d) (-)
4/5 | 120 | 183.912 | 30 | -9.0022 | 30.0334 | 0.14987 | 165 | 20 | -27.9142 | 20.0501 | 0.69611 | 206.61 | 40.74 | 13.6958 | 40.7646 | 0.16799 | 175 | 30 | -17.9142 | 30.0334 | 0.29824 | 235.71 | 27.645 | 42.7958 | 27.6813 | 0.77301 | 198 | 26 | 5.0858 | 26.0386 | 0.097658 | 174.6 | 20.37 | -18.3142 | 20.4192 | 0.44846 | 187 | 27 | -5.9142 | 27.0371 | 0.10937 | [5 empty] | 264.7809 | 47.4736 | 71.8667 | 47.4947 | 0.75658 | 247.35 | 40.74 | 54.4358 | 40.7646 | 0.66768 | 146.664 | 13.386 | -46.2502 | 13.4608 | 1.718 | 192.9142 | 5.8303 | 11.6606 | 6 | 226.98 | 11.64 | 34.0658 | 11.7259 | 1.4526 | 186.35 | 19.0877 | -6.5642 | 19.1402 | 0.17148 | 245 | 36 | 52.0858 | 36.0279 | 0.72285
| 100 | 205.5479 | 30 | -3.759 | 29.8731 | 0.062917 | 180 | 20 | -29.3069 | 19.8092 | 0.73973 | 218.25 | 40.74 | 8.9431 | 40.6467 | 0.11001 | 186 | 30 | -23.3069 | 29.8731 | 0.3901 | 258.99 | 23.28 | 49.6831 | 23.1163 | 1.0746 | 212 | 26 | 2.6931 | 25.8535 | 0.052084 | 203.7 | 20.37 | -5.6069 | 20.1827 | 0.1389 | 199 | 26 | -10.3069 | 25.8535 | 0.19933 | 168.78 | 40.74 | -40.5269 | 40.6467 | 0.49853 | 284.4719 | 47.8189 | 75.165 | 47.7394 | 0.78724 | 241.53 | 40.74 | 32.2231 | 40.6467 | 0.39638 | 158.886 | 13.386 | -50.4209 | 13.0992 | 1.9246 | 209.3069 | 5.7093 | 11.4187 | 5 | 247.35 | 11.64 | 38.0431 | 11.309 | 1.682 | 196.25 | 19.0199 | -13.0569 | 18.8191 | 0.3469 | 261 | 36 | 51.6931 | 35.8943 | 0.72007
| 50 | 262.1473 | 30 | 8.8362 | 29.8406 | 0.14806 | 236 | 20 | -17.3111 | 19.7601 | 0.43803 | 264.81 | 40.74 | 11.4989 | 40.6228 | 0.14153 | 255 | 30 | 1.6889 | 29.8406 | 0.028298 | 328.83 | 93.12 | 75.5189 | 93.0688 | 0.40572 | 268 | 26 | 14.6889 | 25.8159 | 0.28449 | 261.9 | 20.37 | 8.5889 | 20.1345 | 0.21329 | 255 | 26 | 1.6889 | 25.8159 | 0.03271 | 232.8 | 40.74 | -20.5111 | 40.6228 | 0.25246 | 336.4348 | 47.927 | 83.1237 | 47.8274 | 0.869 | 267.72 | 40.74 | 14.4089 | 40.6228 | 0.17735 | 212.721 | 13.386 | -40.5901 | 13.0249 | 1.5582 | 253.3111 | 5.8768 | 11.7536 | 5 | 282.27 | 11.64 | 28.9589 | 11.2229 | 1.2902 | 222.75 | 18.9922 | -30.5611 | 18.7395 | 0.81542 | 300 | 36 | 46.6889 | 35.8673 | 0.65086
| 20 | 322.5735 | 30 | 6.4446 | 29.863 | 0.1079 | 299 | 20 | -17.1289 | 19.7939 | 0.43268 | 320.1 | 40.74 | 3.9711 | 40.6392 | 0.048858 | 320 | 30 | 3.8711 | 29.863 | 0.064814 | 320.1 | 27.645 | 3.9711 | 27.4963 | 0.072212 | 329 | 26 | 12.8711 | 25.8418 | 0.24904 | 320.1 | 20.37 | 3.9711 | 20.1677 | 0.098452 | 320 | 24 | 3.8711 | 23.8285 | 0.081228 | 296.82 | 40.74 | -19.3089 | 40.6392 | 0.23756 | 392.1419 | 47.1843 | 76.013 | 47.0973 | 0.80698 | 323.01 | 40.74 | 6.8811 | 40.6392 | 0.084661 | 284.598 | 13.677 | -31.5309 | 13.3738 | 1.1788 | 316.1289 | 5.7621 | 11.5241 | 5 | 328.83 | 11.64 | 12.7011 | 11.2822 | 0.56288 | 318 | 19.2154 | 1.8711 | 19.0008 | 0.049237 | 357 | 36 | 40.8711 | 35.8859 | 0.56946
| 10 | 358.3083 | 35 | 3.5578 | 34.8209 | 0.051087 | 335 | 20 | -19.7505 | 19.6848 | 0.50167 | 355.02 | 58.2 | 0.26945 | 58.0924 | 0.0023192 | 366 | 30 | 11.2495 | 29.7908 | 0.18881 | 389.94 | 59.655 | 35.1895 | 59.5501 | 0.29546 | 361 | 26 | 6.2495 | 25.7583 | 0.12131 | 378.3 | 20.37 | 23.5495 | 20.0606 | 0.58696 | 354 | 24 | -0.75055 | 23.738 | 0.015809 | 334.65 | 46.56 | -20.1005 | 46.4255 | 0.21648 | 433.9295 | 47.195 | 79.179 | 47.0623 | 0.84121 | 349.2 | 40.74 | -5.5505 | 40.5862 | 0.06838 | 325.047 | 13.677 | -29.7035 | 13.2118 | 1.1241 | 354.7505 | 6.1244 | 12.2488 | 5 | 366.66 | 11.64 | 11.9095 | 11.0897 | 0.53696 | 338.25 | 31.2995 | -16.5005 | 31.099 | 0.26529 | 386 | 36 | 31.2495 | 35.8259 | 0.43613
| 5 | 385.8514 | 35 | 4.9972 | 34.6376 | 0.072135 | 364 | 20 | -16.8543 | 19.3589 | 0.43531 | 381.21 | 81.48 | 0.35572 | 81.325 | 0.002187 | 403 | 30 | 22.1457 | 29.5764 | 0.37438 | 404.49 | 46.56 | 23.6357 | 46.2882 | 0.25531 | 387 | 26 | 6.1457 | 25.5101 | 0.12046 | 378.3 | 32.01 | -2.5543 | 31.6134 | 0.040399 | 386 | 24 | 5.1457 | 23.4684 | 0.10963 | 366.66 | 58.2 | -14.1943 | 57.9828 | 0.1224 | 456.1085 | 56.3368 | 75.2543 | 56.1124 | 0.67057 | 360.84 | 40.74 | -20.0143 | 40.4291 | 0.24752 | 355.893 | 18.042 | -24.9613 | 17.3286 | 0.72023 | 380.8543 | 7.0876 | 14.1752 | 5 | 392.85 | 14.55 | 11.9957 | 13.6554 | 0.43923 | 353.5 | 33.4055 | -27.3543 | 33.0256 | 0.41414 | 412 | 36 | 31.1457 | 35.6478 | 0.43685
| 2 | 412.4925 | 40 | 3.7266 | 39.723 | 0.046908 | 393 | 20 | -15.7659 | 19.4401 | 0.4055 | 407.4 | 81.48 | -1.3659 | 81.3444 | 0.0083955 | 441 | 40 | 32.2341 | 39.723 | 0.40574 | 427.77 | 45.105 | 19.0041 | 44.8595 | 0.21182 | 410 | 26 | 1.2341 | 25.5718 | 0.024131 | 407.4 | 46.56 | -1.3659 | 46.3222 | 0.014743 | 413 | 24 | 4.2341 | 23.5354 | 0.089953 | 424.86 | 98.94 | 16.0941 | 98.8283 | 0.081425 | [5 empty] | 337.56 | 75.66 | -71.2059 | 75.5139 | 0.47148 | 382.956 | 18.333 | -25.8099 | 17.7205 | 0.72825 | 408.7659 | 7.6213 | 15.2426 | 6 | 419.04 | 14.55 | 10.2741 | 13.7702 | 0.37306 | 447.9167 | 45.682 | 39.1508 | 45.4396 | 0.4308 | 437 | 38 | 28.2341 | 37.7083 | 0.37438
| 1 | 428.6866 | 40 | 7.1259 | 39.5987 | 0.089977 | 409 | 20 | -12.5607 | 19.1848 | 0.32736 | 421.95 | 81.48 | 0.38925 | 81.2837 | 0.0023944 | 463 | 40 | 41.4393 | 39.5987 | 0.52324 | 439.41 | 43.65 | 17.8493 | 43.2825 | 0.20619 | 431 | 29 | 9.4393 | 28.4439 | 0.16593 | 407.4 | 46.56 | -14.1607 | 46.2157 | 0.1532 | 427 | 24 | 5.4393 | 23.325 | 0.1166 | 474.33 | 197.88 | 52.7693 | 197.7993 | 0.13339 | [10 empty] | 394.305 | 18.333 | -27.2557 | 17.44 | 0.78141 | 421.5607 | 8.2428 | 16.4856 | 6 | 433.59 | 17.46 | 12.0293 | 16.5199 | 0.36408 | [5 empty] | 450 | 40 | 28.4393 | 39.5987 | 0.35909
5/5 | 120 | 155.605 | 30 | -10.8137 | 30.2859 | 0.17853 | 134 | 20 | -32.4187 | 20.4263 | 0.79355 | 171.69 | 40.74 | 5.2713 | 40.951 | 0.064361 | 139 | 30 | -27.4187 | 30.2859 | 0.45266 | 194.97 | 64.02 | 28.5513 | 64.1545 | 0.22252 | 160 | 26 | -6.4187 | 26.3294 | 0.12189 | 145.5 | 20.37 | -20.9187 | 20.7888 | 0.50312 | 149 | 25 | -17.4187 | 25.3424 | 0.34367 | [10 empty] | 215.34 | 40.74 | 48.9213 | 40.951 | 0.59732 | [5 empty] | 166.4187 | 6.8385 | 13.6769 | 8 | 183.33 | 11.64 | 16.9113 | 12.3582 | 0.68421 | 166.5 | 25.2303 | 0.081318 | 25.5696 | 0.0015901 | 227 | 36 | 60.5813 | 36.2386 | 0.83587
| 100 | 167.2523 | 30 | -1.6804 | 30.2327 | 0.027792 | 144 | 20 | -24.9327 | 20.3474 | 0.61268 | 180.42 | 40.74 | 11.4873 | 40.9117 | 0.14039 | 142 | 30 | -26.9327 | 30.2327 | 0.44542 | 221.16 | 81.48 | 52.2273 | 81.566 | 0.32015 | 172 | 26 | 3.0673 | 26.2681 | 0.058384 | 174.6 | 20.37 | 5.6673 | 20.7112 | 0.13682 | 159 | 25 | -9.9327 | 25.2788 | 0.19646 | 151.32 | 40.74 | -17.6127 | 40.9117 | 0.21525 | 357.9639 | 68.2068 | 189.0313 | 68.3094 | 1.3836 | 197.88 | 40.74 | 28.9473 | 40.9117 | 0.35378 | 133.569 | 13.386 | -35.3637 | 13.8997 | 1.2721 | 168.9327 | 5.9148 | 11.8295 | 7 | 192.06 | 11.64 | 23.1273 | 12.2272 | 0.94573 | 169.5 | 20.1356 | 0.5673 | 20.4807 | 0.01385 | 240 | 36 | 71.0673 | 36.1941 | 0.98175
| 50 | 208.5815 | 30 | 6.4466 | 29.8474 | 0.10799 | 188 | 20 | -14.135 | 19.7704 | 0.35748 | 215.34 | 40.74 | 13.205 | 40.6278 | 0.16251 | 197 | 30 | -5.135 | 29.8474 | 0.08602 | 302.64 | 71.295 | 100.505 | 71.2309 | 0.70549 | 215 | 26 | 12.865 | 25.8238 | 0.24909 | 203.7 | 20.37 | 1.565 | 20.1446 | 0.038845 | 201 | 24 | -1.135 | 23.809 | 0.023835 | 186.24 | 40.74 | -15.895 | 40.6278 | 0.19562 | 364.4436 | 68.1768 | 162.3086 | 68.1098 | 1.1915 | 200.79 | 40.74 | -1.345 | 40.6278 | 0.016552 | 162.96 | 13.386 | -39.175 | 13.0404 | 1.5021 | 202.135 | 5.8423 | 11.6845 | 5 | 221.16 | 11.64 | 19.025 | 11.2409 | 0.84624 | 200.25 | 18.2641 | -1.885 | 18.0124 | 0.052324 | 260 | 36 | 57.865 | 35.8729 | 0.80653
| 20 | 257.4295 | 30 | 6.5346 | 29.8578 | 0.10943 | 238 | 20 | -12.895 | 19.786 | 0.32586 | 256.08 | 40.74 | 5.185 | 40.6354 | 0.0638 | 250 | 30 | -0.89496 | 29.8578 | 0.014987 | 293.91 | 39.285 | 43.015 | 39.1765 | 0.54899 | 262 | 26 | 11.105 | 25.8358 | 0.21492 | 261.9 | 20.37 | 11.005 | 20.1599 | 0.27294 | 250 | 24 | -0.89496 | 23.822 | 0.018784 | 250.26 | 40.74 | -0.63496 | 40.6354 | 0.0078129 | 383.1452 | 68.1732 | 132.2502 | 68.1107 | 0.97085 | 250.26 | 40.74 | -0.63496 | 40.6354 | 0.0078129 | 218.541 | 13.386 | -32.354 | 13.0641 | 1.2383 | 250.895 | 5.7891 | 11.5782 | 5 | 258.99 | 11.64 | 8.095 | 11.2683 | 0.35919 | 254.75 | 17.9982 | 3.855 | 17.7601 | 0.10853 | 296 | 36 | 45.105 | 35.8816 | 0.62853
| 10 | 284.3579 | 35 | 2.815 | 34.8559 | 0.040381 | 264 | 20 | -17.5429 | 19.7468 | 0.4442 | 282.27 | 58.2 | 0.72711 | 58.1135 | 0.006256 | 284 | 30 | 2.4571 | 29.8318 | 0.041183 | 343.38 | 36.375 | 61.8371 | 36.2364 | 0.85325 | 285 | 26 | 3.4571 | 25.8058 | 0.066983 | 291 | 20.37 | 9.4571 | 20.1215 | 0.235 | 276 | 24 | -5.5429 | 23.7894 | 0.1165 | 279.36 | 46.56 | -2.1829 | 46.4518 | 0.023496 | 390.0855 | 68.1814 | 108.5426 | 68.1075 | 0.79685 | 267.72 | 40.74 | -13.8229 | 40.6163 | 0.17016 | 250.26 | 13.386 | -31.2829 | 13.0047 | 1.2028 | 281.5429 | 5.9214 | 11.8428 | 5 | 288.09 | 11.64 | 6.5471 | 11.1994 | 0.2923 | 304.5833 | 20.3934 | 23.0404 | 20.1452 | 0.57186 | 321 | 36 | 39.4571 | 35.86 | 0.55016
| 5 | 305.841 | 35 | 3.3512 | 34.6957 | 0.048295 | 287 | 20 | -15.4898 | 19.4626 | 0.39794 | 302.64 | 81.48 | 0.15024 | 81.3498 | 0.00092344 | 314 | 30 | 11.5102 | 29.6445 | 0.19414 | 325.92 | 46.56 | 23.4302 | 46.3317 | 0.25285 | 308 | 26 | 5.5102 | 25.5889 | 0.10767 | 291 | 32.01 | -11.4898 | 31.677 | 0.18136 | 300 | 24 | -2.4898 | 23.5541 | 0.052852 | 328.83 | 58.2 | 26.3402 | 58.0175 | 0.227 | 388.097 | 81.7859 | 85.6072 | 81.6562 | 0.52419 | 270.63 | 40.74 | -31.8598 | 40.4789 | 0.39354 | 274.704 | 17.751 | -27.7858 | 17.1433 | 0.8104 | 302.4898 | 6.7975 | 13.5951 | 5 | 308.46 | 14.55 | 5.9702 | 13.802 | 0.21628 | 320.25 | 19.262 | 17.7602 | 18.7034 | 0.47479 | 342 | 36 | 39.5102 | 35.7043 | 0.5533
| 2 | 326.7675 | 40 | 4.5468 | 39.7325 | 0.057218 | 309 | 20 | -13.2208 | 19.4594 | 0.3397 | 323.01 | 81.48 | 0.78925 | 81.349 | 0.004851 | 345 | 40 | 22.7792 | 39.7325 | 0.28666 | 343.38 | 43.65 | 21.1592 | 43.405 | 0.24374 | 332 | 26 | 9.7792 | 25.5865 | 0.1911 | 320.1 | 46.56 | -2.1208 | 46.3304 | 0.022887 | 321 | 24 | -1.2208 | 23.5514 | 0.025917 | 384.12 | 98.94 | 61.8992 | 98.8321 | 0.31315 | [5 empty] | 232.8 | 75.66 | -89.4208 | 75.5189 | 0.59204 | 295.656 | 17.751 | -26.5648 | 17.1396 | 0.77495 | 322.2208 | 7.5718 | 15.1436 | 6 | 331.74 | 14.55 | 9.5192 | 13.7975 | 0.34496 | 357.1667 | 45.9535 | 34.9459 | 45.7208 | 0.38217 | 359 | 38 | 36.7792 | 37.7183 | 0.48755
| 1 | 337.5418 | 40 | 4.6152 | 39.6043 | 0.058266 | 322 | 20 | -10.9267 | 19.1964 | 0.2846 | 331.74 | 81.48 | -1.1867 | 81.2865 | 0.0072992 | 364 | 40 | 31.0733 | 39.6043 | 0.3923 | 355.02 | 43.65 | 22.0933 | 43.2877 | 0.25519 | 352 | 29 | 19.0733 | 28.4517 | 0.33519 | 320.1 | 46.56 | -12.8267 | 46.2205 | 0.13876 | 332 | 24 | -0.92666 | 23.3345 | 0.019856 | 462.69 | 197.88 | 129.7633 | 197.8004 | 0.32802 | [10 empty] | 303.222 | 18.042 | -29.7047 | 17.1468 | 0.86619 | 332.9267 | 8.2158 | 16.4317 | 6 | 346.29 | 17.46 | 13.3633 | 16.5333 | 0.40413 | [5 empty] | 369 | 40 | 36.0733 | 39.6043 | 0.45542
6/5 | 120 | 128.8475 | 30 | -11.3014 | 29.9234 | 0.18884 | 114 | 15 | -26.149 | 14.8462 | 0.88066 | 130.95 | 40.74 | -9.199 | 40.6836 | 0.11305 | 127 | 30 | -13.149 | 29.9234 | 0.21971 | 142.59 | 27.645 | 2.441 | 27.5618 | 0.044283 | 140 | 32 | -0.14896 | 31.9282 | 0.0023328 | 145.5 | 20.37 | 5.351 | 20.257 | 0.13208 | 125 | 24 | -15.149 | 23.9042 | 0.31687 | [10 empty] | 174.6 | 81 | 34.451 | 80.9717 | 0.21274 | [5 empty] | 140.149 | 6.3711 | 12.7422 | 6 | 148.41 | 11.64 | 8.261 | 11.4411 | 0.36102 | 156.5 | 19.4955 | 16.351 | 19.3774 | 0.42191 | 205 | 36 | 64.851 | 35.9362 | 0.90231
| 100 | 142.1135 | 30 | -2.7401 | 29.9248 | 0.045784 | 121 | 15 | -23.8536 | 14.8489 | 0.80321 | 133.86 | 40.74 | -10.9936 | 40.6846 | 0.13511 | 117 | 30 | -27.8536 | 29.9248 | 0.46539 | 203.7 | 45.105 | 58.8464 | 45.055 | 0.65305 | 146 | 26 | 1.1464 | 25.9131 | 0.02212 | 145.5 | 20.37 | 0.64637 | 20.259 | 0.015953 | 132 | 24 | -12.8536 | 23.9059 | 0.26884 | [10 empty] | 160.05 | 40.74 | 15.1964 | 40.6846 | 0.18676 | [5 empty] | 144.8536 | 6.3647 | 12.7293 | 6 | 154.23 | 11.64 | 9.3764 | 11.4447 | 0.40964 | 147.75 | 19.8221 | 2.8964 | 19.708 | 0.073482 | 211 | 36 | 66.1464 | 35.9373 | 0.9203
| 50 | 175.8331 | 30 | 8.9352 | 29.846 | 0.14969 | 156 | 15 | -10.8979 | 14.6895 | 0.37094 | 157.14 | 40.74 | -9.7579 | 40.6267 | 0.12009 | 156 | 30 | -10.8979 | 29.846 | 0.18257 | 241.53 | 53.835 | 74.6321 | 53.7493 | 0.69426 | 183 | 26 | 16.1021 | 25.8221 | 0.31179 | 174.6 | 20.37 | 7.7021 | 20.1424 | 0.19119 | 166 | 24 | -0.89795 | 23.8072 | 0.018859 | 157.14 | 40.74 | -9.7579 | 40.6267 | 0.12009 | 198.6269 | 70.4942 | 31.729 | 70.4288 | 0.22526 | 168.78 | 40.74 | 1.8821 | 40.6267 | 0.023163 | 132.114 | 17.169 | -34.7839 | 16.8984 | 1.0292 | 166.8979 | 5.8497 | 11.6994 | 5 | 180.42 | 11.64 | 13.5221 | 11.237 | 0.60167 | 154 | 18.1575 | -12.8979 | 17.9018 | 0.36024 | 226 | 36 | 59.1021 | 35.8717 | 0.8238
| 20 | 215.0999 | 30 | 7.955 | 29.8599 | 0.13321 | 196 | 15 | -11.1449 | 14.7179 | 0.37862 | 192.06 | 40.74 | -15.0849 | 40.637 | 0.18561 | 202 | 30 | -5.1449 | 29.8599 | 0.086151 | 209.52 | 29.1 | 2.3751 | 28.9556 | 0.041012 | 221 | 26 | 13.8551 | 25.8383 | 0.26811 | 203.7 | 20.37 | -3.4449 | 20.1632 | 0.085427 | 208 | 24 | 0.85506 | 23.8247 | 0.017945 | 209.52 | 40.74 | 2.3751 | 40.637 | 0.029223 | 195.7702 | 73.0901 | -11.3747 | 73.0328 | 0.077874 | 203.7 | 40.74 | -3.4449 | 40.637 | 0.042387 | 177.801 | 17.169 | -29.3439 | 16.9231 | 0.86698 | 207.1449 | 5.7779 | 11.5558 | 5 | 221.16 | 11.64 | 14.0151 | 11.2741 | 0.62156 | 208 | 18.4361 | 0.85506 | 18.2073 | 0.023481 | 253 | 36 | 45.8551 | 35.8834 | 0.63895
| 10 | 238.1144 | 35 | 6.6287 | 34.852 | 0.095098 | 219 | 15 | -12.4857 | 14.6513 | 0.42609 | 215.34 | 58.2 | -16.1457 | 58.1111 | 0.13892 | 233 | 30 | 1.5143 | 29.8272 | 0.025385 | 279.36 | 43.65 | 47.8743 | 43.5314 | 0.54988 | 242 | 26 | 10.5143 | 25.8004 | 0.20376 | 232.8 | 20.37 | 1.3143 | 20.1146 | 0.03267 | 229 | 24 | -2.4857 | 23.7836 | 0.052256 | 226.98 | 46.56 | -4.5057 | 46.4488 | 0.048502 | 189.9551 | 76.1178 | -41.5306 | 76.0499 | 0.27305 | 218.25 | 40.74 | -13.2357 | 40.6129 | 0.16295 | 206.028 | 17.46 | -25.4577 | 17.1614 | 0.74171 | 231.4857 | 5.9446 | 11.8892 | 5 | 235.71 | 11.64 | 4.2243 | 11.1871 | 0.1888 | 250.5 | 18.8583 | 19.0143 | 18.5821 | 0.51163 | 270 | 36 | 38.5143 | 35.8561 | 0.53707
| 5 | 255.1488 | 35 | 5.0405 | 34.6282 | 0.072781 | 237 | 15 | -13.1083 | 14.1106 | 0.46448 | 235.71 | 81.48 | -14.3983 | 81.321 | 0.088527 | 257 | 30 | 6.8917 | 29.5653 | 0.11655 | 261.9 | 39.285 | 11.7917 | 38.9541 | 0.15135 | 259 | 26 | 8.8917 | 25.4972 | 0.17437 | 261.9 | 32.01 | 11.7917 | 31.603 | 0.18656 | 248 | 24 | -2.1083 | 23.4544 | 0.044944 | 247.35 | 58.2 | -2.7583 | 57.9771 | 0.023788 | 195.004 | 90.4677 | -55.1043 | 90.3244 | 0.30504 | 215.34 | 40.74 | -34.7683 | 40.421 | 0.43008 | 226.107 | 37.83 | -24.0013 | 37.4862 | 0.32013 | 250.1083 | 7.1338 | 14.2676 | 5 | 253.17 | 14.55 | 3.0617 | 13.6313 | 0.11231 | 270.9167 | 30.5116 | 20.8084 | 30.0843 | 0.34583 | 287 | 36 | 36.8917 | 35.6386 | 0.51758
| 2 | 271.6703 | 40 | 3.5765 | 39.7066 | 0.045037 | 256 | 15 | -12.0938 | 14.1991 | 0.42586 | 253.17 | 81.48 | -14.9238 | 81.3364 | 0.091741 | 286 | 40 | 17.9062 | 39.7066 | 0.22548 | 282.27 | 48.015 | 14.1762 | 47.7709 | 0.14838 | 276 | 26 | 7.9062 | 25.5463 | 0.15474 | 261.9 | 46.56 | -6.1938 | 46.3082 | 0.066876 | 264 | 24 | -4.0938 | 23.5078 | 0.087073 | 296.82 | 98.94 | 28.7262 | 98.8218 | 0.14534 | [5 empty] | 186.24 | 72.75 | -81.8538 | 72.5891 | 0.56382 | 244.149 | 37.83 | -23.9448 | 37.5196 | 0.3191 | 268.0938 | 7.7062 | 15.4123 | 6 | 273.54 | 14.55 | 5.4462 | 13.7229 | 0.19844 | 298.9583 | 45.748 | 30.8646 | 45.4917 | 0.33923 | 302 | 38 | 33.9062 | 37.691 | 0.44979
| 1 | 277.7195 | 40 | 0.83953 | 39.5909 | 0.010603 | 266 | 15 | -10.88 | 13.8722 | 0.39215 | 267.72 | 81.48 | -9.16 | 81.2799 | 0.056348 | 303 | 40 | 26.12 | 39.5909 | 0.32987 | 291 | 48.015 | 14.12 | 47.6747 | 0.14809 | 285 | 26 | 8.12 | 25.3661 | 0.16006 | 261.9 | 46.56 | -14.98 | 46.209 | 0.16209 | 274 | 24 | -2.88 | 23.3117 | 0.06177 | 357.93 | 197.88 | 81.05 | 197.7977 | 0.20488 | [10 empty] | 247.059 | 37.83 | -29.821 | 37.3971 | 0.39871 | 276.88 | 8.2803 | 16.5606 | 6 | 285.18 | 17.46 | 8.3 | 16.5012 | 0.2515 | [5 empty] | 306 | 40 | 29.12 | 39.5909 | 0.36776
8/5 | 120 | 104.1162 | 30 | 0.47655 | 30.1621 | 0.0078999 | 86 | 15 | -17.6396 | 15.3216 | 0.57564 | 110.58 | 40.74 | 6.9404 | 40.8595 | 0.08493 | 113 | 30 | 9.3604 | 30.1621 | 0.15517 | [10 empty] | 87.3 | 20.37 | -16.3396 | 20.608 | 0.39644 | [15 empty] | 139.68 | 40.74 | 36.0404 | 40.8595 | 0.44103 | [5 empty] | 103.6396 | 7.3653 | 14.7306 | 8 | 107.67 | 11.64 | 4.0304 | 12.0516 | 0.16721 | [5 empty] | 170 | 36 | 66.3604 | 36.1352 | 0.91822
| 100 | 113.2936 | 30 | 5.9048 | 30.3471 | 0.097288 | 93 | 15 | -14.3888 | 15.6826 | 0.45875 | 110.58 | 40.74 | 3.1912 | 40.9963 | 0.038921 | 97 | 30 | -10.3888 | 30.3471 | 0.17117 | 119.31 | 26.19 | 11.9212 | 26.5869 | 0.22419 | 125 | 32 | 17.6112 | 32.3256 | 0.2724 | 87.3 | 20.37 | -20.0888 | 20.8778 | 0.4811 | 104 | 24 | -3.3888 | 24.4325 | 0.06935 | [10 empty] | 130.95 | 40.74 | 23.5612 | 40.9963 | 0.28736 | [5 empty] | 107.3888 | 6.5616 | 13.1232 | 8 | 107.67 | 11.64 | 0.28121 | 12.5074 | 0.011242 | 133.0833 | 40.1069 | 25.6945 | 40.3672 | 0.31826 | 178 | 36 | 70.6112 | 36.2897 | 0.97288
| 50 | 137.3884 | 30 | 9.2432 | 30.1964 | 0.15305 | 119 | 15 | -9.1452 | 15.3891 | 0.29713 | 122.22 | 40.74 | -5.9252 | 40.8849 | 0.072462 | 124 | 30 | -4.1452 | 30.1964 | 0.068637 | 189.15 | 62.565 | 61.0048 | 62.6594 | 0.4868 | 143 | 26 | 14.8548 | 26.2264 | 0.2832 | 145.5 | 20.37 | 17.3548 | 20.6582 | 0.42005 | 127 | 24 | -1.1452 | 24.2451 | 0.023617 | 116.4 | 40.74 | -11.7452 | 40.8849 | 0.14364 | 274.3306 | 68.3366 | 146.1853 | 68.423 | 1.0682 | 133.86 | 40.74 | 5.7148 | 40.8849 | 0.069889 | 96.03 | 17.169 | -32.1152 | 17.51 | 0.91705 | 128.1452 | 6.0971 | 12.1943 | 7 | 128.04 | 11.64 | -0.1052 | 12.1373 | 0.0043338 | 130.25 | 35.6702 | 2.1048 | 35.8356 | 0.029367 | 182 | 36 | 53.8548 | 36.1639 | 0.74459
| 20 | 166.8521 | 30 | 8.4328 | 30.0068 | 0.14052 | 149 | 15 | -9.4193 | 15.0136 | 0.31369 | 142.59 | 40.74 | -15.8293 | 40.745 | 0.19425 | 156 | 30 | -2.4193 | 30.0068 | 0.040313 | 160.05 | 26.19 | 1.6307 | 26.1978 | 0.031123 | 169 | 26 | 10.5807 | 26.0078 | 0.20341 | 174.6 | 20.37 | 16.1807 | 20.38 | 0.39698 | 157 | 24 | -1.4193 | 24.0085 | 0.029558 | 151.32 | 40.74 | -7.0993 | 40.745 | 0.087119 | 257.0403 | 69.6031 | 98.621 | 69.606 | 0.70842 | 160.05 | 40.74 | 1.6307 | 40.745 | 0.020011 | 128.331 | 17.169 | -30.0883 | 17.1809 | 0.87563 | 158.4193 | 5.966 | 11.932 | 6 | 165.87 | 11.64 | 7.4507 | 11.6575 | 0.31957 | 145.25 | 35.6481 | -13.1693 | 35.6538 | 0.18468 | 198 | 36 | 39.5807 | 36.0057 | 0.54965
| 10 | 183.6246 | 35 | 11.676 | 34.9918 | 0.16684 | 166 | 15 | -5.9486 | 14.9809 | 0.19854 | 162.96 | 58.2 | -8.9886 | 58.1951 | 0.077228 | 177 | 30 | 5.0514 | 29.9904 | 0.084216 | 174.6 | 26.19 | 2.6514 | 26.179 | 0.050639 | 186 | 26 | 14.0514 | 25.989 | 0.27033 | 174.6 | 20.37 | 2.6514 | 20.3559 | 0.065125 | 173 | 24 | 1.0514 | 23.988 | 0.021914 | 171.69 | 46.56 | -0.25864 | 46.5538 | 0.0027778 | 246.0163 | 71.2437 | 74.0676 | 71.2397 | 0.51985 | 168.78 | 40.74 | -3.1686 | 40.733 | 0.038895 | 150.738 | 17.169 | -21.2106 | 17.1523 | 0.6183 | 171.9486 | 6.0476 | 12.0952 | 6 | 174.6 | 11.64 | 2.6514 | 11.6153 | 0.11413 | 158.25 | 35.6614 | -13.6986 | 35.6534 | 0.19211 | 213 | 36 | 41.0514 | 35.992 | 0.57028
| 5 | 196.876 | 35 | 9.6234 | 34.7721 | 0.13838 | 180 | 15 | -7.2527 | 14.4603 | 0.25078 | 177.51 | 81.48 | -9.7427 | 81.3824 | 0.059857 | 198 | 30 | 10.7473 | 29.7338 | 0.18073 | 194.97 | 39.285 | 7.7173 | 39.0821 | 0.098732 | 198 | 26 | 10.7473 | 25.6924 | 0.20915 | 174.6 | 32.01 | -12.6527 | 31.7607 | 0.19919 | 186 | 24 | -1.2527 | 23.6664 | 0.026465 | 180.42 | 58.2 | -6.8327 | 58.0632 | 0.058838 | 246.4624 | 86.3136 | 59.2098 | 86.2214 | 0.34336 | 165.87 | 40.74 | -21.3827 | 40.5444 | 0.26369 | 166.743 | 37.83 | -20.5097 | 37.6193 | 0.2726 | 187.2527 | 7.2041 | 14.4083 | 6 | 189.15 | 14.55 | 1.8973 | 13.993 | 0.067796 | 174.75 | 38.3432 | -12.5027 | 38.1353 | 0.16393 | 224 | 36 | 36.7473 | 35.7785 | 0.51354
| 2 | 209.942 | 40 | 8.9749 | 39.5811 | 0.11337 | 194 | 15 | -6.967 | 13.8443 | 0.25162 | 192.06 | 81.48 | -8.907 | 81.2752 | 0.054796 | 221 | 40 | 20.033 | 39.5811 | 0.25306 | 209.52 | 48.015 | 8.553 | 47.6666 | 0.089717 | 212 | 26 | 11.033 | 25.3508 | 0.21761 | 203.7 | 46.56 | 2.733 | 46.2006 | 0.029577 | 198 | 24 | -2.967 | 23.2952 | 0.063684 | 224.07 | 98.94 | 23.103 | 98.7714 | 0.11695 | [5 empty] | 139.68 | 72.75 | -61.287 | 72.5205 | 0.42255 | 179.838 | 37.83 | -21.129 | 37.3868 | 0.28257 | 200.967 | 7.6378 | 15.2756 | 5 | 203.7 | 14.55 | 2.733 | 13.3554 | 0.10232 | 183.3333 | 35.7909 | -17.6337 | 35.3222 | 0.24961 | 234 | 38 | 33.033 | 37.5588 | 0.43975
| 1 | 212.8774 | 40 | 2.1515 | 39.5733 | 0.027183 | 201 | 15 | -9.726 | 13.8219 | 0.35183 | 200.79 | 81.48 | -9.936 | 81.2714 | 0.061128 | 236 | 40 | 25.274 | 39.5733 | 0.31933 | 215.34 | 48.015 | 4.614 | 47.6601 | 0.048406 | 227 | 29 | 16.274 | 28.4085 | 0.28643 | 203.7 | 46.56 | -7.026 | 46.1939 | 0.076048 | 205 | 24 | -5.726 | 23.2819 | 0.12297 | 279.36 | 197.88 | 68.634 | 197.7942 | 0.1735 | [10 empty] | 178.965 | 37.83 | -31.761 | 37.3785 | 0.42486 | 210.726 | 8.3639 | 16.7278 | 6 | 218.25 | 17.46 | 7.524 | 16.4589 | 0.22857 | [5 empty] | 235 | 40 | 24.274 | 39.5733 | 0.3067
10/5 | 120 | [5 empty] | 69 | 15 | -8.594 | 13.9142 | 0.30882 | 90.21 | 40.74 | 12.616 | 40.3528 | 0.15632 | [15 empty] | 87.3 | 20.37 | 9.706 | 19.5842 | 0.2478 | [25 empty] | 77.594 | 8.2095 | 16.4191 | 6 | 78.57 | 11.64 | 0.97596 | 10.2026 | 0.047829
| 100 | 84.5064 | 30 | -0.74304 | 29.687 | 0.012515 | 75 | 15 | -10.2494 | 14.3637 | 0.35678 | 90.21 | 40.74 | 4.9606 | 40.5101 | 0.061226 | 99 | 30 | 13.7506 | 29.687 | 0.23159 | [10 empty] | 87.3 | 20.37 | 2.0506 | 19.9061 | 0.051506 | [15 empty] | 119.31 | 40.74 | 34.0606 | 40.5101 | 0.4204 | [5 empty] | 85.2494 | 7.3948 | 14.7896 | 6 | 81.48 | 11.64 | -3.7694 | 10.8077 | 0.17439 | [5 empty] | 134 | 40 | 48.7506 | 39.7658 | 0.61297
| 50 | 103.465 | 30 | 3.7996 | 29.9989 | 0.063329 | 95 | 15 | -4.6654 | 14.9978 | 0.15554 | 98.94 | 40.74 | -0.72542 | 40.7392 | 0.0089033 | 94 | 30 | -5.6654 | 29.9989 | 0.094427 | 104.76 | 27.645 | 5.0946 | 27.6438 | 0.092147 | 119 | 26 | 19.3346 | 25.9987 | 0.37184 | 87.3 | 20.37 | -12.3654 | 20.3683 | 0.30355 | 102 | 24 | 2.3346 | 23.9986 | 0.04864 | 101.85 | 40.74 | 2.1846 | 40.7392 | 0.026812 | 270.2323 | 68.121 | 170.5669 | 68.1205 | 1.2519 | 116.4 | 40.74 | 16.7346 | 40.7392 | 0.20539 | 70.713 | 17.169 | -28.9524 | 17.167 | 0.84326 | 99.6654 | 6.0056 | 12.0112 | 6 | 101.85 | 11.64 | 2.1846 | 11.6371 | 0.093863 | 130.6667 | 37.6832 | 31.0012 | 37.6823 | 0.41135 | 141 | 40 | 41.3346 | 39.9992 | 0.51669
| 20 | 125.2682 | 30 | -0.45359 | 29.8166 | 0.0076064 | 117 | 15 | -8.7218 | 14.6298 | 0.29808 | 116.4 | 40.74 | -9.3218 | 40.6052 | 0.11479 | 126 | 30 | 0.27818 | 29.8166 | 0.0046649 | 151.32 | 27.645 | 25.5982 | 27.4459 | 0.46634 | 140 | 26 | 14.2782 | 25.7882 | 0.27684 | 145.5 | 20.37 | 19.7782 | 20.099 | 0.49202 | 123 | 24 | -2.7218 | 23.7704 | 0.057252 | 125.13 | 40.74 | -0.59182 | 40.6052 | 0.0072875 | 250.6771 | 68.8389 | 124.9553 | 68.7592 | 0.90864 | 125.13 | 40.74 | -0.59182 | 40.6052 | 0.0072875 | 94.284 | 17.169 | -31.4378 | 16.8465 | 0.93307 | 125.7218 | 5.9974 | 11.9948 | 5 | 128.04 | 11.64 | 2.3182 | 11.1589 | 0.10387 | 115.5 | 35.7005 | -10.2218 | 35.5466 | 0.14378 | 156 | 40 | 30.2782 | 39.8627 | 0.37978
| 10 | 138.2541 | 35 | 3.137 | 34.83 | 0.045033 | 130 | 15 | -5.1171 | 14.5989 | 0.17526 | 128.04 | 58.2 | -7.0771 | 58.0979 | 0.060906 | 144 | 30 | 8.8829 | 29.8015 | 0.14903 | 142.59 | 26.19 | 7.4729 | 25.9624 | 0.14392 | 151 | 26 | 15.8829 | 25.7707 | 0.30816 | 145.5 | 20.37 | 10.3829 | 20.0765 | 0.25858 | 134 | 24 | -1.1171 | 23.7514 | 0.023516 | 142.59 | 46.56 | 7.4729 | 46.4323 | 0.080471 | 251.8847 | 69.3717 | 116.7677 | 69.2861 | 0.84265 | 133.86 | 40.74 | -1.2571 | 40.594 | 0.015484 | 112.908 | 17.169 | -22.2091 | 16.8197 | 0.66021 | 135.1171 | 6.0722 | 12.1444 | 5 | 133.86 | 11.64 | -1.2571 | 11.1184 | 0.056532 | 123.4167 | 36.2966 | -11.7004 | 36.1327 | 0.16191 | 173 | 42 | 37.8829 | 41.8584 | 0.45251
| 5 | 148.9847 | 35 | 1.6479 | 34.6026 | 0.023812 | 141 | 15 | -6.3368 | 14.0477 | 0.22555 | 139.68 | 81.48 | -7.6568 | 81.3101 | 0.047084 | 162 | 30 | 14.6632 | 29.5354 | 0.24823 | 165.87 | 40.74 | 18.5332 | 40.3991 | 0.22938 | 160 | 26 | 12.6632 | 25.4625 | 0.24866 | 145.5 | 32.01 | -1.8368 | 31.575 | 0.029086 | 146 | 24 | -1.3368 | 23.4166 | 0.028544 | 151.32 | 58.2 | 3.9832 | 57.9619 | 0.034361 | 266.944 | 83.7801 | 119.6072 | 83.6148 | 0.71523 | 136.77 | 40.74 | -10.5668 | 40.3991 | 0.13078 | 126.294 | 37.83 | -21.0428 | 37.4626 | 0.28085 | 147.3368 | 7.2569 | 14.5138 | 5 | 142.59 | 14.55 | -4.7468 | 13.5661 | 0.17495 | 131.6667 | 37.9738 | -15.6701 | 37.6078 | 0.20834 | 190 | 44 | 42.6632 | 43.6845 | 0.48831
| 2 | 159.3371 | 40 | 3.8154 | 39.6876 | 0.048068 | 152 | 15 | -3.5217 | 14.1457 | 0.12448 | 151.32 | 81.48 | -4.2017 | 81.3271 | 0.025832 | 181 | 40 | 25.4783 | 39.6876 | 0.32099 | 174.6 | 49.47 | 19.0783 | 49.2177 | 0.19382 | 172 | 26 | 16.4783 | 25.5167 | 0.32289 | 145.5 | 46.56 | -10.0217 | 46.2918 | 0.10824 | 156 | 24 | 0.47832 | 23.4756 | 0.010188 | 189.15 | 98.94 | 33.6283 | 98.8141 | 0.17016 | [5 empty] | 107.67 | 72.75 | -47.8517 | 72.5787 | 0.32965 | 135.897 | 37.83 | -19.6247 | 37.4995 | 0.26167 | 155.5217 | 7.8037 | 15.6075 | 6 | 157.14 | 14.55 | 1.6183 | 13.6676 | 0.059203 | 131.4167 | 35.8757 | -24.105 | 35.527 | 0.33925
| 1 | 165.3171 | 40 | 1.6911 | 39.5504 | 0.021379 | 158 | 15 | -5.626 | 13.7564 | 0.20449 | 160.05 | 81.48 | -3.576 | 81.2603 | 0.022003 | 194 | 40 | 30.374 | 39.5504 | 0.38399 | 183.33 | 49.47 | 19.704 | 49.1072 | 0.20062 | 169 | 26 | 5.374 | 25.3029 | 0.10619 | 145.5 | 46.56 | -18.126 | 46.1744 | 0.19628 | 162 | 24 | -1.626 | 23.243 | 0.034978 | 247.35 | 197.88 | 83.724 | 197.7896 | 0.21165 | [10 empty] | 137.352 | 37.83 | -26.274 | 37.3543 | 0.35169 | 163.626 | 8.4712 | 16.9425 | 6 | 168.78 | 17.46 | 5.154 | 16.404 | 0.1571
Annex 4 TAB A4-1 Calculation from results of all labs.
Meaning of individual symbols:
| I1 | Primary current (%)IN
| kI | Transformation ratio I (kA/A)
| B | Burden (VA)
| I1N | Rated primary current (A)
| dL | Phase displacement of individual laboratory (mrad)
| u(dL) | Standard uncertainty of phase displacement of individual laboratory (mrad)
| dr | CRV for phase displacement calculated from results of all laboratories see eqv. (2) (mrad)
| u(dr) | CRV uncertainty calculated from results of all laboratories see eqv. (3) (mrad)
| Dd | Difference between results of individual laboratory dL and dr see eqv. (5) (mrad)
| u(Ddstd) | The uncertainty of the difference Dd according to (8) with transfer standard uncertaity (mrad)
| u(dstd) | Pilot laboratory evaluated uncertainty component of phase displacement of transfer standard (µrad)
| E(d) | Confidence coefficient of phase displacement calculated from results of all laboratories see eqv. (9)
red and underlined results are for E(d) > 1.0
5 VA |  | CMI | CMI | CMI | CMI | CMI | PTB | PTB | PTB | PTB | PTB | METAS | METAS | METAS | METAS | METAS | NPL | NPL | NPL | NPL | NPL | LCOE | LCOE | LCOE | LCOE | LCOE | MIKES | MIKES | MIKES | MIKES | MIKES | BEV | BEV | BEV | BEV | BEV | VSL | VSL | VSL | VSL | VSL | SP | SP | SP | SP | SP | DMDM | DMDM | DMDM | DMDM | DMDM | GUM | GUM | GUM | GUM | GUM | BIM | BIM | BIM | BIM | BIM | weighted mean drC (mrad) | u(drC) (mrad) | U(drC) (mrad) | u(dstd) (mrad) | UME | UME | UME | UME | UME | LNE | LNE | LNE | LNE | LNE | INRIM | INRIM | INRIM | INRIM | INRIM
kI (kA/A) | I % | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) |  |  |  |  | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-)
4/5 | 120 | 159.1334 | 30 | 3.7938 | 31.1796 | 0.060839 | 144 | 10 | -11.3395 | 13.1212 | 0.4321 | 165.87 | 20.37 | 10.5305 | 22.0704 | 0.23857 | 144 | 15 | -11.3395 | 17.2385 | 0.3289 | 224.07 | 91.665 | 68.7305 | 92.0578 | 0.3733 | 165.87 | 26.19 | 10.5305 | 27.5333 | 0.19123 | 145.5 | 20.37 | -9.8395 | 22.0704 | 0.22291 | 156 | 20 | 0.6605 | 21.7294 | 0.015198 | [5 empty] | 215.1605 | 46.4616 | 59.821 | 47.2319 | 0.63327 | 177.51 | 40.74 | 22.1705 | 41.6163 | 0.26637 | 119.892 | 13.386 | -35.4475 | 17.5219 | 1.0115 | 155.3395 | 5.2757 | 10.5514 | 10 | 177.51 | 11.64 | 22.1705 | 14.4103 | 0.76926 | 131.5819 | 18.1683 | -23.7576 | 20.0563 | 0.59227 | 229 | 36 | 73.6605 | 37.7337 | 0.97606
| 100 | 168.8431 | 30 | 10.6307 | 30.6231 | 0.17357 | 151 | 10 | -7.2124 | 11.7378 | 0.30723 | 174.6 | 20.37 | 16.3876 | 21.277 | 0.3851 | 153 | 15 | -5.2124 | 16.2104 | 0.16077 | 235.71 | 91.665 | 77.4976 | 91.8708 | 0.42178 | 177.51 | 26.19 | 19.2976 | 26.9015 | 0.35867 | 174.6 | 20.37 | 16.3876 | 21.277 | 0.3851 | 166 | 20 | 7.7876 | 20.9231 | 0.1861 | 154.23 | 10.185 | -3.9824 | 11.8958 | 0.16739 | 228.7018 | 46.5304 | 70.4894 | 46.9345 | 0.75093 | 171.69 | 40.74 | 13.4776 | 41.201 | 0.16356 | 127.458 | 13.386 | -30.7544 | 16.4137 | 0.93685 | 158.2124 | 5.121 | 10.242 | 8 | 192.06 | 11.64 | 33.8476 | 15.0238 | 1.1265 | 139.8533 | 18.3329 | -18.3591 | 19.3357 | 0.47475 | 245 | 36 | 86.7876 | 37.232 | 1.1655
| 50 | 208.6034 | 30 | 10.9296 | 30.6522 | 0.17828 | 190 | 10 | -7.6737 | 11.8135 | 0.32479 | 206.61 | 20.37 | 8.9363 | 21.3189 | 0.20959 | 201 | 15 | 3.3263 | 16.2653 | 0.10225 | 215.34 | 23.28 | 17.6663 | 24.1147 | 0.3663 | 215.34 | 26.19 | 17.6663 | 26.9347 | 0.32795 | 203.7 | 20.37 | 6.0263 | 21.3189 | 0.14134 | 204 | 19 | 6.3263 | 20.014 | 0.15805 | 192.06 | 10.185 | -5.6137 | 11.9705 | 0.23448 | 259.2034 | 46.5609 | 61.5297 | 46.9838 | 0.6548 | 206.61 | 40.74 | 8.9363 | 41.2227 | 0.10839 | 163.251 | 13.386 | -34.4227 | 16.3593 | 1.0521 | 197.6737 | 4.9437 | 9.8874 | 8 | 226.98 | 11.64 | 29.3063 | 14.9643 | 0.97921 | 167.1198 | 18.2824 | -30.554 | 19.3341 | 0.79016 | 262 | 36 | 64.3263 | 36.5453 | 0.88009
| 20 | 248.3976 | 30 | 9.3667 | 30.1037 | 0.15557 | 232 | 10 | -7.0309 | 10.3069 | 0.34108 | 241.53 | 20.37 | 2.4991 | 20.5224 | 0.060887 | 252 | 15 | 12.9691 | 15.2063 | 0.42644 | 247.35 | 23.28 | 8.3191 | 23.4135 | 0.17766 | 256.08 | 26.19 | 17.0491 | 26.3087 | 0.32402 | 232.8 | 20.37 | -6.2309 | 20.5224 | 0.15181 | 248 | 19 | 8.9691 | 19.1633 | 0.23402 | 235.71 | 10.185 | -3.3209 | 10.4865 | 0.15834 | 292.4162 | 46.5266 | 53.3853 | 46.5936 | 0.57288 | 244.44 | 40.74 | 5.4091 | 40.8164 | 0.066261 | 213.303 | 13.386 | -25.7279 | 13.6168 | 0.94471 | 239.0309 | 4.3322 | 8.6645 | 5 | 253.17 | 11.64 | 14.1391 | 11.9047 | 0.59385 | 217.9166 | 20.4324 | -21.1144 | 20.5844 | 0.51287 | 291 | 36 | 51.9691 | 36.0864 | 0.72006
| 10 | 269.8637 | 35 | 4.4919 | 35.0522 | 0.064075 | 257 | 10 | -8.3718 | 10.1811 | 0.41115 | 264.81 | 29.1 | -0.56179 | 29.1627 | 0.0096321 | 283 | 15 | 17.6282 | 15.1213 | 0.58289 | 288.09 | 42.195 | 22.7182 | 42.2383 | 0.26893 | 276.45 | 26.19 | 11.0782 | 26.2597 | 0.21094 | 291 | 20.37 | 25.6282 | 20.4595 | 0.62632 | 271 | 19 | 5.6282 | 19.0959 | 0.14737 | 256.08 | 11.64 | -9.2918 | 11.7959 | 0.39386 | 314.7505 | 46.5251 | 49.3787 | 46.5644 | 0.53022 | 256.08 | 40.74 | -9.2918 | 40.7848 | 0.11391 | 240.948 | 13.386 | -24.4238 | 13.5218 | 0.90313 | 265.3718 | 4.6202 | 9.2404 | 5 | 276.45 | 11.64 | 11.0782 | 11.7959 | 0.46958 | 231.5622 | 31.1953 | -33.8096 | 31.2538 | 0.54089 | 313 | 36 | 47.6282 | 36.0507 | 0.66057
| 5 | 289.7875 | 35 | 3.962 | 34.9675 | 0.056653 | 275 | 10 | -10.8255 | 9.8857 | 0.54753 | 282.27 | 40.74 | -3.5555 | 40.7121 | 0.043666 | 309 | 15 | 23.1745 | 14.9241 | 0.77641 | 320.1 | 48.015 | 34.2745 | 47.9913 | 0.35709 | 299.73 | 26.19 | 13.9045 | 26.1466 | 0.2659 | 291 | 32.01 | 5.1745 | 31.9745 | 0.080916 | 294 | 19 | 8.1745 | 18.9401 | 0.2158 | 279.36 | 14.55 | -6.4655 | 14.4717 | 0.22338 | 331.3229 | 55.7707 | 45.4974 | 55.7503 | 0.40805 | 258.99 | 40.74 | -26.8355 | 40.7121 | 0.32958 | 262.191 | 17.751 | -23.6345 | 17.6869 | 0.66814 | 285.8255 | 5.2223 | 10.4445 | 5 | 296.82 | 14.55 | 10.9945 | 14.4717 | 0.37986 | 245.2277 | 34.2977 | -40.5978 | 34.2645 | 0.59242 | 326 | 36 | 40.1745 | 35.9684 | 0.55847
| 2 | 304.7425 | 40 | -1.2892 | 39.8931 | 0.016158 | 295 | 10 | -11.0317 | 9.5637 | 0.57675 | 299.73 | 40.74 | -6.3017 | 40.6351 | 0.07754 | 336 | 20 | 29.9683 | 19.7854 | 0.75733 | 328.83 | 43.65 | 22.7983 | 43.5521 | 0.26174 | 317.19 | 26.19 | 11.1583 | 26.0265 | 0.21436 | 320.1 | 46.56 | 14.0683 | 46.4682 | 0.15138 | 313 | 19 | 6.9683 | 18.774 | 0.18558 | 325.92 | 24.735 | 19.8883 | 24.5618 | 0.40486 | [5 empty] | 215.34 | 75.66 | -90.6917 | 75.6036 | 0.59978 | 278.196 | 17.751 | -27.8357 | 17.5089 | 0.7949 | 306.0317 | 5.7911 | 11.5822 | 5 | 317.19 | 14.55 | 11.1583 | 14.2536 | 0.39142 | 264.2293 | 45.7554 | -41.8024 | 45.662 | 0.45774 | 342 | 38 | 35.9683 | 37.8875 | 0.47467
| 1 | 315.9314 | 40 | -0.7507 | 39.8372 | 0.0094221 | 305 | 10 | -11.6821 | 9.3274 | 0.62623 | 308.46 | 40.74 | -8.2221 | 40.5801 | 0.10131 | 353 | 20 | 36.3179 | 19.6723 | 0.92307 | 337.56 | 43.65 | 20.8779 | 43.5008 | 0.23997 | 320.1 | 26.19 | 3.4179 | 25.9406 | 0.065879 | 320.1 | 46.56 | 3.4179 | 46.4202 | 0.036815 | 321 | 19 | 4.3179 | 18.6548 | 0.11573 | 378.3 | 58.2 | 61.6179 | 58.0882 | 0.53038 | [10 empty] | 286.635 | 17.751 | -30.0471 | 17.381 | 0.86437 | 316.6821 | 6.1644 | 12.3288 | 5 | 334.65 | 17.46 | 17.9679 | 17.0837 | 0.52588 | [5 empty] | 349 | 40 | 32.3179 | 39.8372 | 0.40562
10/5 | 120 | [5 empty] | 67 | 7.5 | -4.9281 | 6.9205 | 0.35605 | 81.48 | 20.37 | 9.5519 | 20.1638 | 0.23686 | [15 empty] | 87.3 | 20.37 | 15.3719 | 20.1638 | 0.38118 | [25 empty] | 71.9281 | 5.7756 | 11.5512 | 5 | 75.66 | 11.64 | 3.7319 | 11.2753 | 0.16549
| 100 | 81.1163 | 30 | 1.8877 | 29.9633 | 0.031499 | 71 | 7.5 | -8.2286 | 7.3517 | 0.55964 | 81.48 | 20.37 | 2.2514 | 20.3159 | 0.05541 | 98 | 15 | 18.7714 | 14.9264 | 0.6288 | [10 empty] | 87.3 | 20.37 | 8.0714 | 20.3159 | 0.19865 | [15 empty] | 107.67 | 40.74 | 28.4414 | 40.713 | 0.34929 | [5 empty] | 79.2286 | 5.2156 | 10.4312 | 5 | 78.57 | 11.64 | -0.6586 | 11.545 | 0.028523 | [5 empty] | 117 | 40 | 37.7714 | 39.9725 | 0.47247
| 50 | 97.5723 | 30 | 8.0295 | 30.057 | 0.13357 | 89 | 7.5 | -0.54277 | 7.725 | 0.035131 | 90.21 | 20.37 | 0.66723 | 20.4539 | 0.01631 | 89 | 15 | -0.54277 | 15.1138 | 0.017956 | 96.03 | 29.1 | 6.4872 | 29.1588 | 0.11124 | 113.49 | 26.19 | 23.9472 | 26.2553 | 0.45605 | 87.3 | 20.37 | -2.2428 | 20.4539 | 0.054825 | 93 | 19 | 3.4572 | 19.0899 | 0.090551 | [5 empty] | 258.859 | 68.1058 | 169.3163 | 68.4469 | 1.2368 | 107.67 | 40.74 | 18.1272 | 40.782 | 0.22225 | 64.311 | 17.169 | -25.2318 | 17.2685 | 0.73057 | 89.5428 | 4.6448 | 9.2896 | 5 | 90.21 | 11.64 | 0.66723 | 11.7862 | 0.028305 | 95.723 | 35.5998 | 6.1802 | 35.6479 | 0.086684 | 121 | 40 | 31.4572 | 40.0428 | 0.3928
| 20 | 114.9014 | 30 | 3.3064 | 29.9686 | 0.055165 | 107 | 7.5 | -4.5949 | 7.3734 | 0.31159 | 104.76 | 20.37 | -6.8349 | 20.3237 | 0.16815 | 112 | 15 | 0.40507 | 14.9371 | 0.013559 | 142.59 | 33.465 | 30.9951 | 33.4368 | 0.46349 | 128.04 | 26.19 | 16.4451 | 26.154 | 0.31439 | 116.4 | 20.37 | 4.8051 | 20.3237 | 0.11821 | 110 | 19 | -1.5949 | 18.9504 | 0.042082 | 119.31 | 10.185 | 7.7151 | 10.0921 | 0.38223 | 241.8889 | 68.3626 | 130.294 | 68.3488 | 0.95315 | 116.4 | 40.74 | 4.8051 | 40.7169 | 0.059006 | 84.972 | 17.169 | -26.6229 | 17.1141 | 0.77781 | 111.5949 | 4.2289 | 8.4578 | 4 | 113.49 | 11.64 | 1.8951 | 11.5588 | 0.081975 | 100.2406 | 35.618 | -11.3543 | 35.5915 | 0.15951 | 134 | 40 | 22.4051 | 39.9764 | 0.28023
| 10 | 125.6393 | 35 | 1.6594 | 34.9555 | 0.023736 | 119 | 7.5 | -4.9798 | 7.2896 | 0.34157 | 113.49 | 29.1 | -10.4898 | 29.0465 | 0.18057 | 129 | 15 | 5.0202 | 14.8959 | 0.16851 | 128.04 | 29.1 | 4.0602 | 29.0465 | 0.069891 | 139.68 | 26.19 | 15.7002 | 26.1305 | 0.30042 | 145.5 | 20.37 | 21.5202 | 20.2935 | 0.53022 | 121 | 19 | -2.9798 | 18.9179 | 0.078756 | 133.86 | 11.64 | 9.8802 | 11.5055 | 0.42937 | 254.1594 | 69.247 | 130.1796 | 69.2245 | 0.94027 | 119.31 | 40.74 | -4.6698 | 40.7018 | 0.057366 | 101.268 | 17.169 | -22.7118 | 17.0781 | 0.66494 | 123.9798 | 4.3718 | 8.7436 | 4 | 122.22 | 11.64 | -1.7598 | 11.5055 | 0.076477 | 108.7163 | 36.4661 | -15.2635 | 36.4234 | 0.20953 | 142 | 42 | 18.0202 | 41.9629 | 0.21472
| 5 | 134.5147 | 35 | 0.73411 | 34.8855 | 0.010522 | 128 | 7.5 | -5.7806 | 6.9462 | 0.4161 | 125.13 | 40.74 | -8.6506 | 40.6417 | 0.10643 | 145 | 15 | 11.2194 | 14.7309 | 0.38081 | 145.5 | 39.285 | 11.7194 | 39.183 | 0.14955 | 151.32 | 26.19 | 17.5394 | 26.0368 | 0.33682 | 145.5 | 32.01 | 11.7194 | 31.8848 | 0.18378 | 132 | 19 | -1.7806 | 18.7883 | 0.047387 | 139.68 | 14.55 | 5.8994 | 14.2724 | 0.20667 | 262.0989 | 82.1776 | 128.3182 | 82.1289 | 0.7812 | 119.31 | 40.74 | -14.4706 | 40.6417 | 0.17803 | 112.326 | 37.83 | -21.4546 | 37.7241 | 0.28436 | 133.7806 | 4.899 | 9.7981 | 4 | 130.95 | 14.55 | -2.8306 | 14.2724 | 0.099165 | 115.6842 | 36.2812 | -18.0964 | 36.1707 | 0.25015 | 176 | 44 | 42.2194 | 43.909 | 0.48076
| 2 | 143.2011 | 40 | 0.69229 | 39.8366 | 0.0086891 | 137 | 7.5 | -5.5088 | 6.5732 | 0.41903 | 133.86 | 40.74 | -8.6488 | 40.5796 | 0.10657 | 163 | 20 | 20.4912 | 19.6712 | 0.52084 | 148.41 | 48.015 | 5.9012 | 47.879 | 0.061626 | 160.05 | 26.19 | 17.5412 | 25.9398 | 0.33811 | 145.5 | 46.56 | 2.9912 | 46.4197 | 0.032219 | 140 | 19 | -2.5088 | 18.6536 | 0.067247 | 180.42 | 24.735 | 37.9112 | 24.4699 | 0.77465 | [5 empty] | 93.12 | 75.66 | -49.3888 | 75.5738 | 0.32676 | 120.183 | 37.83 | -22.3258 | 37.6572 | 0.29643 | 142.5088 | 5.3891 | 10.7783 | 4 | 142.59 | 14.55 | 0.081189 | 14.0947 | 0.0028801 | 126.4312 | 41.8072 | -16.0776 | 41.6509 | 0.193
| 1 | 147.6534 | 40 | 0.058197 | 39.788 | 0.00073134 | 142 | 7.5 | -5.5952 | 6.2715 | 0.44608 | 139.68 | 40.74 | -7.9152 | 40.5318 | 0.097642 | 174 | 20 | 26.4048 | 19.5725 | 0.67454 | 154.23 | 48.015 | 6.6348 | 47.8385 | 0.069346 | 174.6 | 34.92 | 27.0048 | 34.6769 | 0.38938 | 145.5 | 46.56 | -2.0952 | 46.378 | 0.022588 | 145 | 19 | -2.5952 | 18.5494 | 0.069954 | 238.62 | 49.47 | 91.0248 | 49.2987 | 0.9232 | [10 empty] | 121.056 | 37.83 | -26.5392 | 37.6057 | 0.35286 | 147.5952 | 5.7375 | 11.4749 | 4 | 148.41 | 17.46 | 0.8148 | 16.9686 | 0.024009
15 VA |  | CMI | CMI | CMI | CMI | CMI | PTB | PTB | PTB | PTB | PTB | METAS | METAS | METAS | METAS | METAS | NPL | NPL | NPL | NPL | NPL | LCOE | LCOE | LCOE | LCOE | LCOE | MIKES | MIKES | MIKES | MIKES | MIKES | BEV | BEV | BEV | BEV | BEV | VSL | VSL | VSL | VSL | VSL | SP | SP | SP | SP | SP | DMDM | DMDM | DMDM | DMDM | DMDM | GUM | GUM | GUM | GUM | GUM | BIM | BIM | BIM | BIM | BIM | weighted mean drC (mrad) | u(drC) (mrad) | U(drC) (mrad) | u(dstd) (mrad) | UME | UME | UME | UME | UME | LNE | LNE | LNE | LNE | LNE | INRIM | INRIM | INRIM | INRIM | INRIM
kI (kA/A) | I % | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) |  |  |  |  | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-) | dL (mrad) | u(dL) (mrad) | DdC (mrad) | u(DdstdC) (mrad) | EC(d) (-)
4/5 | 120 | 183.912 | 30 | -9.4086 | 29.5844 | 0.15901 | 165 | 20 | -28.3206 | 19.371 | 0.73101 | 206.61 | 40.74 | 13.2894 | 40.4349 | 0.16433 | 175 | 30 | -18.3206 | 29.5844 | 0.30963 | 235.71 | 27.645 | 42.3894 | 27.1934 | 0.77941 | 198 | 26 | 4.6794 | 25.5193 | 0.091683 | 174.6 | 20.37 | -18.7206 | 19.7528 | 0.47387 | 187 | 27 | -6.3206 | 26.5374 | 0.11909 | [5 empty] | 264.7809 | 47.4736 | 71.4603 | 47.212 | 0.7568 | 247.35 | 40.74 | 54.0294 | 40.4349 | 0.6681 | 146.664 | 13.386 | -46.6566 | 16.6118 | 1.4043 | 193.3206 | 7.7952 | 15.5904 | 6 | 226.98 | 11.64 | 33.6594 | 15.2399 | 1.1043 | 186.35 | 19.0877 | -6.9706 | 18.4275 | 0.18914 | 245 | 36 | 51.6794 | 35.6544 | 0.72473
| 100 | 205.5479 | 30 | 1.9787 | 29.3597 | 0.033697 | 180 | 20 | -23.5692 | 19.0261 | 0.61939 | 218.25 | 40.74 | 14.6808 | 40.2709 | 0.18228 | 186 | 30 | -17.5692 | 29.3597 | 0.29921 | 258.99 | 23.28 | 55.4208 | 25.0991 | 1.104 | 212 | 26 | 8.4308 | 25.2585 | 0.16689 | 203.7 | 20.37 | 0.13081 | 19.4147 | 0.0033689 | 199 | 26 | -4.5692 | 25.2585 | 0.090448 | 168.78 | 40.74 | -34.7892 | 40.2709 | 0.43194 | 284.4719 | 47.8189 | 80.9027 | 47.4199 | 0.85305 | 241.53 | 40.74 | 37.9608 | 40.2709 | 0.47132 | 158.886 | 13.386 | -44.6832 | 16.346 | 1.3668 | 203.5692 | 7.9376 | 15.8753 | 5 | 247.35 | 11.64 | 43.7808 | 14.9498 | 1.4643 | 196.25 | 19.0199 | -7.3192 | 17.9931 | 0.20339 | 261 | 36 | 57.4308 | 35.4682 | 0.80961
| 50 | 262.1473 | 30 | 8.0364 | 29.3679 | 0.13682 | 236 | 20 | -18.1109 | 19.0387 | 0.47563 | 264.81 | 40.74 | 10.6991 | 40.2768 | 0.13282 | 255 | 30 | 0.88907 | 29.3679 | 0.015137 | 328.83 | 93.12 | 74.7191 | 92.9183 | 0.40207 | 268 | 26 | 13.8891 | 25.268 | 0.27484 | 261.9 | 20.37 | 7.7891 | 19.427 | 0.20047 | 255 | 26 | 0.88907 | 25.268 | 0.017593 | 232.8 | 40.74 | -21.3109 | 40.2768 | 0.26456 | 336.4348 | 47.927 | 82.3239 | 47.5339 | 0.86595 | 267.72 | 40.74 | 13.6091 | 40.2768 | 0.16894 | 212.721 | 13.386 | -41.3899 | 16.3313 | 1.2672 | 254.1109 | 7.9074 | 15.8149 | 5 | 282.27 | 11.64 | 28.1591 | 14.9338 | 0.9428 | 222.75 | 18.9922 | -31.3609 | 17.9771 | 0.87224 | 300 | 36 | 45.8891 | 35.475 | 0.64678
| 20 | 322.5735 | 30 | -0.35949 | 29.7428 | 0.0060433 | 299 | 20 | -23.933 | 19.6121 | 0.61016 | 320.1 | 40.74 | -2.833 | 40.551 | 0.034931 | 320 | 30 | -2.933 | 29.7428 | 0.049306 | 320.1 | 27.645 | -2.833 | 27.3657 | 0.051762 | 329 | 26 | 6.067 | 25.7028 | 0.11802 | 320.1 | 20.37 | -2.833 | 19.9893 | 0.070863 | 320 | 24 | -2.933 | 23.6777 | 0.061936 | 296.82 | 40.74 | -26.113 | 40.551 | 0.32198 | 392.1419 | 47.1843 | 69.2089 | 47.0212 | 0.73593 | 323.01 | 40.74 | 0.077009 | 40.551 | 0.00094953 | 284.598 | 13.677 | -38.335 | 15.8879 | 1.2064 | 322.933 | 6.3534 | 12.7068 | 5 | 328.83 | 11.64 | 5.897 | 10.9601 | 0.26902 | 318 | 19.2154 | -4.933 | 18.8113 | 0.13112 | 357 | 36 | 34.067 | 35.7859 | 0.47598
| 10 | 358.3083 | 35 | -3.8921 | 34.6855 | 0.056105 | 335 | 20 | -27.2004 | 19.4444 | 0.69944 | 355.02 | 58.2 | -7.1804 | 58.0114 | 0.061887 | 366 | 30 | 3.7996 | 29.6325 | 0.064113 | 389.94 | 59.655 | 27.7396 | 59.471 | 0.23322 | 361 | 26 | -1.2004 | 25.5751 | 0.023467 | 378.3 | 20.37 | 16.0996 | 19.8248 | 0.40605 | 354 | 24 | -8.2004 | 23.539 | 0.17419 | 334.65 | 46.56 | -27.5504 | 46.3241 | 0.29737 | 433.9295 | 47.195 | 71.7291 | 46.9623 | 0.76369 | 349.2 | 40.74 | -13.0004 | 40.4701 | 0.16062 | 325.047 | 13.677 | -37.1534 | 16.0927 | 1.1544 | 362.2004 | 6.8495 | 13.699 | 5 | 366.66 | 11.64 | 4.4596 | 10.6571 | 0.20923 | 338.25 | 31.2995 | -23.9504 | 30.9474 | 0.38695 | 386 | 36 | 23.7996 | 35.6943 | 0.33338
| 5 | 385.8514 | 35 | 4.9972 | 34.6376 | 0.072135 | 364 | 20 | -16.8543 | 19.3589 | 0.43531 | 381.21 | 81.48 | 0.35572 | 81.325 | 0.002187 | 403 | 30 | 22.1457 | 29.5764 | 0.37438 | 404.49 | 46.56 | 23.6357 | 46.2882 | 0.25531 | 387 | 26 | 6.1457 | 25.5101 | 0.12046 | 378.3 | 32.01 | -2.5543 | 31.6134 | 0.040399 | 386 | 24 | 5.1457 | 23.4684 | 0.10963 | 366.66 | 58.2 | -14.1943 | 57.9828 | 0.1224 | 456.1085 | 56.3368 | 75.2543 | 56.1124 | 0.67057 | 360.84 | 40.74 | -20.0143 | 40.4291 | 0.24752 | 355.893 | 18.042 | -24.9613 | 17.3286 | 0.72023 | 380.8543 | 7.0876 | 14.1752 | 5 | 392.85 | 14.55 | 11.9957 | 13.6554 | 0.43923 | 353.5 | 33.4055 | -27.3543 | 33.0256 | 0.41414 | 412 | 36 | 31.1457 | 35.6478 | 0.43685
| 2 | 412.4925 | 40 | 3.7266 | 39.723 | 0.046908 | 393 | 20 | -15.7659 | 19.4401 | 0.4055 | 407.4 | 81.48 | -1.3659 | 81.3444 | 0.0083955 | 441 | 40 | 32.2341 | 39.723 | 0.40574 | 427.77 | 45.105 | 19.0041 | 44.8595 | 0.21182 | 410 | 26 | 1.2341 | 25.5718 | 0.024131 | 407.4 | 46.56 | -1.3659 | 46.3222 | 0.014743 | 413 | 24 | 4.2341 | 23.5354 | 0.089953 | 424.86 | 98.94 | 16.0941 | 98.8283 | 0.081425 | [5 empty] | 337.56 | 75.66 | -71.2059 | 75.5139 | 0.47148 | 382.956 | 18.333 | -25.8099 | 17.7205 | 0.72825 | 408.7659 | 7.6213 | 15.2426 | 6 | 419.04 | 14.55 | 10.2741 | 13.7702 | 0.37306 | 447.9167 | 45.682 | 39.1508 | 45.4396 | 0.4308 | 437 | 38 | 28.2341 | 37.7083 | 0.37438
| 1 | 428.6866 | 40 | 7.1259 | 39.5987 | 0.089977 | 409 | 20 | -12.5607 | 19.1848 | 0.32736 | 421.95 | 81.48 | 0.38925 | 81.2837 | 0.0023944 | 463 | 40 | 41.4393 | 39.5987 | 0.52324 | 439.41 | 43.65 | 17.8493 | 43.2825 | 0.20619 | 431 | 29 | 9.4393 | 28.4439 | 0.16593 | 407.4 | 46.56 | -14.1607 | 46.2157 | 0.1532 | 427 | 24 | 5.4393 | 23.325 | 0.1166 | 474.33 | 197.88 | 52.7693 | 197.7993 | 0.13339 | [10 empty] | 394.305 | 18.333 | -27.2557 | 17.44 | 0.78141 | 421.5607 | 8.2428 | 16.4856 | 6 | 433.59 | 17.46 | 12.0293 | 16.5199 | 0.36408 | [5 empty] | 450 | 40 | 28.4393 | 39.5987 | 0.35909
5/5 | 120 | 155.605 | 30 | -10.8137 | 30.2859 | 0.17853 | 134 | 20 | -32.4187 | 20.4263 | 0.79355 | 171.69 | 40.74 | 5.2713 | 40.951 | 0.064361 | 139 | 30 | -27.4187 | 30.2859 | 0.45266 | 194.97 | 64.02 | 28.5513 | 64.1545 | 0.22252 | 160 | 26 | -6.4187 | 26.3294 | 0.12189 | 145.5 | 20.37 | -20.9187 | 20.7888 | 0.50312 | 149 | 25 | -17.4187 | 25.3424 | 0.34367 | [10 empty] | 215.34 | 40.74 | 48.9213 | 40.951 | 0.59732 | [5 empty] | 166.4187 | 6.8385 | 13.6769 | 8 | 183.33 | 11.64 | 16.9113 | 12.3582 | 0.68421 | 166.5 | 25.2303 | 0.081318 | 25.5696 | 0.0015901 | 227 | 36 | 60.5813 | 36.2386 | 0.83587
| 100 | 167.2523 | 30 | -8.5579 | 30.0852 | 0.14223 | 144 | 20 | -31.8101 | 20.1275 | 0.79021 | 180.42 | 40.74 | 4.6099 | 40.8028 | 0.056489 | 142 | 30 | -33.8101 | 30.0852 | 0.56191 | 221.16 | 81.48 | 45.3499 | 81.5114 | 0.27818 | 172 | 26 | -3.8101 | 26.0982 | 0.072996 | 174.6 | 20.37 | -1.2101 | 20.4952 | 0.029523 | 159 | 25 | -16.8101 | 25.1021 | 0.33483 | 151.32 | 40.74 | -24.4901 | 40.8028 | 0.3001 | 357.9639 | 68.2068 | 182.1538 | 68.8843 | 1.3222 | 197.88 | 40.74 | 22.0699 | 40.8028 | 0.27045 | 133.569 | 13.386 | -42.2411 | 16.4945 | 1.2805 | 175.8101 | 6.6244 | 13.2487 | 7 | 192.06 | 11.64 | 16.2499 | 11.8578 | 0.6852 | 169.5 | 20.1356 | -6.3101 | 20.2623 | 0.15571 | 240 | 36 | 64.1899 | 36.071 | 0.88977
| 50 | 208.5815 | 30 | -1.3702 | 29.7061 | 0.023063 | 188 | 20 | -21.9517 | 19.5563 | 0.56124 | 215.34 | 40.74 | 5.3883 | 40.524 | 0.066482 | 197 | 30 | -12.9517 | 29.7061 | 0.218 | 302.64 | 71.295 | 92.6883 | 71.1718 | 0.65116 | 215 | 26 | 5.0483 | 25.6603 | 0.098367 | 203.7 | 20.37 | -6.2517 | 19.9346 | 0.15681 | 201 | 24 | -8.9517 | 23.6316 | 0.1894 | 186.24 | 40.74 | -23.7117 | 40.524 | 0.29256 | 364.4436 | 68.1768 | 154.4918 | 68.6704 | 1.1249 | 200.79 | 40.74 | -9.1617 | 40.524 | 0.11304 | 162.96 | 13.386 | -46.9917 | 15.7078 | 1.4958 | 209.9517 | 6.523 | 13.046 | 5 | 221.16 | 11.64 | 11.2083 | 10.86 | 0.51603 | 200.25 | 18.2641 | -9.7017 | 17.7772 | 0.27287 | 260 | 36 | 50.0483 | 35.7554 | 0.69987
| 20 | 257.4295 | 30 | -0.90898 | 29.7284 | 0.015288 | 238 | 20 | -20.3385 | 19.5902 | 0.5191 | 256.08 | 40.74 | -2.2585 | 40.5404 | 0.027855 | 250 | 30 | -8.3385 | 29.7284 | 0.14024 | 293.91 | 39.285 | 35.5715 | 39.078 | 0.45514 | 262 | 26 | 3.6615 | 25.6861 | 0.071274 | 261.9 | 20.37 | 3.5615 | 19.9678 | 0.089181 | 250 | 24 | -8.3385 | 23.6596 | 0.17622 | 250.26 | 40.74 | -8.0785 | 40.5404 | 0.099635 | 383.1452 | 68.1732 | 124.8067 | 68.0541 | 0.91697 | 250.26 | 40.74 | -8.0785 | 40.5404 | 0.099635 | 218.541 | 13.386 | -39.7975 | 15.6655 | 1.2702 | 258.3385 | 6.4206 | 12.8412 | 5 | 258.99 | 11.64 | 0.65151 | 10.9209 | 0.029829 | 254.75 | 17.9982 | -3.5885 | 17.5417 | 0.10228 | 296 | 36 | 37.6615 | 35.774 | 0.52638
| 10 | 284.3579 | 35 | -4.7956 | 34.7334 | 0.069035 | 264 | 20 | -25.1535 | 19.5296 | 0.64398 | 282.27 | 58.2 | -6.8835 | 58.04 | 0.0593 | 284 | 30 | -5.1535 | 29.6885 | 0.086794 | 343.38 | 36.375 | 54.2265 | 36.1185 | 0.75067 | 285 | 26 | -4.1535 | 25.6399 | 0.080997 | 291 | 20.37 | 1.8465 | 19.9084 | 0.046374 | 276 | 24 | -13.1535 | 23.6095 | 0.27856 | 279.36 | 46.56 | -9.7935 | 46.3599 | 0.10563 | 390.0855 | 68.1814 | 100.932 | 68.0449 | 0.74166 | 267.72 | 40.74 | -21.4335 | 40.5112 | 0.26454 | 250.26 | 13.386 | -38.8935 | 15.741 | 1.2354 | 289.1535 | 6.6025 | 13.205 | 5 | 288.09 | 11.64 | -1.0635 | 10.8119 | 0.049184 | 304.5833 | 20.3934 | 15.4298 | 19.9324 | 0.38705 | 321 | 36 | 31.8465 | 35.7408 | 0.44552
| 5 | 305.841 | 35 | 3.3512 | 34.6957 | 0.048295 | 287 | 20 | -15.4898 | 19.4626 | 0.39794 | 302.64 | 81.48 | 0.15024 | 81.3498 | 0.00092344 | 314 | 30 | 11.5102 | 29.6445 | 0.19414 | 325.92 | 46.56 | 23.4302 | 46.3317 | 0.25285 | 308 | 26 | 5.5102 | 25.5889 | 0.10767 | 291 | 32.01 | -11.4898 | 31.677 | 0.18136 | 300 | 24 | -2.4898 | 23.5541 | 0.052852 | 328.83 | 58.2 | 26.3402 | 58.0175 | 0.227 | 388.097 | 81.7859 | 85.6072 | 81.6562 | 0.52419 | 270.63 | 40.74 | -31.8598 | 40.4789 | 0.39354 | 274.704 | 17.751 | -27.7858 | 17.1433 | 0.8104 | 302.4898 | 6.7975 | 13.5951 | 5 | 308.46 | 14.55 | 5.9702 | 13.802 | 0.21628 | 320.25 | 19.262 | 17.7602 | 18.7034 | 0.47479 | 342 | 36 | 39.5102 | 35.7043 | 0.5533
| 2 | 326.7675 | 40 | 4.5468 | 39.7325 | 0.057218 | 309 | 20 | -13.2208 | 19.4594 | 0.3397 | 323.01 | 81.48 | 0.78925 | 81.349 | 0.004851 | 345 | 40 | 22.7792 | 39.7325 | 0.28666 | 343.38 | 43.65 | 21.1592 | 43.405 | 0.24374 | 332 | 26 | 9.7792 | 25.5865 | 0.1911 | 320.1 | 46.56 | -2.1208 | 46.3304 | 0.022887 | 321 | 24 | -1.2208 | 23.5514 | 0.025917 | 384.12 | 98.94 | 61.8992 | 98.8321 | 0.31315 | [5 empty] | 232.8 | 75.66 | -89.4208 | 75.5189 | 0.59204 | 295.656 | 17.751 | -26.5648 | 17.1396 | 0.77495 | 322.2208 | 7.5718 | 15.1436 | 6 | 331.74 | 14.55 | 9.5192 | 13.7975 | 0.34496 | 357.1667 | 45.9535 | 34.9459 | 45.7208 | 0.38217 | 359 | 38 | 36.7792 | 37.7183 | 0.48755
| 1 | 337.5418 | 40 | 4.6152 | 39.6043 | 0.058266 | 322 | 20 | -10.9267 | 19.1964 | 0.2846 | 331.74 | 81.48 | -1.1867 | 81.2865 | 0.0072992 | 364 | 40 | 31.0733 | 39.6043 | 0.3923 | 355.02 | 43.65 | 22.0933 | 43.2877 | 0.25519 | 352 | 29 | 19.0733 | 28.4517 | 0.33519 | 320.1 | 46.56 | -12.8267 | 46.2205 | 0.13876 | 332 | 24 | -0.92666 | 23.3345 | 0.019856 | 462.69 | 197.88 | 129.7633 | 197.8004 | 0.32802 | [10 empty] | 303.222 | 18.042 | -29.7047 | 17.1468 | 0.86619 | 332.9267 | 8.2158 | 16.4317 | 6 | 346.29 | 17.46 | 13.3633 | 16.5333 | 0.40413 | [5 empty] | 369 | 40 | 36.0733 | 39.6043 | 0.45542
6/5 | 120 | 128.8475 | 30 | -11.3014 | 29.9234 | 0.18884 | 114 | 15 | -26.149 | 14.8462 | 0.88066 | 130.95 | 40.74 | -9.199 | 40.6836 | 0.11305 | 127 | 30 | -13.149 | 29.9234 | 0.21971 | 142.59 | 27.645 | 2.441 | 27.5618 | 0.044283 | 140 | 32 | -0.14896 | 31.9282 | 0.0023328 | 145.5 | 20.37 | 5.351 | 20.257 | 0.13208 | 125 | 24 | -15.149 | 23.9042 | 0.31687 | [10 empty] | 174.6 | 81 | 34.451 | 80.9717 | 0.21274 | [5 empty] | 140.149 | 6.3711 | 12.7422 | 6 | 148.41 | 11.64 | 8.261 | 11.4411 | 0.36102 | 156.5 | 19.4955 | 16.351 | 19.3774 | 0.42191 | 205 | 36 | 64.851 | 35.9362 | 0.90231
| 100 | 142.1135 | 30 | -2.7401 | 29.9248 | 0.045784 | 121 | 15 | -23.8536 | 14.8489 | 0.80321 | 133.86 | 40.74 | -10.9936 | 40.6846 | 0.13511 | 117 | 30 | -27.8536 | 29.9248 | 0.46539 | 203.7 | 45.105 | 58.8464 | 45.055 | 0.65305 | 146 | 26 | 1.1464 | 25.9131 | 0.02212 | 145.5 | 20.37 | 0.64637 | 20.259 | 0.015953 | 132 | 24 | -12.8536 | 23.9059 | 0.26884 | [10 empty] | 160.05 | 40.74 | 15.1964 | 40.6846 | 0.18676 | [5 empty] | 144.8536 | 6.3647 | 12.7293 | 6 | 154.23 | 11.64 | 9.3764 | 11.4447 | 0.40964 | 147.75 | 19.8221 | 2.8964 | 19.708 | 0.073482 | 211 | 36 | 66.1464 | 35.9373 | 0.9203
| 50 | 175.8331 | 30 | 4.367 | 29.7706 | 0.073344 | 156 | 15 | -15.4661 | 14.5357 | 0.532 | 157.14 | 40.74 | -14.3261 | 40.5714 | 0.17655 | 156 | 30 | -15.4661 | 29.7706 | 0.25976 | 241.53 | 53.835 | 70.0639 | 53.7075 | 0.65227 | 183 | 26 | 11.5339 | 25.7349 | 0.22409 | 174.6 | 20.37 | 3.1339 | 20.0306 | 0.078227 | 166 | 24 | -5.4661 | 23.7126 | 0.11526 | 157.14 | 40.74 | -14.3261 | 40.5714 | 0.17655 | 198.6269 | 70.4942 | 27.1608 | 70.3969 | 0.19291 | 168.78 | 40.74 | -2.6861 | 40.5714 | 0.033104 | 132.114 | 17.169 | -39.3521 | 18.9338 | 1.0392 | 171.4661 | 6.222 | 12.4439 | 5 | 180.42 | 11.64 | 8.9539 | 11.0352 | 0.40569 | 154 | 18.1575 | -17.4661 | 17.7759 | 0.49129 | 226 | 36 | 54.5339 | 35.809 | 0.76145
| 20 | 215.0999 | 30 | 7.955 | 29.8599 | 0.13321 | 196 | 15 | -11.1449 | 14.7179 | 0.37862 | 192.06 | 40.74 | -15.0849 | 40.637 | 0.18561 | 202 | 30 | -5.1449 | 29.8599 | 0.086151 | 209.52 | 29.1 | 2.3751 | 28.9556 | 0.041012 | 221 | 26 | 13.8551 | 25.8383 | 0.26811 | 203.7 | 20.37 | -3.4449 | 20.1632 | 0.085427 | 208 | 24 | 0.85506 | 23.8247 | 0.017945 | 209.52 | 40.74 | 2.3751 | 40.637 | 0.029223 | 195.7702 | 73.0901 | -11.3747 | 73.0328 | 0.077874 | 203.7 | 40.74 | -3.4449 | 40.637 | 0.042387 | 177.801 | 17.169 | -29.3439 | 16.9231 | 0.86698 | 207.1449 | 5.7779 | 11.5558 | 5 | 221.16 | 11.64 | 14.0151 | 11.2741 | 0.62156 | 208 | 18.4361 | 0.85506 | 18.2073 | 0.023481 | 253 | 36 | 45.8551 | 35.8834 | 0.63895
| 10 | 238.1144 | 35 | 6.6287 | 34.852 | 0.095098 | 219 | 15 | -12.4857 | 14.6513 | 0.42609 | 215.34 | 58.2 | -16.1457 | 58.1111 | 0.13892 | 233 | 30 | 1.5143 | 29.8272 | 0.025385 | 279.36 | 43.65 | 47.8743 | 43.5314 | 0.54988 | 242 | 26 | 10.5143 | 25.8004 | 0.20376 | 232.8 | 20.37 | 1.3143 | 20.1146 | 0.03267 | 229 | 24 | -2.4857 | 23.7836 | 0.052256 | 226.98 | 46.56 | -4.5057 | 46.4488 | 0.048502 | 189.9551 | 76.1178 | -41.5306 | 76.0499 | 0.27305 | 218.25 | 40.74 | -13.2357 | 40.6129 | 0.16295 | 206.028 | 17.46 | -25.4577 | 17.1614 | 0.74171 | 231.4857 | 5.9446 | 11.8892 | 5 | 235.71 | 11.64 | 4.2243 | 11.1871 | 0.1888 | 250.5 | 18.8583 | 19.0143 | 18.5821 | 0.51163 | 270 | 36 | 38.5143 | 35.8561 | 0.53707
| 5 | 255.1488 | 35 | 5.0405 | 34.6282 | 0.072781 | 237 | 15 | -13.1083 | 14.1106 | 0.46448 | 235.71 | 81.48 | -14.3983 | 81.321 | 0.088527 | 257 | 30 | 6.8917 | 29.5653 | 0.11655 | 261.9 | 39.285 | 11.7917 | 38.9541 | 0.15135 | 259 | 26 | 8.8917 | 25.4972 | 0.17437 | 261.9 | 32.01 | 11.7917 | 31.603 | 0.18656 | 248 | 24 | -2.1083 | 23.4544 | 0.044944 | 247.35 | 58.2 | -2.7583 | 57.9771 | 0.023788 | 195.004 | 90.4677 | -55.1043 | 90.3244 | 0.30504 | 215.34 | 40.74 | -34.7683 | 40.421 | 0.43008 | 226.107 | 37.83 | -24.0013 | 37.4862 | 0.32013 | 250.1083 | 7.1338 | 14.2676 | 5 | 253.17 | 14.55 | 3.0617 | 13.6313 | 0.11231 | 270.9167 | 30.5116 | 20.8084 | 30.0843 | 0.34583 | 287 | 36 | 36.8917 | 35.6386 | 0.51758
| 2 | 271.6703 | 40 | 3.5765 | 39.7066 | 0.045037 | 256 | 15 | -12.0938 | 14.1991 | 0.42586 | 253.17 | 81.48 | -14.9238 | 81.3364 | 0.091741 | 286 | 40 | 17.9062 | 39.7066 | 0.22548 | 282.27 | 48.015 | 14.1762 | 47.7709 | 0.14838 | 276 | 26 | 7.9062 | 25.5463 | 0.15474 | 261.9 | 46.56 | -6.1938 | 46.3082 | 0.066876 | 264 | 24 | -4.0938 | 23.5078 | 0.087073 | 296.82 | 98.94 | 28.7262 | 98.8218 | 0.14534 | [5 empty] | 186.24 | 72.75 | -81.8538 | 72.5891 | 0.56382 | 244.149 | 37.83 | -23.9448 | 37.5196 | 0.3191 | 268.0938 | 7.7062 | 15.4123 | 6 | 273.54 | 14.55 | 5.4462 | 13.7229 | 0.19844 | 298.9583 | 45.748 | 30.8646 | 45.4917 | 0.33923 | 302 | 38 | 33.9062 | 37.691 | 0.44979
| 1 | 277.7195 | 40 | 0.83953 | 39.5909 | 0.010603 | 266 | 15 | -10.88 | 13.8722 | 0.39215 | 267.72 | 81.48 | -9.16 | 81.2799 | 0.056348 | 303 | 40 | 26.12 | 39.5909 | 0.32987 | 291 | 48.015 | 14.12 | 47.6747 | 0.14809 | 285 | 26 | 8.12 | 25.3661 | 0.16006 | 261.9 | 46.56 | -14.98 | 46.209 | 0.16209 | 274 | 24 | -2.88 | 23.3117 | 0.06177 | 357.93 | 197.88 | 81.05 | 197.7977 | 0.20488 | [10 empty] | 247.059 | 37.83 | -29.821 | 37.3971 | 0.39871 | 276.88 | 8.2803 | 16.5606 | 6 | 285.18 | 17.46 | 8.3 | 16.5012 | 0.2515 | [5 empty] | 306 | 40 | 29.12 | 39.5909 | 0.36776
8/5 | 120 | 104.1162 | 30 | 0.47655 | 30.1621 | 0.0078999 | 86 | 15 | -17.6396 | 15.3216 | 0.57564 | 110.58 | 40.74 | 6.9404 | 40.8595 | 0.08493 | 113 | 30 | 9.3604 | 30.1621 | 0.15517 | [10 empty] | 87.3 | 20.37 | -16.3396 | 20.608 | 0.39644 | [15 empty] | 139.68 | 40.74 | 36.0404 | 40.8595 | 0.44103 | [5 empty] | 103.6396 | 7.3653 | 14.7306 | 8 | 107.67 | 11.64 | 4.0304 | 12.0516 | 0.16721 | [5 empty] | 170 | 36 | 66.3604 | 36.1352 | 0.91822
| 100 | 113.2936 | 30 | 5.9048 | 30.3471 | 0.097288 | 93 | 15 | -14.3888 | 15.6826 | 0.45875 | 110.58 | 40.74 | 3.1912 | 40.9963 | 0.038921 | 97 | 30 | -10.3888 | 30.3471 | 0.17117 | 119.31 | 26.19 | 11.9212 | 26.5869 | 0.22419 | 125 | 32 | 17.6112 | 32.3256 | 0.2724 | 87.3 | 20.37 | -20.0888 | 20.8778 | 0.4811 | 104 | 24 | -3.3888 | 24.4325 | 0.06935 | [10 empty] | 130.95 | 40.74 | 23.5612 | 40.9963 | 0.28736 | [5 empty] | 107.3888 | 6.5616 | 13.1232 | 8 | 107.67 | 11.64 | 0.28121 | 12.5074 | 0.011242 | 133.0833 | 40.1069 | 25.6945 | 40.3672 | 0.31826 | 178 | 36 | 70.6112 | 36.2897 | 0.97288
| 50 | 137.3884 | 30 | 10.4162 | 30.1915 | 0.1725 | 119 | 15 | -7.9721 | 15.3794 | 0.25918 | 122.22 | 40.74 | -4.7521 | 40.8812 | 0.058121 | 124 | 30 | -2.9721 | 30.1915 | 0.049222 | 189.15 | 62.565 | 62.1779 | 62.657 | 0.49618 | 143 | 26 | 16.0279 | 26.2207 | 0.30563 | 145.5 | 20.37 | 18.5279 | 20.651 | 0.44859 | 127 | 24 | 0.027857 | 24.2389 | 0.00057463 | 116.4 | 40.74 | -10.5721 | 40.8812 | 0.1293 | 274.3306 | 68.3366 | 147.3584 | 68.9664 | 1.0683 | 133.86 | 40.74 | 6.8879 | 40.8812 | 0.084242 | 96.03 | 17.169 | -30.9421 | 17.5015 | 0.88399 | 126.9721 | 6.1215 | 12.2431 | 7 | 128.04 | 11.64 | 1.0679 | 12.125 | 0.044035 | 130.25 | 35.6702 | 3.2779 | 35.8314 | 0.04574 | 182 | 36 | 55.0279 | 36.1597 | 0.7609
| 20 | 166.8521 | 30 | 8.4328 | 30.0068 | 0.14052 | 149 | 15 | -9.4193 | 15.0136 | 0.31369 | 142.59 | 40.74 | -15.8293 | 40.745 | 0.19425 | 156 | 30 | -2.4193 | 30.0068 | 0.040313 | 160.05 | 26.19 | 1.6307 | 26.1978 | 0.031123 | 169 | 26 | 10.5807 | 26.0078 | 0.20341 | 174.6 | 20.37 | 16.1807 | 20.38 | 0.39698 | 157 | 24 | -1.4193 | 24.0085 | 0.029558 | 151.32 | 40.74 | -7.0993 | 40.745 | 0.087119 | 257.0403 | 69.6031 | 98.621 | 69.606 | 0.70842 | 160.05 | 40.74 | 1.6307 | 40.745 | 0.020011 | 128.331 | 17.169 | -30.0883 | 17.1809 | 0.87563 | 158.4193 | 5.966 | 11.932 | 6 | 165.87 | 11.64 | 7.4507 | 11.6575 | 0.31957 | 145.25 | 35.6481 | -13.1693 | 35.6538 | 0.18468 | 198 | 36 | 39.5807 | 36.0057 | 0.54965
| 10 | 183.6246 | 35 | 11.676 | 34.9918 | 0.16684 | 166 | 15 | -5.9486 | 14.9809 | 0.19854 | 162.96 | 58.2 | -8.9886 | 58.1951 | 0.077228 | 177 | 30 | 5.0514 | 29.9904 | 0.084216 | 174.6 | 26.19 | 2.6514 | 26.179 | 0.050639 | 186 | 26 | 14.0514 | 25.989 | 0.27033 | 174.6 | 20.37 | 2.6514 | 20.3559 | 0.065125 | 173 | 24 | 1.0514 | 23.988 | 0.021914 | 171.69 | 46.56 | -0.25864 | 46.5538 | 0.0027778 | 246.0163 | 71.2437 | 74.0676 | 71.2397 | 0.51985 | 168.78 | 40.74 | -3.1686 | 40.733 | 0.038895 | 150.738 | 17.169 | -21.2106 | 17.1523 | 0.6183 | 171.9486 | 6.0476 | 12.0952 | 6 | 174.6 | 11.64 | 2.6514 | 11.6153 | 0.11413 | 158.25 | 35.6614 | -13.6986 | 35.6534 | 0.19211 | 213 | 36 | 41.0514 | 35.992 | 0.57028
| 5 | 196.876 | 35 | 9.6234 | 34.7721 | 0.13838 | 180 | 15 | -7.2527 | 14.4603 | 0.25078 | 177.51 | 81.48 | -9.7427 | 81.3824 | 0.059857 | 198 | 30 | 10.7473 | 29.7338 | 0.18073 | 194.97 | 39.285 | 7.7173 | 39.0821 | 0.098732 | 198 | 26 | 10.7473 | 25.6924 | 0.20915 | 174.6 | 32.01 | -12.6527 | 31.7607 | 0.19919 | 186 | 24 | -1.2527 | 23.6664 | 0.026465 | 180.42 | 58.2 | -6.8327 | 58.0632 | 0.058838 | 246.4624 | 86.3136 | 59.2098 | 86.2214 | 0.34336 | 165.87 | 40.74 | -21.3827 | 40.5444 | 0.26369 | 166.743 | 37.83 | -20.5097 | 37.6193 | 0.2726 | 187.2527 | 7.2041 | 14.4083 | 6 | 189.15 | 14.55 | 1.8973 | 13.993 | 0.067796 | 174.75 | 38.3432 | -12.5027 | 38.1353 | 0.16393 | 224 | 36 | 36.7473 | 35.7785 | 0.51354
| 2 | 209.942 | 40 | 8.9749 | 39.5811 | 0.11337 | 194 | 15 | -6.967 | 13.8443 | 0.25162 | 192.06 | 81.48 | -8.907 | 81.2752 | 0.054796 | 221 | 40 | 20.033 | 39.5811 | 0.25306 | 209.52 | 48.015 | 8.553 | 47.6666 | 0.089717 | 212 | 26 | 11.033 | 25.3508 | 0.21761 | 203.7 | 46.56 | 2.733 | 46.2006 | 0.029577 | 198 | 24 | -2.967 | 23.2952 | 0.063684 | 224.07 | 98.94 | 23.103 | 98.7714 | 0.11695 | [5 empty] | 139.68 | 72.75 | -61.287 | 72.5205 | 0.42255 | 179.838 | 37.83 | -21.129 | 37.3868 | 0.28257 | 200.967 | 7.6378 | 15.2756 | 5 | 203.7 | 14.55 | 2.733 | 13.3554 | 0.10232 | 183.3333 | 35.7909 | -17.6337 | 35.3222 | 0.24961 | 234 | 38 | 33.033 | 37.5588 | 0.43975
| 1 | 212.8774 | 40 | 2.1515 | 39.5733 | 0.027183 | 201 | 15 | -9.726 | 13.8219 | 0.35183 | 200.79 | 81.48 | -9.936 | 81.2714 | 0.061128 | 236 | 40 | 25.274 | 39.5733 | 0.31933 | 215.34 | 48.015 | 4.614 | 47.6601 | 0.048406 | 227 | 29 | 16.274 | 28.4085 | 0.28643 | 203.7 | 46.56 | -7.026 | 46.1939 | 0.076048 | 205 | 24 | -5.726 | 23.2819 | 0.12297 | 279.36 | 197.88 | 68.634 | 197.7942 | 0.1735 | [10 empty] | 178.965 | 37.83 | -31.761 | 37.3785 | 0.42486 | 210.726 | 8.3639 | 16.7278 | 6 | 218.25 | 17.46 | 7.524 | 16.4589 | 0.22857 | [5 empty] | 235 | 40 | 24.274 | 39.5733 | 0.3067
10/5 | 120 | [5 empty] | 69 | 15 | -8.594 | 13.9142 | 0.30882 | 90.21 | 40.74 | 12.616 | 40.3528 | 0.15632 | [15 empty] | 87.3 | 20.37 | 9.706 | 19.5842 | 0.2478 | [25 empty] | 77.594 | 8.2095 | 16.4191 | 6 | 78.57 | 11.64 | 0.97596 | 10.2026 | 0.047829
| 100 | 84.5064 | 30 | -0.74304 | 29.687 | 0.012515 | 75 | 15 | -10.2494 | 14.3637 | 0.35678 | 90.21 | 40.74 | 4.9606 | 40.5101 | 0.061226 | 99 | 30 | 13.7506 | 29.687 | 0.23159 | [10 empty] | 87.3 | 20.37 | 2.0506 | 19.9061 | 0.051506 | [15 empty] | 119.31 | 40.74 | 34.0606 | 40.5101 | 0.4204 | [5 empty] | 85.2494 | 7.3948 | 14.7896 | 6 | 81.48 | 11.64 | -3.7694 | 10.8077 | 0.17439 | [5 empty] | 134 | 40 | 48.7506 | 39.7658 | 0.61297
| 50 | 103.465 | 30 | 5.1357 | 29.9942 | 0.085612 | 95 | 15 | -3.3293 | 14.9883 | 0.11106 | 98.94 | 40.74 | 0.61067 | 40.7357 | 0.0074955 | 94 | 30 | -4.3293 | 29.9942 | 0.07217 | 104.76 | 27.645 | 6.4307 | 27.6387 | 0.11633 | 119 | 26 | 20.6707 | 25.9933 | 0.39762 | 87.3 | 20.37 | -11.0293 | 20.3614 | 0.27084 | 102 | 24 | 3.6707 | 23.9927 | 0.076496 | 101.85 | 40.74 | 3.5207 | 40.7357 | 0.043214 | 270.2323 | 68.121 | 171.903 | 68.65 | 1.252 | 116.4 | 40.74 | 18.0707 | 40.7357 | 0.2218 | 70.713 | 17.169 | -27.6163 | 17.1588 | 0.80473 | 98.3293 | 6.0291 | 12.0582 | 6 | 101.85 | 11.64 | 3.5207 | 11.625 | 0.15143 | 130.6667 | 37.6832 | 32.3373 | 37.6786 | 0.42912 | 141 | 40 | 42.6707 | 39.9956 | 0.53344
| 20 | 125.2682 | 30 | -0.45359 | 29.8166 | 0.0076064 | 117 | 15 | -8.7218 | 14.6298 | 0.29808 | 116.4 | 40.74 | -9.3218 | 40.6052 | 0.11479 | 126 | 30 | 0.27818 | 29.8166 | 0.0046649 | 151.32 | 27.645 | 25.5982 | 27.4459 | 0.46634 | 140 | 26 | 14.2782 | 25.7882 | 0.27684 | 145.5 | 20.37 | 19.7782 | 20.099 | 0.49202 | 123 | 24 | -2.7218 | 23.7704 | 0.057252 | 125.13 | 40.74 | -0.59182 | 40.6052 | 0.0072875 | 250.6771 | 68.8389 | 124.9553 | 68.7592 | 0.90864 | 125.13 | 40.74 | -0.59182 | 40.6052 | 0.0072875 | 94.284 | 17.169 | -31.4378 | 16.8465 | 0.93307 | 125.7218 | 5.9974 | 11.9948 | 5 | 128.04 | 11.64 | 2.3182 | 11.1589 | 0.10387 | 115.5 | 35.7005 | -10.2218 | 35.5466 | 0.14378 | 156 | 40 | 30.2782 | 39.8627 | 0.37978
| 10 | 138.2541 | 35 | 3.137 | 34.83 | 0.045033 | 130 | 15 | -5.1171 | 14.5989 | 0.17526 | 128.04 | 58.2 | -7.0771 | 58.0979 | 0.060906 | 144 | 30 | 8.8829 | 29.8015 | 0.14903 | 142.59 | 26.19 | 7.4729 | 25.9624 | 0.14392 | 151 | 26 | 15.8829 | 25.7707 | 0.30816 | 145.5 | 20.37 | 10.3829 | 20.0765 | 0.25858 | 134 | 24 | -1.1171 | 23.7514 | 0.023516 | 142.59 | 46.56 | 7.4729 | 46.4323 | 0.080471 | 251.8847 | 69.3717 | 116.7677 | 69.2861 | 0.84265 | 133.86 | 40.74 | -1.2571 | 40.594 | 0.015484 | 112.908 | 17.169 | -22.2091 | 16.8197 | 0.66021 | 135.1171 | 6.0722 | 12.1444 | 5 | 133.86 | 11.64 | -1.2571 | 11.1184 | 0.056532 | 123.4167 | 36.2966 | -11.7004 | 36.1327 | 0.16191 | 173 | 42 | 37.8829 | 41.8584 | 0.45251
| 5 | 148.9847 | 35 | 1.6479 | 34.6026 | 0.023812 | 141 | 15 | -6.3368 | 14.0477 | 0.22555 | 139.68 | 81.48 | -7.6568 | 81.3101 | 0.047084 | 162 | 30 | 14.6632 | 29.5354 | 0.24823 | 165.87 | 40.74 | 18.5332 | 40.3991 | 0.22938 | 160 | 26 | 12.6632 | 25.4625 | 0.24866 | 145.5 | 32.01 | -1.8368 | 31.575 | 0.029086 | 146 | 24 | -1.3368 | 23.4166 | 0.028544 | 151.32 | 58.2 | 3.9832 | 57.9619 | 0.034361 | 266.944 | 83.7801 | 119.6072 | 83.6148 | 0.71523 | 136.77 | 40.74 | -10.5668 | 40.3991 | 0.13078 | 126.294 | 37.83 | -21.0428 | 37.4626 | 0.28085 | 147.3368 | 7.2569 | 14.5138 | 5 | 142.59 | 14.55 | -4.7468 | 13.5661 | 0.17495 | 131.6667 | 37.9738 | -15.6701 | 37.6078 | 0.20834 | 190 | 44 | 42.6632 | 43.6845 | 0.48831
| 2 | 159.3371 | 40 | 3.8154 | 39.6876 | 0.048068 | 152 | 15 | -3.5217 | 14.1457 | 0.12448 | 151.32 | 81.48 | -4.2017 | 81.3271 | 0.025832 | 181 | 40 | 25.4783 | 39.6876 | 0.32099 | 174.6 | 49.47 | 19.0783 | 49.2177 | 0.19382 | 172 | 26 | 16.4783 | 25.5167 | 0.32289 | 145.5 | 46.56 | -10.0217 | 46.2918 | 0.10824 | 156 | 24 | 0.47832 | 23.4756 | 0.010188 | 189.15 | 98.94 | 33.6283 | 98.8141 | 0.17016 | [5 empty] | 107.67 | 72.75 | -47.8517 | 72.5787 | 0.32965 | 135.897 | 37.83 | -19.6247 | 37.4995 | 0.26167 | 155.5217 | 7.8037 | 15.6075 | 6 | 157.14 | 14.55 | 1.6183 | 13.6676 | 0.059203 | 131.4167 | 35.8757 | -24.105 | 35.527 | 0.33925
| 1 | 165.3171 | 40 | 1.6911 | 39.5504 | 0.021379 | 158 | 15 | -5.626 | 13.7564 | 0.20449 | 160.05 | 81.48 | -3.576 | 81.2603 | 0.022003 | 194 | 40 | 30.374 | 39.5504 | 0.38399 | 183.33 | 49.47 | 19.704 | 49.1072 | 0.20062 | 169 | 26 | 5.374 | 25.3029 | 0.10619 | 145.5 | 46.56 | -18.126 | 46.1744 | 0.19628 | 162 | 24 | -1.626 | 23.243 | 0.034978 | 247.35 | 197.88 | 83.724 | 197.7896 | 0.21165 | [10 empty] | 137.352 | 37.83 | -26.274 | 37.3543 | 0.35169 | 163.626 | 8.4712 | 16.9425 | 6 | 168.78 | 17.46 | 5.154 | 16.404 | 0.1571
Annex 4 TAB A4-2. Calculation without outliers.
Meaning of individual symbols:
| I1 | Primary current (%)IN
| kI | Transformation ratio I (kA/A)
| B | Burden (VA)
| I1N | Rated primary current (A)
| dL | Phase displacement of each participating laboratory (mrad)
| u(dL) | Standard uncertainty of phase displacement of each participating laboratory (mrad)
| drC | CRVC for phase displacement according to (2) without outliers (mrad)
| u(drC) | CRVC uncertainty according to (3) without outliers (mrad)
| DdC | Difference between results of individual laboratory dL and drC without outliers, see eq. (10) (mrad)
| u(DdstdC) | The uncertainty of the difference DdC without outliers according to (11) with transfer standard uncertaity (mrad)
| EC(d) | Corrected confidence coefficient whitout outliers, accoring to (12)
5 VA |  |  |  | CMI | CMI |  |  |  | PTB | PTB |  |  |  | METAS | METAS |  |  |  | NPL | NPL |  |  |  | LCOE | LCOE |  |  |  | MIKES | MIKES |  |  |  | BEV | BEV |  |  |  | VSL | VSL |  |  |  | SP | SP |  |  |  | DMDM | DMDM |  |  |  | GUM | GUM |  |  |  | BIM | BIM |  |  |  |  | values for graphs |  |  | UME | UME |  |  |  | LNE | LNE |  |  |  | INRIM | INRIM
kI (kA/A) | I % |  |  | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) |  | -U(drC) (mrad) | U(drC) (mrad) |  | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad) | DdC (mrad) | U(DdstdC) (mrad) | novalue | DdC (mrad) | U(DdstdC) (mrad)
4/5 | 120 |  |  | 3.7938 | 62.3592 | #N/A | 3.7938 | 62.3592 | -11.3395 | 26.2425 | #N/A | -11.3395 | 26.2425 | 10.5305 | 44.1409 | #N/A | 10.5305 | 44.1409 | -11.3395 | 34.4771 | #N/A | -11.3395 | 34.4771 | 68.7305 | 184.1156 | #N/A | 68.7305 | 184.1156 | 10.5305 | 55.0666 | #N/A | 10.5305 | 55.0666 | -9.8395 | 44.1409 | #N/A | -9.8395 | 44.1409 | 0.6605 | 43.4588 | #N/A | 0.6605 | 43.4588 |  |  | 0 | #N/A | #N/A | 59.821 | 94.4637 | #N/A | 59.821 | 94.4637 | 22.1705 | 83.2326 | #N/A | 22.1705 | 83.2326 | -35.4475 | 35.0439 | #N/A | -35.4475 | 35.0439 |  | -10.5514 | 10.5514 |  | 22.1705 | 28.8206 | #N/A | 22.1705 | 28.8206 | -23.7576 | 40.1126 | #N/A | -23.7576 | 40.1126 | 73.6605 | 75.4674 | #N/A | 73.6605 | 75.4674
| 100 |  |  | 10.6307 | 61.2462 | #N/A | 10.6307 | 61.2462 | -7.2124 | 23.4756 | #N/A | -7.2124 | 23.4756 | 16.3876 | 42.5541 | #N/A | 16.3876 | 42.5541 | -5.2124 | 32.4207 | #N/A | -5.2124 | 32.4207 | 77.4976 | 183.7416 | #N/A | 77.4976 | 183.7416 | 19.2976 | 53.803 | #N/A | 19.2976 | 53.803 | 16.3876 | 42.5541 | #N/A | 16.3876 | 42.5541 | 7.7876 | 41.8462 | #N/A | 7.7876 | 41.8462 | -3.9824 | 23.7916 | #N/A | -3.9824 | 23.7916 | 70.4894 | 93.8691 | #N/A | 70.4894 | 93.8691 | 13.4776 | 82.402 | #N/A | 13.4776 | 82.402 | -30.7544 | 32.8274 | #N/A | -30.7544 | 32.8274 |  | -10.242 | 10.242 |  | 33.8476 | 30.0476 | #N/A | 33.8476 | 30.0476 | -18.3591 | 38.6714 | #N/A | -18.3591 | 38.6714 | 86.7876 | 74.4641 | #N/A | 86.7876 | 74.4641
| 50 |  |  | 10.9296 | 61.3045 | #N/A | 10.9296 | 61.3045 | -7.6737 | 23.6271 | #N/A | -7.6737 | 23.6271 | 8.9363 | 42.6378 | #N/A | 8.9363 | 42.6378 | 3.3263 | 32.5306 | #N/A | 3.3263 | 32.5306 | 17.6663 | 48.2294 | #N/A | 17.6663 | 48.2294 | 17.6663 | 53.8693 | #N/A | 17.6663 | 53.8693 | 6.0263 | 42.6378 | #N/A | 6.0263 | 42.6378 | 6.3263 | 40.028 | #N/A | 6.3263 | 40.028 | -5.6137 | 23.9411 | #N/A | -5.6137 | 23.9411 | 61.5297 | 93.9677 | #N/A | 61.5297 | 93.9677 | 8.9363 | 82.4453 | #N/A | 8.9363 | 82.4453 | -34.4227 | 32.7185 | #N/A | -34.4227 | 32.7185 |  | -9.8874 | 9.8874 |  | 29.3063 | 29.9286 | #N/A | 29.3063 | 29.9286 | -30.554 | 38.6681 | #N/A | -30.554 | 38.6681 | 64.3263 | 73.0906 | #N/A | 64.3263 | 73.0906
| 20 |  |  | 9.3667 | 60.2074 | #N/A | 9.3667 | 60.2074 | -7.0309 | 20.6138 | #N/A | -7.0309 | 20.6138 | 2.4991 | 41.0448 | #N/A | 2.4991 | 41.0448 | 12.9691 | 30.4126 | #N/A | 12.9691 | 30.4126 | 8.3191 | 46.8269 | #N/A | 8.3191 | 46.8269 | 17.0491 | 52.6174 | #N/A | 17.0491 | 52.6174 | -6.2309 | 41.0448 | #N/A | -6.2309 | 41.0448 | 8.9691 | 38.3266 | #N/A | 8.9691 | 38.3266 | -3.3209 | 20.9729 | #N/A | -3.3209 | 20.9729 | 53.3853 | 93.1871 | #N/A | 53.3853 | 93.1871 | 5.4091 | 81.6328 | #N/A | 5.4091 | 81.6328 | -25.7279 | 27.2336 | #N/A | -25.7279 | 27.2336 |  | -8.6645 | 8.6645 |  | 14.1391 | 23.8094 | #N/A | 14.1391 | 23.8094 | -21.1144 | 41.1688 | #N/A | -21.1144 | 41.1688 | 51.9691 | 72.1729 | #N/A | 51.9691 | 72.1729
| 10 |  |  | 4.4919 | 70.1043 | #N/A | 4.4919 | 70.1043 | -8.3718 | 20.3621 | #N/A | -8.3718 | 20.3621 | -0.56179 | 58.3254 | #N/A | -0.56179 | 58.3254 | 17.6282 | 30.2426 | #N/A | 17.6282 | 30.2426 | 22.7182 | 84.4766 | #N/A | 22.7182 | 84.4766 | 11.0782 | 52.5193 | #N/A | 11.0782 | 52.5193 | 25.6282 | 40.919 | #N/A | 25.6282 | 40.919 | 5.6282 | 38.1918 | #N/A | 5.6282 | 38.1918 | -9.2918 | 23.5918 | #N/A | -9.2918 | 23.5918 | 49.3787 | 93.1288 | #N/A | 49.3787 | 93.1288 | -9.2918 | 81.5696 | #N/A | -9.2918 | 81.5696 | -24.4238 | 27.0436 | #N/A | -24.4238 | 27.0436 |  | -9.2404 | 9.2404 |  | 11.0782 | 23.5918 | #N/A | 11.0782 | 23.5918 | -33.8096 | 62.5076 | #N/A | -33.8096 | 62.5076 | 47.6282 | 72.1014 | #N/A | 47.6282 | 72.1014
| 5 |  |  | 3.962 | 69.9351 | #N/A | 3.962 | 69.9351 | -10.8255 | 19.7715 | #N/A | -10.8255 | 19.7715 | -3.5555 | 81.4242 | #N/A | -3.5555 | 81.4242 | 23.1745 | 29.8481 | #N/A | 23.1745 | 29.8481 | 34.2745 | 95.9827 | #N/A | 34.2745 | 95.9827 | 13.9045 | 52.2932 | #N/A | 13.9045 | 52.2932 | 5.1745 | 63.949 | #N/A | 5.1745 | 63.949 | 8.1745 | 37.8802 | #N/A | 8.1745 | 37.8802 | -6.4655 | 28.9434 | #N/A | -6.4655 | 28.9434 | 45.4974 | 111.5007 | #N/A | 45.4974 | 111.5007 | -26.8355 | 81.4242 | #N/A | -26.8355 | 81.4242 | -23.6345 | 35.3738 | #N/A | -23.6345 | 35.3738 |  | -10.4445 | 10.4445 |  | 10.9945 | 28.9434 | #N/A | 10.9945 | 28.9434 | -40.5978 | 68.5291 | #N/A | -40.5978 | 68.5291 | 40.1745 | 71.9369 | #N/A | 40.1745 | 71.9369
| 2 |  |  | -1.2892 | 79.7863 | #N/A | -1.2892 | 79.7863 | -11.0317 | 19.1273 | #N/A | -11.0317 | 19.1273 | -6.3017 | 81.2702 | #N/A | -6.3017 | 81.2702 | 29.9683 | 39.5709 | #N/A | 29.9683 | 39.5709 | 22.7983 | 87.1042 | #N/A | 22.7983 | 87.1042 | 11.1583 | 52.053 | #N/A | 11.1583 | 52.053 | 14.0683 | 92.9365 | #N/A | 14.0683 | 92.9365 | 6.9683 | 37.548 | #N/A | 6.9683 | 37.548 | 19.8883 | 49.1237 | #N/A | 19.8883 | 49.1237 |  |  | 0 | #N/A | #N/A | -90.6917 | 151.2071 | #N/A | -90.6917 | 151.2071 | -27.8357 | 35.0178 | #N/A | -27.8357 | 35.0178 |  | -11.5822 | 11.5822 |  | 11.1583 | 28.5073 | #N/A | 11.1583 | 28.5073 | -41.8024 | 91.3241 | #N/A | -41.8024 | 91.3241 | 35.9683 | 75.775 | #N/A | 35.9683 | 75.775
| 1 |  |  | -0.7507 | 79.6743 | #N/A | -0.7507 | 79.6743 | -11.6821 | 18.6548 | #N/A | -11.6821 | 18.6548 | -8.2221 | 81.1603 | #N/A | -8.2221 | 81.1603 | 36.3179 | 39.3446 | #N/A | 36.3179 | 39.3446 | 20.8779 | 87.0017 | #N/A | 20.8779 | 87.0017 | 3.4179 | 51.8813 | #N/A | 3.4179 | 51.8813 | 3.4179 | 92.8404 | #N/A | 3.4179 | 92.8404 | 4.3179 | 37.3095 | #N/A | 4.3179 | 37.3095 | 61.6179 | 116.1764 | #N/A | 61.6179 | 116.1764 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | -30.0471 | 34.7619 | #N/A | -30.0471 | 34.7619 |  | -12.3288 | 12.3288 |  | 17.9679 | 34.1673 | #N/A | 17.9679 | 34.1673 |  |  | 0 | #N/A | #N/A | 32.3179 | 79.6743 | #N/A | 32.3179 | 79.6743
10/5 | 120 |  |  |  |  | 0 | #N/A | #N/A | -4.9281 | 13.8409 | #N/A | -4.9281 | 13.8409 | 9.5519 | 40.3276 | #N/A | 9.5519 | 40.3276 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 15.3719 | 40.3276 | #N/A | 15.3719 | 40.3276 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  | -11.5512 | 11.5512 |  | 3.7319 | 22.5506 | #N/A | 3.7319 | 22.5506 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A
| 100 |  |  | 1.8877 | 59.9265 | #N/A | 1.8877 | 59.9265 | -8.2286 | 14.7034 | #N/A | -8.2286 | 14.7034 | 2.2514 | 40.6317 | #N/A | 2.2514 | 40.6317 | 18.7714 | 29.8528 | #N/A | 18.7714 | 29.8528 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 8.0714 | 40.6317 | #N/A | 8.0714 | 40.6317 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 28.4414 | 81.4259 | #N/A | 28.4414 | 81.4259 |  |  | 0 | #N/A | #N/A |  | -10.4312 | 10.4312 |  | -0.6586 | 23.09 | #N/A | -0.6586 | 23.09 |  |  | 0 | #N/A | #N/A | 37.7714 | 79.9449 | #N/A | 37.7714 | 79.9449
| 50 |  |  | 8.0295 | 60.1141 | #N/A | 8.0295 | 60.1141 | -0.54277 | 15.45 | #N/A | -0.54277 | 15.45 | 0.66723 | 40.9078 | #N/A | 0.66723 | 40.9078 | -0.54277 | 30.2275 | #N/A | -0.54277 | 30.2275 | 6.4872 | 58.3176 | #N/A | 6.4872 | 58.3176 | 23.9472 | 52.5106 | #N/A | 23.9472 | 52.5106 | -2.2428 | 40.9078 | #N/A | -2.2428 | 40.9078 | 3.4572 | 38.1799 | #N/A | 3.4572 | 38.1799 |  |  | 0 | #N/A | #N/A | 169.3163 | 136.8938 | #N/A | 169.3163 | 136.8938 | 18.1272 | 81.564 | #N/A | 18.1272 | 81.564 | -25.2318 | 34.537 | #N/A | -25.2318 | 34.537 |  | -9.2896 | 9.2896 |  | 0.66723 | 23.5725 | #N/A | 0.66723 | 23.5725 | 6.1802 | 71.2958 | #N/A | 6.1802 | 71.2958 | 31.4572 | 80.0856 | #N/A | 31.4572 | 80.0856
| 20 |  |  | 3.3064 | 59.9372 | #N/A | 3.3064 | 59.9372 | -4.5949 | 14.7467 | #N/A | -4.5949 | 14.7467 | -6.8349 | 40.6474 | #N/A | -6.8349 | 40.6474 | 0.40507 | 29.8742 | #N/A | 0.40507 | 29.8742 | 30.9951 | 66.8737 | #N/A | 30.9951 | 66.8737 | 16.4451 | 52.308 | #N/A | 16.4451 | 52.308 | 4.8051 | 40.6474 | #N/A | 4.8051 | 40.6474 | -1.5949 | 37.9007 | #N/A | -1.5949 | 37.9007 | 7.7151 | 20.1842 | #N/A | 7.7151 | 20.1842 | 130.294 | 136.6976 | #N/A | 130.294 | 136.6976 | 4.8051 | 81.4338 | #N/A | 4.8051 | 81.4338 | -26.6229 | 34.2281 | #N/A | -26.6229 | 34.2281 |  | -8.4578 | 8.4578 |  | 1.8951 | 23.1176 | #N/A | 1.8951 | 23.1176 | -11.3543 | 71.1831 | #N/A | -11.3543 | 71.1831 | 22.4051 | 79.9529 | #N/A | 22.4051 | 79.9529
| 10 |  |  | 1.6594 | 69.911 | #N/A | 1.6594 | 69.911 | -4.9798 | 14.5791 | #N/A | -4.9798 | 14.5791 | -10.4898 | 58.0929 | #N/A | -10.4898 | 58.0929 | 5.0202 | 29.7918 | #N/A | 5.0202 | 29.7918 | 4.0602 | 58.0929 | #N/A | 4.0602 | 58.0929 | 15.7002 | 52.261 | #N/A | 15.7002 | 52.261 | 21.5202 | 40.5869 | #N/A | 21.5202 | 40.5869 | -2.9798 | 37.8358 | #N/A | -2.9798 | 37.8358 | 9.8802 | 23.0111 | #N/A | 9.8802 | 23.0111 | 130.1796 | 138.449 | #N/A | 130.1796 | 138.449 | -4.6698 | 81.4036 | #N/A | -4.6698 | 81.4036 | -22.7118 | 34.1562 | #N/A | -22.7118 | 34.1562 |  | -8.7436 | 8.7436 |  | -1.7598 | 23.0111 | #N/A | -1.7598 | 23.0111 | -15.2635 | 72.8469 | #N/A | -15.2635 | 72.8469 | 18.0202 | 83.9259 | #N/A | 18.0202 | 83.9259
| 5 |  |  | 0.73411 | 69.771 | #N/A | 0.73411 | 69.771 | -5.7806 | 13.8924 | #N/A | -5.7806 | 13.8924 | -8.6506 | 81.2834 | #N/A | -8.6506 | 81.2834 | 11.2194 | 29.4618 | #N/A | 11.2194 | 29.4618 | 11.7194 | 78.3661 | #N/A | 11.7194 | 78.3661 | 17.5394 | 52.0736 | #N/A | 17.5394 | 52.0736 | 11.7194 | 63.7696 | #N/A | 11.7194 | 63.7696 | -1.7806 | 37.5766 | #N/A | -1.7806 | 37.5766 | 5.8994 | 28.5448 | #N/A | 5.8994 | 28.5448 | 128.3182 | 164.2577 | #N/A | 128.3182 | 164.2577 | -14.4706 | 81.2834 | #N/A | -14.4706 | 81.2834 | -21.4546 | 75.4482 | #N/A | -21.4546 | 75.4482 |  | -9.7981 | 9.7981 |  | -2.8306 | 28.5448 | #N/A | -2.8306 | 28.5448 | -18.0964 | 72.3415 | #N/A | -18.0964 | 72.3415 | 42.2194 | 87.818 | #N/A | 42.2194 | 87.818
| 2 |  |  | 0.69229 | 79.6733 | #N/A | 0.69229 | 79.6733 | -5.5088 | 13.1464 | #N/A | -5.5088 | 13.1464 | -8.6488 | 81.1592 | #N/A | -8.6488 | 81.1592 | 20.4912 | 39.3425 | #N/A | 20.4912 | 39.3425 | 5.9012 | 95.758 | #N/A | 5.9012 | 95.758 | 17.5412 | 51.8796 | #N/A | 17.5412 | 51.8796 | 2.9912 | 92.8394 | #N/A | 2.9912 | 92.8394 | -2.5088 | 37.3072 | #N/A | -2.5088 | 37.3072 | 37.9112 | 48.9399 | #N/A | 37.9112 | 48.9399 |  |  | 0 | #N/A | #N/A | -49.3888 | 151.1475 | #N/A | -49.3888 | 151.1475 | -22.3258 | 75.3144 | #N/A | -22.3258 | 75.3144 |  | -10.7783 | 10.7783 |  | 0.081189 | 28.1893 | #N/A | 0.081189 | 28.1893 | -16.0776 | 83.3018 | #N/A | -16.0776 | 83.3018 |  |  | 0 | #N/A | #N/A
| 1 |  |  | 0.058197 | 79.5759 | #N/A | 0.058197 | 79.5759 | -5.5952 | 12.543 | #N/A | -5.5952 | 12.543 | -7.9152 | 81.0637 | #N/A | -7.9152 | 81.0637 | 26.4048 | 39.1449 | #N/A | 26.4048 | 39.1449 | 6.6348 | 95.677 | #N/A | 6.6348 | 95.677 | 27.0048 | 69.3538 | #N/A | 27.0048 | 69.3538 | -2.0952 | 92.7559 | #N/A | -2.0952 | 92.7559 | -2.5952 | 37.0989 | #N/A | -2.5952 | 37.0989 | 91.0248 | 98.5974 | #N/A | 91.0248 | 98.5974 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | -26.5392 | 75.2114 | #N/A | -26.5392 | 75.2114 |  | -11.4749 | 11.4749 |  | 0.8148 | 33.9372 | #N/A | 0.8148 | 33.9372 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A
15 VA |  |  |  | CMI | CMI |  |  |  | PTB | PTB |  |  |  | METAS | METAS |  |  |  | NPL | NPL |  |  |  | LCOE | LCOE |  |  |  | MIKES | MIKES |  |  |  | BEV | BEV |  |  |  | VSL | VSL |  |  |  | SP | SP |  |  |  | DMDM | DMDM |  |  |  | GUM | GUM |  |  |  | BIM | BIM |  |  |  |  | -U(drC) (mrad) | U(drC) (mrad) |  | UME | UME |  |  |  | LNE | LNE |  |  |  | INRIM | INRIM
kI (kA/A) | I % |  |  | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad) | [7 empty] | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad) |  |  |  | DdC (mrad) | U(DdstdC) (mrad)
4/5 | 120 |  |  | -9.4086 | 59.1687 | #N/A | -9.4086 | 59.1687 | -28.3206 | 38.7419 | #N/A | -28.3206 | 38.7419 | 13.2894 | 80.8698 | #N/A | 13.2894 | 80.8698 | -18.3206 | 59.1687 | #N/A | -18.3206 | 59.1687 | 42.3894 | 54.3868 | #N/A | 42.3894 | 54.3868 | 4.6794 | 51.0386 | #N/A | 4.6794 | 51.0386 | -18.7206 | 39.5055 | #N/A | -18.7206 | 39.5055 | -6.3206 | 53.0748 | #N/A | -6.3206 | 53.0748 |  |  | 0 | #N/A | #N/A | 71.4603 | 94.4241 | #N/A | 71.4603 | 94.4241 | 54.0294 | 80.8698 | #N/A | 54.0294 | 80.8698 | -46.6566 | 33.2235 | #N/A | -46.6566 | 33.2235 |  | -15.5904 | 15.5904 |  | 33.6594 | 30.4798 | #N/A | 33.6594 | 30.4798 | -6.9706 | 36.8551 | #N/A | -6.9706 | 36.8551 | 51.6794 | 71.3088 | #N/A | 51.6794 | 71.3088
| 100 |  |  | 1.9787 | 58.7195 | #N/A | 1.9787 | 58.7195 | -23.5692 | 38.0523 | #N/A | -23.5692 | 38.0523 | 14.6808 | 80.5417 | #N/A | 14.6808 | 80.5417 | -17.5692 | 58.7195 | #N/A | -17.5692 | 58.7195 | 55.4208 | 50.1982 | #N/A | 55.4208 | 50.1982 | 8.4308 | 50.5171 | #N/A | 8.4308 | 50.5171 | 0.13081 | 38.8294 | #N/A | 0.13081 | 38.8294 | -4.5692 | 50.5171 | #N/A | -4.5692 | 50.5171 | -34.7892 | 80.5417 | #N/A | -34.7892 | 80.5417 | 80.9027 | 94.8398 | #N/A | 80.9027 | 94.8398 | 37.9608 | 80.5417 | #N/A | 37.9608 | 80.5417 | -44.6832 | 32.6919 | #N/A | -44.6832 | 32.6919 |  | -15.8753 | 15.8753 |  | 43.7808 | 29.8995 | #N/A | 43.7808 | 29.8995 | -7.3192 | 35.9861 | #N/A | -7.3192 | 35.9861 | 57.4308 | 70.9364 | #N/A | 57.4308 | 70.9364
| 50 |  |  | 8.0364 | 58.7358 | #N/A | 8.0364 | 58.7358 | -18.1109 | 38.0774 | #N/A | -18.1109 | 38.0774 | 10.6991 | 80.5536 | #N/A | 10.6991 | 80.5536 | 0.88907 | 58.7358 | #N/A | 0.88907 | 58.7358 | 74.7191 | 185.8366 | #N/A | 74.7191 | 185.8366 | 13.8891 | 50.536 | #N/A | 13.8891 | 50.536 | 7.7891 | 38.854 | #N/A | 7.7891 | 38.854 | 0.88907 | 50.536 | #N/A | 0.88907 | 50.536 | -21.3109 | 80.5536 | #N/A | -21.3109 | 80.5536 | 82.3239 | 95.0678 | #N/A | 82.3239 | 95.0678 | 13.6091 | 80.5536 | #N/A | 13.6091 | 80.5536 | -41.3899 | 32.6627 | #N/A | -41.3899 | 32.6627 |  | -15.8149 | 15.8149 |  | 28.1591 | 29.8675 | #N/A | 28.1591 | 29.8675 | -31.3609 | 35.9543 | #N/A | -31.3609 | 35.9543 | 45.8891 | 70.9499 | #N/A | 45.8891 | 70.9499
| 20 |  |  | -0.35949 | 59.4856 | #N/A | -0.35949 | 59.4856 | -23.933 | 39.2242 | #N/A | -23.933 | 39.2242 | -2.833 | 81.102 | #N/A | -2.833 | 81.102 | -2.933 | 59.4856 | #N/A | -2.933 | 59.4856 | -2.833 | 54.7313 | #N/A | -2.833 | 54.7313 | 6.067 | 51.4056 | #N/A | 6.067 | 51.4056 | -2.833 | 39.9785 | #N/A | -2.833 | 39.9785 | -2.933 | 47.3554 | #N/A | -2.933 | 47.3554 | -26.113 | 81.102 | #N/A | -26.113 | 81.102 | 69.2089 | 94.0424 | #N/A | 69.2089 | 94.0424 | 0.077009 | 81.102 | #N/A | 0.077009 | 81.102 | -38.335 | 31.7759 | #N/A | -38.335 | 31.7759 |  | -12.7068 | 12.7068 |  | 5.897 | 21.9202 | #N/A | 5.897 | 21.9202 | -4.933 | 37.6226 | #N/A | -4.933 | 37.6226 | 34.067 | 71.5719 | #N/A | 34.067 | 71.5719
| 10 |  |  | -3.8921 | 69.371 | #N/A | -3.8921 | 69.371 | -27.2004 | 38.8888 | #N/A | -27.2004 | 38.8888 | -7.1804 | 116.0228 | #N/A | -7.1804 | 116.0228 | 3.7996 | 59.265 | #N/A | 3.7996 | 59.265 | 27.7396 | 118.9421 | #N/A | 27.7396 | 118.9421 | -1.2004 | 51.1501 | #N/A | -1.2004 | 51.1501 | 16.0996 | 39.6495 | #N/A | 16.0996 | 39.6495 | -8.2004 | 47.078 | #N/A | -8.2004 | 47.078 | -27.5504 | 92.6481 | #N/A | -27.5504 | 92.6481 | 71.7291 | 93.9246 | #N/A | 71.7291 | 93.9246 | -13.0004 | 80.9403 | #N/A | -13.0004 | 80.9403 | -37.1534 | 32.1855 | #N/A | -37.1534 | 32.1855 |  | -13.699 | 13.699 |  | 4.4596 | 21.3142 | #N/A | 4.4596 | 21.3142 | -23.9504 | 61.8948 | #N/A | -23.9504 | 61.8948 | 23.7996 | 71.3886 | #N/A | 23.7996 | 71.3886
| 5 |  |  | 4.9972 | 69.2753 | #N/A | 4.9972 | 69.2753 | -16.8543 | 38.7178 | #N/A | -16.8543 | 38.7178 | 0.35572 | 162.65 | #N/A | 0.35572 | 162.65 | 22.1457 | 59.1529 | #N/A | 22.1457 | 59.1529 | 23.6357 | 92.5765 | #N/A | 23.6357 | 92.5765 | 6.1457 | 51.0202 | #N/A | 6.1457 | 51.0202 | -2.5543 | 63.2268 | #N/A | -2.5543 | 63.2268 | 5.1457 | 46.9368 | #N/A | 5.1457 | 46.9368 | -14.1943 | 115.9656 | #N/A | -14.1943 | 115.9656 | 75.2543 | 112.2248 | #N/A | 75.2543 | 112.2248 | -20.0143 | 80.8582 | #N/A | -20.0143 | 80.8582 | -24.9613 | 34.6572 | #N/A | -24.9613 | 34.6572 |  | -14.1752 | 14.1752 |  | 11.9957 | 27.3107 | #N/A | 11.9957 | 27.3107 | -27.3543 | 66.0512 | #N/A | -27.3543 | 66.0512 | 31.1457 | 71.2956 | #N/A | 31.1457 | 71.2956
| 2 |  |  | 3.7266 | 79.446 | #N/A | 3.7266 | 79.446 | -15.7659 | 38.8801 | #N/A | -15.7659 | 38.8801 | -1.3659 | 162.6887 | #N/A | -1.3659 | 162.6887 | 32.2341 | 79.446 | #N/A | 32.2341 | 79.446 | 19.0041 | 89.719 | #N/A | 19.0041 | 89.719 | 1.2341 | 51.1435 | #N/A | 1.2341 | 51.1435 | -1.3659 | 92.6445 | #N/A | -1.3659 | 92.6445 | 4.2341 | 47.0708 | #N/A | 4.2341 | 47.0708 | 16.0941 | 197.6567 | #N/A | 16.0941 | 197.6567 |  |  | 0 | #N/A | #N/A | -71.2059 | 151.0278 | #N/A | -71.2059 | 151.0278 | -25.8099 | 35.4409 | #N/A | -25.8099 | 35.4409 |  | -15.2426 | 15.2426 |  | 10.2741 | 27.5404 | #N/A | 10.2741 | 27.5404 | 39.1508 | 90.8792 | #N/A | 39.1508 | 90.8792 | 28.2341 | 75.4166 | #N/A | 28.2341 | 75.4166
| 1 |  |  | 7.1259 | 79.1974 | #N/A | 7.1259 | 79.1974 | -12.5607 | 38.3696 | #N/A | -12.5607 | 38.3696 | 0.38925 | 162.5675 | #N/A | 0.38925 | 162.5675 | 41.4393 | 79.1974 | #N/A | 41.4393 | 79.1974 | 17.8493 | 86.5651 | #N/A | 17.8493 | 86.5651 | 9.4393 | 56.8878 | #N/A | 9.4393 | 56.8878 | -14.1607 | 92.4314 | #N/A | -14.1607 | 92.4314 | 5.4393 | 46.65 | #N/A | 5.4393 | 46.65 | 52.7693 | 395.5985 | #N/A | 52.7693 | 395.5985 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | -27.2557 | 34.8801 | #N/A | -27.2557 | 34.8801 |  | -16.4856 | 16.4856 |  | 12.0293 | 33.0398 | #N/A | 12.0293 | 33.0398 |  |  | 0 | #N/A | #N/A | 28.4393 | 79.1974 | #N/A | 28.4393 | 79.1974
5/5 | 120 |  |  | -10.8137 | 60.5718 | #N/A | -10.8137 | 60.5718 | -32.4187 | 40.8527 | #N/A | -32.4187 | 40.8527 | 5.2713 | 81.902 | #N/A | 5.2713 | 81.902 | -27.4187 | 60.5718 | #N/A | -27.4187 | 60.5718 | 28.5513 | 128.3089 | #N/A | 28.5513 | 128.3089 | -6.4187 | 52.6587 | #N/A | -6.4187 | 52.6587 | -20.9187 | 41.5775 | #N/A | -20.9187 | 41.5775 | -17.4187 | 50.6847 | #N/A | -17.4187 | 50.6847 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 48.9213 | 81.902 | #N/A | 48.9213 | 81.902 |  |  | 0 | #N/A | #N/A |  | -13.6769 | 13.6769 |  | 16.9113 | 24.7164 | #N/A | 16.9113 | 24.7164 | 0.081318 | 51.1392 | #N/A | 0.081318 | 51.1392 | 60.5813 | 72.4772 | #N/A | 60.5813 | 72.4772
| 100 |  |  | -8.5579 | 60.1704 | #N/A | -8.5579 | 60.1704 | -31.8101 | 40.2551 | #N/A | -31.8101 | 40.2551 | 4.6099 | 81.6055 | #N/A | 4.6099 | 81.6055 | -33.8101 | 60.1704 | #N/A | -33.8101 | 60.1704 | 45.3499 | 163.0228 | #N/A | 45.3499 | 163.0228 | -3.8101 | 52.1965 | #N/A | -3.8101 | 52.1965 | -1.2101 | 40.9905 | #N/A | -1.2101 | 40.9905 | -16.8101 | 50.2043 | #N/A | -16.8101 | 50.2043 | -24.4901 | 81.6055 | #N/A | -24.4901 | 81.6055 | 182.1538 | 137.7686 | #N/A | 182.1538 | 137.7686 | 22.0699 | 81.6055 | #N/A | 22.0699 | 81.6055 | -42.2411 | 32.9889 | #N/A | -42.2411 | 32.9889 |  | -13.2487 | 13.2487 |  | 16.2499 | 23.7156 | #N/A | 16.2499 | 23.7156 | -6.3101 | 40.5247 | #N/A | -6.3101 | 40.5247 | 64.1899 | 72.142 | #N/A | 64.1899 | 72.142
| 50 |  |  | -1.3702 | 59.4121 | #N/A | -1.3702 | 59.4121 | -21.9517 | 39.1127 | #N/A | -21.9517 | 39.1127 | 5.3883 | 81.0481 | #N/A | 5.3883 | 81.0481 | -12.9517 | 59.4121 | #N/A | -12.9517 | 59.4121 | 92.6883 | 142.3436 | #N/A | 92.6883 | 142.3436 | 5.0483 | 51.3206 | #N/A | 5.0483 | 51.3206 | -6.2517 | 39.8691 | #N/A | -6.2517 | 39.8691 | -8.9517 | 47.2631 | #N/A | -8.9517 | 47.2631 | -23.7117 | 81.0481 | #N/A | -23.7117 | 81.0481 | 154.4918 | 137.3408 | #N/A | 154.4918 | 137.3408 | -9.1617 | 81.0481 | #N/A | -9.1617 | 81.0481 | -46.9917 | 31.4156 | #N/A | -46.9917 | 31.4156 |  | -13.046 | 13.046 |  | 11.2083 | 21.72 | #N/A | 11.2083 | 21.72 | -9.7017 | 35.5544 | #N/A | -9.7017 | 35.5544 | 50.0483 | 71.5108 | #N/A | 50.0483 | 71.5108
| 20 |  |  | -0.90898 | 59.4567 | #N/A | -0.90898 | 59.4567 | -20.3385 | 39.1804 | #N/A | -20.3385 | 39.1804 | -2.2585 | 81.0808 | #N/A | -2.2585 | 81.0808 | -8.3385 | 59.4567 | #N/A | -8.3385 | 59.4567 | 35.5715 | 78.1559 | #N/A | 35.5715 | 78.1559 | 3.6615 | 51.3722 | #N/A | 3.6615 | 51.3722 | 3.5615 | 39.9356 | #N/A | 3.5615 | 39.9356 | -8.3385 | 47.3192 | #N/A | -8.3385 | 47.3192 | -8.0785 | 81.0808 | #N/A | -8.0785 | 81.0808 | 124.8067 | 136.1082 | #N/A | 124.8067 | 136.1082 | -8.0785 | 81.0808 | #N/A | -8.0785 | 81.0808 | -39.7975 | 31.3311 | #N/A | -39.7975 | 31.3311 |  | -12.8412 | 12.8412 |  | 0.65151 | 21.8417 | #N/A | 0.65151 | 21.8417 | -3.5885 | 35.0834 | #N/A | -3.5885 | 35.0834 | 37.6615 | 71.5479 | #N/A | 37.6615 | 71.5479
| 10 |  |  | -4.7956 | 69.4667 | #N/A | -4.7956 | 69.4667 | -25.1535 | 39.0593 | #N/A | -25.1535 | 39.0593 | -6.8835 | 116.0801 | #N/A | -6.8835 | 116.0801 | -5.1535 | 59.377 | #N/A | -5.1535 | 59.377 | 54.2265 | 72.237 | #N/A | 54.2265 | 72.237 | -4.1535 | 51.2799 | #N/A | -4.1535 | 51.2799 | 1.8465 | 39.8168 | #N/A | 1.8465 | 39.8168 | -13.1535 | 47.2189 | #N/A | -13.1535 | 47.2189 | -9.7935 | 92.7198 | #N/A | -9.7935 | 92.7198 | 100.932 | 136.0898 | #N/A | 100.932 | 136.0898 | -21.4335 | 81.0223 | #N/A | -21.4335 | 81.0223 | -38.8935 | 31.4819 | #N/A | -38.8935 | 31.4819 |  | -13.205 | 13.205 |  | -1.0635 | 21.6237 | #N/A | -1.0635 | 21.6237 | 15.4298 | 39.8648 | #N/A | 15.4298 | 39.8648 | 31.8465 | 71.4817 | #N/A | 31.8465 | 71.4817
| 5 |  |  | 3.3512 | 69.3915 | #N/A | 3.3512 | 69.3915 | -15.4898 | 38.9252 | #N/A | -15.4898 | 38.9252 | 0.15024 | 162.6995 | #N/A | 0.15024 | 162.6995 | 11.5102 | 59.2889 | #N/A | 11.5102 | 59.2889 | 23.4302 | 92.6634 | #N/A | 23.4302 | 92.6634 | 5.5102 | 51.1779 | #N/A | 5.5102 | 51.1779 | -11.4898 | 63.354 | #N/A | -11.4898 | 63.354 | -2.4898 | 47.1081 | #N/A | -2.4898 | 47.1081 | 26.3402 | 116.0351 | #N/A | 26.3402 | 116.0351 | 85.6072 | 163.3123 | #N/A | 85.6072 | 163.3123 | -31.8598 | 80.9578 | #N/A | -31.8598 | 80.9578 | -27.7858 | 34.2865 | #N/A | -27.7858 | 34.2865 |  | -13.5951 | 13.5951 |  | 5.9702 | 27.6041 | #N/A | 5.9702 | 27.6041 | 17.7602 | 37.4068 | #N/A | 17.7602 | 37.4068 | 39.5102 | 71.4085 | #N/A | 39.5102 | 71.4085
| 2 |  |  | 4.5468 | 79.4649 | #N/A | 4.5468 | 79.4649 | -13.2208 | 38.9188 | #N/A | -13.2208 | 38.9188 | 0.78925 | 162.698 | #N/A | 0.78925 | 162.698 | 22.7792 | 79.4649 | #N/A | 22.7792 | 79.4649 | 21.1592 | 86.8099 | #N/A | 21.1592 | 86.8099 | 9.7792 | 51.173 | #N/A | 9.7792 | 51.173 | -2.1208 | 92.6607 | #N/A | -2.1208 | 92.6607 | -1.2208 | 47.1028 | #N/A | -1.2208 | 47.1028 | 61.8992 | 197.6643 | #N/A | 61.8992 | 197.6643 |  |  | 0 | #N/A | #N/A | -89.4208 | 151.0378 | #N/A | -89.4208 | 151.0378 | -26.5648 | 34.2792 | #N/A | -26.5648 | 34.2792 |  | -15.1436 | 15.1436 |  | 9.5192 | 27.595 | #N/A | 9.5192 | 27.595 | 34.9459 | 91.4416 | #N/A | 34.9459 | 91.4416 | 36.7792 | 75.4365 | #N/A | 36.7792 | 75.4365
| 1 |  |  | 4.6152 | 79.2086 | #N/A | 4.6152 | 79.2086 | -10.9267 | 38.3927 | #N/A | -10.9267 | 38.3927 | -1.1867 | 162.5729 | #N/A | -1.1867 | 162.5729 | 31.0733 | 79.2086 | #N/A | 31.0733 | 79.2086 | 22.0933 | 86.5753 | #N/A | 22.0933 | 86.5753 | 19.0733 | 56.9034 | #N/A | 19.0733 | 56.9034 | -12.8267 | 92.441 | #N/A | -12.8267 | 92.441 | -0.92666 | 46.669 | #N/A | -0.92666 | 46.669 | 129.7633 | 395.6008 | #N/A | 129.7633 | 395.6008 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | -29.7047 | 34.2937 | #N/A | -29.7047 | 34.2937 |  | -16.4317 | 16.4317 |  | 13.3633 | 33.0667 | #N/A | 13.3633 | 33.0667 |  |  | 0 | #N/A | #N/A | 36.0733 | 79.2086 | #N/A | 36.0733 | 79.2086
6/5 | 120 |  |  | -11.3014 | 59.8468 | #N/A | -11.3014 | 59.8468 | -26.149 | 29.6924 | #N/A | -26.149 | 29.6924 | -9.199 | 81.3672 | #N/A | -9.199 | 81.3672 | -13.149 | 59.8468 | #N/A | -13.149 | 59.8468 | 2.441 | 55.1237 | #N/A | 2.441 | 55.1237 | -0.14896 | 63.8564 | #N/A | -0.14896 | 63.8564 | 5.351 | 40.514 | #N/A | 5.351 | 40.514 | -15.149 | 47.8083 | #N/A | -15.149 | 47.8083 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 34.451 | 161.9433 | #N/A | 34.451 | 161.9433 |  |  | 0 | #N/A | #N/A |  | -12.7422 | 12.7422 |  | 8.261 | 22.8822 | #N/A | 8.261 | 22.8822 | 16.351 | 38.7547 | #N/A | 16.351 | 38.7547 | 64.851 | 71.8724 | #N/A | 64.851 | 71.8724
| 100 |  |  | -2.7401 | 59.8495 | #N/A | -2.7401 | 59.8495 | -23.8536 | 29.6979 | #N/A | -23.8536 | 29.6979 | -10.9936 | 81.3692 | #N/A | -10.9936 | 81.3692 | -27.8536 | 59.8495 | #N/A | -27.8536 | 59.8495 | 58.8464 | 90.11 | #N/A | 58.8464 | 90.11 | 1.1464 | 51.8263 | #N/A | 1.1464 | 51.8263 | 0.64637 | 40.518 | #N/A | 0.64637 | 40.518 | -12.8536 | 47.8118 | #N/A | -12.8536 | 47.8118 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 15.1964 | 81.3692 | #N/A | 15.1964 | 81.3692 |  |  | 0 | #N/A | #N/A |  | -12.7293 | 12.7293 |  | 9.3764 | 22.8894 | #N/A | 9.3764 | 22.8894 | 2.8964 | 39.4161 | #N/A | 2.8964 | 39.4161 | 66.1464 | 71.8746 | #N/A | 66.1464 | 71.8746
| 50 |  |  | 4.367 | 59.5411 | #N/A | 4.367 | 59.5411 | -15.4661 | 29.0714 | #N/A | -15.4661 | 29.0714 | -14.3261 | 81.1427 | #N/A | -14.3261 | 81.1427 | -15.4661 | 59.5411 | #N/A | -15.4661 | 59.5411 | 70.0639 | 107.415 | #N/A | 70.0639 | 107.415 | 11.5339 | 51.4699 | #N/A | 11.5339 | 51.4699 | 3.1339 | 40.0612 | #N/A | 3.1339 | 40.0612 | -5.4661 | 47.4252 | #N/A | -5.4661 | 47.4252 | -14.3261 | 81.1427 | #N/A | -14.3261 | 81.1427 | 27.1608 | 140.7937 | #N/A | 27.1608 | 140.7937 | -2.6861 | 81.1427 | #N/A | -2.6861 | 81.1427 | -39.3521 | 37.8675 | #N/A | -39.3521 | 37.8675 |  | -12.4439 | 12.4439 |  | 8.9539 | 22.0705 | #N/A | 8.9539 | 22.0705 | -17.4661 | 35.5517 | #N/A | -17.4661 | 35.5517 | 54.5339 | 71.6181 | #N/A | 54.5339 | 71.6181
| 20 |  |  | 7.955 | 59.7199 | #N/A | 7.955 | 59.7199 | -11.1449 | 29.4358 | #N/A | -11.1449 | 29.4358 | -15.0849 | 81.2739 | #N/A | -15.0849 | 81.2739 | -5.1449 | 59.7199 | #N/A | -5.1449 | 59.7199 | 2.3751 | 57.9112 | #N/A | 2.3751 | 57.9112 | 13.8551 | 51.6765 | #N/A | 13.8551 | 51.6765 | -3.4449 | 40.3263 | #N/A | -3.4449 | 40.3263 | 0.85506 | 47.6494 | #N/A | 0.85506 | 47.6494 | 2.3751 | 81.2739 | #N/A | 2.3751 | 81.2739 | -11.3747 | 146.0655 | #N/A | -11.3747 | 146.0655 | -3.4449 | 81.2739 | #N/A | -3.4449 | 81.2739 | -29.3439 | 33.8462 | #N/A | -29.3439 | 33.8462 |  | -11.5558 | 11.5558 |  | 14.0151 | 22.5482 | #N/A | 14.0151 | 22.5482 | 0.85506 | 36.4146 | #N/A | 0.85506 | 36.4146 | 45.8551 | 71.7667 | #N/A | 45.8551 | 71.7667
| 10 |  |  | 6.6287 | 69.704 | #N/A | 6.6287 | 69.704 | -12.4857 | 29.3027 | #N/A | -12.4857 | 29.3027 | -16.1457 | 116.2222 | #N/A | -16.1457 | 116.2222 | 1.5143 | 59.6544 | #N/A | 1.5143 | 59.6544 | 47.8743 | 87.0628 | #N/A | 47.8743 | 87.0628 | 10.5143 | 51.6009 | #N/A | 10.5143 | 51.6009 | 1.3143 | 40.2293 | #N/A | 1.3143 | 40.2293 | -2.4857 | 47.5673 | #N/A | -2.4857 | 47.5673 | -4.5057 | 92.8977 | #N/A | -4.5057 | 92.8977 | -41.5306 | 152.0998 | #N/A | -41.5306 | 152.0998 | -13.2357 | 81.2258 | #N/A | -13.2357 | 81.2258 | -25.4577 | 34.3228 | #N/A | -25.4577 | 34.3228 |  | -11.8892 | 11.8892 |  | 4.2243 | 22.3742 | #N/A | 4.2243 | 22.3742 | 19.0143 | 37.1643 | #N/A | 19.0143 | 37.1643 | 38.5143 | 71.7123 | #N/A | 38.5143 | 71.7123
| 5 |  |  | 5.0405 | 69.2563 | #N/A | 5.0405 | 69.2563 | -13.1083 | 28.2212 | #N/A | -13.1083 | 28.2212 | -14.3983 | 162.6419 | #N/A | -14.3983 | 162.6419 | 6.8917 | 59.1307 | #N/A | 6.8917 | 59.1307 | 11.7917 | 77.9082 | #N/A | 11.7917 | 77.9082 | 8.8917 | 50.9945 | #N/A | 8.8917 | 50.9945 | 11.7917 | 63.206 | #N/A | 11.7917 | 63.206 | -2.1083 | 46.9088 | #N/A | -2.1083 | 46.9088 | -2.7583 | 115.9543 | #N/A | -2.7583 | 115.9543 | -55.1043 | 180.6489 | #N/A | -55.1043 | 180.6489 | -34.7683 | 80.842 | #N/A | -34.7683 | 80.842 | -24.0013 | 74.9725 | #N/A | -24.0013 | 74.9725 |  | -14.2676 | 14.2676 |  | 3.0617 | 27.2625 | #N/A | 3.0617 | 27.2625 | 20.8084 | 60.1686 | #N/A | 20.8084 | 60.1686 | 36.8917 | 71.2772 | #N/A | 36.8917 | 71.2772
| 2 |  |  | 3.5765 | 79.4132 | #N/A | 3.5765 | 79.4132 | -12.0938 | 28.3983 | #N/A | -12.0938 | 28.3983 | -14.9238 | 162.6727 | #N/A | -14.9238 | 162.6727 | 17.9062 | 79.4132 | #N/A | 17.9062 | 79.4132 | 14.1762 | 95.5417 | #N/A | 14.1762 | 95.5417 | 7.9062 | 51.0927 | #N/A | 7.9062 | 51.0927 | -6.1938 | 92.6164 | #N/A | -6.1938 | 92.6164 | -4.0938 | 47.0155 | #N/A | -4.0938 | 47.0155 | 28.7262 | 197.6435 | #N/A | 28.7262 | 197.6435 |  |  | 0 | #N/A | #N/A | -81.8538 | 145.1782 | #N/A | -81.8538 | 145.1782 | -23.9448 | 75.0393 | #N/A | -23.9448 | 75.0393 |  | -15.4123 | 15.4123 |  | 5.4462 | 27.4458 | #N/A | 5.4462 | 27.4458 | 30.8646 | 90.9834 | #N/A | 30.8646 | 90.9834 | 33.9062 | 75.3821 | #N/A | 33.9062 | 75.3821
| 1 |  |  | 0.83953 | 79.1817 | #N/A | 0.83953 | 79.1817 | -10.88 | 27.7443 | #N/A | -10.88 | 27.7443 | -9.16 | 162.5599 | #N/A | -9.16 | 162.5599 | 26.12 | 79.1817 | #N/A | 26.12 | 79.1817 | 14.12 | 95.3494 | #N/A | 14.12 | 95.3494 | 8.12 | 50.7321 | #N/A | 8.12 | 50.7321 | -14.98 | 92.418 | #N/A | -14.98 | 92.418 | -2.88 | 46.6235 | #N/A | -2.88 | 46.6235 | 81.05 | 395.5954 | #N/A | 81.05 | 395.5954 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | -29.821 | 74.7943 | #N/A | -29.821 | 74.7943 |  | -16.5606 | 16.5606 |  | 8.3 | 33.0023 | #N/A | 8.3 | 33.0023 |  |  | 0 | #N/A | #N/A | 29.12 | 79.1817 | #N/A | 29.12 | 79.1817
8/5 | 120 |  |  | 0.47655 | 60.3242 | #N/A | 0.47655 | 60.3242 | -17.6396 | 30.6432 | #N/A | -17.6396 | 30.6432 | 6.9404 | 81.719 | #N/A | 6.9404 | 81.719 | 9.3604 | 60.3242 | #N/A | 9.3604 | 60.3242 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | -16.3396 | 41.216 | #N/A | -16.3396 | 41.216 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 36.0404 | 81.719 | #N/A | 36.0404 | 81.719 |  |  | 0 | #N/A | #N/A |  | -14.7306 | 14.7306 |  | 4.0304 | 24.1033 | #N/A | 4.0304 | 24.1033 |  |  | 0 | #N/A | #N/A | 66.3604 | 72.2704 | #N/A | 66.3604 | 72.2704
| 100 |  |  | 5.9048 | 60.6942 | #N/A | 5.9048 | 60.6942 | -14.3888 | 31.3653 | #N/A | -14.3888 | 31.3653 | 3.1912 | 81.9925 | #N/A | 3.1912 | 81.9925 | -10.3888 | 60.6942 | #N/A | -10.3888 | 60.6942 | 11.9212 | 53.1737 | #N/A | 11.9212 | 53.1737 | 17.6112 | 64.6512 | #N/A | 17.6112 | 64.6512 | -20.0888 | 41.7556 | #N/A | -20.0888 | 41.7556 | -3.3888 | 48.8649 | #N/A | -3.3888 | 48.8649 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 23.5612 | 81.9925 | #N/A | 23.5612 | 81.9925 |  |  | 0 | #N/A | #N/A |  | -13.1232 | 13.1232 |  | 0.28121 | 25.0148 | #N/A | 0.28121 | 25.0148 | 25.6945 | 80.7344 | #N/A | 25.6945 | 80.7344 | 70.6112 | 72.5795 | #N/A | 70.6112 | 72.5795
| 50 |  |  | 10.4162 | 60.383 | #N/A | 10.4162 | 60.383 | -7.9721 | 30.7588 | #N/A | -7.9721 | 30.7588 | -4.7521 | 81.7624 | #N/A | -4.7521 | 81.7624 | -2.9721 | 60.383 | #N/A | -2.9721 | 60.383 | 62.1779 | 125.3141 | #N/A | 62.1779 | 125.3141 | 16.0279 | 52.4415 | #N/A | 16.0279 | 52.4415 | 18.5279 | 41.302 | #N/A | 18.5279 | 41.302 | 0.027857 | 48.4779 | #N/A | 0.027857 | 48.4779 | -10.5721 | 81.7624 | #N/A | -10.5721 | 81.7624 | 147.3584 | 137.9327 | #N/A | 147.3584 | 137.9327 | 6.8879 | 81.7624 | #N/A | 6.8879 | 81.7624 | -30.9421 | 35.0029 | #N/A | -30.9421 | 35.0029 |  | -12.2431 | 12.2431 |  | 1.0679 | 24.2501 | #N/A | 1.0679 | 24.2501 | 3.2779 | 71.6629 | #N/A | 3.2779 | 71.6629 | 55.0279 | 72.3195 | #N/A | 55.0279 | 72.3195
| 20 |  |  | 8.4328 | 60.0136 | #N/A | 8.4328 | 60.0136 | -9.4193 | 30.0271 | #N/A | -9.4193 | 30.0271 | -15.8293 | 81.49 | #N/A | -15.8293 | 81.49 | -2.4193 | 60.0136 | #N/A | -2.4193 | 60.0136 | 1.6307 | 52.3955 | #N/A | 1.6307 | 52.3955 | 10.5807 | 52.0157 | #N/A | 10.5807 | 52.0157 | 16.1807 | 40.76 | #N/A | 16.1807 | 40.76 | -1.4193 | 48.017 | #N/A | -1.4193 | 48.017 | -7.0993 | 81.49 | #N/A | -7.0993 | 81.49 | 98.621 | 139.212 | #N/A | 98.621 | 139.212 | 1.6307 | 81.49 | #N/A | 1.6307 | 81.49 | -30.0883 | 34.3617 | #N/A | -30.0883 | 34.3617 |  | -11.932 | 11.932 |  | 7.4507 | 23.3149 | #N/A | 7.4507 | 23.3149 | -13.1693 | 71.3076 | #N/A | -13.1693 | 71.3076 | 39.5807 | 72.0113 | #N/A | 39.5807 | 72.0113
| 10 |  |  | 11.676 | 69.9836 | #N/A | 11.676 | 69.9836 | -5.9486 | 29.9617 | #N/A | -5.9486 | 29.9617 | -8.9886 | 116.3901 | #N/A | -8.9886 | 116.3901 | 5.0514 | 59.9809 | #N/A | 5.0514 | 59.9809 | 2.6514 | 52.3581 | #N/A | 2.6514 | 52.3581 | 14.0514 | 51.9779 | #N/A | 14.0514 | 51.9779 | 2.6514 | 40.7118 | #N/A | 2.6514 | 40.7118 | 1.0514 | 47.9761 | #N/A | 1.0514 | 47.9761 | -0.25864 | 93.1077 | #N/A | -0.25864 | 93.1077 | 74.0676 | 142.4794 | #N/A | 74.0676 | 142.4794 | -3.1686 | 81.4659 | #N/A | -3.1686 | 81.4659 | -21.2106 | 34.3046 | #N/A | -21.2106 | 34.3046 |  | -12.0952 | 12.0952 |  | 2.6514 | 23.2307 | #N/A | 2.6514 | 23.2307 | -13.6986 | 71.3068 | #N/A | -13.6986 | 71.3068 | 41.0514 | 71.9841 | #N/A | 41.0514 | 71.9841
| 5 |  |  | 9.6234 | 69.5442 | #N/A | 9.6234 | 69.5442 | -7.2527 | 28.9206 | #N/A | -7.2527 | 28.9206 | -9.7427 | 162.7647 | #N/A | -9.7427 | 162.7647 | 10.7473 | 59.4676 | #N/A | 10.7473 | 59.4676 | 7.7173 | 78.1642 | #N/A | 7.7173 | 78.1642 | 10.7473 | 51.3848 | #N/A | 10.7473 | 51.3848 | -12.6527 | 63.5213 | #N/A | -12.6527 | 63.5213 | -1.2527 | 47.3329 | #N/A | -1.2527 | 47.3329 | -6.8327 | 116.1265 | #N/A | -6.8327 | 116.1265 | 59.2098 | 172.4429 | #N/A | 59.2098 | 172.4429 | -21.3827 | 81.0888 | #N/A | -21.3827 | 81.0888 | -20.5097 | 75.2385 | #N/A | -20.5097 | 75.2385 |  | -14.4083 | 14.4083 |  | 1.8973 | 27.9859 | #N/A | 1.8973 | 27.9859 | -12.5027 | 76.2706 | #N/A | -12.5027 | 76.2706 | 36.7473 | 71.557 | #N/A | 36.7473 | 71.557
| 2 |  |  | 8.9749 | 79.1622 | #N/A | 8.9749 | 79.1622 | -6.967 | 27.6886 | #N/A | -6.967 | 27.6886 | -8.907 | 162.5504 | #N/A | -8.907 | 162.5504 | 20.033 | 79.1622 | #N/A | 20.033 | 79.1622 | 8.553 | 95.3332 | #N/A | 8.553 | 95.3332 | 11.033 | 50.7016 | #N/A | 11.033 | 50.7016 | 2.733 | 92.4013 | #N/A | 2.733 | 92.4013 | -2.967 | 46.5903 | #N/A | -2.967 | 46.5903 | 23.103 | 197.5428 | #N/A | 23.103 | 197.5428 |  |  | 0 | #N/A | #N/A | -61.287 | 145.0411 | #N/A | -61.287 | 145.0411 | -21.129 | 74.7736 | #N/A | -21.129 | 74.7736 |  | -15.2756 | 15.2756 |  | 2.733 | 26.7108 | #N/A | 2.733 | 26.7108 | -17.6337 | 70.6443 | #N/A | -17.6337 | 70.6443 | 33.033 | 75.1176 | #N/A | 33.033 | 75.1176
| 1 |  |  | 2.1515 | 79.1466 | #N/A | 2.1515 | 79.1466 | -9.726 | 27.6438 | #N/A | -9.726 | 27.6438 | -9.936 | 162.5427 | #N/A | -9.936 | 162.5427 | 25.274 | 79.1466 | #N/A | 25.274 | 79.1466 | 4.614 | 95.3202 | #N/A | 4.614 | 95.3202 | 16.274 | 56.8171 | #N/A | 16.274 | 56.8171 | -7.026 | 92.3878 | #N/A | -7.026 | 92.3878 | -5.726 | 46.5637 | #N/A | -5.726 | 46.5637 | 68.634 | 395.5884 | #N/A | 68.634 | 395.5884 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | -31.761 | 74.757 | #N/A | -31.761 | 74.757 |  | -16.7278 | 16.7278 |  | 7.524 | 32.9179 | #N/A | 7.524 | 32.9179 |  |  | 0 | #N/A | #N/A | 24.274 | 79.1466 | #N/A | 24.274 | 79.1466
10/5 | 120 |  |  |  |  | 0 | #N/A | #N/A | -8.594 | 27.8283 | #N/A | -8.594 | 27.8283 | 12.616 | 80.7057 | #N/A | 12.616 | 80.7057 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 9.706 | 39.1684 | #N/A | 9.706 | 39.1684 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  | -16.4191 | 16.4191 |  | 0.97596 | 20.4052 | #N/A | 0.97596 | 20.4052 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A
| 100 |  |  | -0.74304 | 59.374 | #N/A | -0.74304 | 59.374 | -10.2494 | 28.7275 | #N/A | -10.2494 | 28.7275 | 4.9606 | 81.0201 | #N/A | 4.9606 | 81.0201 | 13.7506 | 59.374 | #N/A | 13.7506 | 59.374 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 2.0506 | 39.8123 | #N/A | 2.0506 | 39.8123 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | 34.0606 | 81.0201 | #N/A | 34.0606 | 81.0201 |  |  | 0 | #N/A | #N/A |  | -14.7896 | 14.7896 |  | -3.7694 | 21.6154 | #N/A | -3.7694 | 21.6154 |  |  | 0 | #N/A | #N/A | 48.7506 | 79.5316 | #N/A | 48.7506 | 79.5316
| 50 |  |  | 5.1357 | 59.9883 | #N/A | 5.1357 | 59.9883 | -3.3293 | 29.9767 | #N/A | -3.3293 | 29.9767 | 0.61067 | 81.4714 | #N/A | 0.61067 | 81.4714 | -4.3293 | 59.9883 | #N/A | -4.3293 | 59.9883 | 6.4307 | 55.2773 | #N/A | 6.4307 | 55.2773 | 20.6707 | 51.9865 | #N/A | 20.6707 | 51.9865 | -11.0293 | 40.7228 | #N/A | -11.0293 | 40.7228 | 3.6707 | 47.9854 | #N/A | 3.6707 | 47.9854 | 3.5207 | 81.4714 | #N/A | 3.5207 | 81.4714 | 171.903 | 137.3 | #N/A | 171.903 | 137.3 | 18.0707 | 81.4714 | #N/A | 18.0707 | 81.4714 | -27.6163 | 34.3176 | #N/A | -27.6163 | 34.3176 |  | -12.0582 | 12.0582 |  | 3.5207 | 23.2499 | #N/A | 3.5207 | 23.2499 | 32.3373 | 75.3572 | #N/A | 32.3373 | 75.3572 | 42.6707 | 79.9913 | #N/A | 42.6707 | 79.9913
| 20 |  |  | -0.45359 | 59.6332 | #N/A | -0.45359 | 59.6332 | -8.7218 | 29.2596 | #N/A | -8.7218 | 29.2596 | -9.3218 | 81.2103 | #N/A | -9.3218 | 81.2103 | 0.27818 | 59.6332 | #N/A | 0.27818 | 59.6332 | 25.5982 | 54.8918 | #N/A | 25.5982 | 54.8918 | 14.2782 | 51.5764 | #N/A | 14.2782 | 51.5764 | 19.7782 | 40.1979 | #N/A | 19.7782 | 40.1979 | -2.7218 | 47.5408 | #N/A | -2.7218 | 47.5408 | -0.59182 | 81.2103 | #N/A | -0.59182 | 81.2103 | 124.9553 | 137.5184 | #N/A | 124.9553 | 137.5184 | -0.59182 | 81.2103 | #N/A | -0.59182 | 81.2103 | -31.4378 | 33.6931 | #N/A | -31.4378 | 33.6931 |  | -11.9948 | 11.9948 |  | 2.3182 | 22.3178 | #N/A | 2.3182 | 22.3178 | -10.2218 | 71.0932 | #N/A | -10.2218 | 71.0932 | 30.2782 | 79.7253 | #N/A | 30.2782 | 79.7253
| 10 |  |  | 3.137 | 69.66 | #N/A | 3.137 | 69.66 | -5.1171 | 29.1978 | #N/A | -5.1171 | 29.1978 | -7.0771 | 116.1958 | #N/A | -7.0771 | 116.1958 | 8.8829 | 59.603 | #N/A | 8.8829 | 59.603 | 7.4729 | 51.9247 | #N/A | 7.4729 | 51.9247 | 15.8829 | 51.5414 | #N/A | 15.8829 | 51.5414 | 10.3829 | 40.153 | #N/A | 10.3829 | 40.153 | -1.1171 | 47.5028 | #N/A | -1.1171 | 47.5028 | 7.4729 | 92.8647 | #N/A | 7.4729 | 92.8647 | 116.7677 | 138.5723 | #N/A | 116.7677 | 138.5723 | -1.2571 | 81.1881 | #N/A | -1.2571 | 81.1881 | -22.2091 | 33.6394 | #N/A | -22.2091 | 33.6394 |  | -12.1444 | 12.1444 |  | -1.2571 | 22.2367 | #N/A | -1.2571 | 22.2367 | -11.7004 | 72.2654 | #N/A | -11.7004 | 72.2654 | 37.8829 | 83.7169 | #N/A | 37.8829 | 83.7169
| 5 |  |  | 1.6479 | 69.2051 | #N/A | 1.6479 | 69.2051 | -6.3368 | 28.0954 | #N/A | -6.3368 | 28.0954 | -7.6568 | 162.6201 | #N/A | -7.6568 | 162.6201 | 14.6632 | 59.0707 | #N/A | 14.6632 | 59.0707 | 18.5332 | 80.7982 | #N/A | 18.5332 | 80.7982 | 12.6632 | 50.925 | #N/A | 12.6632 | 50.925 | -1.8368 | 63.1499 | #N/A | -1.8368 | 63.1499 | -1.3368 | 46.8332 | #N/A | -1.3368 | 46.8332 | 3.9832 | 115.9237 | #N/A | 3.9832 | 115.9237 | 119.6072 | 167.2297 | #N/A | 119.6072 | 167.2297 | -10.5668 | 80.7982 | #N/A | -10.5668 | 80.7982 | -21.0428 | 74.9252 | #N/A | -21.0428 | 74.9252 |  | -14.5138 | 14.5138 |  | -4.7468 | 27.1323 | #N/A | -4.7468 | 27.1323 | -15.6701 | 75.2156 | #N/A | -15.6701 | 75.2156 | 42.6632 | 87.3691 | #N/A | 42.6632 | 87.3691
| 2 |  |  | 3.8154 | 79.3751 | #N/A | 3.8154 | 79.3751 | -3.5217 | 28.2915 | #N/A | -3.5217 | 28.2915 | -4.2017 | 162.6541 | #N/A | -4.2017 | 162.6541 | 25.4783 | 79.3751 | #N/A | 25.4783 | 79.3751 | 19.0783 | 98.4354 | #N/A | 19.0783 | 98.4354 | 16.4783 | 51.0334 | #N/A | 16.4783 | 51.0334 | -10.0217 | 92.5837 | #N/A | -10.0217 | 92.5837 | 0.47832 | 46.9511 | #N/A | 0.47832 | 46.9511 | 33.6283 | 197.6282 | #N/A | 33.6283 | 197.6282 |  |  | 0 | #N/A | #N/A | -47.8517 | 145.1574 | #N/A | -47.8517 | 145.1574 | -19.6247 | 74.9989 | #N/A | -19.6247 | 74.9989 |  | -15.6075 | 15.6075 |  | 1.6183 | 27.3353 | #N/A | 1.6183 | 27.3353 | -24.105 | 71.054 | #N/A | -24.105 | 71.054 |  |  | 0 | #N/A | #N/A
| 1 |  |  | 1.6911 | 79.1009 | #N/A | 1.6911 | 79.1009 | -5.626 | 27.5128 | #N/A | -5.626 | 27.5128 | -3.576 | 162.5205 | #N/A | -3.576 | 162.5205 | 30.374 | 79.1009 | #N/A | 30.374 | 79.1009 | 19.704 | 98.2144 | #N/A | 19.704 | 98.2144 | 5.374 | 50.6059 | #N/A | 5.374 | 50.6059 | -18.126 | 92.3487 | #N/A | -18.126 | 92.3487 | -1.626 | 46.486 | #N/A | -1.626 | 46.486 | 83.724 | 395.5792 | #N/A | 83.724 | 395.5792 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A | -26.274 | 74.7087 | #N/A | -26.274 | 74.7087 |  | -16.9425 | 16.9425 |  | 5.154 | 32.8079 | #N/A | 5.154 | 32.8079 |  |  | 0 | #N/A | #N/A |  |  | 0 | #N/A | #N/A
Annex 4. TAB A4-3 auxiliary calculation for graphs