## Slide 1
Development of FPGA-based electronic modules for digital counting system at KRISS
Kyoung Beom Lee, Agung Agusbudiman
Korea Research Institute of Standards and Science
Daejeon, Republic of Korea
BIPM Workshop on Digital Electronics for the RMO SIRTIs 8th June 2022

## Slide 2
Development of FPGA-based TDCR counting system
1
2
3
4
Design and manufacture of optical chamber
Optical Chamber
Development of FPGA-based coincidence module
DAQ
Development of voltage divider
PMT characterization
Detector
Validation of measurement result
Validation
Method

## Slide 3
Fast amp
Coincidence and dead-time module
Counter
A
B
C
A
B
C
A
B
C
AB
BC
CA
T
D
Simplified schematic diagram of a TDCR counting system

## Slide 4
TDCR system at KRISS[1]
FPGA-BASED TDCR system
Hamamatsu R329-02 PMT
Coincidence module:
  NI cRIO system (Zynq-7020 FPGA)
  NI 9401 C-series module + NI 9924 screw-terminal block
KRISS TDCR counting system:
  Burle-RCA 8850 PMT (discontinued)
  MAC3 module is not a commercial product
  > 15 years old
[1] Lee, K.B., et al, 2004. Implementation of TDCR method in KRISS. Nucl. Instrum. Methods Phys. Res. Sect. A534, 496-502. https://doi.org/10.1016/j.nima.2004.06.150.

## Slide 5
Conventional setup for TDCR system
New setup for TDCR system
Purposes

## Slide 6
Modules specification

## Slide 7
Method
Development of FPGA-based coincidence module
Analog system
Network
Quad CFD
PC
LabView Real-Time System
FPGA
FPGA vi
Counter
Dead time
Coincidence logic
Real-Time Processor
RT Target vi
FPGA control
Data Acquisition
DAQ Setting
NI cRIO 9068
NI 9401
C-series DIO module
Notes: NI 9401 C series module
  8 DIO
  5V/ TTL

## Slide 8
Realization of coincidence module, dead-time module, and counter on FPGA
A single-cycle timed loop (SCTL) is used to contain the code implemented in the FPGA.vi.
FPGA clock is set at 80 MHz (derived clock)
Notes: When used with an FPGA target SCTL executes all functions inside within one tick of the FPGA clock that have been selected.
Limitation of SCTL
  Some function can not be implemented, e.g. while loop, divide function, etc
  higher clock = fewer function available

## Slide 9
FPGA coincidence module
Dead-time module & Counters
Trigger
Resolving time = 3 x SCTL = 37.5 ns
Coincidence module

## Slide 10
Extendable dead-time
Adjustable base dead-time
  Base dead-time 20 s
Extending-type dead-time
Triggering based on Common Dead-Time (CDT)
Resolving time: 37.5 ns
Dead-time
Notes: Dead-time used to prevent spurious after-pulse from the PMT
Extendable dead-type provide a cleaner result and simpler

## Slide 11
FPGA dead-time (DT) module and counters
Other event detected within DT range
1st event trigger DT counter and counters
Compared with base DT
FPGA VI

## Slide 12
RT vi

## Slide 13
Front Panel of the RT vi
Typical output file from the DAQ

## Slide 14
Validation
Comparison between MAC3 and FPGA-based system:
Output count rates
Activity measurement of 3H and 14C
[2] Agusbudiman, A., et al, 2022. Development of FPGA-based coincidence module for TDCR counting system. Radiat. Phys. Chem. https://doi.org/10.1016/j.radphyschem.2022.110377
Schematic diagram of TDCR system used to evaluate the performance of FPGA-based coincidence module. [2]
Preliminary Test

## Slide 15
Count distribution comparison between MAC3 and FPGA
 | MAC3-FPGA (%)
AB | 0.18
BC | 0.13
CA | 0.21
Preliminary Results

## Slide 16
Preliminary Results
Preliminary result of the FPGA-based module[2] :
TDCR value comparison with base dead-time variation
Double efficiency comparison with different level of count rates
[2] Agusbudiman, A., et al, 2022. Development of FPGA-based coincidence module for TDCR counting system. Radiat. Phys. Chem. https://doi.org/10.1016/j.radphyschem.2022.110377

## Slide 17
Evaluation – measurement result of 3H and 14C
The preliminary evaluation results show a good agreement between the FPGA-based and the MAC3 system.
For both 3H and 14C cases, the discrepancies of activity result between the FPGA-based and the MAC3 system is around 0.20%, reflecting the difference obtained on the count rates comparison.
Preliminary Results

## Slide 18
Application on SIRTI
[3] Michotte, C., et al, 2013. The SIRTI: a new tool developed at the BIMP for comparing activity measurements of short-lived radionuclides world-wide. Rapport BIPM-2013/02.
Block diagram of the SIRTI. [3]
The development of FPGA-based electronic modules for the TDCR system above does not use any digitizer; therefore, this technique can only be applied to the SIRTI to replace the functionality of the MTR2 and NIM scalers module.

## Slide 19
Thank you