## Introduction
APMP.AUV.A-K5 Final Report Tables of Data
This Microsoft Excel spreadsheet contains supplementary data for Report on Key Comparison APMP.AUV.A-K5
The size and number of these tables made it impractical to include them within the body of the report.
The 'Parameters' tab lists all declared microphone paramaters used by each participant.
For each of the two reference microphones (4160.2652765 and 4160.2652762) there are the following tabs:
Level | Tables of declared sensitivity level measurements and associated expanded uncertainties for all participating laboratories
Phase | Tables of declared sensitivity phase measurements and associated expanded uncertainties for all participating laboratories
Additionally for 4160.2652765 there are the following tabs:
DoE Level | Degrees of equivalence for sensitivity level relative to the KCRV value derived through CCAUV.A-K5
DoE Phase | Degrees of equivalence for sensitivity phase, relative to the KCRV value derived through CCAUV.A-K5
Spreadsheet protection
This spreadsheet is password protected. The password can be requested from fxj@nim.ac.cn

## Parameters
Declared Parameters
2652762 | NIM | NMIA | SCL | NMIJ | KRISS | NIMT | NMIM | CMS-ITRI
total volume (mm3)
front volume (mm3) | 537 | 533 | 542 | 538 | 539.3 | 544 | 544 | 531.6
cavity depth (mm) | 1.95 | 1.96 | 1.99 | 1.96 | 1.965 | 1.979 | 1.97 | 1.96
acoustic mass (kg m-4)
acoustic compliance (kg-1 m4 s2)
acoustic resistance (kg m-4 s-1)
loss factor | 1.07 | 1.01 | 1.05 | 1.05 | 1.05 | 1 | 1.05 | 1.05
res. frequency (kHz) | 8.46 | 8.3 | 8.2 | 8.2 | 8.2 | 8.3 | 8.2 | 8.2
equivalent volume (mm3) | 140 | 136 | 128.7 | 129 | 129.1 | 124 | 130 | 140.7
temp. coeff. at 250 Hz (dB/K)(Level) | -0.002 |  | -0.0149 |  | -0.015 |  | -0.002 | -0.016
press. coeff. at 250 Hz (dB/kPa)(Level) | -0.016 |  | -0.003 |  | -0.0025 |  | -0.015 | -0.002
temp. coeff. at 250 Hz (deg/K)(Phase)
press. coeff. at 250 Hz (deg/kPa)(Phase)
diaphragm diameter (mm) |  | 17.9 | 17.9 |  |  | 17.9 |  | 17.9
damping
front cavity excessive surface/added thread area (mm2) |  | 0 | 0
2652765 | NIM | NMIA | SCL | NMIJ | KRISS | NIMT | NMIM | CMS-ITRI
total volume (mm3)
front volume (mm3) | 534 | 533 | 542 | 536 | 539.6 | 524.7 | 542 | 532.6
cavity depth (mm) | 1.94 | 1.96 | 1.99 | 1.95 | 1.967 | 1.903 | 1.97 | 1.95
acoustic mass (kg m-4)
acoustic compliance (kg-1 m4 s2)
acoustic resistance (kg m-4 s-1)
loss factor | 1.09 | 1.01 | 1.05 | 1.05 | 1.05 | 1.05 | 1.05 | 1.05
res. frequency (kHz) | 8.58 | 8.25 | 8.2 | 8.2 | 8.2 | 8.2 | 8.2 | 8.2
equivalent volume (mm3) | 140 | 139 | 133.6 | 141 | 134.5 | 151.9 | 135 | 143.8
temp. coeff. at 250 Hz (dB/K)(Level) | -0.016 |  | -0.0156 |  | -0.015 |  | -0.002 | -0.016
press. coeff. at 250 Hz (dB/kPa)(Level) | -0.002 |  | -0.003 |  | -0.0025 |  | -0.015 | -0.002
temp. coeff. at 250 Hz (deg/K)(Phase)
press. coeff. at 250 Hz (deg/kPa)(Phase)
diaphragm diameter (mm) |  | 17.9 | 17.9 |  |  | 17.9 |  | 17.9
damping
front cavity excessive surface/added thread area (mm2) |  | 0 | 0

## 2652762 Level
4160. 2652762 | Sensitivity Level (dB re 1 V/Pa) | [7 empty] | Expanded Uncertainties (dB) | [7 empty] | Unweighted mean of sensitivity level (dB re 1 V/Pa) | Level difference relative to unweighted mean of sensitivity level (dB)
Frequency (Hz) | NIM | NMIA | SCL | NMIJ | KRISS | NIMT | NMIM | CMS-ITRI | NIM | NMIA | SCL | NMIJ | KRISS | NIMT | NMIM | CMS-ITRI |  | NIM | NMIA | SCL | NMIJ | KRISS | NIMT | NMIM | CMS-ITRI
1.9953 | -26.2502 | -26.349 |  | -26.46 | -26.492 | -26.37 |  | -26.3777 | 0.2 | 0.2 |  | 0.23 | 0.09 | 0.2 |  | 0.05 | -26.3831 | 0.13292 | 0.034146 |  | -0.076854 | -0.10885 | 0.013146 |  | 0.005493 | Frequency | NMIM
2.5119 | -26.3833 | -26.482 |  | -26.58 | -26.619 | -26.5 |  | -26.5089 | 0.18 | 0.15 |  | 0.22 | 0.08 | 0.2 |  | 0.05 | -26.5122 | 0.1289 | 0.030197 |  | -0.067803 | -0.1068 | 0.012197 |  | 0.0033147 | 19.9526 | 0.018771
3.1623 | -26.5004 | -26.597 |  | -26.6 | -26.714 | -26.63 |  | -26.6334 | 0.16 | 0.13 |  | 0.22 | 0.07 | 0.2 |  | 0.05 | -26.6125 | 0.11202 | 0.01547 |  | 0.01247 | -0.10153 | -0.01753 |  | -0.020899 | 31.6228 | 0.012335
3.9811 | -26.6093 | -26.713 |  | -26.74 | -26.819 | -26.74 |  | -26.7455 | 0.16 | 0.12 |  | 0.19 | 0.07 | 0.2 |  | 0.05 | -26.7278 | 0.11848 | 0.014795 |  | -0.012205 | -0.091205 | -0.012205 |  | -0.017662 | 63.0957 | 0.021086
5.0119 | -26.7056 | -26.812 |  | -26.85 | -26.902 | -26.85 |  | -26.8453 | 0.1 | 0.11 |  | 0.09 | 0.06 | 0.2 |  | 0.05 | -26.8275 | 0.12185 | 0.015481 |  | -0.022519 | -0.074519 | -0.022519 |  | -0.017771 | 125.8925 | 0.019878
6.3096 | -26.7904 | -26.896 |  | -26.94 | -26.985 | -26.93 |  | -26.9339 | 0.1 | 0.1 |  | 0.08 | 0.06 | 0.13 |  | 0.05 | -26.9126 | 0.12211 | 0.016554 |  | -0.027446 | -0.072446 | -0.017446 |  | -0.021327 | 251.1886 | 0.022448
7.9433 | -26.8629 | -26.977 |  | -27.01 | -27.053 | -27.01 |  | -27.009 | 0.08 | 0.09 |  | 0.05 | 0.06 | 0.09 |  | 0.05 | -26.987 | 0.12407 | 0.0099835 |  | -0.023017 | -0.066017 | -0.023017 |  | -0.022008 | 501.1872 | 0.027872
10 | -26.9228 | -27.037 |  | -27.08 | -27.11 | -27.08 |  | -27.072 | 0.08 | 0.08 |  | 0.05 | 0.06 | 0.06 |  | 0.05 | -27.0503 | 0.12748 | 0.013308 |  | -0.029692 | -0.059692 | -0.029692 |  | -0.021713 | 1000 | 0.026294
12.5893 | -26.9722 | -27.084 |  | -27.13 | -27.157 | -27.13 |  | -27.1243 | 0.08 | 0.08 |  | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.0996 | 0.12742 | 0.015586 |  | -0.030414 | -0.057414 | -0.030414 |  | -0.024763 | 1258.9254 | 0.025091
15.8489 | -27.0142 | -27.134 |  | -27.17 | -27.197 | -27.17 |  | -27.1674 | 0.08 | 0.07 |  | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.1421 | 0.12791 | 0.0080978 |  | -0.027902 | -0.054902 | -0.027902 |  | -0.025304 | 1584.8932 | 0.033069
19.9526 | -27.0481 | -27.166 | -27.2029 | -27.21 | -27.229 | -27.21 | -27.16 | -27.2042 | 0.07 | 0.07 | 0.06 | 0.04 | 0.06 | 0.05 | 0.07 | 0.05 | -27.1788 | 0.13068 | 0.012762 | -0.024092 | -0.031238 | -0.050238 | -0.031238 | 0.018771 | -0.025411 | 1995.2623 | 0.027494
21.1349 | -27.0554 | -27.173 | -27.2114 | -27.21 | -27.239 | -27.22 |  | -27.2119 | 0.06 | 0.07 | 0.06 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.1887 | 0.13324 | 0.015678 | -0.022698 | -0.021322 | -0.050322 | -0.031322 |  | -0.023252 | 2511.8864 | 0.024434
22.3872 | -27.0625 | -27.181 | -27.2207 | -27.22 | -27.245 | -27.23 |  | -27.2198 | 0.06 | 0.07 | 0.06 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.197 | 0.13448 | 0.016001 | -0.023694 | -0.022999 | -0.047999 | -0.032999 |  | -0.022792 | 3162.2777 | 0.027512
23.7137 | -27.0693 | -27.188 | -27.2263 | -27.23 | -27.252 | -27.24 |  | -27.2273 | 0.06 | 0.07 | 0.06 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.2047 | 0.13538 | 0.016698 | -0.021591 | -0.025302 | -0.047302 | -0.035302 |  | -0.022583 | 3981.0717 | 0.03254
25.1189 | -27.0758 | -27.195 | -27.2337 | -27.24 | -27.26 | -27.24 |  | -27.2347 | 0.06 | 0.06 | 0.05 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.2113 | 0.1355 | 0.01632 | -0.022408 | -0.02868 | -0.04868 | -0.02868 |  | -0.023375 | 5011.8723 | 0.028189
26.6073 | -27.082 | -27.202 | -27.2405 | -27.24 | -27.266 | -27.25 |  | -27.2415 | 0.06 | 0.06 | 0.05 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.2174 | 0.13541 | 0.015439 | -0.023074 | -0.022561 | -0.048561 | -0.032561 |  | -0.024092 | 6309.5734 | 0.028836
28.1838 | -27.088 | -27.208 | -27.2468 | -27.25 | -27.273 | -27.26 |  | -27.2482 | 0.06 | 0.06 | 0.05 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.2249 | 0.1369 | 0.016854 | -0.02194 | -0.025146 | -0.048146 | -0.035146 |  | -0.023374 | 7943.2823 | -0.0053821
29.8538 | -27.0936 | -27.215 | -27.2539 | -27.26 | -27.277 | -27.26 |  | -27.2548 | 0.06 | 0.06 | 0.05 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.2306 | 0.13701 | 0.015626 | -0.023298 | -0.029374 | -0.046374 | -0.029374 |  | -0.024213 | 10000 | 0.093391
31.6228 | -27.099 | -27.221 | -27.2595 | -27.27 | -27.283 | -27.27 | -27.2236 | -27.2612 | 0.06 | 0.06 | 0.05 | 0.04 | 0.06 | 0.05 | 0.05 | 0.05 | -27.2359 | 0.13689 | 0.014917 | -0.023631 | -0.034083 | -0.047083 | -0.034083 | 0.012335 | -0.025262
33.4965 | -27.1042 | -27.227 | -27.2657 | -27.27 | -27.289 | -27.28 |  | -27.267 | 0.05 | 0.06 | 0.05 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.2433 | 0.13906 | 0.01627 | -0.022449 | -0.02673 | -0.04573 | -0.03673 |  | -0.023691
35.4813 | -27.1092 | -27.232 | -27.2702 | -27.28 | -27.295 | -27.28 |  | -27.2729 | 0.05 | 0.06 | 0.05 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.2485 | 0.13928 | 0.016466 | -0.02172 | -0.031534 | -0.046534 | -0.031534 |  | -0.024424
37.5837 | -27.114 | -27.238 | -27.276 | -27.28 | -27.299 | -27.29 |  | -27.2785 | 0.05 | 0.06 | 0.05 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.2536 | 0.13964 | 0.015635 | -0.022358 | -0.026365 | -0.045365 | -0.036365 |  | -0.024826
39.8107 | -27.1186 | -27.243 | -27.2817 | -27.29 | -27.304 | -27.29 |  | -27.284 | 0.05 | 0.06 | 0.05 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.2588 | 0.14012 | 0.015763 | -0.022926 | -0.031237 | -0.045237 | -0.031237 |  | -0.025248
42.1697 | -27.1232 | -27.248 | -27.286 | -27.29 | -27.309 | -27.3 |  | -27.2891 | 0.05 | 0.06 | 0.05 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.2636 | 0.14045 | 0.015607 | -0.022345 | -0.026393 | -0.045393 | -0.036393 |  | -0.025527
44.6684 | -27.1276 | -27.253 | -27.2921 | -27.3 | -27.312 | -27.3 |  | -27.2942 | 0.05 | 0.06 | 0.05 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.2684 | 0.14083 | 0.015405 | -0.023666 | -0.031595 | -0.043595 | -0.031595 |  | -0.025788
47.3151 | -27.1319 | -27.258 | -27.297 | -27.31 | -27.317 | -27.31 |  | -27.2994 | 0.05 | 0.06 | 0.05 | 0.04 | 0.06 | 0.05 |  | 0.05 | -27.2748 | 0.14288 | 0.016765 | -0.022268 | -0.035235 | -0.042235 | -0.035235 |  | -0.024673
50.1187 | -27.1361 | -27.262 | -27.3015 | -27.31 | -27.322 | -27.31 |  | -27.3041 | 0.05 | 0.06 | 0.04 | 0.04 | 0.05 | 0.05 |  | 0.05 | -27.278 | 0.14186 | 0.015965 | -0.023549 | -0.032035 | -0.044035 | -0.032035 |  | -0.026175
53.0884 | -27.1402 | -27.266 | -27.3062 | -27.32 | -27.326 | -27.32 |  | -27.3088 | 0.05 | 0.06 | 0.04 | 0.04 | 0.05 | 0.05 |  | 0.05 | -27.2839 | 0.14368 | 0.017896 | -0.022341 | -0.036104 | -0.042104 | -0.036104 |  | -0.024922
56.2341 | -27.1442 | -27.27 | -27.3107 | -27.32 | -27.331 | -27.32 |  | -27.3135 | 0.05 | 0.06 | 0.04 | 0.04 | 0.05 | 0.05 |  | 0.05 | -27.2871 | 0.14284 | 0.017063 | -0.023637 | -0.032937 | -0.043937 | -0.032937 |  | -0.026455
59.5662 | -27.1481 | -27.274 | -27.315 | -27.32 | -27.334 | -27.33 |  | -27.318 | 0.05 | 0.06 | 0.04 | 0.04 | 0.05 | 0.05 |  | 0.05 | -27.2913 | 0.14321 | 0.017307 | -0.023715 | -0.028693 | -0.042693 | -0.038693 |  | -0.026726
63.0957 | -27.1518 | -27.279 | -27.3193 | -27.33 | -27.338 | -27.33 | -27.2717 | -27.3221 | 0.04 | 0.06 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | -27.2927 | 0.14093 | 0.013736 | -0.026561 | -0.037264 | -0.045264 | -0.037264 | 0.021086 | -0.029401
66.8344 | -27.1553 | -27.282 | -27.3234 | -27.33 | -27.342 | -27.34 |  | -27.3263 | 0.04 | 0.06 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.2999 | 0.14452 | 0.017856 | -0.023501 | -0.030144 | -0.042144 | -0.040144 |  | -0.026445
70.7946 | -27.1587 | -27.286 | -27.3274 | -27.34 | -27.346 | -27.34 |  | -27.3304 | 0.04 | 0.06 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3041 | 0.14541 | 0.018072 | -0.023326 | -0.035928 | -0.041928 | -0.035928 |  | -0.026371
74.9894 | -27.1618 | -27.29 | -27.3312 | -27.35 | -27.348 | -27.34 |  | -27.3343 | 0.04 | 0.06 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3079 | 0.14612 | 0.017901 | -0.023292 | -0.042099 | -0.040099 | -0.032099 |  | -0.026432
79.4328 | -27.1647 | -27.293 | -27.3351 | -27.35 | -27.352 | -27.35 |  | -27.3383 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3119 | 0.14718 | 0.018875 | -0.023222 | -0.038125 | -0.040125 | -0.038125 |  | -0.026459
84.1395 | -27.1674 | -27.297 | -27.3387 | -27.35 | -27.356 | -27.35 |  | -27.3418 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3144 | 0.14699 | 0.017416 | -0.024267 | -0.035584 | -0.041584 | -0.035584 |  | -0.02739
89.1251 | -27.17 | -27.3 | -27.3423 | -27.36 | -27.358 | -27.35 |  | -27.3455 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.318 | 0.14795 | 0.017971 | -0.024356 | -0.042029 | -0.040029 | -0.032029 |  | -0.027481
94.4061 | -27.1725 | -27.303 | -27.3458 | -27.36 | -27.362 | -27.36 |  | -27.349 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3218 | 0.14922 | 0.018756 | -0.024029 | -0.038244 | -0.040244 | -0.038244 |  | -0.027214
100 | -27.175 | -27.306 | -27.3491 | -27.36 | -27.365 | -27.36 |  | -27.3524 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3239 | 0.1489 | 0.017941 | -0.025159 | -0.036059 | -0.041059 | -0.036059 |  | -0.028508
105.9254 | -27.1776 | -27.309 | -27.3525 | -27.36 | -27.369 | -27.36 |  | -27.3558 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3263 | 0.14869 | 0.017274 | -0.026238 | -0.033726 | -0.042726 | -0.033726 |  | -0.029552
112.2018 | -27.1802 | -27.313 | -27.3559 | -27.37 | -27.372 | -27.37 |  | -27.359 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3315 | 0.15124 | 0.018451 | -0.024459 | -0.038549 | -0.040549 | -0.038549 |  | -0.02759
118.8502 | -27.1829 | -27.316 | -27.3589 | -27.37 | -27.373 | -27.37 |  | -27.3622 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3333 | 0.15039 | 0.017299 | -0.025637 | -0.036701 | -0.039701 | -0.036701 |  | -0.028949
125.8925 | -27.1856 | -27.319 | -27.3619 | -27.38 | -27.376 | -27.37 | -27.3141 | -27.3651 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | -27.334 | 0.14833 | 0.014965 | -0.027907 | -0.046035 | -0.042035 | -0.036035 | 0.019878 | -0.031159
133.3521 | -27.1883 | -27.321 | -27.3646 | -27.37 | -27.38 | -27.38 |  | -27.3682 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3389 | 0.15059 | 0.017878 | -0.02577 | -0.031122 | -0.041122 | -0.041122 |  | -0.029328
141.2538 | -27.1908 | -27.324 | -27.3674 | -27.38 | -27.383 | -27.38 |  | -27.3711 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3423 | 0.15156 | 0.018319 | -0.025097 | -0.037681 | -0.040681 | -0.037681 |  | -0.028744
149.6236 | -27.1929 | -27.326 | -27.3702 | -27.38 | -27.384 | -27.38 |  | -27.3739 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3439 | 0.15094 | 0.017861 | -0.02638 | -0.036139 | -0.040139 | -0.036139 |  | -0.030004
158.4893 | -27.1948 | -27.329 | -27.3728 | -27.38 | -27.388 | -27.39 |  | -27.3766 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3473 | 0.15254 | 0.0183 | -0.025476 | -0.0327 | -0.0407 | -0.0427 |  | -0.029265
167.8804 | -27.1964 | -27.331 | -27.3754 | -27.39 | -27.387 | -27.39 |  | -27.3793 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3499 | 0.1535 | 0.01887 | -0.025539 | -0.04013 | -0.03713 | -0.04013 |  | -0.029443
177.8279 | -27.1988 | -27.333 | -27.3779 | -27.39 | -27.394 | -27.39 |  | -27.3819 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.3522 | 0.15342 | 0.019232 | -0.025678 | -0.037768 | -0.041768 | -0.037768 |  | -0.029665
188.3649 | -27.2007 | -27.336 | -27.3802 | -27.39 | -27.397 | -27.39 |  | -27.3841 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 |  | 0.05 | -27.354 | 0.15328 | 0.018005 | -0.026152 | -0.035995 | -0.042995 | -0.035995 |  | -0.030144
199.5262 | -27.2026 | -27.338 | -27.3824 | -27.4 | -27.398 | -27.4 |  | -27.3867 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3582 | 0.15565 | 0.020242 | -0.024186 | -0.041758 | -0.039758 | -0.041758 |  | -0.028431
211.3489 | -27.2046 | -27.34 | -27.3846 | -27.4 | -27.4 | -27.4 |  | -27.3891 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3598 | 0.15514 | 0.019751 | -0.024819 | -0.040249 | -0.040249 | -0.040249 |  | -0.02932
223.8721 | -27.2063 | -27.342 | -27.3866 | -27.4 | -27.401 | -27.4 |  | -27.3912 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.361 | 0.15471 | 0.019021 | -0.025627 | -0.038979 | -0.039979 | -0.038979 |  | -0.030166
237.1374 | -27.2079 | -27.344 | -27.3885 | -27.4 | -27.405 | -27.4 |  | -27.3935 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3627 | 0.1548 | 0.018703 | -0.02584 | -0.037297 | -0.042297 | -0.037297 |  | -0.030767
251.1886 | -27.2094 | -27.345 | -27.3903 | -27.4 | -27.408 | -27.4 | -27.3383 | -27.3953 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -27.3608 | 0.15139 | 0.015793 | -0.0295 | -0.039207 | -0.047207 | -0.039207 | 0.022448 | -0.034509
266.0725 | -27.2108 | -27.347 | -27.3923 | -27.4 | -27.409 | -27.41 |  | -27.3975 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3666 | 0.15584 | 0.019646 | -0.025605 | -0.033354 | -0.042354 | -0.043354 |  | -0.030818
281.8383 | -27.2121 | -27.349 | -27.3939 | -27.41 | -27.411 | -27.41 |  | -27.3992 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3693 | 0.1572 | 0.020315 | -0.024563 | -0.040685 | -0.041685 | -0.040685 |  | -0.0299
298.5383 | -27.2133 | -27.35 | -27.3955 | -27.41 | -27.412 | -27.41 |  | -27.4012 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3703 | 0.15697 | 0.020291 | -0.025224 | -0.039709 | -0.041709 | -0.039709 |  | -0.030907
316.2278 | -27.2144 | -27.352 | -27.3971 | -27.41 | -27.414 | -27.41 |  | -27.4027 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3715 | 0.15702 | 0.01946 | -0.025658 | -0.03854 | -0.04254 | -0.03854 |  | -0.031203
334.9654 | -27.2155 | -27.353 | -27.3985 | -27.41 | -27.415 | -27.41 |  | -27.4041 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3723 | 0.15685 | 0.019301 | -0.026208 | -0.037699 | -0.042699 | -0.037699 |  | -0.031845
354.8134 | -27.2164 | -27.354 | -27.3998 | -27.41 | -27.416 | -27.41 |  | -27.4054 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3731 | 0.15672 | 0.019082 | -0.026702 | -0.036918 | -0.042918 | -0.036918 |  | -0.032345
375.8374 | -27.2172 | -27.355 | -27.401 | -27.42 | -27.417 | -27.42 |  | -27.4066 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3767 | 0.15952 | 0.02168 | -0.024316 | -0.04332 | -0.04032 | -0.04332 |  | -0.029924
398.1072 | -27.2178 | -27.356 | -27.4021 | -27.42 | -27.418 | -27.42 |  | -27.4078 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3774 | 0.15956 | 0.021397 | -0.024713 | -0.042603 | -0.040603 | -0.042603 |  | -0.03044
421.6965 | -27.2184 | -27.357 | -27.4029 | -27.41 | -27.419 | -27.42 |  | -27.4087 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3766 | 0.1582 | 0.019572 | -0.026365 | -0.033428 | -0.042428 | -0.043428 |  | -0.032126
446.6836 | -27.2188 | -27.357 | -27.4038 | -27.41 | -27.42 | -27.42 |  | -27.4095 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.377 | 0.15825 | 0.02 | -0.026783 | -0.033 | -0.043 | -0.043 |  | -0.032464
473.1513 | -27.219 | -27.357 | -27.4044 | -27.42 | -27.42 | -27.42 |  | -27.4102 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3786 | 0.15968 | 0.021639 | -0.025714 | -0.041361 | -0.041361 | -0.041361 |  | -0.031526
501.1872 | -27.2188 | -27.358 | -27.4048 | -27.42 | -27.42 | -27.42 | -27.347 | -27.4106 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -27.3749 | 0.15608 | 0.016905 | -0.029846 | -0.045095 | -0.045095 | -0.045095 | 0.027872 | -0.035731
530.8844 | -27.2186 | -27.358 | -27.4048 | -27.42 | -27.42 | -27.42 |  | -27.4108 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3789 | 0.16033 | 0.020882 | -0.025957 | -0.041118 | -0.041118 | -0.041118 |  | -0.031901
562.3413 | -27.2181 | -27.357 | -27.4049 | -27.42 | -27.42 | -27.42 |  | -27.4106 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3786 | 0.16057 | 0.021649 | -0.026237 | -0.041351 | -0.041351 | -0.041351 |  | -0.031934
595.6621 | -27.2174 | -27.357 | -27.4046 | -27.42 | -27.419 | -27.42 |  | -27.4105 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3784 | 0.16099 | 0.021354 | -0.026271 | -0.041646 | -0.040646 | -0.041646 |  | -0.03214
630.9573 | -27.2164 | -27.356 | -27.404 | -27.42 | -27.418 | -27.42 |  | -27.4098 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3777 | 0.16136 | 0.021738 | -0.026237 | -0.042262 | -0.040262 | -0.042262 |  | -0.03207
668.3439 | -27.2151 | -27.355 | -27.403 | -27.42 | -27.418 | -27.42 |  | -27.4092 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3772 | 0.16208 | 0.022186 | -0.025853 | -0.042814 | -0.040814 | -0.042814 |  | -0.031966
707.9458 | -27.2135 | -27.354 | -27.4018 | -27.41 | -27.417 | -27.42 |  | -27.4079 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3749 | 0.16138 | 0.020884 | -0.026904 | -0.035116 | -0.042116 | -0.045116 |  | -0.033006
749.8942 | -27.2115 | -27.352 | -27.4002 | -27.41 | -27.415 | -27.42 |  | -27.4064 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3736 | 0.16204 | 0.02159 | -0.026597 | -0.03641 | -0.04141 | -0.04641 |  | -0.032807
794.3282 | -27.2092 | -27.35 | -27.3981 | -27.41 | -27.412 | -27.41 |  | -27.4043 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3705 | 0.16134 | 0.020521 | -0.027615 | -0.039479 | -0.041479 | -0.039479 |  | -0.033809
841.3951 | -27.2064 | -27.347 | -27.3958 | -27.41 | -27.41 | -27.41 |  | -27.4021 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3687 | 0.16238 | 0.021748 | -0.027063 | -0.041252 | -0.041252 | -0.041252 |  | -0.033306
891.2509 | -27.2031 | -27.344 | -27.3929 | -27.41 | -27.407 | -27.41 |  | -27.399 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3666 | 0.16351 | 0.022576 | -0.026366 | -0.043424 | -0.040424 | -0.043424 |  | -0.032448
944.0609 | -27.1992 | -27.34 | -27.3892 | -27.4 | -27.402 | -27.4 |  | -27.3953 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3608 | 0.16161 | 0.020825 | -0.028395 | -0.039175 | -0.041175 | -0.039175 |  | -0.034513
1000 | -27.1948 | -27.336 | -27.385 | -27.4 | -27.398 | -27.4 | -27.3278 | -27.3913 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -27.3541 | 0.15936 | 0.018119 | -0.0309 | -0.045881 | -0.043881 | -0.045881 | 0.026294 | -0.037224
1059.2537 | -27.1896 | -27.331 | -27.3802 | -27.39 | -27.394 | -27.4 |  | -27.3864 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.353 | 0.16339 | 0.022033 | -0.02719 | -0.036967 | -0.040967 | -0.046967 |  | -0.033333
1122.0185 | -27.1838 | -27.326 | -27.3747 | -27.39 | -27.389 | -27.39 |  | -27.3809 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3478 | 0.16399 | 0.021759 | -0.026931 | -0.042241 | -0.041241 | -0.042241 |  | -0.033092
1188.5022 | -27.1771 | -27.319 | -27.3682 | -27.38 | -27.381 | -27.38 |  | -27.3741 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3399 | 0.16283 | 0.020905 | -0.028293 | -0.040095 | -0.041095 | -0.040095 |  | -0.034155
1258.9254 | -27.1695 | -27.312 | -27.3607 | -27.38 | -27.374 | -27.38 | -27.3061 | -27.367 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -27.3312 | 0.16169 | 0.019158 | -0.02958 | -0.048842 | -0.042842 | -0.048842 | 0.025091 | -0.035831
1333.5214 | -27.1609 | -27.303 | -27.3523 | -27.37 | -27.366 | -27.37 |  | -27.3585 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3258 | 0.16496 | 0.022809 | -0.026521 | -0.044191 | -0.040191 | -0.044191 |  | -0.032674
1412.5375 | -27.1511 | -27.294 | -27.343 | -27.36 | -27.355 | -27.36 |  | -27.3494 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3161 | 0.16497 | 0.022076 | -0.026917 | -0.043924 | -0.038924 | -0.043924 |  | -0.033353
1496.2357 | -27.1401 | -27.283 | -27.3321 | -27.35 | -27.342 | -27.35 |  | -27.3385 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.3051 | 0.16499 | 0.022111 | -0.027007 | -0.044889 | -0.036889 | -0.044889 |  | -0.033424
1584.8932 | -27.1278 | -27.27 | -27.3199 | -27.34 | -27.331 | -27.34 | -27.2558 | -27.3264 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -27.2889 | 0.16109 | 0.018864 | -0.031059 | -0.051136 | -0.042136 | -0.051136 | 0.033069 | -0.037555
1678.804 | -27.1139 | -27.257 | -27.3061 | -27.32 | -27.318 | -27.32 |  | -27.3127 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.2782 | 0.16432 | 0.021245 | -0.027864 | -0.041755 | -0.039755 | -0.041755 |  | -0.034442
1778.2794 | -27.0984 | -27.241 | -27.2905 | -27.31 | -27.302 | -27.31 |  | -27.2967 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.2641 | 0.16567 | 0.023087 | -0.026408 | -0.045913 | -0.037913 | -0.045913 |  | -0.032605
1883.6491 | -27.0811 | -27.224 | -27.2731 | -27.29 | -27.282 | -27.29 |  | -27.2795 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.2457 | 0.16455 | 0.021665 | -0.027398 | -0.044335 | -0.036335 | -0.044335 |  | -0.033811
1995.2623 | -27.0618 | -27.205 | -27.2536 | -27.26 | -27.264 | -27.27 | -27.1935 | -27.2604 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -27.221 | 0.15921 | 0.016042 | -0.032511 | -0.038958 | -0.042958 | -0.048958 | 0.027494 | -0.039356
2113.489 | -27.0404 | -27.184 | -27.2316 | -27.25 | -27.245 | -27.25 |  | -27.238 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.2056 | 0.1652 | 0.021575 | -0.026061 | -0.044425 | -0.039425 | -0.044425 |  | -0.032436
2238.7211 | -27.0165 | -27.16 | -27.2073 | -27.22 | -27.216 | -27.22 |  | -27.2139 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.1791 | 0.16264 | 0.0191 | -0.028228 | -0.0409 | -0.0369 | -0.0409 |  | -0.034807
2371.3737 | -26.9897 | -27.133 | -27.1803 | -27.2 | -27.19 | -27.2 |  | -27.1869 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.1543 | 0.16458 | 0.021268 | -0.026019 | -0.045732 | -0.035732 | -0.045732 |  | -0.032632
2511.8864 | -26.9594 | -27.104 | -27.1508 | -27.17 | -27.16 | -27.17 | -27.0965 | -27.1566 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -27.1209 | 0.1615 | 0.016903 | -0.029872 | -0.049097 | -0.039097 | -0.049097 | 0.024434 | -0.035675
2660.7251 | -26.9267 | -27.071 | -27.117 | -27.14 | -27.124 | -27.13 |  | -27.1234 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0903 | 0.16358 | 0.019302 | -0.026652 | -0.049698 | -0.033698 | -0.039698 |  | -0.033139
2818.3829 | -26.8907 | -27.035 | -27.0805 | -27.1 | -27.086 | -27.1 |  | -27.087 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0542 | 0.16348 | 0.019167 | -0.026345 | -0.045833 | -0.031833 | -0.045833 |  | -0.032804
2985.3826 | -26.8512 | -26.997 | -27.0404 | -27.06 | -27.048 | -27.06 |  | -27.0467 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0148 | 0.16355 | 0.01775 | -0.025624 | -0.04525 | -0.03325 | -0.04525 |  | -0.031927
3162.2777 | -26.8082 | -26.954 | -26.9971 | -27.02 | -27.006 | -27.01 | -26.9396 | -27.0024 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -26.9672 | 0.15899 | 0.013154 | -0.029921 | -0.052846 | -0.038846 | -0.042846 | 0.027512 | -0.035197
3349.6544 | -26.7616 | -26.912 | -26.9498 | -26.97 | -26.955 | -26.97 |  | -26.9548 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.9247 | 0.16318 | 0.01275 | -0.0251 | -0.04525 | -0.03025 | -0.04525 |  | -0.030081
3548.1339 | -26.7115 | -26.86 | -26.8988 | -26.92 | -26.906 | -26.92 |  | -26.9032 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.8742 | 0.16268 | 0.014223 | -0.024584 | -0.045777 | -0.031777 | -0.045777 |  | -0.028984
3758.374 | -26.6585 | -26.807 | -26.8447 | -26.87 | -26.853 | -26.86 |  | -26.8485 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.8202 | 0.16175 | 0.013233 | -0.02446 | -0.049767 | -0.032767 | -0.039767 |  | -0.028225
3981.0717 | -26.6031 | -26.756 | -26.7885 | -26.82 | -26.797 | -26.81 | -26.7294 | -26.7913 | 0.04 | 0.05 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -26.7619 | 0.15882 | 0.0059195 | -0.026614 | -0.05808 | -0.03508 | -0.04808 | 0.03254 | -0.029423
4216.965 | -26.5466 | -26.698 | -26.7301 | -26.76 | -26.736 | -26.75 |  | -26.7324 | 0.05 | 0.05 | 0.05 | 0.04 | 0.03 | 0.05 |  | 0.05 | -26.7076 | 0.16097 | 0.0095889 | -0.022551 | -0.052411 | -0.028411 | -0.042411 |  | -0.024773
4466.8359 | -26.4909 | -26.643 | -26.6721 | -26.71 | -26.673 | -26.69 |  | -26.672 | 0.05 | 0.05 | 0.05 | 0.04 | 0.03 | 0.05 |  | 0.05 | -26.6501 | 0.15921 | 0.0071351 | -0.021917 | -0.059865 | -0.022865 | -0.039865 |  | -0.02183
4731.5126 | -26.4388 | -26.592 | -26.6159 | -26.65 | -26.622 | -26.64 |  | -26.6132 | 0.05 | 0.05 | 0.05 | 0.04 | 0.03 | 0.05 |  | 0.06 | -26.596 | 0.15722 | 0.0039681 | -0.019898 | -0.054032 | -0.026032 | -0.044032 |  | -0.01719
5011.8723 | -26.3938 | -26.541 | -26.5656 | -26.61 | -26.573 | -26.59 | -26.5155 | -26.5605 | 0.06 | 0.05 | 0.05 | 0.04 | 0.03 | 0.05 | 0.05 | 0.06 | -26.5437 | 0.14985 | 0.0026724 | -0.021933 | -0.066328 | -0.029328 | -0.046328 | 0.028189 | -0.016795
5308.8444 | -26.3597 | -26.507 | -26.5253 | -26.57 | -26.533 | -26.55 |  | -26.5153 | 0.06 | 0.05 | 0.05 | 0.04 | 0.03 | 0.05 |  | 0.07 | -26.5086 | 0.14893 | 0.0016048 | -0.016649 | -0.061395 | -0.024395 | -0.041395 |  | -0.0066994
5623.4133 | -26.3434 | -26.489 | -26.5025 | -26.55 | -26.512 | -26.52 |  | -26.4897 | 0.06 | 0.06 | 0.05 | 0.04 | 0.03 | 0.05 |  | 0.08 | -26.4866 | 0.14327 | -0.0023507 | -0.01584 | -0.063351 | -0.025351 | -0.033351 |  | -0.0030289
5956.6214 | -26.3534 | -26.493 | -26.5039 | -26.55 | -26.516 | -26.52 |  | -26.4827 | 0.06 | 0.06 | 0.05 | 0.04 | 0.03 | 0.05 |  | 0.09 | -26.4884 | 0.13499 | -0.004559 | -0.015468 | -0.061559 | -0.027559 | -0.031559 |  | 0.0057101
6309.5734 | -26.4009 | -26.535 | -26.5438 | -26.59 | -26.554 | -26.56 | -26.4947 | -26.51 | 0.06 | 0.06 | 0.05 | 0.04 | 0.03 | 0.05 | 0.06 | 0.12 | -26.5236 | 0.12265 | -0.011441 | -0.020233 | -0.066441 | -0.030441 | -0.036441 | 0.028836 | 0.013516
6683.4392 | -26.4997 | -26.635 | -26.6339 | -26.68 | -26.638 | -26.65 |  | -26.6356 | 0.06 | 0.06 | 0.06 | 0.04 | 0.04 | 0.05 |  | 0.12 | -26.6246 | 0.12492 | -0.010407 | -0.0093129 | -0.055407 | -0.013407 | -0.025407 |  | -0.010979
7079.4578 | -26.6667 | -26.799 | -26.7878 | -26.83 | -26.79 | -26.8 |  | -26.7883 | 0.06 | 0.07 | 0.06 | 0.04 | 0.04 | 0.05 |  | 0.13 | -26.7803 | 0.11358 | -0.018746 | -0.0075308 | -0.049746 | -0.0097464 | -0.019746 |  | -0.0080654
7498.9421 | -26.9216 | -27.047 | -27.0255 | -27.07 | -27.03 | -27.03 |  | -27.0236 | 0.06 | 0.07 | 0.06 | 0.04 | 0.04 | 0.05 |  | 0.13 | -27.0211 | 0.099481 | -0.025899 | -0.0044205 | -0.048899 | -0.0088991 | -0.0088991 |  | -0.002464
7943.2823 | -27.2859 | -27.381 | -27.3672 | -27.4 | -27.375 | -27.37 | -27.3691 | -27.3617 | 0.06 | 0.07 | 0.06 | 0.04 | 0.04 | 0.05 | 0.07 | 0.13 | -27.3637 | 0.077843 | -0.017265 | -0.0034423 | -0.036265 | -0.011265 | -0.006265 | -0.0053821 | 0.0020411
8413.9514 | -27.7783 | -27.861 | -27.834 | -27.86 | -27.834 | -27.83 |  | -27.8304 | 0.08 | 0.07 | 0.09 | 0.09 | 0.07 | 0.08 |  | 0.13 | -27.8325 | 0.05428 | -0.028469 | -0.0014953 | -0.027469 | -0.0014694 | 0.0025306 |  | 0.0020928
8912.5094 | -28.4115 | -28.465 | -28.4461 | -28.47 | -28.445 | -28.43 |  | -28.4545 | 0.08 | 0.08 | 0.09 | 0.09 | 0.07 | 0.08 |  | 0.13 | -28.446 | 0.034538 | -0.01899 | -6.7717e-05 | -0.02399 | 0.0010105 | 0.01601 |  | -0.0085118
9440.6088 | -29.1806 | -29.189 | -29.2063 | -29.25 | -29.209 | -29.15 |  | -29.1632 | 0.08 | 0.1 | 0.09 | 0.09 | 0.07 | 0.08 |  | 0.13 | -29.1926 | 0.012021 | 0.0035749 | -0.013707 | -0.057425 | -0.016425 | 0.042575 |  | 0.029387
10000 | -30.0453 | -30.029 | -30.0699 | -30.12 | -30.083 | -29.97 | -29.9332 | -29.9627 | 0.08 | 0.15 | 0.09 | 0.09 | 0.07 | 0.08 | 0.1 | 0.13 | -30.0266 | -0.01862 | -0.002362 | -0.043289 | -0.093362 | -0.056362 | 0.056638 | 0.093391 | 0.063966

## 2652762 Phase
4160. 2652762 | Sensitivity Phase (degrees) |  |  |  | Expanded Uncertainties (degrees) |  |  |  | Unweighted mean of sensitivity phase (degrees) | Phase difference relative to unweighted mean of sensitivity phase (degrees)
Frequency (Hz) | NIM | NMIJ | KRISS | CMS/ITRI | NIM | NMIJ | KRISS | CMS/ITRI |  | NIM | NMIJ | KRISS | CMS/ITRI
1.9953 | 175.996 | 176.86 | 176.75 | 175.5887 | 1 | 2.5 | 1.5 | 1.3 | 176.2987 | -0.30268 | 0.56132 | 0.45132 | -0.70997
2.5119 | 176.0564 | 177.36 | 176.503 | 175.6257 | 0.9 | 2.43 | 1.5 | 1.1 | 176.3863 | -0.32985 | 0.97371 | 0.11671 | -0.76056
3.1623 | 176.1972 | 176.8 | 176.513 | 175.775 | 0.8 | 2.22 | 1 | 0.9 | 176.3213 | -0.12408 | 0.47868 | 0.19168 | -0.54628
3.9811 | 176.3759 | 177.27 | 176.585 | 175.9979 | 0.7 | 1.44 | 0.8 | 0.8 | 176.5572 | -0.18127 | 0.71279 | 0.027786 | -0.5593
5.0119 | 176.5338 | 177.19 | 176.746 | 176.2671 | 0.6 | 1.18 | 0.8 | 0.7 | 176.6842 | -0.15041 | 0.50576 | 0.061756 | -0.41711
6.3096 | 176.8007 | 177.56 | 176.942 | 176.5731 | 0.6 | 1.01 | 0.7 | 0.6 | 176.969 | -0.16823 | 0.59103 | -0.026971 | -0.39583
7.9433 | 177.154 | 177.52 | 177.202 | 176.8928 | 0.5 | 0.76 | 0.6 | 0.6 | 177.1922 | -0.038163 | 0.32779 | 0.0097924 | -0.29942
10 | 177.4091 | 177.76 | 177.461 | 177.2096 | 0.5 | 0.6 | 0.5 | 0.5 | 177.4599 | -0.050839 | 0.30007 | 0.0010672 | -0.2503
12.5893 | 177.6829 | 177.95 | 177.696 | 177.5 | 0.4 | 0.46 | 0.4 | 0.5 | 177.7072 | -0.024321 | 0.24278 | -0.011216 | -0.20725
15.8489 | 177.9284 | 178.04 | 177.915 | 177.7593 | 0.35 | 0.37 | 0.4 | 0.5 | 177.9107 | 0.017724 | 0.12934 | 0.004337 | -0.1514
19.9526 | 178.1473 | 178.16 | 178.11 | 177.9837 | 0.3 | 0.29 | 0.4 | 0.5 | 178.1003 | 0.047068 | 0.059745 | 0.0097448 | -0.11656
21.1349 | 178.1967 | 178.22 | 178.154 | 178.0366 | 0.3 | 0.29 | 0.4 | 0.5 | 178.1518 | 0.044835 | 0.068176 | 0.0021763 | -0.11519
22.3872 | 178.2433 | 178.23 | 178.2 | 178.0837 | 0.3 | 0.29 | 0.4 | 0.5 | 178.1892 | 0.054071 | 0.040759 | 0.010759 | -0.10559
23.7137 | 178.2872 | 178.3 | 178.238 | 178.129 | 0.3 | 0.29 | 0.4 | 0.5 | 178.2386 | 0.048675 | 0.06143 | -0.0005697 | -0.10954
25.1189 | 178.3284 | 178.32 | 178.279 | 178.1732 | 0.25 | 0.25 | 0.4 | 0.5 | 178.2751 | 0.053254 | 0.04486 | 0.0038601 | -0.10197
26.6073 | 178.3667 | 178.35 | 178.307 | 178.2129 | 0.25 | 0.25 | 0.4 | 0.5 | 178.3092 | 0.057585 | 0.040836 | -0.0021638 | -0.096258
28.1838 | 178.4024 | 178.38 | 178.341 | 178.251 | 0.25 | 0.25 | 0.4 | 0.5 | 178.3436 | 0.058759 | 0.036407 | -0.0025926 | -0.092573
29.8538 | 178.4353 | 178.42 | 178.372 | 178.2873 | 0.25 | 0.25 | 0.4 | 0.5 | 178.3787 | 0.05663 | 0.041342 | -0.0066581 | -0.091314
31.6228 | 178.4657 | 178.45 | 178.399 | 178.3213 | 0.2 | 0.19 | 0.4 | 0.5 | 178.409 | 0.056702 | 0.041011 | -0.0099889 | -0.087724
33.4965 | 178.4937 | 178.47 | 178.428 | 178.3533 | 0.2 | 0.19 | 0.4 | 0.5 | 178.4363 | 0.057456 | 0.033735 | -0.008265 | -0.082926
35.4813 | 178.5196 | 178.51 | 178.456 | 178.383 | 0.2 | 0.19 | 0.4 | 0.5 | 178.4671 | 0.052413 | 0.042855 | -0.011145 | -0.084122
37.5837 | 178.5434 | 178.52 | 178.479 | 178.4103 | 0.2 | 0.19 | 0.4 | 0.5 | 178.4882 | 0.055199 | 0.031823 | -0.0091765 | -0.077846
39.8107 | 178.5653 | 178.54 | 178.498 | 178.4341 | 0.2 | 0.16 | 0.4 | 0.5 | 178.5094 | 0.055965 | 0.030646 | -0.011354 | -0.075257
42.1697 | 178.5855 | 178.55 | 178.517 | 178.4557 | 0.2 | 0.16 | 0.4 | 0.5 | 178.5271 | 0.058424 | 0.022945 | -0.010055 | -0.071314
44.6684 | 178.6039 | 178.57 | 178.535 | 178.4769 | 0.2 | 0.16 | 0.4 | 0.5 | 178.5464 | 0.057434 | 0.023569 | -0.011431 | -0.069572
47.3151 | 178.6204 | 178.58 | 178.552 | 178.4958 | 0.2 | 0.16 | 0.4 | 0.5 | 178.5621 | 0.058338 | 0.01795 | -0.01005 | -0.066237
50.1187 | 178.6349 | 178.6 | 178.556 | 178.5105 | 0.2 | 0.14 | 0.4 | 0.5 | 178.5753 | 0.059525 | 0.02467 | -0.01933 | -0.064865
53.0884 | 178.647 | 178.62 | 178.578 | 178.5229 | 0.15 | 0.14 | 0.4 | 0.5 | 178.592 | 0.055015 | 0.028034 | -0.013966 | -0.069083
56.2341 | 178.6564 | 178.62 | 178.592 | 178.5349 | 0.15 | 0.14 | 0.4 | 0.5 | 178.6008 | 0.055595 | 0.019153 | -0.0088468 | -0.065901
59.5662 | 178.663 | 178.63 | 178.587 | 178.5425 | 0.15 | 0.14 | 0.4 | 0.5 | 178.6056 | 0.05734 | 0.02438 | -0.01862 | -0.0631
63.0957 | 178.6664 | 178.63 | 178.59 | 178.5495 | 0.15 | 0.12 | 0.4 | 0.5 | 178.609 | 0.057449 | 0.021012 | -0.018988 | -0.059473
66.8344 | 178.6671 | 178.63 | 178.593 | 178.5538 | 0.15 | 0.12 | 0.4 | 0.5 | 178.611 | 0.05613 | 0.019017 | -0.017983 | -0.057165
70.7946 | 178.6657 | 178.62 | 178.593 | 178.5533 | 0.15 | 0.12 | 0.4 | 0.5 | 178.608 | 0.057728 | 0.011999 | -0.015001 | -0.054726
74.9894 | 178.6636 | 178.63 | 178.591 | 178.5528 | 0.15 | 0.12 | 0.4 | 0.5 | 178.6094 | 0.054222 | 0.020645 | -0.018355 | -0.056511
79.4328 | 178.6623 | 178.62 | 178.583 | 178.5504 | 0.15 | 0.1 | 0.4 | 0.5 | 178.6039 | 0.058345 | 0.016084 | -0.020916 | -0.053513
84.1395 | 178.6628 | 178.61 | 178.578 | 178.5436 | 0.15 | 0.1 | 0.4 | 0.5 | 178.5986 | 0.064203 | 0.011392 | -0.020608 | -0.054986
89.1251 | 178.6635 | 178.62 | 178.564 | 178.5334 | 0.15 | 0.1 | 0.4 | 0.5 | 178.5952 | 0.068287 | 0.024771 | -0.031229 | -0.06183
94.4061 | 178.6554 | 178.58 | 178.552 | 178.5238 | 0.15 | 0.1 | 0.4 | 0.5 | 178.5778 | 0.077616 | 0.0022107 | -0.025789 | -0.054037
100 | 178.6136 | 178.56 | 178.536 | 178.5083 | 0.15 | 0.09 | 0.4 | 0.5 | 178.5545 | 0.059117 | 0.0055295 | -0.018471 | -0.046176
105.9254 | 178.5943 | 178.56 | 178.518 | 178.4908 | 0.15 | 0.09 | 0.4 | 0.5 | 178.5408 | 0.053496 | 0.01922 | -0.02278 | -0.049935
112.2018 | 178.5719 | 178.53 | 178.496 | 178.4699 | 0.15 | 0.09 | 0.4 | 0.5 | 178.517 | 0.054965 | 0.013039 | -0.020961 | -0.047043
118.8502 | 178.5464 | 178.5 | 178.471 | 178.4463 | 0.15 | 0.09 | 0.4 | 0.5 | 178.4909 | 0.055496 | 0.0090533 | -0.019947 | -0.044603
125.8925 | 178.5176 | 178.47 | 178.444 | 178.4186 | 0.15 | 0.07 | 0.4 | 0.5 | 178.4626 | 0.055059 | 0.0074366 | -0.018563 | -0.043932
133.3521 | 178.4853 | 178.44 | 178.413 | 178.389 | 0.15 | 0.07 | 0.4 | 0.5 | 178.4318 | 0.053444 | 0.0081776 | -0.018822 | -0.0428
141.2538 | 178.4492 | 178.42 | 178.377 | 178.3546 | 0.15 | 0.07 | 0.4 | 0.5 | 178.4002 | 0.048993 | 0.01981 | -0.02319 | -0.045612
149.6236 | 178.4092 | 178.35 | 178.339 | 178.316 | 0.15 | 0.07 | 0.4 | 0.5 | 178.3535 | 0.055649 | -0.0035346 | -0.014535 | -0.03758
158.4893 | 178.3651 | 178.33 | 178.297 | 178.2755 | 0.15 | 0.07 | 0.3 | 0.5 | 178.3169 | 0.048193 | 0.013102 | -0.019898 | -0.041397
167.8804 | 178.3167 | 178.28 | 178.25 | 178.2297 | 0.15 | 0.07 | 0.3 | 0.5 | 178.2691 | 0.047619 | 0.010886 | -0.019114 | -0.03939
177.8279 | 178.2639 | 178.23 | 178.2 | 178.1794 | 0.15 | 0.07 | 0.3 | 0.5 | 178.2183 | 0.045598 | 0.011653 | -0.018347 | -0.038904
188.3649 | 178.2066 | 178.17 | 178.144 | 178.125 | 0.15 | 0.07 | 0.3 | 0.5 | 178.1614 | 0.045169 | 0.0086042 | -0.017396 | -0.036378
199.5262 | 178.1444 | 178.11 | 178.085 | 178.0656 | 0.15 | 0.06 | 0.3 | 0.5 | 178.1013 | 0.043177 | 0.0087478 | -0.016252 | -0.035673
211.3489 | 178.0774 | 178.04 | 178.02 | 178.0024 | 0.15 | 0.06 | 0.3 | 0.5 | 178.0349 | 0.042439 | 0.0050695 | -0.01493 | -0.032578
223.8721 | 178.0052 | 177.97 | 177.951 | 177.9338 | 0.15 | 0.06 | 0.3 | 0.5 | 177.965 | 0.040207 | 0.0050084 | -0.013992 | -0.031224
237.1374 | 177.9277 | 177.89 | 177.875 | 177.8581 | 0.15 | 0.06 | 0.3 | 0.5 | 177.8877 | 0.039995 | 0.0022867 | -0.012713 | -0.029568
251.1886 | 177.8447 | 177.82 | 177.795 | 177.7769 | 0.15 | 0.05 | 0.3 | 0.5 | 177.8091 | 0.035519 | 0.010868 | -0.014132 | -0.032255
266.0725 | 177.7557 | 177.72 | 177.709 | 177.6909 | 0.15 | 0.05 | 0.3 | 0.5 | 177.7189 | 0.036813 | 0.0010817 | -0.0099183 | -0.027976
281.8383 | 177.6606 | 177.63 | 177.616 | 177.5982 | 0.15 | 0.05 | 0.3 | 0.5 | 177.6262 | 0.034408 | 0.0038126 | -0.010187 | -0.028034
298.5383 | 177.5588 | 177.53 | 177.514 | 177.4984 | 0.15 | 0.05 | 0.3 | 0.5 | 177.5253 | 0.033517 | 0.004692 | -0.011308 | -0.026901
316.2278 | 177.4499 | 177.42 | 177.41 | 177.392 | 0.15 | 0.05 | 0.3 | 0.5 | 177.418 | 0.031942 | 0.0020098 | -0.0079902 | -0.025961
334.9654 | 177.3334 | 177.31 | 177.296 | 177.2777 | 0.15 | 0.05 | 0.3 | 0.5 | 177.3043 | 0.029098 | 0.0057237 | -0.0082763 | -0.026545
354.8134 | 177.2086 | 177.18 | 177.175 | 177.1556 | 0.15 | 0.05 | 0.3 | 0.5 | 177.1798 | 0.028793 | 0.00021753 | -0.0047825 | -0.024228
375.8374 | 177.075 | 177.05 | 177.044 | 177.0256 | 0.15 | 0.05 | 0.3 | 0.5 | 177.0487 | 0.026323 | 0.0013484 | -0.0046516 | -0.02302
398.1072 | 176.9321 | 176.91 | 176.906 | 176.8869 | 0.15 | 0.05 | 0.3 | 0.5 | 176.9088 | 0.023332 | 0.0012473 | -0.0027527 | -0.021826
421.6965 | 176.7796 | 176.76 | 176.757 | 176.7398 | 0.15 | 0.05 | 0.3 | 0.5 | 176.7591 | 0.020479 | 0.00090353 | -0.0020965 | -0.019286
446.6836 | 176.6174 | 176.6 | 176.6 | 176.5808 | 0.15 | 0.05 | 0.3 | 0.5 | 176.5995 | 0.01782 | 0.00046454 | 0.00046454 | -0.018749
473.1513 | 176.4456 | 176.43 | 176.432 | 176.413 | 0.15 | 0.05 | 0.3 | 0.5 | 176.4301 | 0.015469 | -0.0001414 | 0.0018586 | -0.017187
501.1872 | 176.2647 | 176.25 | 176.254 | 176.2332 | 0.15 | 0.06 | 0.3 | 0.5 | 176.2505 | 0.01424 | -0.00049009 | 0.0035099 | -0.017259
530.8844 | 176.0691 | 176.06 | 176.063 | 176.0429 | 0.15 | 0.06 | 0.3 | 0.5 | 176.0587 | 0.010308 | 0.0012534 | 0.0042534 | -0.015814
562.3413 | 175.861 | 175.85 | 175.86 | 175.8403 | 0.15 | 0.06 | 0.3 | 0.5 | 175.8528 | 0.008163 | -0.0028093 | 0.0071907 | -0.012544
595.6621 | 175.6396 | 175.64 | 175.643 | 175.6229 | 0.15 | 0.06 | 0.3 | 0.5 | 175.6364 | 0.0032607 | 0.0036252 | 0.0066252 | -0.013511
630.9573 | 175.4042 | 175.41 | 175.413 | 175.3929 | 0.15 | 0.06 | 0.3 | 0.5 | 175.405 | -0.00081584 | 0.0049612 | 0.0079612 | -0.012107
668.3439 | 175.1539 | 175.17 | 175.171 | 175.1486 | 0.15 | 0.07 | 0.3 | 0.5 | 175.1609 | -0.007005 | 0.0091322 | 0.010132 | -0.012259
707.9458 | 174.8876 | 174.9 | 174.911 | 174.8887 | 0.15 | 0.07 | 0.3 | 0.5 | 174.8968 | -0.0092015 | 0.0031715 | 0.014171 | -0.0081414
749.8942 | 174.6045 | 174.63 | 174.635 | 174.6107 | 0.15 | 0.07 | 0.3 | 0.5 | 174.6201 | -0.015528 | 0.0099468 | 0.014947 | -0.009366
794.3282 | 174.3035 | 174.32 | 174.342 | 174.3184 | 0.15 | 0.07 | 0.3 | 0.5 | 174.321 | -0.01748 | -0.00098 | 0.02102 | -0.0025596
841.3951 | 173.9834 | 174.01 | 174.029 | 174.0038 | 0.15 | 0.08 | 0.3 | 0.5 | 174.0066 | -0.02315 | 0.0034371 | 0.022437 | -0.0027238
891.2509 | 173.643 | 173.68 | 173.698 | 173.6708 | 0.15 | 0.08 | 0.3 | 0.5 | 173.673 | -0.029911 | 0.007047 | 0.025047 | -0.0021827
944.0609 | 173.2811 | 173.32 | 173.343 | 173.3167 | 0.15 | 0.08 | 0.3 | 0.5 | 173.3152 | -0.034126 | 0.004808 | 0.027808 | 0.0015103
1000 | 172.8961 | 172.95 | 172.967 | 172.9395 | 0.15 | 0.08 | 0.3 | 0.5 | 172.9381 | -0.042078 | 0.011869 | 0.028869 | 0.0013402
1059.2537 | 172.4865 | 172.54 | 172.568 | 172.5391 | 0.15 | 0.09 | 0.3 | 0.5 | 172.5334 | -0.046939 | 0.0066103 | 0.03461 | 0.0057185
1122.0185 | 172.0506 | 172.11 | 172.142 | 172.1106 | 0.15 | 0.09 | 0.3 | 0.5 | 172.1033 | -0.052734 | 0.0067026 | 0.038703 | 0.0073289
1188.5022 | 171.5865 | 171.65 | 171.688 | 171.6577 | 0.15 | 0.09 | 0.3 | 0.5 | 171.6456 | -0.059011 | 0.0044451 | 0.042445 | 0.01212
1258.9254 | 171.0924 | 171.17 | 171.207 | 171.1694 | 0.15 | 0.09 | 0.3 | 0.5 | 171.1597 | -0.067327 | 0.010295 | 0.047295 | 0.0097373
1333.5214 | 170.5659 | 170.65 | 170.69 | 170.6545 | 0.15 | 0.11 | 0.3 | 0.5 | 170.6401 | -0.074215 | 0.0099158 | 0.049916 | 0.014384
1412.5375 | 170.0046 | 170.1 | 170.139 | 170.1037 | 0.15 | 0.11 | 0.3 | 0.5 | 170.0868 | -0.082209 | 0.013169 | 0.052169 | 0.016871
1496.2357 | 169.406 | 169.51 | 169.556 | 169.5169 | 0.15 | 0.11 | 0.3 | 0.5 | 169.4972 | -0.091202 | 0.012757 | 0.058757 | 0.019688
1584.8932 | 168.7673 | 168.88 | 168.932 | 168.8929 | 0.15 | 0.11 | 0.3 | 0.5 | 168.8681 | -0.10075 | 0.011938 | 0.063938 | 0.024876
1678.804 | 168.0854 | 168.22 | 168.265 | 168.2251 | 0.15 | 0.13 | 0.3 | 0.5 | 168.1989 | -0.11347 | 0.02114 | 0.06614 | 0.026195
1778.2794 | 167.357 | 167.5 | 167.554 | 167.5087 | 0.15 | 0.13 | 0.3 | 0.5 | 167.4799 | -0.12296 | 0.020076 | 0.074076 | 0.028804
1883.6491 | 166.5784 | 166.74 | 166.791 | 166.7458 | 0.15 | 0.13 | 0.3 | 0.5 | 166.7138 | -0.13537 | 0.026189 | 0.077189 | 0.03199
1995.2623 | 165.7458 | 165.89 | 165.978 | 165.9297 | 0.2 | 0.13 | 0.3 | 0.5 | 165.8859 | -0.14011 | 0.0041258 | 0.092126 | 0.043859
2113.489 | 164.8542 | 165.03 | 165.09 | 165.053 | 0.2 | 0.13 | 0.3 | 0.5 | 165.0068 | -0.15261 | 0.023194 | 0.083194 | 0.046227
2238.7211 | 163.8978 | 164.08 | 164.149 | 164.1118 | 0.2 | 0.13 | 0.3 | 0.5 | 164.0597 | -0.16184 | 0.020336 | 0.089336 | 0.052165
2371.3737 | 162.8687 | 163.07 | 163.141 | 163.0993 | 0.2 | 0.13 | 0.3 | 0.5 | 163.0447 | -0.17603 | 0.025251 | 0.096251 | 0.054527
2511.8864 | 161.7553 | 161.98 | 162.052 | 162.0136 | 0.2 | 0.13 | 0.3 | 0.5 | 161.9502 | -0.19493 | 0.029771 | 0.10177 | 0.063388
2660.7251 | 160.5598 | 160.81 | 160.88 | 160.8408 | 0.2 | 0.14 | 0.3 | 0.5 | 160.7727 | -0.21284 | 0.037331 | 0.10733 | 0.068175
2818.3829 | 159.2693 | 159.55 | 159.611 | 159.5725 | 0.2 | 0.14 | 0.3 | 0.5 | 159.5007 | -0.23141 | 0.049321 | 0.11032 | 0.071772
2985.3826 | 157.8733 | 158.18 | 158.245 | 158.2005 | 0.2 | 0.14 | 0.3 | 0.5 | 158.1247 | -0.25144 | 0.055298 | 0.1203 | 0.075844
3162.2777 | 156.3599 | 156.69 | 156.758 | 156.7172 | 0.2 | 0.14 | 0.3 | 0.5 | 156.6313 | -0.27138 | 0.058722 | 0.12672 | 0.085931
3349.6544 | 154.7154 | 155.06 | 155.134 | 155.0986 | 0.2 | 0.15 | 0.3 | 0.5 | 155.002 | -0.28659 | 0.057997 | 0.132 | 0.09659
3548.1339 | 152.9239 | 153.32 | 153.372 | 153.3411 | 0.2 | 0.15 | 0.3 | 0.5 | 153.2393 | -0.31532 | 0.080739 | 0.13274 | 0.10185
3758.374 | 150.9672 | 151.4 | 151.446 | 151.4143 | 0.2 | 0.15 | 0.3 | 0.5 | 151.3069 | -0.33967 | 0.093113 | 0.13911 | 0.10745
3981.0717 | 148.8244 | 149.32 | 149.332 | 149.3115 | 0.25 | 0.15 | 0.3 | 0.5 | 149.197 | -0.3726 | 0.12303 | 0.13503 | 0.11454
4216.965 | 146.4718 | 147.01 | 147.016 | 147.0074 | 0.25 | 0.15 | 0.3 | 0.5 | 146.8763 | -0.40454 | 0.13371 | 0.13971 | 0.13113
4466.8359 | 143.8828 | 144.52 | 144.462 | 144.4691 | 0.25 | 0.15 | 0.3 | 0.5 | 144.3335 | -0.45068 | 0.18654 | 0.12854 | 0.1356
4731.5126 | 141.0281 | 141.77 | 141.643 | 141.6623 | 0.25 | 0.15 | 0.3 | 0.5 | 141.5258 | -0.49779 | 0.24415 | 0.11715 | 0.13648
5011.8723 | 137.8758 | 138.74 | 138.535 | 138.5673 | 0.25 | 0.15 | 0.3 | 0.6 | 138.4295 | -0.5537 | 0.31046 | 0.10546 | 0.13778
5308.8444 | 134.3927 | 135.38 | 135.047 | 135.133 | 0.25 | 0.27 | 0.3 | 0.6 | 134.9882 | -0.59549 | 0.39183 | 0.058834 | 0.14483
5623.4133 | 130.5449 | 131.7 | 131.27 | 131.3467 | 0.25 | 0.27 | 0.3 | 0.6 | 131.2154 | -0.67049 | 0.4846 | 0.054596 | 0.13129
5956.6214 | 126.3009 | 127.57 | 127.038 | 127.1609 | 0.25 | 0.27 | 0.3 | 0.7 | 127.0175 | -0.71653 | 0.55253 | 0.020528 | 0.14348
6309.5734 | 121.6345 | 123.09 | 122.389 | 122.5465 | 0.25 | 0.27 | 0.3 | 0.8 | 122.415 | -0.78047 | 0.675 | -0.025995 | 0.13146
6683.4392 | 116.5298 | 118.13 | 117.277 | 117.2308 | 0.3 | 0.62 | 0.4 | 1 | 117.2919 | -0.76212 | 0.83809 | -0.014907 | -0.061069
7079.4578 | 110.9881 | 112.73 | 111.724 | 111.7016 | 0.35 | 0.62 | 0.4 | 1.1 | 111.7859 | -0.79781 | 0.94406 | -0.061936 | -0.084323
7498.9421 | 105.0376 | 106.89 | 105.755 | 105.7643 | 0.4 | 0.62 | 0.4 | 1.1 | 105.8617 | -0.82414 | 1.0283 | -0.10672 | -0.09741
7943.2823 | 98.7453 | 100.67 | 99.448 | 99.4571 | 0.45 | 0.62 | 0.4 | 1.1 | 99.5801 | -0.83484 | 1.0899 | -0.1321 | -0.12297
8413.9514 | 92.2332 | 94.14 | 92.887 | 92.9074 | 0.55 | 1.07 | 0.7 | 1.1 | 93.0419 | -0.80869 | 1.0981 | -0.15489 | -0.13452
8912.5094 | 85.6979 | 87.5 | 86.281 | 86.4122 | 0.65 | 1.07 | 0.7 | 1.1 | 86.4728 | -0.77487 | 1.0272 | -0.19179 | -0.060559
9440.6088 | 79.4333 | 78.76 | 79.906 | 80.4888 | 0.7 | 1.07 | 0.7 | 1.1 | 79.647 | -0.2137 | -0.88703 | 0.25897 | 0.84177
10000 | 73.8559 | 72.91 | 74.083 | 73.7777 | 0.72 | 1.07 | 0.7 | 1.1 | 73.6567 | 0.19927 | -0.74665 | 0.42635 | 0.12104

## 2652762 Phase-Revised
Results highlighted in blue were corrected during the circulation of APMP.AUV.A-K5 Draft A.
4160. 2652762 | Sensitivity Phase (degrees) |  |  |  | Expanded Uncertainties (degrees) |  |  |  | Unweighted mean of sensitivity phase (degrees) | Phase difference relative to unweighted mean of sensitivity phase (degrees)
Frequency (Hz) | NIM | NMIJ | KRISS | CMS/ITRI | NIM | NMIJ | KRISS | CMS/ITRI |  | NIM | NMIJ | KRISS | CMS/ITRI
1.9953 | 175.996 | 176.3 | 176.75 | 175.5887 | 1 | 2.5 | 1.5 | 1.3 | 176.1587 | -0.16268 | 0.14132 | 0.59132 | -0.56997
2.5119 | 176.0564 | 176.84 | 176.503 | 175.6257 | 0.9 | 2.43 | 1.5 | 1.1 | 176.2563 | -0.19985 | 0.58371 | 0.24671 | -0.63056
3.1623 | 176.1972 | 176.32 | 176.513 | 175.775 | 0.8 | 2.22 | 1 | 0.9 | 176.2013 | -0.0040766 | 0.11868 | 0.31168 | -0.42628
3.9811 | 176.3759 | 176.83 | 176.585 | 175.9979 | 0.7 | 1.44 | 0.8 | 0.8 | 176.4472 | -0.071274 | 0.38279 | 0.13779 | -0.4493
5.0119 | 176.5338 | 176.79 | 176.746 | 176.2671 | 0.6 | 1.18 | 0.8 | 0.7 | 176.5842 | -0.050407 | 0.20576 | 0.16176 | -0.31711
6.3096 | 176.8007 | 177.2 | 176.942 | 176.5731 | 0.6 | 1.01 | 0.7 | 0.6 | 176.879 | -0.07823 | 0.32103 | 0.063029 | -0.30583
7.9433 | 177.154 | 177.19 | 177.202 | 176.8928 | 0.5 | 0.76 | 0.6 | 0.6 | 177.1097 | 0.044337 | 0.080292 | 0.092292 | -0.21692
10 | 177.4091 | 177.46 | 177.461 | 177.2096 | 0.5 | 0.6 | 0.5 | 0.5 | 177.3849 | 0.024161 | 0.075067 | 0.076067 | -0.1753
12.5893 | 177.6829 | 177.68 | 177.696 | 177.5 | 0.4 | 0.46 | 0.4 | 0.5 | 177.6397 | 0.043179 | 0.040284 | 0.056284 | -0.13975
15.8489 | 177.9284 | 177.88 | 177.915 | 177.7593 | 0.35 | 0.37 | 0.4 | 0.5 | 177.8707 | 0.057724 | 0.009337 | 0.044337 | -0.1114
19.9526 | 178.1473 | 178.1 | 178.11 | 177.9837 | 0.3 | 0.29 | 0.4 | 0.5 | 178.0853 | 0.062068 | 0.014745 | 0.024745 | -0.10156
21.1349 | 178.1967 | 178.16 | 178.154 | 178.0366 | 0.3 | 0.29 | 0.4 | 0.5 | 178.1368 | 0.059835 | 0.023176 | 0.017176 | -0.10019
22.3872 | 178.2433 | 178.17 | 178.2 | 178.0837 | 0.3 | 0.29 | 0.4 | 0.5 | 178.1742 | 0.069071 | -0.004241 | 0.025759 | -0.090589
23.7137 | 178.2872 | 178.24 | 178.238 | 178.129 | 0.3 | 0.29 | 0.4 | 0.5 | 178.2236 | 0.063675 | 0.01643 | 0.01443 | -0.094535
25.1189 | 178.3284 | 178.26 | 178.279 | 178.1732 | 0.25 | 0.25 | 0.4 | 0.5 | 178.2601 | 0.068254 | -0.00013987 | 0.01886 | -0.086974
26.6073 | 178.3667 | 178.3 | 178.307 | 178.2129 | 0.25 | 0.25 | 0.4 | 0.5 | 178.2967 | 0.070085 | 0.0033362 | 0.010336 | -0.083758
28.1838 | 178.4024 | 178.33 | 178.341 | 178.251 | 0.25 | 0.25 | 0.4 | 0.5 | 178.3311 | 0.071259 | -0.0010926 | 0.0099074 | -0.080073
29.8538 | 178.4353 | 178.37 | 178.372 | 178.2873 | 0.25 | 0.25 | 0.4 | 0.5 | 178.3662 | 0.06913 | 0.0038419 | 0.0058419 | -0.078814
31.6228 | 178.4657 | 178.4 | 178.399 | 178.3213 | 0.2 | 0.19 | 0.4 | 0.5 | 178.3965 | 0.069202 | 0.0035111 | 0.0025111 | -0.075224
33.4965 | 178.4937 | 178.42 | 178.428 | 178.3533 | 0.2 | 0.19 | 0.4 | 0.5 | 178.4238 | 0.069956 | -0.003765 | 0.004235 | -0.070426
35.4813 | 178.5196 | 178.46 | 178.456 | 178.383 | 0.2 | 0.19 | 0.4 | 0.5 | 178.4546 | 0.064913 | 0.0053546 | 0.0013546 | -0.071622
37.5837 | 178.5434 | 178.47 | 178.479 | 178.4103 | 0.2 | 0.19 | 0.4 | 0.5 | 178.4757 | 0.067699 | -0.0056765 | 0.0033235 | -0.065346
39.8107 | 178.5653 | 178.5 | 178.498 | 178.4341 | 0.2 | 0.16 | 0.4 | 0.5 | 178.4994 | 0.065965 | 0.00064587 | -0.0013541 | -0.065257
42.1697 | 178.5855 | 178.51 | 178.517 | 178.4557 | 0.2 | 0.16 | 0.4 | 0.5 | 178.5171 | 0.068424 | -0.0070548 | -5.4798e-05 | -0.061314
44.6684 | 178.6039 | 178.53 | 178.535 | 178.4769 | 0.2 | 0.16 | 0.4 | 0.5 | 178.5364 | 0.067434 | -0.0064308 | -0.0014308 | -0.059572
47.3151 | 178.6204 | 178.54 | 178.552 | 178.4958 | 0.2 | 0.16 | 0.4 | 0.5 | 178.5521 | 0.068338 | -0.01205 | -5.0253e-05 | -0.056237
50.1187 | 178.6349 | 178.56 | 178.556 | 178.5105 | 0.2 | 0.14 | 0.4 | 0.5 | 178.5653 | 0.069525 | -0.0053297 | -0.0093297 | -0.054865
53.0884 | 178.647 | 178.58 | 178.578 | 178.5229 | 0.15 | 0.14 | 0.4 | 0.5 | 178.582 | 0.065015 | -0.0019656 | -0.0039656 | -0.059083
56.2341 | 178.6564 | 178.58 | 178.592 | 178.5349 | 0.15 | 0.14 | 0.4 | 0.5 | 178.5908 | 0.065595 | -0.010847 | 0.0011532 | -0.055901
59.5662 | 178.663 | 178.59 | 178.587 | 178.5425 | 0.15 | 0.14 | 0.4 | 0.5 | 178.5956 | 0.06734 | -0.0056201 | -0.0086201 | -0.0531
63.0957 | 178.6664 | 178.6 | 178.59 | 178.5495 | 0.15 | 0.12 | 0.4 | 0.5 | 178.6015 | 0.064949 | -0.0014878 | -0.011488 | -0.051973
66.8344 | 178.6671 | 178.6 | 178.593 | 178.5538 | 0.15 | 0.12 | 0.4 | 0.5 | 178.6035 | 0.06363 | -0.0034827 | -0.010483 | -0.049665
70.7946 | 178.6657 | 178.59 | 178.593 | 178.5533 | 0.15 | 0.12 | 0.4 | 0.5 | 178.6005 | 0.065228 | -0.010501 | -0.0075009 | -0.047226
74.9894 | 178.6636 | 178.6 | 178.591 | 178.5528 | 0.15 | 0.12 | 0.4 | 0.5 | 178.6019 | 0.061722 | -0.0018553 | -0.010855 | -0.049011
79.4328 | 178.6623 | 178.59 | 178.583 | 178.5504 | 0.15 | 0.1 | 0.4 | 0.5 | 178.5964 | 0.065845 | -0.0064159 | -0.013416 | -0.046013
84.1395 | 178.6628 | 178.58 | 178.578 | 178.5436 | 0.15 | 0.1 | 0.4 | 0.5 | 178.5911 | 0.071703 | -0.011108 | -0.013108 | -0.047486
89.1251 | 178.6635 | 178.59 | 178.564 | 178.5334 | 0.15 | 0.1 | 0.4 | 0.5 | 178.5877 | 0.075787 | 0.0022714 | -0.023729 | -0.05433
94.4061 | 178.6554 | 178.55 | 178.552 | 178.5238 | 0.15 | 0.1 | 0.4 | 0.5 | 178.5703 | 0.085116 | -0.020289 | -0.018289 | -0.046537
100 | 178.6136 | 178.54 | 178.536 | 178.5083 | 0.15 | 0.09 | 0.4 | 0.5 | 178.5495 | 0.064117 | -0.0094705 | -0.013471 | -0.041176
105.9254 | 178.5943 | 178.53 | 178.518 | 178.4908 | 0.15 | 0.09 | 0.4 | 0.5 | 178.5333 | 0.060996 | -0.0032802 | -0.01528 | -0.042435
112.2018 | 178.5719 | 178.5 | 178.496 | 178.4699 | 0.15 | 0.09 | 0.4 | 0.5 | 178.5095 | 0.062465 | -0.0094614 | -0.013461 | -0.039543
118.8502 | 178.5464 | 178.48 | 178.471 | 178.4463 | 0.15 | 0.09 | 0.4 | 0.5 | 178.4859 | 0.060496 | -0.0059467 | -0.014947 | -0.039603
125.8925 | 178.5176 | 178.45 | 178.444 | 178.4186 | 0.15 | 0.07 | 0.4 | 0.5 | 178.4576 | 0.060059 | -0.0075634 | -0.013563 | -0.038932
133.3521 | 178.4853 | 178.42 | 178.413 | 178.389 | 0.15 | 0.07 | 0.4 | 0.5 | 178.4268 | 0.058444 | -0.0068224 | -0.013822 | -0.0378
141.2538 | 178.4492 | 178.4 | 178.377 | 178.3546 | 0.15 | 0.07 | 0.4 | 0.5 | 178.3952 | 0.053993 | 0.0048099 | -0.01819 | -0.040612
149.6236 | 178.4092 | 178.33 | 178.339 | 178.316 | 0.15 | 0.07 | 0.4 | 0.5 | 178.3485 | 0.060649 | -0.018535 | -0.0095346 | -0.03258
158.4893 | 178.3651 | 178.31 | 178.297 | 178.2755 | 0.15 | 0.07 | 0.3 | 0.5 | 178.3119 | 0.053193 | -0.0018979 | -0.014898 | -0.036397
167.8804 | 178.3167 | 178.26 | 178.25 | 178.2297 | 0.15 | 0.07 | 0.3 | 0.5 | 178.2641 | 0.052619 | -0.0041144 | -0.014114 | -0.03439
177.8279 | 178.2639 | 178.21 | 178.2 | 178.1794 | 0.15 | 0.07 | 0.3 | 0.5 | 178.2133 | 0.050598 | -0.0033472 | -0.013347 | -0.033904
188.3649 | 178.2066 | 178.15 | 178.144 | 178.125 | 0.15 | 0.07 | 0.3 | 0.5 | 178.1564 | 0.050169 | -0.0063958 | -0.012396 | -0.031378
199.5262 | 178.1444 | 178.09 | 178.085 | 178.0656 | 0.15 | 0.06 | 0.3 | 0.5 | 178.0963 | 0.048177 | -0.0062522 | -0.011252 | -0.030673
211.3489 | 178.0774 | 178.03 | 178.02 | 178.0024 | 0.15 | 0.06 | 0.3 | 0.5 | 178.0324 | 0.044939 | -0.0024305 | -0.01243 | -0.030078
223.8721 | 178.0052 | 177.96 | 177.951 | 177.9338 | 0.15 | 0.06 | 0.3 | 0.5 | 177.9625 | 0.042707 | -0.0024916 | -0.011492 | -0.028724
237.1374 | 177.9277 | 177.88 | 177.875 | 177.8581 | 0.15 | 0.06 | 0.3 | 0.5 | 177.8852 | 0.042495 | -0.0052133 | -0.010213 | -0.027068
251.1886 | 177.8447 | 177.81 | 177.795 | 177.7769 | 0.15 | 0.05 | 0.3 | 0.5 | 177.8066 | 0.038019 | 0.0033682 | -0.011632 | -0.029755
266.0725 | 177.7557 | 177.71 | 177.709 | 177.6909 | 0.15 | 0.05 | 0.3 | 0.5 | 177.7164 | 0.039313 | -0.0064183 | -0.0074183 | -0.025476
281.8383 | 177.6606 | 177.62 | 177.616 | 177.5982 | 0.15 | 0.05 | 0.3 | 0.5 | 177.6237 | 0.036908 | -0.0036874 | -0.0076874 | -0.025534
298.5383 | 177.5588 | 177.53 | 177.514 | 177.4984 | 0.15 | 0.05 | 0.3 | 0.5 | 177.5253 | 0.033517 | 0.004692 | -0.011308 | -0.026901
316.2278 | 177.4499 | 177.42 | 177.41 | 177.392 | 0.15 | 0.05 | 0.3 | 0.5 | 177.418 | 0.031942 | 0.0020098 | -0.0079902 | -0.025961
334.9654 | 177.3334 | 177.3 | 177.296 | 177.2777 | 0.15 | 0.05 | 0.3 | 0.5 | 177.3018 | 0.031598 | -0.0017763 | -0.0057763 | -0.024045
354.8134 | 177.2086 | 177.18 | 177.175 | 177.1556 | 0.15 | 0.05 | 0.3 | 0.5 | 177.1798 | 0.028793 | 0.00021753 | -0.0047825 | -0.024228
375.8374 | 177.075 | 177.05 | 177.044 | 177.0256 | 0.15 | 0.05 | 0.3 | 0.5 | 177.0487 | 0.026323 | 0.0013484 | -0.0046516 | -0.02302
398.1072 | 176.9321 | 176.91 | 176.906 | 176.8869 | 0.15 | 0.05 | 0.3 | 0.5 | 176.9088 | 0.023332 | 0.0012473 | -0.0027527 | -0.021826
421.6965 | 176.7796 | 176.76 | 176.757 | 176.7398 | 0.15 | 0.05 | 0.3 | 0.5 | 176.7591 | 0.020479 | 0.00090353 | -0.0020965 | -0.019286
446.6836 | 176.6174 | 176.6 | 176.6 | 176.5808 | 0.15 | 0.05 | 0.3 | 0.5 | 176.5995 | 0.01782 | 0.00046454 | 0.00046454 | -0.018749
473.1513 | 176.4456 | 176.43 | 176.432 | 176.413 | 0.15 | 0.05 | 0.3 | 0.5 | 176.4301 | 0.015469 | -0.0001414 | 0.0018586 | -0.017187
501.1872 | 176.2647 | 176.25 | 176.254 | 176.2332 | 0.15 | 0.06 | 0.3 | 0.5 | 176.2505 | 0.01424 | -0.00049009 | 0.0035099 | -0.017259
530.8844 | 176.0691 | 176.07 | 176.063 | 176.0429 | 0.15 | 0.06 | 0.3 | 0.5 | 176.0612 | 0.0078077 | 0.0087534 | 0.0017534 | -0.018314
562.3413 | 175.861 | 175.86 | 175.86 | 175.8403 | 0.15 | 0.06 | 0.3 | 0.5 | 175.8553 | 0.005663 | 0.0046907 | 0.0046907 | -0.015044
595.6621 | 175.6396 | 175.64 | 175.643 | 175.6229 | 0.15 | 0.06 | 0.3 | 0.5 | 175.6364 | 0.0032607 | 0.0036252 | 0.0066252 | -0.013511
630.9573 | 175.4042 | 175.42 | 175.413 | 175.3929 | 0.15 | 0.06 | 0.3 | 0.5 | 175.4075 | -0.0033158 | 0.012461 | 0.0054612 | -0.014607
668.3439 | 175.1539 | 175.18 | 175.171 | 175.1486 | 0.15 | 0.07 | 0.3 | 0.5 | 175.1634 | -0.009505 | 0.016632 | 0.0076322 | -0.014759
707.9458 | 174.8876 | 174.91 | 174.911 | 174.8887 | 0.15 | 0.07 | 0.3 | 0.5 | 174.8993 | -0.011701 | 0.010671 | 0.011671 | -0.010641
749.8942 | 174.6045 | 174.64 | 174.635 | 174.6107 | 0.15 | 0.07 | 0.3 | 0.5 | 174.6226 | -0.018028 | 0.017447 | 0.012447 | -0.011866
794.3282 | 174.3035 | 174.34 | 174.342 | 174.3184 | 0.15 | 0.07 | 0.3 | 0.5 | 174.326 | -0.02248 | 0.01402 | 0.01602 | -0.0075596
841.3951 | 173.9834 | 174.03 | 174.029 | 174.0038 | 0.15 | 0.08 | 0.3 | 0.5 | 174.0116 | -0.02815 | 0.018437 | 0.017437 | -0.0077238
891.2509 | 173.643 | 173.7 | 173.698 | 173.6708 | 0.15 | 0.08 | 0.3 | 0.5 | 173.678 | -0.034911 | 0.022047 | 0.020047 | -0.0071827
944.0609 | 173.2811 | 173.34 | 173.343 | 173.3167 | 0.15 | 0.08 | 0.3 | 0.5 | 173.3202 | -0.039126 | 0.019808 | 0.022808 | -0.0034897
1000 | 172.8961 | 172.97 | 172.967 | 172.9395 | 0.15 | 0.08 | 0.3 | 0.5 | 172.9431 | -0.047078 | 0.026869 | 0.023869 | -0.0036598
1059.2537 | 172.4865 | 172.57 | 172.568 | 172.5391 | 0.15 | 0.09 | 0.3 | 0.5 | 172.5409 | -0.054439 | 0.02911 | 0.02711 | -0.0017815
1122.0185 | 172.0506 | 172.14 | 172.142 | 172.1106 | 0.15 | 0.09 | 0.3 | 0.5 | 172.1108 | -0.060234 | 0.029203 | 0.031203 | -0.00017113
1188.5022 | 171.5865 | 171.69 | 171.688 | 171.6577 | 0.15 | 0.09 | 0.3 | 0.5 | 171.6556 | -0.069011 | 0.034445 | 0.032445 | 0.0021203
1258.9254 | 171.0924 | 171.2 | 171.207 | 171.1694 | 0.15 | 0.09 | 0.3 | 0.5 | 171.1672 | -0.074827 | 0.032795 | 0.039795 | 0.0022373
1333.5214 | 170.5659 | 170.69 | 170.69 | 170.6545 | 0.15 | 0.11 | 0.3 | 0.5 | 170.6501 | -0.084215 | 0.039916 | 0.039916 | 0.0043836
1412.5375 | 170.0046 | 170.14 | 170.139 | 170.1037 | 0.15 | 0.11 | 0.3 | 0.5 | 170.0968 | -0.092209 | 0.043169 | 0.042169 | 0.0068708
1496.2357 | 169.406 | 169.55 | 169.556 | 169.5169 | 0.15 | 0.11 | 0.3 | 0.5 | 169.5072 | -0.1012 | 0.042757 | 0.048757 | 0.0096883
1584.8932 | 168.7673 | 168.92 | 168.932 | 168.8929 | 0.15 | 0.11 | 0.3 | 0.5 | 168.8781 | -0.11075 | 0.041938 | 0.053938 | 0.014876
1678.804 | 168.0854 | 168.26 | 168.265 | 168.2251 | 0.15 | 0.13 | 0.3 | 0.5 | 168.2089 | -0.12347 | 0.05114 | 0.05614 | 0.016195
1778.2794 | 167.357 | 167.55 | 167.554 | 167.5087 | 0.15 | 0.13 | 0.3 | 0.5 | 167.4924 | -0.13546 | 0.057576 | 0.061576 | 0.016304
1883.6491 | 166.5784 | 166.79 | 166.791 | 166.7458 | 0.15 | 0.13 | 0.3 | 0.5 | 166.7263 | -0.14787 | 0.063689 | 0.064689 | 0.01949
1995.2623 | 165.7458 | 165.95 | 165.978 | 165.9297 | 0.2 | 0.13 | 0.3 | 0.5 | 165.9009 | -0.15511 | 0.049126 | 0.077126 | 0.028859
2113.489 | 164.8542 | 165.09 | 165.09 | 165.053 | 0.2 | 0.13 | 0.3 | 0.5 | 165.0218 | -0.16761 | 0.068194 | 0.068194 | 0.031227
2238.7211 | 163.8978 | 164.14 | 164.149 | 164.1118 | 0.2 | 0.13 | 0.3 | 0.5 | 164.0747 | -0.17684 | 0.065336 | 0.074336 | 0.037165
2371.3737 | 162.8687 | 163.13 | 163.141 | 163.0993 | 0.2 | 0.13 | 0.3 | 0.5 | 163.0597 | -0.19103 | 0.070251 | 0.081251 | 0.039527
2511.8864 | 161.7553 | 162.05 | 162.052 | 162.0136 | 0.2 | 0.13 | 0.3 | 0.5 | 161.9677 | -0.21243 | 0.082271 | 0.084271 | 0.045888
2660.7251 | 160.5598 | 160.88 | 160.88 | 160.8408 | 0.2 | 0.14 | 0.3 | 0.5 | 160.7902 | -0.23034 | 0.089831 | 0.089831 | 0.050675
2818.3829 | 159.2693 | 159.62 | 159.611 | 159.5725 | 0.2 | 0.14 | 0.3 | 0.5 | 159.5182 | -0.24891 | 0.10182 | 0.092821 | 0.054272
2985.3826 | 157.8733 | 158.25 | 158.245 | 158.2005 | 0.2 | 0.14 | 0.3 | 0.5 | 158.1422 | -0.26894 | 0.1078 | 0.1028 | 0.058344
3162.2777 | 156.3599 | 156.76 | 156.758 | 156.7172 | 0.2 | 0.14 | 0.3 | 0.5 | 156.6488 | -0.28888 | 0.11122 | 0.10922 | 0.068431
3349.6544 | 154.7154 | 155.13 | 155.134 | 155.0986 | 0.2 | 0.15 | 0.3 | 0.5 | 155.0195 | -0.30409 | 0.1105 | 0.1145 | 0.07909
3548.1339 | 152.9239 | 153.37 | 153.372 | 153.3411 | 0.2 | 0.15 | 0.3 | 0.5 | 153.2518 | -0.32782 | 0.11824 | 0.12024 | 0.089346
3758.374 | 150.9672 | 151.45 | 151.446 | 151.4143 | 0.2 | 0.15 | 0.3 | 0.5 | 151.3194 | -0.35217 | 0.13061 | 0.12661 | 0.094947
3981.0717 | 148.8244 | 149.35 | 149.332 | 149.3115 | 0.25 | 0.15 | 0.3 | 0.5 | 149.2045 | -0.3801 | 0.14553 | 0.12753 | 0.10704
4216.965 | 146.4718 | 147.02 | 147.016 | 147.0074 | 0.25 | 0.15 | 0.3 | 0.5 | 146.8788 | -0.40704 | 0.14121 | 0.13721 | 0.12863
4466.8359 | 143.8828 | 144.5 | 144.462 | 144.4691 | 0.25 | 0.15 | 0.3 | 0.5 | 144.3285 | -0.44568 | 0.17154 | 0.13354 | 0.1406
4731.5126 | 141.0281 | 141.71 | 141.643 | 141.6623 | 0.25 | 0.15 | 0.3 | 0.5 | 141.5108 | -0.48279 | 0.19915 | 0.13215 | 0.15148
5011.8723 | 137.8758 | 138.62 | 138.535 | 138.5673 | 0.25 | 0.15 | 0.3 | 0.6 | 138.3995 | -0.5237 | 0.22046 | 0.13546 | 0.16778
5308.8444 | 134.3927 | 135.2 | 135.047 | 135.133 | 0.25 | 0.27 | 0.3 | 0.6 | 134.9432 | -0.55049 | 0.25683 | 0.10383 | 0.18983
5623.4133 | 130.5449 | 131.43 | 131.27 | 131.3467 | 0.25 | 0.27 | 0.3 | 0.6 | 131.1479 | -0.60299 | 0.2821 | 0.1221 | 0.19879
5956.6214 | 126.3009 | 127.2 | 127.038 | 127.1609 | 0.25 | 0.27 | 0.3 | 0.7 | 126.925 | -0.62403 | 0.27503 | 0.11303 | 0.23598
6309.5734 | 121.6345 | 122.62 | 122.389 | 122.5465 | 0.25 | 0.27 | 0.3 | 0.8 | 122.2975 | -0.66297 | 0.3225 | 0.091505 | 0.24896
6683.4392 | 116.5298 | 117.57 | 117.277 | 117.2308 | 0.3 | 0.62 | 0.4 | 1 | 117.1519 | -0.62212 | 0.41809 | 0.12509 | 0.078931
7079.4578 | 110.9881 | 112.08 | 111.724 | 111.7016 | 0.35 | 0.62 | 0.4 | 1.1 | 111.6234 | -0.63531 | 0.45656 | 0.10056 | 0.078177
7498.9421 | 105.0376 | 106.18 | 105.755 | 105.7643 | 0.4 | 0.62 | 0.4 | 1.1 | 105.6842 | -0.64664 | 0.49578 | 0.070776 | 0.08009
7943.2823 | 98.7453 | 99.92 | 99.448 | 99.4571 | 0.45 | 0.62 | 0.4 | 1.1 | 99.3926 | -0.64734 | 0.5274 | 0.055403 | 0.06453
8413.9514 | 92.2332 | 93.38 | 92.887 | 92.9074 | 0.55 | 1.07 | 0.7 | 1.1 | 92.8519 | -0.61869 | 0.52811 | 0.035106 | 0.055482
8912.5094 | 85.6979 | 86.75 | 86.281 | 86.4122 | 0.65 | 1.07 | 0.7 | 1.1 | 86.2853 | -0.58737 | 0.46471 | -0.0042855 | 0.12694
9440.6088 | 79.4333 | 80.35 | 79.906 | 80.4888 | 0.7 | 1.07 | 0.7 | 1.1 | 80.0445 | -0.6112 | 0.30547 | -0.13853 | 0.44427
10000 | 73.8559 | 74.87 | 74.083 | 73.7777 | 0.72 | 1.07 | 0.7 | 1.1 | 74.1467 | -0.29073 | 0.72335 | -0.063655 | -0.36896

## 2652765 Level
4160. 2652765 | Sensitivity Level (dB re 1 V/Pa) | [7 empty] | Expanded Uncertainties (dB) | [7 empty] | Estimated reference sensitivity level | Std Unc (ERSL) | Level difference relative to estimated reference sensitivity level (dB)
Frequency (Hz) | NIM | NMIA | SCL | NMIJ | KRISS | NIMT | NMIM | CMS-ITRI | NIM | NMIA | SCL | NMIJ | KRISS | NIMT | NMIM | CMS-ITRI | dB re 1 V/Pa | dB | NIM | NMIA | SCL | NMIJ | KRISS | NIMT | NMIM | CMS-ITRI
1.9953 | -26.1722 | -26.139 |  | -26.22 | -26.246 | -26.11 |  | -26.1205 | 0.2 | 0.2 |  | 0.23 | 0.1 | 0.2 |  | 0.05 | -26.1473 | 0.020517 | -0.024889 | 0.008295 |  | -0.072705 | -0.098705 | 0.037295 |  | 0.026818 | Frequency (Hz) | NMIM | 0
2.5119 | -26.2934 | -26.262 |  | -26.32 | -26.34 | -26.24 |  | -26.2496 | 0.18 | 0.15 |  | 0.22 | 0.1 | 0.2 |  | 0.05 | -26.2692 | 0.020065 | -0.024147 | 0.0072071 |  | -0.050793 | -0.070793 | 0.029207 |  | 0.019559 | 19.9526 | -0.0034181
3.1623 | -26.4038 | -26.375 |  | -26.41 | -26.443 | -26.35 |  | -26.3647 | 0.16 | 0.13 |  | 0.22 | 0.07 | 0.2 |  | 0.05 | -26.3897 | 0.018282 | -0.014138 | 0.014671 |  | -0.020329 | -0.053329 | 0.039671 |  | 0.02499 | 31.6228 | -7.0542e-05
3.9811 | -26.5025 | -26.484 |  | -26.45 | -26.54 | -26.46 |  | -26.4685 | 0.16 | 0.12 |  | 0.19 | 0.07 | 0.2 |  | 0.05 | -26.4898 | 0.018074 | -0.012746 | 0.0057633 |  | 0.039763 | -0.050237 | 0.029763 |  | 0.021261 | 63.0957 | 0.0087216
5.0119 | -26.5873 | -26.573 |  | -26.56 | -26.617 | -26.55 |  | -26.5594 | 0.1 | 0.11 |  | 0.09 | 0.06 | 0.2 |  | 0.05 | -26.579 | 0.015742 | -0.0083755 | 0.0059739 |  | 0.018974 | -0.038026 | 0.028974 |  | 0.0196 | 125.8925 | 0.0079885
6.3096 | -26.665 | -26.652 |  | -26.65 | -26.684 | -26.63 |  | -26.6396 | 0.1 | 0.1 |  | 0.08 | 0.06 | 0.13 |  | 0.05 | -26.6553 | 0.015122 | -0.0097131 | 0.0033123 |  | 0.0053123 | -0.028688 | 0.025312 |  | 0.015703 | 251.1886 | 0.0090501
7.9433 | -26.7309 | -26.723 |  | -26.72 | -26.759 | -26.7 |  | -26.7066 | 0.08 | 0.09 |  | 0.05 | 0.06 | 0.09 |  | 0.05 | -26.7234 | 0.012993 | -0.0074946 | 0.00043645 |  | 0.0034364 | -0.035564 | 0.023436 |  | 0.01682 | 501.1872 | 0.014202
10 | -26.7852 | -26.778 |  | -26.76 | -26.809 | -26.76 |  | -26.7631 | 0.08 | 0.08 |  | 0.05 | 0.06 | 0.06 |  | 0.05 | -26.773 | 0.012242 | -0.012257 | -0.005044 |  | 0.012956 | -0.036044 | 0.012956 |  | 0.0098357 | 1000 | 0.012805
12.5893 | -26.8307 | -26.824 |  | -26.8 | -26.851 | -26.8 |  | -26.8106 | 0.08 | 0.07 |  | 0.04 | 0.06 | 0.05 |  | 0.05 | -26.8136 | 0.011008 | -0.017105 | -0.01037 |  | 0.01363 | -0.03737 | 0.01363 |  | 0.0029951 | 1258.9254 | 0.010242
15.8489 | -26.8692 | -26.866 |  | -26.84 | -26.885 | -26.84 |  | -26.8485 | 0.08 | 0.06 |  | 0.04 | 0.06 | 0.05 |  | 0.05 | -26.853 | 0.010817 | -0.016195 | -0.013042 |  | 0.012958 | -0.032042 | 0.012958 |  | 0.0044306 | 1584.8932 | 0.019529
19.9526 | -26.9005 | -26.897 | -26.8823 | -26.87 | -26.911 | -26.87 | -26.8884 | -26.8809 | 0.07 | 0.06 | 0.06 | 0.04 | 0.05 | 0.05 | 0.07 | 0.05 | -26.8849 | 0.0094682 | -0.015552 | -0.012066 | 0.0026004 | 0.014934 | -0.026066 | 0.014934 | -0.0034181 | 0.0040514 | 1995.2623 | 0.016677
21.1349 | -26.9075 | -26.904 | -26.8898 | -26.88 | -26.918 | -26.88 |  | -26.8881 | 0.06 | 0.06 | 0.06 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.8934 | 0.0096976 | -0.014147 | -0.010649 | 0.0035157 | 0.013351 | -0.024649 | 0.013351 |  | 0.0052142 | 2511.8864 | 0.012629
22.3872 | -26.9143 | -26.911 | -26.8959 | -26.89 | -26.927 | -26.89 |  | -26.8949 | 0.06 | 0.06 | 0.06 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9017 | 0.0096976 | -0.012596 | -0.0093426 | 0.0057559 | 0.011657 | -0.025343 | 0.011657 |  | 0.0067084 | 3162.2777 | 0.018375
23.7137 | -26.9208 | -26.918 | -26.9036 | -26.89 | -26.932 | -26.9 |  | -26.9013 | 0.06 | 0.06 | 0.06 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9071 | 0.0096976 | -0.013681 | -0.010906 | 0.0034735 | 0.017094 | -0.024906 | 0.0070944 |  | 0.0057631 | 3981.0717 | 0.027968
25.1189 | -26.9271 | -26.924 | -26.9085 | -26.9 | -26.938 | -26.9 |  | -26.9078 | 0.06 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9134 | 0.0092803 | -0.013675 | -0.010621 | 0.0048974 | 0.013379 | -0.024621 | 0.013379 |  | 0.0055583 | 5011.8723 | 0.031845
26.6073 | -26.9331 | -26.93 | -26.9149 | -26.9 | -26.943 | -26.91 |  | -26.914 | 0.06 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9186 | 0.0092803 | -0.014501 | -0.011422 | 0.0037073 | 0.018578 | -0.024422 | 0.0085781 |  | 0.0046003 | 6309.5734 | 0.048285
28.1838 | -26.9389 | -26.936 | -26.9215 | -26.91 | -26.947 | -26.91 |  | -26.92 | 0.06 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9244 | 0.0092803 | -0.014457 | -0.011605 | 0.0029272 | 0.014395 | -0.022605 | 0.014395 |  | 0.0044372 | 7943.2823 | 0.0072229
29.8538 | -26.9444 | -26.942 | -26.9269 | -26.92 | -26.955 | -26.92 |  | -26.9258 | 0.06 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9319 | 0.0092803 | -0.012448 | -0.010074 | 0.0050671 | 0.011926 | -0.023074 | 0.011926 |  | 0.0061631 | 10000 | 0.13014
31.6228 | -26.9497 | -26.948 | -26.9326 | -26.92 | -26.96 | -26.93 | -26.9369 | -26.9311 | 0.06 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | -26.9369 | 0.0087002 | -0.012806 | -0.011148 | 0.0042836 | 0.016852 | -0.023148 | 0.006852 | -7.0542e-05 | 0.0057931
33.4965 | -26.9547 | -26.953 | -26.9388 | -26.93 | -26.963 | -26.93 |  | -26.9365 | 0.05 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9427 | 0.0090909 | -0.012023 | -0.010304 | 0.0039358 | 0.012696 | -0.020304 | 0.012696 |  | 0.0061605
35.4813 | -26.9596 | -26.958 | -26.9437 | -26.94 | -26.967 | -26.94 |  | -26.9415 | 0.05 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9492 | 0.0090909 | -0.010349 | -0.0087725 | 0.005497 | 0.0092275 | -0.017772 | 0.0092275 |  | 0.0077519
37.5837 | -26.9643 | -26.963 | -26.9483 | -26.94 | -26.973 | -26.94 |  | -26.9465 | 0.05 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9526 | 0.0090909 | -0.011689 | -0.010432 | 0.0042796 | 0.012568 | -0.020432 | 0.012568 |  | 0.0060693
39.8107 | -26.9688 | -26.968 | -26.9535 | -26.95 | -26.977 | -26.95 |  | -26.9512 | 0.05 | 0.04 | 0.05 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9597 | 0.0087706 | -0.0091023 | -0.0083169 | 0.0061579 | 0.0096831 | -0.017317 | 0.0096831 |  | 0.0084437
42.1697 | -26.9732 | -26.972 | -26.9573 | -26.95 | -26.981 | -26.95 |  | -26.956 | 0.05 | 0.04 | 0.05 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9625 | 0.0087706 | -0.010644 | -0.009457 | 0.0052161 | 0.012543 | -0.018457 | 0.012543 |  | 0.0065196
44.6684 | -26.9775 | -26.977 | -26.9618 | -26.96 | -26.985 | -26.95 |  | -26.9604 | 0.05 | 0.04 | 0.05 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9675 | 0.0087706 | -0.0099515 | -0.0094689 | 0.0057614 | 0.0075311 | -0.017469 | 0.017531 |  | 0.0071555
47.3151 | -26.9817 | -26.981 | -26.9656 | -26.96 | -26.989 | -26.96 |  | -26.9648 | 0.05 | 0.04 | 0.05 | 0.04 | 0.05 | 0.05 |  | 0.05 | -26.9716 | 0.0087706 | -0.010128 | -0.0094395 | 0.0059221 | 0.01156 | -0.01744 | 0.01156 |  | 0.0067709
50.1187 | -26.9858 | -26.986 | -26.9699 | -26.97 | -26.992 | -26.96 |  | -26.9688 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 |  | 0.05 | -26.9769 | 0.0082199 | -0.0089065 | -0.0090966 | 0.0069697 | 0.0069034 | -0.015097 | 0.016903 |  | 0.0081286
53.0884 | -26.9898 | -26.99 | -26.974 | -26.97 | -26.994 | -26.97 |  | -26.9727 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 |  | 0.05 | -26.9805 | 0.0082199 | -0.0093029 | -0.0094608 | 0.0065855 | 0.010539 | -0.013461 | 0.010539 |  | 0.0078215
56.2341 | -26.9938 | -26.994 | -26.9778 | -26.97 | -26.997 | -26.97 |  | -26.9767 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 |  | 0.05 | -26.9832 | 0.0082199 | -0.010543 | -0.010774 | 0.0054296 | 0.013226 | -0.013774 | 0.013226 |  | 0.0065251
59.5662 | -26.9976 | -26.998 | -26.9816 | -26.98 | -27.002 | -26.97 |  | -26.9804 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 |  | 0.05 | -26.9879 | 0.0082199 | -0.0096793 | -0.010116 | 0.0063068 | 0.0078843 | -0.014116 | 0.017884 |  | 0.0074831
63.0957 | -27.0012 | -27.002 | -26.9853 | -26.98 | -27.006 | -26.98 | -26.9832 | -26.9839 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -26.9919 | 0.0068934 | -0.0092631 | -0.010064 | 0.0066085 | 0.011936 | -0.014064 | 0.011936 | 0.0087216 | 0.0080092
66.8344 | -27.0047 | -27.005 | -26.989 | -26.98 | -27.008 | -26.98 |  | -26.9876 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.9953 | 0.0073433 | -0.009386 | -0.0097314 | 0.0062767 | 0.015269 | -0.012731 | 0.015269 |  | 0.0077144
70.7946 | -27.0079 | -27.009 | -26.9924 | -26.98 | -27.009 | -26.99 |  | -26.9911 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.9986 | 0.0073433 | -0.0093278 | -0.0104 | 0.0062311 | 0.0186 | -0.0104 | 0.0086002 |  | 0.0074759
74.9894 | -27.0111 | -27.012 | -26.996 | -26.99 | -27.014 | -26.99 |  | -26.9944 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0027 | 0.0073433 | -0.0083057 | -0.0092524 | 0.0067391 | 0.012748 | -0.011252 | 0.012748 |  | 0.0083254
79.4328 | -27.0141 | -27.015 | -26.9991 | -26.99 | -27.016 | -26.99 |  | -26.9977 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0047 | 0.0073433 | -0.0093603 | -0.010254 | 0.0056471 | 0.014746 | -0.011254 | 0.014746 |  | 0.0070046
84.1395 | -27.0172 | -27.019 | -27.0021 | -27 | -27.019 | -27 |  | -27.0007 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0098 | 0.0073433 | -0.0074116 | -0.00922 | 0.0076575 | 0.00978 | -0.00922 | 0.00978 |  | 0.0090705
89.1251 | -27.0204 | -27.022 | -27.0052 | -27 | -27.023 | -27 |  | -27.0039 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0123 | 0.0073433 | -0.0081129 | -0.0097346 | 0.0070665 | 0.012265 | -0.010735 | 0.012265 |  | 0.0083343
94.4061 | -27.0235 | -27.025 | -27.0082 | -27 | -27.025 | -27 |  | -27.0068 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0142 | 0.0073433 | -0.0093068 | -0.010776 | 0.006029 | 0.014224 | -0.010776 | 0.014224 |  | 0.0074389
100 | -27.026 | -27.028 | -27.0112 | -27 | -27.028 | -27 |  | -27.0097 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0163 | 0.0073433 | -0.0096193 | -0.011667 | 0.0051347 | 0.016333 | -0.011667 | 0.016333 |  | 0.0066078
105.9254 | -27.0286 | -27.03 | -27.014 | -27.01 | -27.029 | -27.01 |  | -27.0125 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0214 | 0.0069862 | -0.0072063 | -0.0085834 | 0.0074247 | 0.011417 | -0.0075834 | 0.011417 |  | 0.0088895
112.2018 | -27.0312 | -27.033 | -27.0168 | -27.01 | -27.031 | -27.01 |  | -27.0152 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0234 | 0.0069862 | -0.0078644 | -0.0096198 | 0.0065352 | 0.01338 | -0.0076198 | 0.01338 |  | 0.0081519
118.8502 | -27.0338 | -27.036 | -27.0195 | -27.01 | -27.035 | -27.01 |  | -27.0179 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0257 | 0.0069862 | -0.0080592 | -0.010251 | 0.0062131 | 0.015749 | -0.0092512 | 0.015749 |  | 0.007843
125.8925 | -27.0363 | -27.039 | -27.022 | -27.02 | -27.036 | -27.02 | -27.0209 | -27.0204 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -27.0289 | 0.0065954 | -0.0074215 | -0.010117 | 0.0068668 | 0.0088835 | -0.0071165 | 0.0088835 | 0.0079885 | 0.0084935
133.3521 | -27.0387 | -27.041 | -27.0246 | -27.02 | -27.038 | -27.02 |  | -27.023 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0315 | 0.0069862 | -0.0071899 | -0.0094616 | 0.00695 | 0.011538 | -0.0064616 | 0.011538 |  | 0.0085487
141.2538 | -27.0411 | -27.043 | -27.0269 | -27.02 | -27.039 | -27.02 |  | -27.0252 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0329 | 0.0069862 | -0.0081437 | -0.010072 | 0.0060523 | 0.012928 | -0.006072 | 0.012928 |  | 0.0077124
149.6236 | -27.0433 | -27.046 | -27.0291 | -27.02 | -27.043 | -27.02 |  | -27.0275 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0352 | 0.0069862 | -0.0081614 | -0.010831 | 0.0060918 | 0.015169 | -0.0078315 | 0.015169 |  | 0.0076737
158.4893 | -27.0455 | -27.048 | -27.0315 | -27.02 | -27.044 | -27.02 |  | -27.0296 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0365 | 0.0069862 | -0.0089596 | -0.01146 | 0.0050459 | 0.01654 | -0.0074601 | 0.01654 |  | 0.0069844
167.8804 | -27.0476 | -27.05 | -27.0336 | -27.03 | -27.049 | -27.03 |  | -27.0318 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0412 | 0.0069862 | -0.0064013 | -0.0088234 | 0.0076097 | 0.011177 | -0.0078234 | 0.011177 |  | 0.0094261
177.8279 | -27.0496 | -27.052 | -27.0355 | -27.03 | -27.049 | -27.03 |  | -27.0336 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0422 | 0.0069862 | -0.007327 | -0.0097628 | 0.0067325 | 0.012237 | -0.0067628 | 0.012237 |  | 0.008592
188.3649 | -27.0515 | -27.054 | -27.0376 | -27.03 | -27.05 | -27.03 |  | -27.0355 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0435 | 0.0069862 | -0.0079392 | -0.010481 | 0.0059437 | 0.013519 | -0.0064806 | 0.013519 |  | 0.007984
199.5262 | -27.0533 | -27.056 | -27.0394 | -27.03 | -27.052 | -27.03 |  | -27.0375 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.045 | 0.0069862 | -0.008276 | -0.011014 | 0.0055714 | 0.014986 | -0.0070136 | 0.014986 |  | 0.0074692
211.3489 | -27.055 | -27.058 | -27.0412 | -27.04 | -27.054 | -27.03 |  | -27.0393 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0476 | 0.0069862 | -0.0073339 | -0.010357 | 0.0064449 | 0.0076431 | -0.0063569 | 0.017643 |  | 0.008307
223.8721 | -27.0566 | -27.059 | -27.0429 | -27.04 | -27.055 | -27.04 |  | -27.0411 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0498 | 0.0069862 | -0.0067618 | -0.0091574 | 0.0069084 | 0.0098426 | -0.0051574 | 0.0098426 |  | 0.0087762
237.1374 | -27.0581 | -27.061 | -27.0446 | -27.04 | -27.057 | -27.04 |  | -27.0427 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0512 | 0.0069862 | -0.0069168 | -0.0097711 | 0.0066255 | 0.011229 | -0.0057711 | 0.011229 |  | 0.0085378
251.1886 | -27.0596 | -27.062 | -27.0462 | -27.04 | -27.059 | -27.04 | -27.0422 | -27.0442 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -27.0513 | 0.0065954 | -0.0083346 | -0.010734 | 0.0050692 | 0.011266 | -0.0077337 | 0.011266 | 0.0090501 | 0.0070521
266.0725 | -27.061 | -27.063 | -27.0477 | -27.04 | -27.061 | -27.04 |  | -27.0457 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0535 | 0.0069862 | -0.0074841 | -0.0095159 | 0.0058237 | 0.013484 | -0.0075159 | 0.013484 |  | 0.0077669
281.8383 | -27.0622 | -27.065 | -27.0489 | -27.04 | -27.062 | -27.04 |  | -27.047 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0545 | 0.0069862 | -0.0076999 | -0.010457 | 0.0056127 | 0.014543 | -0.0074565 | 0.014543 |  | 0.0075712
298.5383 | -27.0634 | -27.066 | -27.0501 | -27.05 | -27.063 | -27.04 |  | -27.0483 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0566 | 0.0069862 | -0.0068284 | -0.0094083 | 0.0064444 | 0.0065917 | -0.0064083 | 0.016592 |  | 0.0083113
316.2278 | -27.0645 | -27.067 | -27.0511 | -27.05 | -27.064 | -27.04 |  | -27.0492 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0573 | 0.0069862 | -0.007141 | -0.0096514 | 0.0062174 | 0.0073486 | -0.0066514 | 0.017349 |  | 0.008139
334.9654 | -27.0654 | -27.068 | -27.0523 | -27.05 | -27.065 | -27.04 |  | -27.0501 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0581 | 0.0069862 | -0.0073395 | -0.0098968 | 0.005846 | 0.0081032 | -0.0068968 | 0.018103 |  | 0.0080353
354.8134 | -27.0663 | -27.069 | -27.053 | -27.05 | -27.066 | -27.05 |  | -27.0512 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.06 | 0.0069862 | -0.006231 | -0.0089605 | 0.0069925 | 0.01004 | -0.0059605 | 0.01004 |  | 0.0088837
375.8374 | -27.067 | -27.07 | -27.0539 | -27.05 | -27.066 | -27.05 |  | -27.0519 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0605 | 0.0069862 | -0.0064689 | -0.0095023 | 0.0066437 | 0.010498 | -0.0055023 | 0.010498 |  | 0.0086007
398.1072 | -27.0675 | -27.07 | -27.0545 | -27.05 | -27.067 | -27.05 |  | -27.0526 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0609 | 0.0069862 | -0.0066166 | -0.0090889 | 0.0064449 | 0.010911 | -0.0060889 | 0.010911 |  | 0.0083316
421.6965 | -27.068 | -27.07 | -27.0549 | -27.05 | -27.068 | -27.05 |  | -27.053 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0613 | 0.0069862 | -0.0066856 | -0.0087344 | 0.0063523 | 0.011266 | -0.0067344 | 0.011266 |  | 0.0082848
446.6836 | -27.0682 | -27.071 | -27.0551 | -27.05 | -27.068 | -27.05 |  | -27.0532 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0616 | 0.0069862 | -0.0066597 | -0.0094365 | 0.0064142 | 0.011564 | -0.0064365 | 0.011564 |  | 0.0083393
473.1513 | -27.0683 | -27.071 | -27.0552 | -27.06 | -27.067 | -27.05 |  | -27.0533 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0626 | 0.0069862 | -0.0057027 | -0.0084121 | 0.0074011 | 0.0025879 | -0.0044121 | 0.012588 |  | 0.0092566
501.1872 | -27.0681 | -27.071 | -27.055 | -27.05 | -27.067 | -27.05 | -27.0454 | -27.0533 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -27.0596 | 0.0065954 | -0.0084992 | -0.011415 | 0.0045494 | 0.0095854 | -0.0074146 | 0.0095854 | 0.014202 | 0.0063299
530.8844 | -27.0676 | -27.071 | -27.0546 | -27.05 | -27.067 | -27.05 |  | -27.0527 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0612 | 0.0069862 | -0.006432 | -0.0098305 | 0.0065222 | 0.011169 | -0.0058305 | 0.011169 |  | 0.0084572
562.3413 | -27.0669 | -27.07 | -27.054 | -27.05 | -27.067 | -27.05 |  | -27.0523 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0608 | 0.0069862 | -0.0061571 | -0.0092428 | 0.006763 | 0.010757 | -0.0062428 | 0.010757 |  | 0.0084524
595.6621 | -27.066 | -27.069 | -27.0532 | -27.05 | -27.065 | -27.05 |  | -27.0514 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0598 | 0.0069862 | -0.0061758 | -0.009181 | 0.0066515 | 0.009819 | -0.005181 | 0.009819 |  | 0.0084664
630.9573 | -27.0648 | -27.068 | -27.0522 | -27.05 | -27.064 | -27.04 |  | -27.0503 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0578 | 0.0069862 | -0.0069826 | -0.010171 | 0.005581 | 0.007829 | -0.006171 | 0.017829 |  | 0.007494
668.3439 | -27.0633 | -27.066 | -27.0508 | -27.05 | -27.062 | -27.04 |  | -27.0489 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0565 | 0.0069862 | -0.0068386 | -0.0095085 | 0.0057213 | 0.0064915 | -0.0055085 | 0.016491 |  | 0.0075493
707.9458 | -27.0615 | -27.065 | -27.0491 | -27.05 | -27.06 | -27.04 |  | -27.0472 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0553 | 0.0069862 | -0.0062354 | -0.009723 | 0.0061875 | 0.005277 | -0.004723 | 0.015277 |  | 0.0080872
749.8942 | -27.0593 | -27.062 | -27.0469 | -27.04 | -27.057 | -27.04 |  | -27.0448 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.052 | 0.0069862 | -0.0072768 | -0.0099604 | 0.0051171 | 0.01204 | -0.0049604 | 0.01204 |  | 0.0071975
794.3282 | -27.0567 | -27.06 | -27.0443 | -27.04 | -27.054 | -27.04 |  | -27.0424 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0501 | 0.0069862 | -0.0065691 | -0.0098726 | 0.0057903 | 0.010127 | -0.0038726 | 0.010127 |  | 0.0077499
841.3951 | -27.0536 | -27.056 | -27.0414 | -27.04 | -27.052 | -27.03 |  | -27.0393 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0466 | 0.0069862 | -0.0069769 | -0.0093749 | 0.0052599 | 0.0066251 | -0.0053749 | 0.016625 |  | 0.0073261
891.2509 | -27.05 | -27.053 | -27.0376 | -27.03 | -27.047 | -27.03 |  | -27.0355 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0425 | 0.0069862 | -0.0074989 | -0.010521 | 0.004835 | 0.012479 | -0.0045212 | 0.012479 |  | 0.006951
944.0609 | -27.0458 | -27.049 | -27.0335 | -27.03 | -27.042 | -27.03 |  | -27.0312 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0392 | 0.0069862 | -0.0065852 | -0.0098222 | 0.0056349 | 0.0091778 | -0.0028222 | 0.0091778 |  | 0.0079278
1000 | -27.0409 | -27.044 | -27.0287 | -27.03 | -27.038 | -27.02 | -27.0201 | -27.0264 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -27.0329 | 0.0065954 | -0.0080106 | -0.011119 | 0.004178 | 0.0028806 | -0.0051194 | 0.012881 | 0.012805 | 0.0064619
1059.2537 | -27.0353 | -27.038 | -27.023 | -27.02 | -27.033 | -27.02 |  | -27.0209 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.029 | 0.0069862 | -0.0062677 | -0.0089783 | 0.0060534 | 0.0090217 | -0.0039783 | 0.0090217 |  | 0.0081328
1122.0185 | -27.0289 | -27.032 | -27.0165 | -27.01 | -27.026 | -27.01 |  | -27.0147 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0217 | 0.0069862 | -0.0071664 | -0.010289 | 0.0051745 | 0.011711 | -0.0042886 | 0.011711 |  | 0.0070065
1188.5022 | -27.0216 | -27.025 | -27.0095 | -27.01 | -27.017 | -27 |  | -27.0073 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -27.0147 | 0.0069862 | -0.0068785 | -0.01031 | 0.0052355 | 0.0046898 | -0.0023102 | 0.01469 |  | 0.0073431
1258.9254 | -27.0133 | -27.017 | -27.0014 | -27 | -27.01 | -26.99 | -26.9949 | -26.999 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -27.0051 | 0.0065954 | -0.0081689 | -0.011904 | 0.0037035 | 0.0050957 | -0.0049043 | 0.015096 | 0.010242 | 0.0061161
1333.5214 | -27.0039 | -27.007 | -26.9921 | -26.99 | -27.002 | -26.99 |  | -26.99 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.9982 | 0.0069862 | -0.005625 | -0.0087644 | 0.0061807 | 0.0082356 | -0.0037644 | 0.0082356 |  | 0.0081977
1412.5375 | -26.9932 | -26.996 | -26.9817 | -26.98 | -26.988 | -26.97 |  | -26.9797 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.9858 | 0.0069862 | -0.007457 | -0.010218 | 0.0040805 | 0.0057823 | -0.0022177 | 0.015782 |  | 0.0061235
1496.2357 | -26.9813 | -26.984 | -26.97 | -26.97 | -26.975 | -26.96 |  | -26.9678 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.9741 | 0.0069862 | -0.0071667 | -0.0098943 | 0.0040747 | 0.0041057 | -0.00089429 | 0.014106 |  | 0.0063441
1584.8932 | -26.9678 | -26.971 | -26.9569 | -26.96 | -26.962 | -26.95 | -26.9398 | -26.9545 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -26.9594 | 0.0065954 | -0.0084495 | -0.011632 | 0.0024398 | -0.00063155 | -0.0026316 | 0.0093684 | 0.019529 | 0.0048446
1678.804 | -26.9527 | -26.956 | -26.942 | -26.94 | -26.948 | -26.94 |  | -26.9394 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.9469 | 0.0069862 | -0.0057702 | -0.0090503 | 0.0049709 | 0.0069497 | -0.0010503 | 0.0069497 |  | 0.0075885
1778.2794 | -26.9358 | -26.939 | -26.9252 | -26.92 | -26.931 | -26.92 |  | -26.9226 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.9293 | 0.0069862 | -0.0065349 | -0.0097281 | 0.0040579 | 0.0092719 | -0.0017281 | 0.0092719 |  | 0.0067179
1883.6491 | -26.9169 | -26.92 | -26.9066 | -26.91 | -26.91 | -26.9 |  | -26.9036 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.9109 | 0.0069862 | -0.006022 | -0.0091316 | 0.0043056 | 0.00086837 | 0.00086837 | 0.010868 |  | 0.0072967
1995.2623 | -26.8958 | -26.899 | -26.8855 | -26.89 | -26.89 | -26.88 | -26.8716 | -26.8829 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -26.8883 | 0.0065954 | -0.0074613 | -0.010696 | 0.0027556 | -0.0016964 | -0.0016964 | 0.0083036 | 0.016677 | 0.0053946
2113.489 | -26.8722 | -26.875 | -26.8626 | -26.86 | -26.869 | -26.85 |  | -26.8589 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.8657 | 0.0069862 | -0.0064932 | -0.0092851 | 0.0030738 | 0.0057149 | -0.0032851 | 0.015715 |  | 0.0067759
2238.7211 | -26.846 | -26.848 | -26.8368 | -26.84 | -26.84 | -26.83 |  | -26.8326 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.8403 | 0.0069862 | -0.0057095 | -0.0077284 | 0.0035212 | 0.00027162 | 0.00027162 | 0.010272 |  | 0.0076586
2371.3737 | -26.8168 | -26.819 | -26.8079 | -26.81 | -26.812 | -26.8 |  | -26.8039 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.8113 | 0.0069862 | -0.0055649 | -0.0077321 | 0.003377 | 0.0012679 | -0.00073206 | 0.011268 |  | 0.0073426
2511.8864 | -26.7844 | -26.786 | -26.776 | -26.78 | -26.781 | -26.77 | -26.7655 | -26.771 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -26.7781 | 0.0065954 | -0.0062993 | -0.0078899 | 0.0020769 | -0.0018899 | -0.0028899 | 0.0081101 | 0.012629 | 0.0070903
2660.7251 | -26.7483 | -26.751 | -26.7406 | -26.74 | -26.742 | -26.73 |  | -26.7359 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.7424 | 0.0069862 | -0.0059761 | -0.008627 | 0.0017409 | 0.002373 | 0.00037297 | 0.012373 |  | 0.0065049
2818.3829 | -26.7089 | -26.712 | -26.7017 | -26.71 | -26.701 | -26.69 |  | -26.6964 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.7038 | 0.0069862 | -0.0050704 | -0.0081669 | 0.0021173 | -0.0061669 | 0.0028331 | 0.013833 |  | 0.0074514
2985.3826 | -26.6658 | -26.668 | -26.6593 | -26.67 | -26.66 | -26.65 |  | -26.6535 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.6618 | 0.0069862 | -0.003923 | -0.0061529 | 0.0025398 | -0.0081529 | 0.0018471 | 0.011847 |  | 0.0083498
3162.2777 | -26.6187 | -26.622 | -26.6126 | -26.62 | -26.614 | -26.6 | -26.5939 | -26.6062 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -26.6123 | 0.0065954 | -0.0063986 | -0.0096864 | -0.00027166 | -0.0076864 | -0.0016864 | 0.012314 | 0.018375 | 0.0060731
3349.6544 | -26.5676 | -26.573 | -26.5624 | -26.57 | -26.56 | -26.55 |  | -26.5557 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.5637 | 0.0069862 | -0.0039032 | -0.0092949 | 0.0013298 | -0.0062949 | 0.0037051 | 0.013705 |  | 0.00797
3548.1339 | -26.5125 | -26.516 | -26.5088 | -26.52 | -26.507 | -26.5 |  | -26.5009 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.5101 | 0.0069862 | -0.0024205 | -0.0059037 | 0.0013023 | -0.0099037 | 0.0030963 | 0.010096 |  | 0.0092448
3758.374 | -26.4538 | -26.458 | -26.4517 | -26.46 | -26.45 | -26.45 |  | -26.4432 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 |  | 0.05 | -26.4531 | 0.0069862 | -0.0006689 | -0.0049062 | 0.0013783 | -0.0069062 | 0.0030938 | 0.0030938 |  | 0.0098832
3981.0717 | -26.392 | -26.401 | -26.392 | -26.41 | -26.391 | -26.39 | -26.3636 | -26.3823 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | -26.3915 | 0.0065954 | -0.00052632 | -0.0094776 | -0.00044496 | -0.018478 | 0.00052243 | 0.0015224 | 0.027968 | 0.0091852
4216.965 | -26.3286 | -26.335 | -26.3309 | -26.35 | -26.328 | -26.33 |  | -26.3199 | 0.05 | 0.03 | 0.05 | 0.04 | 0.03 | 0.05 |  | 0.05 | -26.3326 | 0.0074977 | 0.0040125 | -0.0023907 | 0.0017018 | -0.017391 | 0.0046093 | 0.0026093 |  | 0.012687
4466.8359 | -26.2653 | -26.273 | -26.27 | -26.29 | -26.262 | -26.27 |  | -26.2567 | 0.05 | 0.03 | 0.05 | 0.04 | 0.03 | 0.05 |  | 0.05 | -26.2699 | 0.0074977 | 0.0046193 | -0.0030577 | -6.0969e-05 | -0.020058 | 0.0079423 | -5.765e-05 |  | 0.013271
4731.5126 | -26.2051 | -26.216 | -26.2123 | -26.23 | -26.207 | -26.21 |  | -26.1959 | 0.05 | 0.03 | 0.05 | 0.04 | 0.03 | 0.05 |  | 0.06 | -26.2125 | 0.0076029 | 0.0074545 | -0.0034933 | 0.00025598 | -0.017493 | 0.0055067 | 0.0025067 |  | 0.016594
5011.8723 | -26.1518 | -26.158 | -26.1606 | -26.19 | -26.155 | -26.16 | -26.1261 | -26.1427 | 0.06 | 0.03 | 0.05 | 0.04 | 0.03 | 0.05 | 0.05 | 0.06 | -26.1579 | 0.0073699 | 0.0061794 | -5.0058e-05 | -0.0026305 | -0.03205 | 0.0029499 | -0.0020501 | 0.031845 | 0.015216
5308.8444 | -26.1109 | -26.119 | -26.1208 | -26.15 | -26.115 | -26.12 |  | -26.0986 | 0.06 | 0.03 | 0.05 | 0.04 | 0.03 | 0.05 |  | 0.07 | -26.1213 | 0.0077811 | 0.010411 | 0.0023294 | 0.00056768 | -0.028671 | 0.0063294 | 0.0013294 |  | 0.022773
5623.4133 | -26.0897 | -26.097 | -26.1008 | -26.13 | -26.099 | -26.1 |  | -26.0756 | 0.06 | 0.04 | 0.05 | 0.04 | 0.03 | 0.05 |  | 0.08 | -26.1026 | 0.008339 | 0.012908 | 0.0056187 | 0.0017962 | -0.027381 | 0.0036187 | 0.0026187 |  | 0.027069
5956.6214 | -26.0975 | -26.095 | -26.1088 | -26.13 | -26.106 | -26.11 |  | -26.0767 | 0.06 | 0.04 | 0.05 | 0.04 | 0.03 | 0.05 |  | 0.09 | -26.1074 | 0.0083773 | 0.0098928 | 0.012367 | -0.0014667 | -0.022633 | 0.0013666 | -0.0026334 |  | 0.030701
6309.5734 | -26.146 | -26.137 | -26.1597 | -26.19 | -26.155 | -26.16 | -26.1067 | -26.1177 | 0.06 | 0.05 | 0.05 | 0.04 | 0.03 | 0.05 | 0.06 | 0.12 | -26.155 | 0.0083787 | 0.0090593 | 0.01801 | -0.0046774 | -0.03499 | 9.8478e-06 | -0.0049902 | 0.048285 | 0.037324
6683.4392 | -26.2495 | -26.261 | -26.2679 | -26.3 | -26.255 | -26.26 |  | -26.2518 | 0.06 | 0.05 | 0.06 | 0.04 | 0.04 | 0.05 |  | 0.12 | -26.2678 | 0.0096674 | 0.018321 | 0.006849 | -5.5895e-05 | -0.032151 | 0.012849 | 0.007849 |  | 0.015999
7079.4578 | -26.4253 | -26.434 | -26.4456 | -26.48 | -26.431 | -26.44 |  | -26.4267 | 0.06 | 0.06 | 0.06 | 0.04 | 0.04 | 0.05 |  | 0.13 | -26.4457 | 0.0099161 | 0.02038 | 0.011656 | 8.6972e-05 | -0.034344 | 0.014656 | 0.0056558 |  | 0.018925
7498.9421 | -26.6926 | -26.699 | -26.7135 | -26.74 | -26.699 | -26.71 |  | -26.6918 | 0.06 | 0.06 | 0.06 | 0.04 | 0.04 | 0.05 |  | 0.13 | -26.7115 | 0.0099161 | 0.018897 | 0.012533 | -0.0019849 | -0.028467 | 0.012533 | 0.0015331 |  | 0.019708
7943.2823 | -27.0725 | -27.076 | -27.0914 | -27.12 | -27.081 | -27.08 | -27.0821 | -27.0711 | 0.06 | 0.06 | 0.06 | 0.04 | 0.04 | 0.05 | 0.07 | 0.13 | -27.0893 | 0.0095405 | 0.016761 | 0.013287 | -0.0021463 | -0.030713 | 0.0082872 | 0.0092872 | 0.0072229 | 0.018196
8413.9514 | -27.59 | -27.584 | -27.6023 | -27.62 | -27.586 | -27.58 |  | -27.5859 | 0.08 | 0.06 | 0.09 | 0.09 | 0.07 | 0.08 |  | 0.13 | -27.5908 | 0.015073 | 0.00084813 | 0.0068481 | -0.011468 | -0.029152 | 0.0048481 | 0.010848 |  | 0.0049946
8912.5094 | -28.25 | -28.238 | -28.2591 | -28.27 | -28.247 | -28.23 |  | -28.2538 | 0.08 | 0.07 | 0.09 | 0.09 | 0.07 | 0.08 |  | 0.13 | -28.2477 | 0.015604 | -0.0023062 | 0.0096938 | -0.011394 | -0.022306 | 0.00069376 | 0.017694 |  | -0.0061465
9440.6088 | -29.07 | -29.019 | -29.0605 | -29.08 | -29.062 | -28.99 |  | -29.0038 | 0.08 | 0.07 | 0.09 | 0.09 | 0.07 | 0.08 |  | 0.13 | -29.0423 | 0.015604 | -0.027661 | 0.023339 | -0.018119 | -0.037661 | -0.019661 | 0.052339 |  | 0.038529
10000 | -29.9886 | -29.89 | -29.9796 | -30 | -29.987 | -29.87 | -29.8053 | -29.8775 | 0.08 | 0.09 | 0.09 | 0.09 | 0.07 | 0.08 | 0.1 | 0.13 | -29.9354 | 0.015459 | -0.053163 | 0.045414 | -0.044214 | -0.064586 | -0.051586 | 0.065414 | 0.13014 | 0.057871

## 2652765 DoE Level
4160. 2652765 | Linking DoEs from CCAUV.A-K5 |  |  | 4160 2652765 | [7 empty] | Linking U(DoE)s from CCAUV.A-K5 |  |  | U(DoE) (dB)
DoE (dB)
Frequency (Hz) | NIM | NMIJ | KRISS | NIM | NMIJ | KRISS | NMIA | SCL | NIMT | NIMM | CMS-ITRI | NIM | NMIJ | KRISS | NIM | NMIJ | KRISS | NMIA | SCL | NIMT | NIMM | CMS-ITRI
1.9953 | -0.002047 | -0.012047 | -0.032047 | 0.032932 | -0.014884 | -0.040884 | 0.066116 |  | 0.095116 |  | 0.08464 | 0.1961 | 0.18589 | 0.069688 | 0.19352 | 0.22266 | 0.084932 | 0.21708 |  | 0.21708 |  | 0.098107
2.5119 | 0.0038185 | -0.016182 | -0.026182 | 0.013973 | -0.012673 | -0.032673 | 0.045327 |  | 0.067327 |  | 0.057679 | 0.1763 | 0.16608 | 0.071283 | 0.17252 | 0.21228 | 0.079617 | 0.1704 |  | 0.21572 |  | 0.095057
3.1623 | -0.00055047 | 0.01945 | -0.01055 | 0.024318 | 0.018127 | -0.014873 | 0.053127 |  | 0.078127 |  | 0.063446 | 0.13705 | 0.13705 | 0.041009 | 0.15045 | 0.21162 | 0.049508 | 0.13809 |  | 0.20535 |  | 0.068323
3.9811 | 0.001366 | 0.011366 | -0.008634 | 0.018977 | 0.071486 | -0.018514 | 0.037486 |  | 0.061486 |  | 0.052984 | 0.11773 | 0.097261 | 0.032554 | 0.14911 | 0.17932 | 0.052848 | 0.12884 |  | 0.20542 |  | 0.06855
5.0119 | 0.010488 | 0.00048772 | -0.0095123 | 0.0042259 | 0.031575 | -0.025425 | 0.018575 |  | 0.041575 |  | 0.032201 | 0.0978 | 0.077232 | 0.034128 | 0.090494 | 0.0757 | 0.054473 | 0.11301 |  | 0.20167 |  | 0.056319
6.3096 | 0.008607 | 0.008607 | -0.011393 | 0.0022773 | 0.017303 | -0.016697 | 0.015303 |  | 0.037303 |  | 0.027693 | 0.077912 | 0.057187 | 0.035641 | 0.094289 | 0.066008 | 0.05225 | 0.10279 |  | 0.13216 |  | 0.055373
7.9433 | 0.0049767 | 0.0049767 | -0.0050233 | 0.0042852 | 0.015216 | -0.023784 | 0.012216 |  | 0.035216 |  | 0.0286 | 0.057953 | 0.047524 | 0.025662 | 0.077096 | 0.034196 | 0.061258 | 0.093209 |  | 0.093209 |  | 0.05557
10 | 0.0072267 | -0.0027733 | -0.0027733 | -0.0048111 | 0.020402 | -0.028598 | 0.0024021 |  | 0.020402 |  | 0.017282 | 0.058109 | 0.047715 | 0.026013 | 0.077407 | 0.033857 | 0.06175 | 0.083603 |  | 0.064725 |  | 0.055582
12.5893 | 0.0078617 | -0.012138 | -0.0021383 | -0.014962 | 0.015773 | -0.035227 | -0.0082265 |  | 0.015773 |  | 0.0051381 | 0.048146 | 0.048146 | 0.026796 | 0.082781 | 0.023811 | 0.06365 | 0.072907 |  | 0.053994 |  | 0.053994
15.8489 | 0.011741 | -0.0082591 | 0.0017409 | -0.012361 | 0.016792 | -0.028208 | -0.009208 |  | 0.016792 |  | 0.0082649 | 0.048301 | 0.037854 | 0.027073 | 0.073363 | 0.023137 | 0.064157 | 0.063447 |  | 0.054089 |  | 0.054089
19.9526 | 0.0080582 | -0.011942 | -0.0019418 | -0.015076 | 0.01541 | -0.02559 | -0.01159 | 0.0030771 | 0.01541 | -0.0029415 | 0.004528 | 0.048742 | 0.038416 | 0.027853 | 0.064464 | 0.02712 | 0.052796 | 0.063044 | 0.063044 | 0.053614 | 0.072626 | 0.053614
21.1349 | 0.0072489 | -0.0027511 | -0.0027511 | -0.010847 | 0.016652 | -0.021348 | -0.0073479 | 0.0068165 | 0.016652 |  | 0.008515 | 0.048694 | 0.038355 | 0.027769 | 0.054 | 0.028241 | 0.053233 | 0.062848 | 0.062848 | 0.053384 |  | 0.053384
22.3872 | 0.012729 | -0.0072713 | 0.0027287 | -0.0060463 | 0.018207 | -0.018793 | -0.002793 | 0.012306 | 0.018207 |  | 0.013258 | 0.048694 | 0.038355 | 0.027769 | 0.054 | 0.028241 | 0.053233 | 0.062848 | 0.062848 | 0.053384 |  | 0.053384
23.7137 | 0.011722 | -0.0082775 | 0.0017225 | -0.010649 | 0.020127 | -0.021873 | -0.0078733 | 0.0065058 | 0.010127 |  | 0.0087954 | 0.048694 | 0.038355 | 0.027769 | 0.054 | 0.028241 | 0.053233 | 0.062848 | 0.062848 | 0.053384 |  | 0.053384
25.1189 | 0.010652 | -0.0093484 | 0.00065159 | -0.0089663 | 0.018087 | -0.019913 | -0.0059127 | 0.0096059 | 0.018087 |  | 0.010267 | 0.03859 | 0.03859 | 0.028092 | 0.056068 | 0.026955 | 0.051962 | 0.053449 | 0.053449 | 0.053449 |  | 0.053449
26.6073 | 0.0061544 | -0.0038456 | 0.0061544 | -0.0093164 | 0.023763 | -0.019237 | -0.0062369 | 0.0088923 | 0.013763 |  | 0.0097853 | 0.03859 | 0.03859 | 0.028092 | 0.056068 | 0.026955 | 0.051962 | 0.053449 | 0.053449 | 0.053449 |  | 0.053449
28.1838 | 0.011273 | -0.008727 | 0.001273 | -0.010067 | 0.018785 | -0.018215 | -0.0072149 | 0.0073177 | 0.018785 |  | 0.0088278 | 0.03859 | 0.03859 | 0.028092 | 0.056068 | 0.026955 | 0.051962 | 0.053449 | 0.053449 | 0.053449 |  | 0.053449
29.8538 | 0.010723 | -0.0092771 | 0.0007229 | -0.0077082 | 0.016666 | -0.018334 | -0.0053337 | 0.009807 | 0.016666 |  | 0.010903 | 0.03859 | 0.03859 | 0.028092 | 0.056068 | 0.026955 | 0.051962 | 0.053449 | 0.053449 | 0.053449 |  | 0.053449
31.6228 | 0.0063173 | -0.0036827 | 0.0063173 | -0.0073828 | 0.022275 | -0.017725 | -0.0057246 | 0.0097071 | 0.012275 | 0.0053529 | 0.011217 | 0.038742 | 0.038742 | 0.0283 | 0.054682 | 0.02773 | 0.052713 | 0.053821 | 0.053821 | 0.053821 | 0.053821 | 0.053821
33.4965 | 0.013431 | -0.0065691 | 0.0034309 | -0.0050861 | 0.019633 | -0.013367 | -0.0033673 | 0.010872 | 0.019633 |  | 0.013097 | 0.038801 | 0.038801 | 0.028381 | 0.044415 | 0.029127 | 0.053608 | 0.053293 | 0.053293 | 0.053293 |  | 0.053293
35.4813 | 0.0090061 | -0.00099394 | 0.0090061 | -0.00014378 | 0.019433 | -0.0075667 | 0.0014333 | 0.015703 | 0.019433 |  | 0.017958 | 0.038801 | 0.038801 | 0.028381 | 0.044415 | 0.029127 | 0.053608 | 0.053293 | 0.053293 | 0.053293 |  | 0.053293
37.5837 | 0.015115 | -0.0048847 | 0.0051153 | -0.0030698 | 0.021187 | -0.011813 | -0.0018125 | 0.012899 | 0.021187 |  | 0.014689 | 0.038801 | 0.038801 | 0.028381 | 0.044415 | 0.029127 | 0.053608 | 0.053293 | 0.053293 | 0.053293 |  | 0.053293
39.8107 | 0.011167 | -0.0088332 | 0.0011668 | -0.0047665 | 0.014019 | -0.012981 | -0.0039812 | 0.010494 | 0.014019 |  | 0.012779 | 0.038804 | 0.038804 | 0.028386 | 0.043762 | 0.028494 | 0.054675 | 0.043384 | 0.052746 | 0.052746 |  | 0.052746
42.1697 | 0.015864 | -0.0041356 | 0.0058644 | -0.0021119 | 0.021075 | -0.0099253 | -0.00092532 | 0.013748 | 0.021075 |  | 0.015051 | 0.038804 | 0.038804 | 0.028386 | 0.043762 | 0.028494 | 0.054675 | 0.043384 | 0.052746 | 0.052746 |  | 0.052746
44.6684 | 0.012636 | -0.0073636 | 0.0026364 | -0.0029212 | 0.014561 | -0.010439 | -0.0024386 | 0.012792 | 0.024561 |  | 0.014186 | 0.038804 | 0.038804 | 0.028386 | 0.043762 | 0.028494 | 0.054675 | 0.043384 | 0.052746 | 0.052746 |  | 0.052746
47.3151 | 0.01653 | -0.0034696 | 0.0065304 | -0.00092971 | 0.020759 | -0.0082413 | -0.00024133 | 0.01512 | 0.020759 |  | 0.015969 | 0.038804 | 0.038804 | 0.028386 | 0.043762 | 0.028494 | 0.054675 | 0.043384 | 0.052746 | 0.052746 |  | 0.052746
50.1187 | 0.021243 | -0.008757 | 0.001243 | 0.00024619 | 0.016056 | -0.0059439 | 5.6095e-05 | 0.016122 | 0.026056 |  | 0.017281 | 0.03869 | 0.03869 | 0.02823 | 0.04715 | 0.03324 | 0.038966 | 0.044868 | 0.044868 | 0.053973 |  | 0.053973
53.0884 | 0.016533 | -0.0034666 | 0.0065334 | 0.00059251 | 0.020435 | -0.0035655 | 0.00043453 | 0.016481 | 0.020435 |  | 0.017717 | 0.038804 | 0.038804 | 0.028386 | 0.047246 | 0.033285 | 0.039151 | 0.044808 | 0.044808 | 0.053924 |  | 0.053924
56.2341 | 0.010409 | -0.0095907 | 0.010409 | -0.0026638 | 0.021105 | -0.0058955 | -0.0028955 | 0.013308 | 0.021105 |  | 0.014404 | 0.038804 | 0.038804 | 0.028386 | 0.047246 | 0.033285 | 0.039151 | 0.044808 | 0.044808 | 0.053924 |  | 0.053924
59.5662 | 0.015291 | -0.0047089 | 0.0052911 | 0.00031874 | 0.017882 | -0.0041177 | -0.00011768 | 0.016305 | 0.027882 |  | 0.017481 | 0.038804 | 0.038804 | 0.028386 | 0.047246 | 0.033285 | 0.039151 | 0.044808 | 0.044808 | 0.053924 |  | 0.053924
63.0957 | 0.018196 | -0.0018036 | 0.0081964 | 0.0048439 | 0.026043 | 4.3362e-05 | 0.0040434 | 0.020715 | 0.026043 | 0.022829 | 0.022116 | 0.038804 | 0.038804 | 0.028386 | 0.038199 | 0.039006 | 0.02869 | 0.04509 | 0.04509 | 0.04509 | 0.04509 | 0.054158
66.8344 | 0.013821 | -0.0061789 | 0.0038211 | -0.0011308 | 0.023524 | -0.0044762 | -0.0014762 | 0.014532 | 0.023524 |  | 0.01597 | 0.038804 | 0.038804 | 0.028386 | 0.038199 | 0.039006 | 0.02869 | 0.04509 | 0.04509 | 0.04509 |  | 0.054158
70.7946 | 0.008187 | -0.001813 | 0.008187 | -0.0013483 | 0.02658 | -0.0024203 | -0.0024203 | 0.014211 | 0.01658 |  | 0.015455 | 0.038899 | 0.038899 | 0.028515 | 0.038282 | 0.039105 | 0.028825 | 0.045092 | 0.045092 | 0.045092 |  | 0.05416
74.9894 | 0.0095611 | -0.00043893 | 0.0095611 | 0.0027278 | 0.023781 | -0.00021896 | 0.001781 | 0.017773 | 0.023781 |  | 0.019359 | 0.038899 | 0.038899 | 0.028515 | 0.038282 | 0.039105 | 0.028825 | 0.045092 | 0.045092 | 0.045092 |  | 0.05416
79.4328 | 0.01546 | -0.0045401 | 0.0054599 | -3.0001e-05 | 0.024076 | -0.0019238 | -0.00092379 | 0.014977 | 0.024076 |  | 0.016335 | 0.038899 | 0.038899 | 0.028515 | 0.038487 | 0.038949 | 0.02879 | 0.045267 | 0.045267 | 0.045267 |  | 0.054305
84.1395 | 0.0081959 | -0.0018041 | 0.0081959 | 0.0018491 | 0.019041 | 4.0731e-05 | 4.0731e-05 | 0.016918 | 0.019041 |  | 0.018331 | 0.038899 | 0.038899 | 0.028515 | 0.038487 | 0.038949 | 0.02879 | 0.045267 | 0.045267 | 0.045267 |  | 0.054305
89.1251 | 0.013623 | -0.0063774 | 0.0036226 | -0.00053784 | 0.01984 | -0.0031596 | -0.0021596 | 0.014642 | 0.01984 |  | 0.015909 | 0.038899 | 0.038899 | 0.028515 | 0.038487 | 0.038949 | 0.02879 | 0.045267 | 0.045267 | 0.045267 |  | 0.054305
94.4061 | 0.014813 | -0.0051873 | 0.0048127 | -0.00072496 | 0.022806 | -0.0021937 | -0.0021937 | 0.014611 | 0.022806 |  | 0.016021 | 0.038899 | 0.038899 | 0.028515 | 0.038487 | 0.038949 | 0.02879 | 0.045267 | 0.045267 | 0.045267 |  | 0.054305
100 | 0.0086347 | -0.011365 | 0.0086347 | -0.0025652 | 0.023387 | -0.0046133 | -0.0046133 | 0.012189 | 0.023387 |  | 0.013662 | 0.038899 | 0.038899 | 0.028515 | 0.038345 | 0.039406 | 0.028545 | 0.045147 | 0.045147 | 0.045147 |  | 0.054205
105.9254 | 0.01218 | -0.0078199 | 0.01218 | 0.0021338 | 0.020757 | 0.0017567 | 0.00075675 | 0.016765 | 0.020757 |  | 0.01823 | 0.038904 | 0.038904 | 0.028522 | 0.038349 | 0.039411 | 0.028552 | 0.036582 | 0.045147 | 0.045147 |  | 0.054205
112.2018 | 0.0047087 | -0.0052913 | 0.014709 | 0.0010289 | 0.022273 | 0.0012735 | -0.0007265 | 0.015429 | 0.022273 |  | 0.017045 | 0.038904 | 0.038904 | 0.028522 | 0.038349 | 0.039411 | 0.028552 | 0.036582 | 0.045147 | 0.045147 |  | 0.054205
118.8502 | 0.0083578 | -0.011642 | 0.0083578 | -0.0026775 | 0.021131 | -0.0038695 | -0.0048695 | 0.011595 | 0.021131 |  | 0.013225 | 0.038948 | 0.038948 | 0.028583 | 0.038389 | 0.039461 | 0.028613 | 0.036584 | 0.045148 | 0.045148 |  | 0.054207
125.8925 | 0.01192 | -0.0080799 | 0.01192 | 0.0021665 | 0.018471 | 0.0024715 | -0.00052852 | 0.016455 | 0.018471 | 0.017577 | 0.018082 | 0.038948 | 0.038948 | 0.028583 | 0.038574 | 0.039009 | 0.028821 | 0.036777 | 0.045305 | 0.045305 | 0.045305 | 0.054338
133.3521 | 0.014113 | -0.005887 | 0.004113 | -0.0011872 | 0.017541 | -0.00045889 | -0.0034589 | 0.012953 | 0.017541 |  | 0.014551 | 0.038948 | 0.038948 | 0.028583 | 0.038574 | 0.039009 | 0.028821 | 0.036777 | 0.045305 | 0.045305 |  | 0.054338
141.2538 | 0.006248 | -0.003752 | 0.006248 | -0.0029517 | 0.01812 | -0.0008801 | -0.0048801 | 0.011244 | 0.01812 |  | 0.012904 | 0.039058 | 0.039058 | 0.018043 | 0.038675 | 0.03912 | 0.028976 | 0.036787 | 0.045313 | 0.045313 |  | 0.054344
149.6236 | 0.0078274 | -0.0021726 | 0.0078274 | -0.0011505 | 0.022179 | -0.00082058 | -0.0038206 | 0.013103 | 0.022179 |  | 0.014685 | 0.039058 | 0.039058 | 0.018043 | 0.038675 | 0.03912 | 0.028976 | 0.036787 | 0.045313 | 0.045313 |  | 0.054344
158.4893 | 0.011268 | -0.0087322 | 0.011268 | -0.0014818 | 0.024018 | 1.7632e-05 | -0.0039824 | 0.012524 | 0.024018 |  | 0.014462 | 0.039058 | 0.039058 | 0.018043 | 0.038715 | 0.039252 | 0.028846 | 0.036829 | 0.045347 | 0.045347 |  | 0.054373
167.8804 | 0.0033772 | -0.0066228 | 0.013377 | 0.0014523 | 0.01903 | 3.0191e-05 | -0.00096981 | 0.015463 | 0.01903 |  | 0.01728 | 0.039058 | 0.039058 | 0.018043 | 0.038715 | 0.039252 | 0.028846 | 0.036829 | 0.045347 | 0.045347 |  | 0.054373
177.8279 | 0.0064627 | -0.0035373 | 0.0064627 | -0.0016286 | 0.017936 | -0.0010644 | -0.0040644 | 0.012431 | 0.017936 |  | 0.01429 | 0.039058 | 0.039058 | 0.018043 | 0.038715 | 0.039252 | 0.028846 | 0.036829 | 0.045347 | 0.045347 |  | 0.054373
188.3649 | 0.0077518 | -0.0022482 | 0.0077518 | -0.001262 | 0.020197 | 0.00019669 | -0.0038033 | 0.012621 | 0.020197 |  | 0.014661 | 0.039058 | 0.039058 | 0.018043 | 0.038715 | 0.039252 | 0.028846 | 0.036829 | 0.045347 | 0.045347 |  | 0.054373
199.5262 | 0.010695 | -0.0093053 | 0.010695 | -0.00044725 | 0.022815 | 0.00081508 | -0.0031849 | 0.0134 | 0.022815 |  | 0.015298 | 0.028778 | 0.039092 | 0.018115 | 0.039612 | 0.037615 | 0.026743 | 0.036114 | 0.044769 | 0.044769 |  | 0.053891
211.3489 | 0.01149 | -0.0085103 | 0.01149 | 0.0026752 | 0.017652 | 0.0036522 | -0.00034776 | 0.016454 | 0.027652 |  | 0.018316 | 0.028778 | 0.039092 | 0.018115 | 0.039612 | 0.037615 | 0.026743 | 0.036114 | 0.044769 | 0.044769 |  | 0.053891
223.8721 | 0.0058013 | -0.0041987 | 0.0058013 | -0.0015441 | 0.01506 | 6.0301e-05 | -0.0039397 | 0.012126 | 0.01506 |  | 0.013994 | 0.028778 | 0.039092 | 0.018115 | 0.039612 | 0.037615 | 0.026743 | 0.036114 | 0.044769 | 0.044769 |  | 0.053891
237.1374 | 0.0070694 | -0.0029306 | 0.0070694 | -0.00046815 | 0.017678 | 0.00067757 | -0.0033224 | 0.013074 | 0.017678 |  | 0.014987 | 0.028778 | 0.039092 | 0.018115 | 0.039612 | 0.037615 | 0.026743 | 0.036114 | 0.044769 | 0.044769 |  | 0.053891
251.1886 | 0.0082476 | -0.011752 | 0.0082476 | -0.0016139 | 0.017987 | -0.001013 | -0.004013 | 0.01179 | 0.017987 | 0.015771 | 0.013773 | 0.028778 | 0.039092 | 0.018115 | 0.039537 | 0.037663 | 0.026767 | 0.036135 | 0.044786 | 0.044786 | 0.044786 | 0.053905
266.0725 | 0.010027 | -0.0099728 | 0.010027 | 0.00010153 | 0.02107 | 6.974e-05 | -0.0019303 | 0.013409 | 0.02107 |  | 0.015353 | 0.028778 | 0.039092 | 0.018115 | 0.039537 | 0.037663 | 0.026767 | 0.036135 | 0.044786 | 0.044786 |  | 0.053905
281.8383 | 0.01177 | -0.0082304 | 0.01177 | 0.0013769 | 0.02362 | 0.0016203 | -0.0013797 | 0.01469 | 0.02362 |  | 0.016648 | 0.02893 | 0.039204 | 0.018357 | 0.039691 | 0.037764 | 0.026908 | 0.036128 | 0.04478 | 0.04478 |  | 0.0539
298.5383 | 0.012526 | -0.0074743 | 0.012526 | 0.0043855 | 0.017806 | 0.0048055 | 0.0018055 | 0.017658 | 0.027806 |  | 0.019525 | 0.02893 | 0.039204 | 0.018357 | 0.039691 | 0.037764 | 0.026908 | 0.036128 | 0.04478 | 0.04478 |  | 0.0539
316.2278 | 0.013452 | -0.0065479 | 0.013452 | 0.004996 | 0.019486 | 0.0054856 | 0.0024856 | 0.018354 | 0.029486 |  | 0.020276 | 0.02893 | 0.039204 | 0.018357 | 0.039687 | 0.037783 | 0.026897 | 0.036124 | 0.044776 | 0.044776 |  | 0.053897
334.9654 | 0.0057286 | -0.0042714 | 0.0057286 | -0.00076261 | 0.01468 | -0.00031991 | -0.0033199 | 0.012423 | 0.02468 |  | 0.014612 | 0.02893 | 0.039204 | 0.018357 | 0.039687 | 0.037783 | 0.026897 | 0.036124 | 0.044776 | 0.044776 |  | 0.053897
354.8134 | 0.0060637 | -0.0039363 | 0.0060637 | -0.00056565 | 0.015705 | -0.00029516 | -0.0032952 | 0.012658 | 0.015705 |  | 0.014549 | 0.02893 | 0.039204 | 0.018357 | 0.039687 | 0.037783 | 0.026897 | 0.036124 | 0.044776 | 0.044776 |  | 0.053897
375.8374 | 0.0063987 | -0.0036013 | 0.0063987 | -0.00074243 | 0.016224 | 0.00022415 | -0.0037759 | 0.01237 | 0.016224 |  | 0.014327 | 0.02893 | 0.039204 | 0.018357 | 0.039687 | 0.037783 | 0.026897 | 0.036124 | 0.044776 | 0.044776 |  | 0.053897
398.1072 | 0.0071169 | -0.0028831 | 0.0071169 | 3.3775e-05 | 0.017562 | 0.00056152 | -0.0024385 | 0.013095 | 0.017562 |  | 0.014982 | 0.02893 | 0.039204 | 0.018357 | 0.039718 | 0.037769 | 0.026882 | 0.036157 | 0.044803 | 0.044803 |  | 0.05392
421.6965 | 0.0071325 | -0.0028675 | 0.0071325 | 0.00020853 | 0.01816 | 0.0001597 | -0.0018403 | 0.013246 | 0.01816 |  | 0.015179 | 0.02893 | 0.039204 | 0.018357 | 0.039718 | 0.037769 | 0.026882 | 0.036157 | 0.044803 | 0.044803 |  | 0.05392
446.6836 | 0.0066295 | -0.0033705 | 0.0066295 | -0.00049453 | 0.017729 | -0.00027128 | -0.0032713 | 0.012579 | 0.017729 |  | 0.014505 | 0.02893 | 0.039204 | 0.018357 | 0.039718 | 0.037769 | 0.026882 | 0.036157 | 0.044803 | 0.044803 |  | 0.05392
473.1513 | 0.0065598 | -0.0034402 | 0.0065598 | 0.0015627 | 0.0098534 | 0.0028534 | -0.0011466 | 0.014667 | 0.019853 |  | 0.016522 | 0.02893 | 0.039204 | 0.018357 | 0.039718 | 0.037769 | 0.026882 | 0.036157 | 0.044803 | 0.044803 |  | 0.05392
501.1872 | 0.0057218 | -0.0042782 | 0.0057218 | -0.0017708 | 0.016314 | -0.00068628 | -0.0046863 | 0.011278 | 0.016314 | 0.02093 | 0.013058 | 0.02893 | 0.039204 | 0.018357 | 0.039676 | 0.037842 | 0.026859 | 0.036136 | 0.044787 | 0.044787 | 0.044787 | 0.053906
530.8844 | 0.0053326 | -0.0046674 | 0.0053326 | -0.0018048 | 0.015797 | -0.0012034 | -0.0052034 | 0.011149 | 0.015797 |  | 0.013084 | 0.02893 | 0.039204 | 0.018357 | 0.039676 | 0.037842 | 0.026859 | 0.036136 | 0.044787 | 0.044787 |  | 0.053906
562.3413 | 0.0051686 | -0.0048314 | 0.0051686 | -0.0014635 | 0.015451 | -0.0015492 | -0.0045492 | 0.011457 | 0.015451 |  | 0.013146 | 0.02893 | 0.039204 | 0.018357 | 0.039676 | 0.037842 | 0.026859 | 0.036136 | 0.044787 | 0.044787 |  | 0.053906
595.6621 | 0.0028141 | -0.0071859 | 0.012814 | -0.00018069 | 0.015814 | 0.00081411 | -0.0031859 | 0.012647 | 0.015814 |  | 0.014461 | 0.02893 | 0.039204 | 0.018357 | 0.039676 | 0.037842 | 0.026859 | 0.036136 | 0.044787 | 0.044787 |  | 0.053906
630.9573 | 0.011265 | -0.0087351 | 0.011265 | 0.002572 | 0.017384 | 0.0033836 | -0.00061639 | 0.015136 | 0.027384 |  | 0.017049 | 0.02893 | 0.039204 | 0.018357 | 0.039635 | 0.03798 | 0.026783 | 0.036091 | 0.04475 | 0.04475 |  | 0.053876
668.3439 | 0.010322 | -0.0096778 | 0.010322 | 0.0017775 | 0.015108 | 0.0031076 | -0.00089244 | 0.014337 | 0.025108 |  | 0.016165 | 0.02893 | 0.039204 | 0.018357 | 0.039635 | 0.03798 | 0.026783 | 0.036091 | 0.04475 | 0.04475 |  | 0.053876
707.9458 | 0.0076656 | -0.0023344 | 0.0076656 | 0.0017114 | 0.013224 | 0.0032237 | -0.0017763 | 0.014134 | 0.023224 |  | 0.016034 | 0.02893 | 0.039204 | 0.018357 | 0.039635 | 0.03798 | 0.026783 | 0.036091 | 0.04475 | 0.04475 |  | 0.053876
749.8942 | 0.0033521 | -0.0066479 | 0.013352 | -0.0011439 | 0.018173 | 0.0011726 | -0.0038274 | 0.01125 | 0.018173 |  | 0.01333 | 0.02893 | 0.039204 | 0.018357 | 0.039635 | 0.03798 | 0.026783 | 0.036091 | 0.04475 | 0.04475 |  | 0.053876
794.3282 | 0.011094 | -0.0089062 | 0.011094 | 0.001015 | 0.017711 | 0.0037115 | -0.0022885 | 0.013374 | 0.017711 |  | 0.015334 | 0.02893 | 0.039204 | 0.018357 | 0.039666 | 0.037897 | 0.026824 | 0.036126 | 0.044778 | 0.044778 |  | 0.053899
841.3951 | 0.0072379 | -0.0027621 | 0.0072379 | 0.00068836 | 0.01429 | 0.0022903 | -0.0017097 | 0.012925 | 0.02429 |  | 0.014991 | 0.02893 | 0.039204 | 0.018357 | 0.039666 | 0.037897 | 0.026824 | 0.036126 | 0.044778 | 0.044778 |  | 0.053899
891.2509 | 0.011756 | -0.008244 | 0.011756 | 0.00067629 | 0.020654 | 0.003654 | -0.002346 | 0.01301 | 0.020654 |  | 0.015126 | 0.02893 | 0.039204 | 0.018357 | 0.039666 | 0.037897 | 0.026824 | 0.036126 | 0.044778 | 0.044778 |  | 0.053899
944.0609 | 0.0074998 | -0.0025002 | 0.0074998 | -0.00063868 | 0.015124 | 0.0031242 | -0.0038758 | 0.011581 | 0.015124 |  | 0.013874 | 0.02893 | 0.039204 | 0.018357 | 0.039666 | 0.037897 | 0.026824 | 0.036126 | 0.044778 | 0.044778 |  | 0.053899
1000 | 0.010304 | 0.00030365 | 0.010304 | 0.003865 | 0.014756 | 0.0067562 | 0.00075618 | 0.016054 | 0.024756 | 0.024681 | 0.018337 | 0.02893 | 0.039204 | 0.018357 | 0.039687 | 0.037849 | 0.026845 | 0.036148 | 0.044796 | 0.044796 | 0.044796 | 0.053914
1059.2537 | 0.0032595 | -0.0067405 | 0.013259 | -0.00015776 | 0.015132 | 0.0021317 | -0.0028683 | 0.012163 | 0.015132 |  | 0.014243 | 0.02893 | 0.039204 | 0.018357 | 0.039687 | 0.037849 | 0.026845 | 0.036148 | 0.044796 | 0.044796 |  | 0.053914
1122.0185 | 0.0055101 | -0.0044899 | 0.0055101 | -0.0030363 | 0.015841 | -0.00015851 | -0.0061585 | 0.0093046 | 0.015841 |  | 0.011137 | 0.028928 | 0.039203 | 0.018353 | 0.039685 | 0.037847 | 0.026843 | 0.036148 | 0.044796 | 0.044796 |  | 0.053914
1188.5022 | 0.0075625 | -0.0024375 | 0.0075625 | 0.0001554 | 0.011724 | 0.0047237 | -0.0032763 | 0.012269 | 0.021724 |  | 0.014377 | 0.028926 | 0.039201 | 0.01835 | 0.039683 | 0.037846 | 0.026841 | 0.036148 | 0.044796 | 0.044796 |  | 0.053914
1258.9254 | 0.0072695 | -0.0027305 | 0.0072695 | 2.6891e-05 | 0.013292 | 0.0032915 | -0.0037085 | 0.011899 | 0.023292 | 0.018437 | 0.014312 | 0.028925 | 0.0392 | 0.018348 | 0.03963 | 0.038105 | 0.026674 | 0.036092 | 0.044751 | 0.044751 | 0.044751 | 0.053876
1333.5214 | 0.0042242 | -0.0057758 | 0.014224 | 0.001387 | 0.015248 | 0.0032476 | -0.0017524 | 0.013193 | 0.015248 |  | 0.01521 | 0.028925 | 0.0392 | 0.018348 | 0.03963 | 0.038105 | 0.026674 | 0.036092 | 0.044751 | 0.044751 |  | 0.053876
1412.5375 | 0.01004 | 4.0473e-05 | 0.01004 | 0.0018752 | 0.015115 | 0.0071146 | -0.00088541 | 0.013413 | 0.025115 |  | 0.015456 | 0.028925 | 0.0392 | 0.018348 | 0.03963 | 0.038105 | 0.026674 | 0.036092 | 0.044751 | 0.044751 |  | 0.053876
1496.2357 | 0.0054928 | -0.0045072 | 0.015493 | 0.0012646 | 0.012537 | 0.0075369 | -0.0014631 | 0.012506 | 0.022537 |  | 0.014775 | 0.028925 | 0.0392 | 0.018348 | 0.03963 | 0.038105 | 0.026674 | 0.036092 | 0.044751 | 0.044751 |  | 0.053876
1584.8932 | 0.0084528 | -0.0015472 | 0.0084528 | 0.0014504 | 0.0092683 | 0.0072683 | -0.0017317 | 0.01234 | 0.019268 | 0.029429 | 0.014744 | 0.028926 | 0.039201 | 0.01835 | 0.039747 | 0.037824 | 0.026805 | 0.036219 | 0.044853 | 0.044853 | 0.044853 | 0.053961
1678.804 | 0.0098498 | -0.01015 | 0.0098498 | -0.00012861 | 0.012591 | 0.0045913 | -0.0034087 | 0.010612 | 0.012591 |  | 0.01323 | 0.028926 | 0.039201 | 0.01835 | 0.039747 | 0.037824 | 0.026805 | 0.036219 | 0.044853 | 0.044853 |  | 0.053961
1778.2794 | 0.0089232 | -0.011077 | 0.0089232 | -0.0019132 | 0.013894 | 0.0028936 | -0.0051064 | 0.0086796 | 0.013894 |  | 0.01134 | 0.028926 | 0.039201 | 0.01835 | 0.039747 | 0.037824 | 0.026805 | 0.036219 | 0.044853 | 0.044853 |  | 0.053961
1883.6491 | 0.005003 | -0.004997 | 0.015003 | 0.0016439 | 0.0085342 | 0.0085342 | -0.0014658 | 0.011971 | 0.018534 |  | 0.014963 | 0.028926 | 0.039201 | 0.01835 | 0.039747 | 0.037824 | 0.026805 | 0.036219 | 0.044853 | 0.044853 |  | 0.053961
1995.2623 | 0.0093808 | -0.010619 | 0.0093808 | 0.00075165 | 0.0065166 | 0.0065166 | -0.0024834 | 0.010969 | 0.016517 | 0.02489 | 0.013608 | 0.028926 | 0.039201 | 0.01835 | 0.039652 | 0.037763 | 0.026932 | 0.036114 | 0.044769 | 0.044769 | 0.044769 | 0.053891
2113.489 | 0.0095914 | -0.010409 | 0.0095914 | -0.00044398 | 0.011764 | 0.0027641 | -0.0032359 | 0.009123 | 0.021764 |  | 0.012825 | 0.039176 | 0.039176 | 0.018296 | 0.038765 | 0.03933 | 0.029085 | 0.036869 | 0.04538 | 0.04538 |  | 0.0544
2238.7211 | 0.0078522 | -0.0021478 | 0.0078522 | 0.00080724 | 0.0067884 | 0.0067884 | -0.0012116 | 0.010038 | 0.016788 |  | 0.014175 | 0.039176 | 0.039176 | 0.018296 | 0.038765 | 0.03933 | 0.029085 | 0.036869 | 0.04538 | 0.04538 |  | 0.0544
2371.3737 | 0.0094638 | -0.00053617 | 0.0094638 | 0.0027338 | 0.0095666 | 0.0075666 | 0.00056663 | 0.011676 | 0.019567 |  | 0.015641 | 0.039176 | 0.039176 | 0.018296 | 0.038765 | 0.03933 | 0.029085 | 0.036869 | 0.04538 | 0.04538 |  | 0.0544
2511.8864 | 0.010486 | -0.0095139 | 0.010486 | 0.0024204 | 0.0068298 | 0.0058298 | 0.00082975 | 0.010797 | 0.01683 | 0.021348 | 0.01581 | 0.039176 | 0.039176 | 0.018296 | 0.038919 | 0.038975 | 0.029237 | 0.037031 | 0.045511 | 0.045511 | 0.045511 | 0.054509
2660.7251 | 0.0039896 | -0.0060104 | 0.01399 | 0.0008506 | 0.0091997 | 0.0071997 | -0.0018003 | 0.0085676 | 0.0192 |  | 0.013332 | 0.039119 | 0.039119 | 0.018174 | 0.038865 | 0.038921 | 0.029157 | 0.037023 | 0.045505 | 0.045505 |  | 0.054504
2818.3829 | 0.0060488 | -0.0039512 | 0.016049 | 0.0046786 | 0.0035821 | 0.012582 | 0.0015821 | 0.011866 | 0.023582 |  | 0.0172 | 0.039024 | 0.039024 | 0.028685 | 0.038775 | 0.038829 | 0.029021 | 0.037009 | 0.045494 | 0.045494 |  | 0.054495
2985.3826 | 0.013126 | -0.0068744 | 0.013126 | 0.0062358 | 0.0020059 | 0.012006 | 0.0040059 | 0.012699 | 0.022006 |  | 0.018509 | 0.039024 | 0.039024 | 0.028685 | 0.038775 | 0.038829 | 0.029021 | 0.037009 | 0.045494 | 0.045494 |  | 0.054495
3162.2777 | 0.014168 | -0.0058321 | 0.014168 | 0.0069971 | 0.0057093 | 0.011709 | 0.0037093 | 0.013124 | 0.025709 | 0.03177 | 0.019469 | 0.039018 | 0.039018 | 0.028677 | 0.038712 | 0.038799 | 0.029075 | 0.036949 | 0.045444 | 0.045444 | 0.045444 | 0.054454
3349.6544 | 0.012076 | -0.0079236 | 0.012076 | 0.003835 | 0.0014433 | 0.011443 | -0.0015567 | 0.009068 | 0.021443 |  | 0.015708 | 0.038998 | 0.038998 | 0.028651 | 0.038694 | 0.03878 | 0.029046 | 0.036946 | 0.045442 | 0.045442 |  | 0.054452
3548.1339 | 0.013164 | -0.0068356 | 0.013164 | 0.007255 | -0.00022825 | 0.012772 | 0.0037717 | 0.010978 | 0.019772 |  | 0.01892 | 0.038991 | 0.038991 | 0.028641 | 0.038687 | 0.038774 | 0.029036 | 0.036945 | 0.045442 | 0.045442 |  | 0.054451
3758.374 | 0.012714 | -0.0072856 | 0.012714 | 0.0073007 | 0.0010633 | 0.011063 | 0.0030633 | 0.0093479 | 0.011063 |  | 0.017853 | 0.038922 | 0.038922 | 0.028547 | 0.038622 | 0.038708 | 0.028938 | 0.036936 | 0.045435 | 0.045435 |  | 0.054445
3981.0717 | 0.0097154 | -0.010285 | 0.0097154 | 0.0086796 | -0.0092717 | 0.0097283 | -0.00027166 | 0.0087609 | 0.010728 | 0.037174 | 0.018391 | 0.038914 | 0.038914 | 0.028536 | 0.038538 | 0.038518 | 0.029119 | 0.036772 | 0.045301 | 0.045301 | 0.045301 | 0.054334
4216.965 | 0.015452 | -0.0045475 | 0.015452 | 0.01611 | -0.005293 | 0.016707 | 0.009707 | 0.0138 | 0.014707 |  | 0.024784 | 0.038631 | 0.038631 | 0.028148 | 0.048142 | 0.037632 | 0.027862 | 0.037311 | 0.0547 | 0.0547 |  | 0.0547
4466.8359 | 0.013358 | -0.006642 | 0.013358 | 0.013556 | -0.011121 | 0.016879 | 0.0058786 | 0.0088753 | 0.0088786 |  | 0.022207 | 0.038618 | 0.038618 | 0.028131 | 0.048133 | 0.03762 | 0.027843 | 0.03731 | 0.054699 | 0.054699 |  | 0.054699
4731.5126 | 0.016087 | -0.0039134 | 0.0060866 | 0.014399 | -0.010549 | 0.012451 | 0.0034513 | 0.0072006 | 0.0094513 |  | 0.023538 | 0.038569 | 0.038569 | 0.028064 | 0.048097 | 0.037575 | 0.027769 | 0.037306 | 0.054697 | 0.054697 |  | 0.063967
5011.8723 | 0.0087966 | -0.011203 | 0.0087966 | 0.016591 | -0.021639 | 0.013361 | 0.010361 | 0.007781 | 0.0083615 | 0.042257 | 0.025628 | 0.048824 | 0.03852 | 0.027996 | 0.059051 | 0.037011 | 0.028206 | 0.037433 | 0.054783 | 0.054783 | 0.054783 | 0.064041
5308.8444 | 0.014169 | -0.0058313 | 0.0041687 | 0.01855 | -0.020532 | 0.014468 | 0.010468 | 0.0087061 | 0.0094678 |  | 0.030911 | 0.048782 | 0.048782 | 0.027923 | 0.059801 | 0.036525 | 0.029765 | 0.037886 | 0.055094 | 0.055094 |  | 0.073725
5623.4133 | 0.0093135 | -0.010686 | -0.00068646 | 0.016942 | -0.023347 | 0.007653 | 0.009653 | 0.0058305 | 0.006653 |  | 0.031103 | 0.048753 | 0.048753 | 0.027873 | 0.059773 | 0.036503 | 0.029708 | 0.046207 | 0.055092 | 0.055092 |  | 0.083277
5956.6214 | 0.0077228 | -0.012277 | -0.0022772 | 0.012521 | -0.020005 | 0.003995 | 0.014995 | 0.0011617 | -4.9611e-06 |  | 0.03333 | 0.048663 | 0.048663 | 0.027714 | 0.059687 | 0.036433 | 0.029529 | 0.046198 | 0.055084 | 0.055084 |  | 0.092921
6309.5734 | 0.017882 | -0.012118 | -0.0021183 | 0.018638 | -0.025411 | 0.0095886 | 0.027589 | 0.0049013 | 0.0045886 | 0.057864 | 0.046903 | 0.048732 | 0.048732 | 0.027835 | 0.060085 | 0.036752 | 0.029311 | 0.05545 | 0.05545 | 0.05545 | 0.064612 | 0.12237
6683.4392 | 0.014863 | -0.015137 | 0.0048626 | 0.025656 | -0.024816 | 0.020184 | 0.014184 | 0.0072789 | 0.015184 |  | 0.023334 | 0.048725 | 0.048725 | 0.027822 | 0.058903 | 0.034794 | 0.037665 | 0.057016 | 0.065961 | 0.057016 |  | 0.12309
7079.4578 | 0.012347 | -0.0076534 | 0.0023466 | 0.027076 | -0.027649 | 0.021351 | 0.018351 | 0.0067824 | 0.012351 |  | 0.025621 | 0.048591 | 0.048591 | 0.027588 | 0.058757 | 0.034669 | 0.037459 | 0.065937 | 0.065937 | 0.056988 |  | 0.13284
7498.9421 | 0.010842 | -0.019158 | 0.00084154 | 0.018813 | -0.028551 | 0.012449 | 0.012449 | -0.0020693 | 0.0014487 |  | 0.019624 | 0.058803 | 0.048557 | 0.027528 | 0.060895 | 0.038812 | 0.039677 | 0.067451 | 0.067451 | 0.058734 |  | 0.1336
7943.2823 | 0.0083064 | -0.021694 | 0.0083064 | 0.021983 | -0.025491 | 0.013509 | 0.018509 | 0.0030757 | 0.014509 | 0.012445 | 0.023418 | 0.058677 | 0.048405 | 0.027258 | 0.061272 | 0.03882 | 0.038949 | 0.067402 | 0.067402 | 0.058678 | 0.07644 | 0.13358
8413.9514 | -0.0027903 | -0.02279 | 0.01721 | 0.013649 | -0.016351 | 0.017649 | 0.019649 | 0.0013331 | 0.023649 |  | 0.017795 | 0.058376 | 0.099034 | 0.026604 | 0.073635 | 0.080784 | 0.060077 | 0.077911 | 0.10281 | 0.094181 |  | 0.13918
8912.5094 | -0.016468 | -0.006468 | 0.023532 | 0.016874 | -0.0031262 | 0.019874 | 0.028874 | 0.0077863 | 0.036874 |  | 0.013034 | 0.07825 | 0.098605 | 0.036374 | 0.072643 | 0.081619 | 0.061346 | 0.086602 | 0.10344 | 0.094867 |  | 0.13964
9440.6088 | -0.019077 | -0.029077 | 0.010923 | -0.003138 | -0.013138 | 0.004862 | 0.047862 | 0.0064035 | 0.076862 |  | 0.063052 | 0.078108 | 0.098493 | 0.046914 | 0.07594 | 0.083766 | 0.063586 | 0.088397 | 0.10495 | 0.096509 |  | 0.14076
10000 | -0.031569 | -0.041569 | -0.0015686 | -0.018191 | -0.029614 | -0.016614 | 0.080386 | -0.0092418 | 0.10039 | 0.16512 | 0.092843 | 0.07712 | 0.097711 | 0.056103 | 0.081769 | 0.084557 | 0.066579 | 0.10513 | 0.10513 | 0.09671 | 0.11381 | 0.1409

## 2652765 Phase
Results highlighted in gray have been excluded from the calculation of the estimated reference sensitivity phase.
4160. 2652765 | Sensitivity Phase (4160 2652765) degrees |  |  |  | Expanded Uncertainties (degrees) |  |  |  | Estimated reference sensitivity phase | Std Unc (ERSP) | Phase difference relative to estimated reference sensitivity phase (degrees)
Frequency (Hz) | NIM | NMIJ | KRISS | CMS/ITRI | NIM | NMIJ | KRISS | CMS/ITRI | (degrees) | (degrees) | NIM | NMIJ | KRISS | CMS/ITRI
1.9953 | 176.3269 | 177.36 | 176.531 | 175.5648 | 1 | 2.5 | 2.5 | 1.3 | 176.1946 | 0.36162 | 0.1323 | 1.1654 | 0.33645 | -0.6297
2.5119 | 176.3712 | 177.2 | 176.437 | 175.704 | 0.9 | 2.43 | 2 | 1.1 | 176.2121 | 0.31748 | 0.15906 | 0.9879 | 0.2249 | -0.50807
3.1623 | 176.4582 | 176.53 | 176.504 | 175.9194 | 0.8 | 2.22 | 1.5 | 0.9 | 176.2752 | 0.26941 | 0.18298 | 0.25478 | 0.22878 | -0.35582
3.9811 | 176.6084 | 176.77 | 176.672 | 176.1862 | 0.7 | 1.44 | 1.5 | 0.8 | 176.4862 | 0.23492 | 0.12218 | 0.28377 | 0.18577 | -0.30001
5.0119 | 176.8252 | 177.35 | 176.886 | 176.4785 | 0.6 | 1.18 | 1.2 | 0.7 | 176.7789 | 0.2003 | 0.046303 | 0.57108 | 0.10708 | -0.30043
6.3096 | 177.1457 | 177.59 | 177.235 | 176.791 | 0.6 | 1.01 | 1.2 | 0.6 | 177.0783 | 0.18595 | 0.067441 | 0.51173 | 0.15673 | -0.28722
7.9433 | 177.3736 | 177.7 | 177.36 | 177.1107 | 0.5 | 0.76 | 1 | 0.6 | 177.3548 | 0.16214 | 0.018771 | 0.34521 | 0.0052066 | -0.24406
10 | 177.655 | 177.97 | 177.616 | 177.4178 | 0.5 | 0.6 | 0.8 | 0.5 | 177.6441 | 0.14233 | 0.010927 | 0.32593 | -0.028067 | -0.22631
12.5893 | 177.8633 | 178.16 | 177.856 | 177.6979 | 0.4 | 0.46 | 0.6 | 0.5 | 177.9039 | 0.11867 | -0.040595 | 0.25614 | -0.047858 | -0.20596
15.8489 | 178.0732 | 178.25 | 178.058 | 177.941 | 0.35 | 0.37 | 0.5 | 0.5 | 178.1031 | 0.10321 | -0.029903 | 0.14689 | -0.045114 | -0.1621
19.9526 | 178.2578 | 178.35 | 178.254 | 178.1512 | 0.3 | 0.29 | 0.5 | 0.5 | 178.2789 | 0.0898 | -0.021101 | 0.071065 | -0.024935 | -0.1277
21.1349 | 178.2997 | 178.38 | 178.294 | 178.1979 | 0.3 | 0.29 | 0.5 | 0.5 | 178.3166 | 0.0898 | -0.016926 | 0.063371 | -0.022629 | -0.11873
22.3872 | 178.339 | 178.43 | 178.336 | 178.2425 | 0.3 | 0.29 | 0.5 | 0.5 | 178.3611 | 0.0898 | -0.022071 | 0.068938 | -0.025062 | -0.11856
23.7137 | 178.3759 | 178.45 | 178.367 | 178.2851 | 0.3 | 0.29 | 0.5 | 0.5 | 178.3915 | 0.0898 | -0.015567 | 0.058542 | -0.024458 | -0.10633
25.1189 | 178.4106 | 178.48 | 178.408 | 178.3251 | 0.25 | 0.25 | 0.4 | 0.5 | 178.4284 | 0.076923 | -0.01779 | 0.05159 | -0.02041 | -0.10331
26.6073 | 178.4434 | 178.51 | 178.439 | 178.3619 | 0.25 | 0.25 | 0.4 | 0.5 | 178.4602 | 0.076923 | -0.016859 | 0.049751 | -0.021249 | -0.098366
28.1838 | 178.4744 | 178.54 | 178.47 | 178.3991 | 0.25 | 0.25 | 0.4 | 0.5 | 178.4915 | 0.076923 | -0.017067 | 0.048534 | -0.021466 | -0.092327
29.8538 | 178.5038 | 178.56 | 178.509 | 178.4312 | 0.25 | 0.25 | 0.4 | 0.5 | 178.519 | 0.076923 | -0.015181 | 0.041026 | -0.0099745 | -0.087792
31.6228 | 178.5317 | 178.6 | 178.525 | 178.4633 | 0.2 | 0.19 | 0.4 | 0.5 | 178.5567 | 0.063019 | -0.025074 | 0.043276 | -0.031724 | -0.093413
33.4965 | 178.558 | 178.61 | 178.553 | 178.4929 | 0.2 | 0.19 | 0.4 | 0.5 | 178.5762 | 0.063019 | -0.018271 | 0.03377 | -0.02323 | -0.083374
35.4813 | 178.5826 | 178.64 | 178.579 | 178.5183 | 0.2 | 0.19 | 0.4 | 0.5 | 178.6034 | 0.063019 | -0.020812 | 0.036586 | -0.024414 | -0.085138
37.5837 | 178.6054 | 178.66 | 178.599 | 178.5437 | 0.2 | 0.19 | 0.4 | 0.5 | 178.6249 | 0.063019 | -0.019489 | 0.035143 | -0.025857 | -0.081171
39.8107 | 178.626 | 178.68 | 178.62 | 178.5667 | 0.2 | 0.16 | 0.4 | 0.5 | 178.6507 | 0.058001 | -0.024712 | 0.029321 | -0.030679 | -0.083955
42.1697 | 178.6441 | 178.69 | 178.635 | 178.5852 | 0.2 | 0.16 | 0.4 | 0.5 | 178.6643 | 0.058001 | -0.020208 | 0.025714 | -0.029286 | -0.079048
44.6684 | 178.6594 | 178.7 | 178.652 | 178.6027 | 0.2 | 0.16 | 0.4 | 0.5 | 178.6771 | 0.058001 | -0.017647 | 0.022922 | -0.025078 | -0.074368
47.3151 | 178.6719 | 178.71 | 178.661 | 178.6193 | 0.2 | 0.16 | 0.4 | 0.5 | 178.6882 | 0.058001 | -0.016286 | 0.021823 | -0.027177 | -0.068865
50.1187 | 178.6816 | 178.72 | 178.675 | 178.6326 | 0.2 | 0.14 | 0.4 | 0.5 | 178.7015 | 0.053832 | -0.01998 | 0.018451 | -0.026549 | -0.06899
53.0884 | 178.6889 | 178.73 | 178.686 | 178.6449 | 0.15 | 0.14 | 0.4 | 0.5 | 178.7069 | 0.04863 | -0.018003 | 0.023106 | -0.020894 | -0.062039
56.2341 | 178.6946 | 178.73 | 178.683 | 178.6527 | 0.15 | 0.14 | 0.4 | 0.5 | 178.7094 | 0.04863 | -0.014811 | 0.020582 | -0.026418 | -0.056675
59.5662 | 178.6996 | 178.74 | 178.694 | 178.659 | 0.15 | 0.14 | 0.4 | 0.5 | 178.7173 | 0.04863 | -0.017605 | 0.022749 | -0.023251 | -0.058218
63.0957 | 178.7047 | 178.74 | 178.702 | 178.6645 | 0.15 | 0.12 | 0.4 | 0.5 | 178.723 | 0.044876 | -0.018286 | 0.016966 | -0.021034 | -0.058516
66.8344 | 178.7098 | 178.74 | 178.699 | 178.6659 | 0.15 | 0.12 | 0.4 | 0.5 | 178.7247 | 0.044876 | -0.014938 | 0.015265 | -0.025735 | -0.058826
70.7946 | 178.7129 | 178.73 | 178.695 | 178.6658 | 0.15 | 0.12 | 0.4 | 0.5 | 178.72 | 0.044876 | -0.0071512 | 0.0099577 | -0.025042 | -0.054289
74.9894 | 178.7094 | 178.73 | 178.691 | 178.662 | 0.15 | 0.12 | 0.4 | 0.5 | 178.7185 | 0.044876 | -0.0090658 | 0.011528 | -0.027472 | -0.056482
79.4328 | 178.6919 | 178.72 | 178.684 | 178.6569 | 0.15 | 0.1 | 0.4 | 0.5 | 178.7088 | 0.040201 | -0.016959 | 0.011167 | -0.024833 | -0.051939
84.1395 | 178.6526 | 178.71 | 178.672 | 178.6483 | 0.15 | 0.1 | 0.4 | 0.5 | 178.6904 | 0.040201 | -0.03777 | 0.01962 | -0.01838 | -0.042105
89.1251 | 178.5932 | 178.72 | 178.662 | 178.6378 | 0.15 | 0.1 | 0.4 | 0.5 | 178.6791 | 0.040201 | -0.085904 | 0.0409 | -0.0171 | -0.041277
94.4061 | 178.5489 | 178.68 | 178.648 | 178.6241 | 0.15 | 0.1 | 0.4 | 0.5 | 178.6396 | 0.040201 | -0.090717 | 0.040413 | 0.0084133 | -0.015517
100 | 178.6431 | 178.65 | 178.63 | 178.6078 | 0.15 | 0.09 | 0.4 | 0.5 | 178.6466 | 0.037461 | -0.0035471 | 0.0033767 | -0.016623 | -0.038832
105.9254 | 178.6224 | 178.64 | 178.611 | 178.5886 | 0.15 | 0.09 | 0.4 | 0.5 | 178.6334 | 0.037461 | -0.011073 | 0.0065749 | -0.022425 | -0.04486
112.2018 | 178.5985 | 178.62 | 178.588 | 178.5659 | 0.15 | 0.09 | 0.4 | 0.5 | 178.6123 | 0.037461 | -0.013776 | 0.0076929 | -0.024307 | -0.046386
118.8502 | 178.5716 | 178.58 | 178.56 | 178.5416 | 0.15 | 0.09 | 0.4 | 0.5 | 178.5763 | 0.037461 | -0.004748 | 0.0036623 | -0.016338 | -0.034751
125.8925 | 178.5415 | 178.54 | 178.532 | 178.5121 | 0.15 | 0.07 | 0.4 | 0.5 | 178.5396 | 0.031082 | 0.0018699 | 0.00036634 | -0.0076337 | -0.02754
133.3521 | 178.5082 | 178.53 | 178.498 | 178.4799 | 0.15 | 0.07 | 0.4 | 0.5 | 178.5247 | 0.031082 | -0.016483 | 0.0052863 | -0.026714 | -0.044825
141.2538 | 178.4717 | 178.5 | 178.462 | 178.4443 | 0.15 | 0.07 | 0.4 | 0.5 | 178.4934 | 0.031082 | -0.021658 | 0.0066396 | -0.03136 | -0.049107
149.6236 | 178.4318 | 178.44 | 178.423 | 178.4048 | 0.15 | 0.07 | 0.4 | 0.5 | 178.4376 | 0.031082 | -0.0058332 | 0.0023615 | -0.014639 | -0.032798
158.4893 | 178.3884 | 178.39 | 178.379 | 178.3612 | 0.15 | 0.07 | 0.3 | 0.5 | 178.3888 | 0.030794 | -0.00045117 | 0.0011741 | -0.0098259 | -0.027598
167.8804 | 178.3412 | 178.34 | 178.33 | 178.3133 | 0.15 | 0.07 | 0.3 | 0.5 | 178.3394 | 0.030794 | 0.0018079 | 0.00062847 | -0.0093715 | -0.026121
177.8279 | 178.2899 | 178.3 | 178.279 | 178.2619 | 0.15 | 0.07 | 0.3 | 0.5 | 178.2968 | 0.030794 | -0.0069183 | 0.0031624 | -0.017838 | -0.03493
188.3649 | 178.2342 | 178.24 | 178.22 | 178.2045 | 0.15 | 0.07 | 0.3 | 0.5 | 178.2376 | 0.030794 | -0.0034188 | 0.0023542 | -0.017646 | -0.033108
199.5262 | 178.1737 | 178.17 | 178.159 | 178.144 | 0.15 | 0.06 | 0.3 | 0.5 | 178.1698 | 0.027223 | 0.0038678 | 0.00018558 | -0.010814 | -0.025823
211.3489 | 178.1078 | 178.11 | 178.093 | 178.0773 | 0.15 | 0.06 | 0.3 | 0.5 | 178.1088 | 0.027223 | -0.00092916 | 0.0012319 | -0.015768 | -0.031427
223.8721 | 178.0363 | 178.03 | 178.021 | 178.0071 | 0.15 | 0.06 | 0.3 | 0.5 | 178.0303 | 0.027223 | 0.0060087 | -0.00025706 | -0.0092571 | -0.023198
237.1374 | 177.9586 | 177.95 | 177.945 | 177.9295 | 0.15 | 0.06 | 0.3 | 0.5 | 177.9507 | 0.027223 | 0.0078749 | -0.00072475 | -0.0057248 | -0.021266
251.1886 | 177.8746 | 177.87 | 177.861 | 177.8466 | 0.15 | 0.05 | 0.3 | 0.5 | 177.87 | 0.023324 | 0.0045771 | -2.4083e-05 | -0.0090241 | -0.023381
266.0725 | 177.7842 | 177.78 | 177.772 | 177.758 | 0.15 | 0.05 | 0.3 | 0.5 | 177.78 | 0.023324 | 0.0041787 | -2.1169e-05 | -0.0080212 | -0.022032
281.8383 | 177.6875 | 177.68 | 177.679 | 177.6634 | 0.15 | 0.05 | 0.3 | 0.5 | 177.6806 | 0.023324 | 0.0069435 | -0.00055696 | -0.001557 | -0.017129
298.5383 | 177.5847 | 177.58 | 177.575 | 177.5615 | 0.15 | 0.05 | 0.3 | 0.5 | 177.5802 | 0.023324 | 0.0045413 | -0.00017414 | -0.0051741 | -0.018672
316.2278 | 177.476 | 177.47 | 177.468 | 177.4524 | 0.15 | 0.05 | 0.3 | 0.5 | 177.4704 | 0.023324 | 0.0056032 | -0.00037693 | -0.0023769 | -0.017963
334.9654 | 177.361 | 177.35 | 177.35 | 177.3346 | 0.15 | 0.05 | 0.3 | 0.5 | 177.3509 | 0.023324 | 0.010083 | -0.00093104 | -0.00093104 | -0.016345
354.8134 | 177.2386 | 177.22 | 177.225 | 177.2102 | 0.15 | 0.05 | 0.3 | 0.5 | 177.2218 | 0.023324 | 0.0168 | -0.0018382 | 0.0031618 | -0.011639
375.8374 | 177.1063 | 177.09 | 177.093 | 177.0772 | 0.15 | 0.05 | 0.3 | 0.5 | 177.0915 | 0.023324 | 0.014739 | -0.0015354 | 0.0014646 | -0.014289
398.1072 | 176.9598 | 176.95 | 176.952 | 176.9358 | 0.15 | 0.05 | 0.3 | 0.5 | 176.9509 | 0.023324 | 0.0089488 | -0.00087467 | 0.0011253 | -0.01509
421.6965 | 176.7949 | 176.79 | 176.8 | 176.7849 | 0.15 | 0.05 | 0.3 | 0.5 | 176.7907 | 0.023324 | 0.0042168 | -0.0006697 | 0.0093303 | -0.0058003
446.6836 | 176.6113 | 176.63 | 176.639 | 176.6227 | 0.15 | 0.05 | 0.3 | 0.5 | 176.6283 | 0.023324 | -0.017015 | 0.0016507 | 0.010651 | -0.0056045
473.1513 | 176.4252 | 176.46 | 176.467 | 176.4509 | 0.15 | 0.05 | 0.3 | 0.5 | 176.4567 | 0.023324 | -0.031488 | 0.0032719 | 0.010272 | -0.0058644
501.1872 | 176.2957 | 176.27 | 176.287 | 176.2688 | 0.15 | 0.06 | 0.3 | 0.5 | 176.2739 | 0.027223 | 0.021769 | -0.0039319 | 0.013068 | -0.0051361
530.8844 | 176.1016 | 176.08 | 176.091 | 176.0737 | 0.15 | 0.06 | 0.3 | 0.5 | 176.0831 | 0.027223 | 0.018453 | -0.0031312 | 0.0078688 | -0.0094496
562.3413 | 175.8949 | 175.87 | 175.884 | 175.8665 | 0.15 | 0.06 | 0.3 | 0.5 | 175.8737 | 0.027223 | 0.021171 | -0.0036955 | 0.010304 | -0.0072282
595.6621 | 175.6749 | 175.65 | 175.666 | 175.6462 | 0.15 | 0.06 | 0.3 | 0.5 | 175.6538 | 0.027223 | 0.021109 | -0.0037588 | 0.012241 | -0.0075186
630.9573 | 175.4408 | 175.41 | 175.43 | 175.4119 | 0.15 | 0.06 | 0.3 | 0.5 | 175.4147 | 0.027223 | 0.026057 | -0.0047388 | 0.015261 | -0.0028369
668.3439 | 175.1918 | 175.16 | 175.183 | 175.1626 | 0.15 | 0.07 | 0.3 | 0.5 | 175.1664 | 0.030794 | 0.025441 | -0.0063719 | 0.016628 | -0.0037712
707.9458 | 174.927 | 174.89 | 174.918 | 174.8985 | 0.15 | 0.07 | 0.3 | 0.5 | 174.8976 | 0.030794 | 0.029476 | -0.0075506 | 0.020449 | 0.00091437
749.8942 | 174.6455 | 174.61 | 174.638 | 174.6153 | 0.15 | 0.07 | 0.3 | 0.5 | 174.6172 | 0.030794 | 0.028244 | -0.0072427 | 0.020757 | -0.0019504
794.3282 | 174.3462 | 174.3 | 174.338 | 174.315 | 0.15 | 0.07 | 0.3 | 0.5 | 174.3096 | 0.030794 | 0.036557 | -0.009613 | 0.028387 | 0.00542
841.3951 | 174.028 | 173.98 | 174.02 | 173.9968 | 0.15 | 0.08 | 0.3 | 0.5 | 173.9923 | 0.034036 | 0.035724 | -0.012252 | 0.027748 | 0.0045804
891.2509 | 173.6897 | 173.64 | 173.681 | 173.6579 | 0.15 | 0.08 | 0.3 | 0.5 | 173.6527 | 0.034036 | 0.03703 | -0.01268 | 0.02832 | 0.0051817
944.0609 | 173.3301 | 173.28 | 173.323 | 173.2982 | 0.15 | 0.08 | 0.3 | 0.5 | 173.2929 | 0.034036 | 0.037209 | -0.012863 | 0.030137 | 0.0053033
1000 | 172.9476 | 172.89 | 172.938 | 172.9118 | 0.15 | 0.08 | 0.3 | 0.5 | 172.9047 | 0.034036 | 0.042888 | -0.014746 | 0.033254 | 0.0070922
1059.2537 | 172.5408 | 172.48 | 172.53 | 172.5031 | 0.15 | 0.09 | 0.3 | 0.5 | 172.4983 | 0.03696 | 0.042518 | -0.018312 | 0.031688 | 0.0047546
1122.0185 | 172.1079 | 172.04 | 172.098 | 172.0679 | 0.15 | 0.09 | 0.3 | 0.5 | 172.0606 | 0.03696 | 0.047308 | -0.020629 | 0.037371 | 0.0072511
1188.5022 | 171.6471 | 171.58 | 171.636 | 171.6045 | 0.15 | 0.09 | 0.3 | 0.5 | 171.6002 | 0.03696 | 0.046835 | -0.02022 | 0.03578 | 0.0043026
1258.9254 | 171.1561 | 171.09 | 171.143 | 171.1105 | 0.15 | 0.09 | 0.3 | 0.5 | 171.1097 | 0.03696 | 0.046379 | -0.019717 | 0.033283 | 0.00075708
1333.5214 | 170.6328 | 170.55 | 170.62 | 170.5854 | 0.15 | 0.11 | 0.3 | 0.5 | 170.5823 | 0.04193 | 0.050438 | -0.032338 | 0.037662 | 0.003099
1412.5375 | 170.0746 | 169.99 | 170.063 | 170.0264 | 0.15 | 0.11 | 0.3 | 0.5 | 170.0232 | 0.04193 | 0.05144 | -0.033174 | 0.039826 | 0.003218
1496.2357 | 169.4789 | 169.39 | 169.465 | 169.4269 | 0.15 | 0.11 | 0.3 | 0.5 | 169.4247 | 0.04193 | 0.05424 | -0.034693 | 0.040307 | 0.0021773
1584.8932 | 168.8429 | 168.75 | 168.828 | 168.7894 | 0.15 | 0.11 | 0.3 | 0.5 | 168.7862 | 0.04193 | 0.056648 | -0.036232 | 0.041768 | 0.003154
1678.804 | 168.1634 | 168.07 | 168.15 | 168.1091 | 0.15 | 0.13 | 0.3 | 0.5 | 168.1138 | 0.045888 | 0.049649 | -0.04378 | 0.03622 | -0.0046366
1778.2794 | 167.4374 | 167.33 | 167.424 | 167.3823 | 0.15 | 0.13 | 0.3 | 0.5 | 167.3808 | 0.045888 | 0.056624 | -0.050758 | 0.043242 | 0.0015789
1883.6491 | 166.6613 | 166.56 | 166.648 | 166.6031 | 0.15 | 0.13 | 0.3 | 0.5 | 166.6076 | 0.045888 | 0.053696 | -0.047613 | 0.040387 | -0.0044701
1995.2623 | 165.8314 | 165.72 | 165.817 | 165.7682 | 0.2 | 0.13 | 0.3 | 0.5 | 165.7608 | 0.05018 | 0.070537 | -0.040844 | 0.056156 | 0.0073538
2113.489 | 164.943 | 164.82 | 164.91 | 164.8745 | 0.2 | 0.13 | 0.3 | 0.5 | 164.8632 | 0.05018 | 0.079724 | -0.043226 | 0.046774 | 0.01124
2238.7211 | 163.9897 | 163.85 | 163.951 | 163.9138 | 0.2 | 0.13 | 0.3 | 0.5 | 163.8991 | 0.05018 | 0.090666 | -0.049057 | 0.051943 | 0.014749
2371.3737 | 162.9621 | 162.82 | 162.92 | 162.8816 | 0.2 | 0.13 | 0.3 | 0.5 | 162.8695 | 0.05018 | 0.092676 | -0.049466 | 0.050534 | 0.012153
2511.8864 | 161.8445 | 161.71 | 161.807 | 161.7704 | 0.2 | 0.13 | 0.3 | 0.5 | 161.7571 | 0.05018 | 0.087313 | -0.047146 | 0.049854 | 0.013234
2660.7251 | 160.6504 | 160.51 | 160.61 | 160.5682 | 0.2 | 0.14 | 0.3 | 0.5 | 160.5633 | 0.052376 | 0.087119 | -0.053258 | 0.046742 | 0.0049736
2818.3829 | 159.3567 | 159.21 | 159.322 | 159.274 | 0.2 | 0.14 | 0.3 | 0.5 | 159.2667 | 0.052376 | 0.089976 | -0.056703 | 0.055297 | 0.007305
2985.3826 | 157.9536 | 157.81 | 157.916 | 157.8668 | 0.2 | 0.14 | 0.3 | 0.5 | 157.8648 | 0.052376 | 0.088811 | -0.054816 | 0.051184 | 0.0019382
3162.2777 | 156.4298 | 156.28 | 156.39 | 156.3411 | 0.2 | 0.14 | 0.3 | 0.5 | 156.3372 | 0.052376 | 0.092602 | -0.057184 | 0.052816 | 0.0039053
3349.6544 | 154.7717 | 154.63 | 154.719 | 154.6821 | 0.2 | 0.15 | 0.3 | 0.5 | 154.686 | 0.054375 | 0.085631 | -0.056047 | 0.032953 | -0.0039823
3548.1339 | 152.9634 | 152.83 | 152.9 | 152.8741 | 0.2 | 0.15 | 0.3 | 0.5 | 152.8807 | 0.054375 | 0.082672 | -0.050725 | 0.019275 | -0.0066371
3758.374 | 150.9862 | 150.85 | 150.907 | 150.8913 | 0.2 | 0.15 | 0.3 | 0.5 | 150.8997 | 0.054375 | 0.086498 | -0.049717 | 0.0072827 | -0.0084271
3981.0717 | 148.8182 | 148.72 | 148.738 | 148.7257 | 0.25 | 0.15 | 0.3 | 0.5 | 148.7437 | 0.057522 | 0.074466 | -0.023748 | -0.0057479 | -0.018034
4216.965 | 146.434 | 146.36 | 146.347 | 146.3426 | 0.25 | 0.15 | 0.3 | 0.5 | 146.3728 | 0.057522 | 0.061158 | -0.012833 | -0.025833 | -0.030283
4466.8359 | 143.8045 | 143.78 | 143.728 | 143.7173 | 0.25 | 0.15 | 0.3 | 0.5 | 143.7742 | 0.057522 | 0.030296 | 0.0057727 | -0.046227 | -0.056915
4731.5126 | 140.8973 | 140.93 | 140.826 | 140.8084 | 0.25 | 0.15 | 0.3 | 0.5 | 140.9013 | 0.057522 | -0.0040477 | 0.028656 | -0.075344 | -0.092927
5011.8723 | 137.6769 | 137.78 | 137.597 | 137.5952 | 0.25 | 0.15 | 0.3 | 0.6 | 137.7235 | 0.057993 | -0.04668 | 0.056462 | -0.12654 | -0.12835
5308.8444 | 134.106 | 134.33 | 134.006 | 134.0282 | 0.25 | 0.27 | 0.3 | 0.6 | 134.146 | 0.075717 | -0.040035 | 0.18398 | -0.14002 | -0.11785
5623.4133 | 130.1477 | 130.52 | 130.067 | 130.0742 | 0.25 | 0.27 | 0.3 | 0.6 | 130.1109 | 0.091457 | 0.036838 | 0.40913 | -0.043869 | -0.036711
5956.6214 | 125.7685 | 126.25 | 125.66 | 125.7239 | 0.25 | 0.27 | 0.3 | 0.7 | 125.724 | 0.092605 | 0.044469 | 0.52599 | -0.064013 | -0.00012404
6309.5734 | 120.9428 | 121.52 | 120.842 | 120.9165 | 0.25 | 0.27 | 0.3 | 0.8 | 120.9023 | 0.093375 | 0.040511 | 0.61768 | -0.060324 | 0.01414
6683.4392 | 115.6597 | 116.41 | 115.541 | 115.4825 | 0.3 | 0.62 | 0.4 | 1 | 115.6097 | 0.11669 | 0.050061 | 0.80034 | -0.068656 | -0.12714
7079.4578 | 109.9308 | 110.84 | 109.82 | 109.7861 | 0.35 | 0.62 | 0.4 | 1.1 | 109.8775 | 0.12808 | 0.053279 | 0.9625 | -0.0575 | -0.091426
7498.9421 | 103.8004 | 104.88 | 103.685 | 103.6632 | 0.4 | 0.62 | 0.4 | 1.1 | 103.7378 | 0.13697 | 0.062624 | 1.1422 | -0.052762 | -0.074587
7943.2823 | 97.3567 | 98.55 | 97.24 | 97.2355 | 0.45 | 0.62 | 0.4 | 1.1 | 97.2876 | 0.14425 | 0.06902 | 1.2624 | -0.047646 | -0.052097
8413.9514 | 90.7418 | 91.98 | 90.635 | 90.6298 | 0.55 | 1.07 | 0.7 | 1.1 | 90.6915 | 0.20124 | 0.050283 | 1.2885 | -0.056476 | -0.061672
8912.5094 | 84.1576 | 85.4 | 84.007 | 84.1361 | 0.65 | 1.07 | 0.7 | 1.1 | 84.0955 | 0.21855 | 0.062085 | 1.3045 | -0.088475 | 0.040675
9440.6088 | 77.8601 | 76.57 | 77.724 | 78.1892 | 0.7 | 1.07 | 0.7 | 1.1 | 77.8589 | 0.22569 | 0.001157 | -1.2889 | -0.13492 | 0.33032
10000 | 72.1303 | 70.66 | 72.037 | 71.6052 | 0.72 | 1.07 | 0.7 | 1.1 | 72.0001 | 0.22831 | 0.13021 | -1.3401 | 0.036859 | -0.39494

## 2652765 DoE Phase
4160. 2652765 | Linking DoEs from CCAUV.A-K5 |  |  | 4160 2652765 DoE (degrees) |  |  |  | Linking U(DoE)s from CCAUV.A-K5 |  |  | U(DoE) (degrees)
Frequency (Hz) | NIM | NMIJ | KRISS | NIM | NMIJ | KRISS | CMS-ITRI | NIM | NMIJ | KRISS | NIM | NMIJ | KRISS | CMS-ITRI
1.9953 | 0.12068 | 0.48068 |  | 0.1006 | 1.1338 | 0.30475 | -0.6614 | 1.9012 | 2.7229 |  | 1.52 | 2.8379 | 2.9462 | 2.0298
2.5119 | 0.0041898 | 0.45419 |  | 0.086451 | 0.9153 | 0.1523 | -0.58068 | 1.8942 | 2.3297 |  | 1.4156 | 2.8066 | 2.4622 | 1.809
3.1623 | 0.20273 | 0.31273 |  | 0.23568 | 0.30748 | 0.28148 | -0.30313 | 1.7076 | 2.2444 |  | 1.3152 | 2.5791 | 1.9692 | 1.5613
3.9811 | -0.012347 | 0.25765 |  | 0.074267 | 0.23586 | 0.13786 | -0.34792 | 1.5277 | 1.7543 |  | 1.1722 | 1.6687 | 1.8993 | 1.4133
5.0119 | -0.014873 | 0.15513 |  | -0.068673 | 0.45611 | -0.0078921 | -0.41541 | 1.1072 | 1.6644 |  | 0.96384 | 1.2985 | 1.4838 | 1.1187
6.3096 | -0.15021 | -0.030212 |  | -0.21928 | 0.22502 | -0.12998 | -0.57394 | 0.92822 | 1.1747 |  | 0.80928 | 1.0433 | 1.3852 | 0.91586
7.9433 | -0.12801 | -0.018008 |  | -0.18239 | 0.14404 | -0.19596 | -0.44522 | 0.79655 | 0.98086 |  | 0.68905 | 0.82832 | 1.1864 | 0.87616
10 | -0.064805 | -0.0048051 |  | -0.16874 | 0.14627 | -0.20773 | -0.40597 | 0.6608 | 0.78374 |  | 0.62583 | 0.62306 | 0.95695 | 0.72509
12.5893 | -0.080102 | -0.0001022 |  | -0.17207 | 0.12467 | -0.17933 | -0.33743 | 0.52414 | 0.58577 |  | 0.51499 | 0.44253 | 0.70408 | 0.62107
15.8489 | -0.053307 | -0.023307 |  | -0.12725 | 0.049544 | -0.14246 | -0.25944 | 0.39127 | 0.58736 |  | 0.43283 | 0.37203 | 0.6044 | 0.6044
19.9526 | -0.031477 | -0.031477 |  | -0.079121 | 0.013044 | -0.082956 | -0.18572 | 0.37348 | 0.48723 |  | 0.39572 | 0.31034 | 0.57605 | 0.57605
21.1349 | -0.10436 | -0.034358 | 0.46564 | -0.11939 | -0.039092 | -0.12509 | -0.2212 | 0.37581 | 0.48902 | 0.28131 | 0.39867 | 0.31225 | 0.56493 | 0.57675
22.3872 | -0.091011 | -0.071011 | 0.42899 | -0.13098 | -0.039972 | -0.13397 | -0.22747 | 0.35566 | 0.48948 | 0.28212 | 0.39405 | 0.29999 | 0.55982 | 0.57126
23.7137 | -0.092555 | -0.012555 | 0.38745 | -0.10557 | -0.031461 | -0.11446 | -0.19633 | 0.32457 | 0.48964 | 0.28239 | 0.38567 | 0.27962 | 0.55112 | 0.56257
25.1189 | -0.099613 | -0.069613 | 0.43039 | -0.12482 | -0.055437 | -0.12744 | -0.21033 | 0.32494 | 0.48988 | 0.28281 | 0.36721 | 0.25636 | 0.45086 | 0.5526
26.6073 | -0.090545 | -0.10055 | 0.39945 | -0.12676 | -0.060146 | -0.13115 | -0.20826 | 0.30413 | 0.48999 | 0.28301 | 0.36063 | 0.24067 | 0.44334 | 0.54636
28.1838 | -0.079244 | -0.019244 | 0.38076 | -0.095384 | -0.029784 | -0.099784 | -0.17064 | 0.29371 | 0.49007 | 0.28314 | 0.35725 | 0.23247 | 0.43958 | 0.54323
29.8538 | -0.081785 | -0.051785 | 0.34821 | -0.10156 | -0.045352 | -0.096352 | -0.17417 | 0.2941 | 0.4903 | 0.28353 | 0.35776 | 0.23265 | 0.43995 | 0.54328
31.6228 | -0.071152 | 0.018848 | 0.31885 | -0.074678 | -0.0063285 | -0.081329 | -0.14302 | 0.28439 | 0.38843 | 0.28439 | 0.2739 | 0.24621 | 0.45004 | 0.55073
33.4965 | -0.071151 | -0.0011506 | 0.29885 | -0.078842 | -0.0268 | -0.0838 | -0.14394 | 0.23461 | 0.39057 | 0.2873 | 0.25749 | 0.21814 | 0.43994 | 0.54069
35.4813 | -0.061411 | -0.021411 | 0.27859 | -0.080992 | -0.023594 | -0.084594 | -0.14532 | 0.22409 | 0.39066 | 0.28743 | 0.25333 | 0.21146 | 0.43731 | 0.53847
37.5837 | -0.058085 | -0.048085 | 0.35191 | -0.084317 | -0.029685 | -0.090685 | -0.146 | 0.21357 | 0.39079 | 0.2876 | 0.2491 | 0.20463 | 0.43473 | 0.53627
39.8107 | -0.069055 | 0.020945 | 0.32094 | -0.081101 | -0.027067 | -0.087067 | -0.14034 | 0.20397 | 0.39141 | 0.28845 | 0.25757 | 0.18149 | 0.43218 | 0.53353
42.1697 | -0.06004 | 0.0099602 | 0.20996 | -0.028078 | 0.017844 | -0.037156 | -0.086918 | 0.18279 | 0.39168 | 0.28881 | 0.23275 | 0.15498 | 0.4232 | 0.52372
44.6684 | -0.054422 | -0.0044225 | 0.19558 | -0.02966 | 0.01091 | -0.03709 | -0.08638 | 0.17213 | 0.39183 | 0.28901 | 0.23015 | 0.1494 | 0.41913 | 0.5223
47.3151 | -0.043572 | -0.013572 | 0.18643 | -0.023354 | 0.014755 | -0.034245 | -0.075933 | 0.17219 | 0.39185 | 0.28905 | 0.23022 | 0.14944 | 0.41917 | 0.52231
50.1187 | -0.036237 | 0.083763 | 0.18376 | -0.010568 | 0.027864 | -0.017136 | -0.059577 | 0.16149 | 0.39202 | 0.28927 | 0.23524 | 0.13253 | 0.41903 | 0.52037
53.0884 | -0.050777 | -0.050777 | 0.14922 | -0.044682 | -0.0035725 | -0.047573 | -0.088718 | 0.15235 | 0.39283 | 0.29037 | 0.19361 | 0.12778 | 0.41604 | 0.51479
56.2341 | -0.065779 | 0.034221 | 0.13422 | -0.032358 | 0.003035 | -0.043965 | -0.074221 | 0.1435 | 0.39369 | 0.29154 | 0.18288 | 0.11635 | 0.42263 | 0.51231
59.5662 | -0.059114 | 0.030886 | 0.13089 | -0.031624 | 0.0087294 | -0.037271 | -0.072237 | 0.14357 | 0.39371 | 0.29157 | 0.18296 | 0.11637 | 0.42268 | 0.51232
63.0957 | -0.063705 | 0.016295 | 0.11629 | -0.039631 | -0.0043792 | -0.042379 | -0.079862 | 0.14551 | 0.39443 | 0.29253 | 0.16859 | 0.12921 | 0.43353 | 0.51759
66.8344 | -0.062945 | 0.017055 | 0.11706 | -0.039392 | -0.0091889 | -0.050189 | -0.08328 | 0.13481 | 0.39456 | 0.2927 | 0.16505 | 0.12382 | 0.42873 | 0.51638
70.7946 | -0.060883 | 0.019117 | 0.11912 | -0.037176 | -0.020067 | -0.055067 | -0.084314 | 0.11525 | 0.39545 | 0.29391 | 0.15931 | 0.11398 | 0.42051 | 0.51429
74.9894 | -0.057789 | 0.022211 | 0.12221 | -0.039935 | -0.019341 | -0.058341 | -0.087351 | 0.10446 | 0.39562 | 0.29413 | 0.15545 | 0.10807 | 0.41567 | 0.51313
79.4328 | -0.048931 | 0.031069 | 0.13107 | -0.042271 | -0.014144 | -0.050144 | -0.07725 | 0.093628 | 0.39581 | 0.29439 | 0.15884 | 0.090229 | 0.41358 | 0.51093
84.1395 | -0.055145 | 0.034855 | 0.13485 | -0.067298 | -0.0099084 | -0.047908 | -0.071633 | 0.094911 | 0.39612 | 0.2948 | 0.1599 | 0.091184 | 0.41436 | 0.51108
89.1251 | -0.041551 | -0.051551 | 0.14845 | -0.10807 | 0.018738 | -0.039262 | -0.063439 | 0.095126 | 0.39617 | 0.29487 | 0.16008 | 0.091345 | 0.41449 | 0.51111
94.4061 | -0.046945 | -0.036945 | 0.063055 | -0.12713 | 0.0040034 | -0.027997 | -0.051927 | 0.087471 | 0.39705 | 0.29606 | 0.15879 | 0.086621 | 0.41165 | 0.51033
100 | -0.046418 | -0.016418 | 0.083582 | -0.038654 | -0.03173 | -0.05173 | -0.073939 | 0.076796 | 0.29631 | 0.29631 | 0.15699 | 0.074092 | 0.40768 | 0.50839
105.9254 | -0.047461 | 0.0025389 | 0.10254 | -0.04431 | -0.026662 | -0.055662 | -0.078097 | 0.079342 | 0.29698 | 0.29698 | 0.15912 | 0.075873 | 0.40892 | 0.50861
112.2018 | -0.036067 | 0.033933 | 0.13393 | -0.035968 | -0.014498 | -0.046498 | -0.068577 | 0.068527 | 0.29715 | 0.29715 | 0.1556 | 0.068438 | 0.40491 | 0.50766
118.8502 | -0.03014 | 0.05986 | 0.15986 | -0.018255 | -0.0098443 | -0.029844 | -0.048258 | 0.068527 | 0.29715 | 0.29715 | 0.1556 | 0.068438 | 0.40491 | 0.50766
125.8925 | -0.029474 | -0.0094735 | 0.090526 | -0.01932 | -0.020824 | -0.028824 | -0.048729 | 0.068536 | 0.29715 | 0.29715 | 0.16258 | 0.06179 | 0.40788 | 0.50649
133.3521 | -0.029087 | 0.020913 | 0.12091 | -0.036454 | -0.014685 | -0.046685 | -0.064796 | 0.057667 | 0.29736 | 0.29736 | 0.15897 | 0.053329 | 0.4042 | 0.50559
141.2538 | -0.02845 | 0.06155 | 0.06155 | -0.042042 | -0.013744 | -0.051744 | -0.069491 | 0.057667 | 0.29736 | 0.29736 | 0.15897 | 0.053329 | 0.4042 | 0.50559
149.6236 | -0.032149 | -0.0021493 | 0.097851 | -0.032571 | -0.024377 | -0.041377 | -0.059537 | 0.049947 | 0.29815 | 0.29815 | 0.15807 | 0.046824 | 0.40193 | 0.50498
158.4893 | -0.026561 | 0.043439 | 0.043439 | -0.022252 | -0.020627 | -0.031627 | -0.0494 | 0.049947 | 0.29815 | 0.29815 | 0.15836 | 0.047453 | 0.30066 | 0.50493
167.8804 | -0.028351 | -0.0083512 | 0.091649 | -0.021761 | -0.02294 | -0.03294 | -0.049689 | 0.049947 | 0.29815 | 0.29815 | 0.15836 | 0.047453 | 0.30066 | 0.50493
177.8279 | -0.022709 | 0.047291 | 0.047291 | -0.026481 | -0.0164 | -0.0374 | -0.054492 | 0.039233 | 0.29839 | 0.19758 | 0.15472 | 0.039436 | 0.29762 | 0.50428
188.3649 | -0.024744 | 0.005256 | 0.10526 | -0.024462 | -0.018689 | -0.038689 | -0.054151 | 0.039233 | 0.29839 | 0.19758 | 0.15472 | 0.039436 | 0.29762 | 0.50428
199.5262 | -0.021023 | -0.031023 | -0.031023 | -0.017991 | -0.021673 | -0.032673 | -0.047682 | 0.041352 | 0.29868 | 0.19802 | 0.16017 | 0.037305 | 0.29991 | 0.50371
211.3489 | -0.02349 | 0.03651 | -0.06349 | -0.02427 | -0.022109 | -0.039109 | -0.054769 | 0.041352 | 0.29868 | 0.19802 | 0.16017 | 0.037305 | 0.29991 | 0.50371
223.8721 | -0.02049 | 0.0095104 | 0.0095104 | -0.01318 | -0.019446 | -0.028446 | -0.042387 | 0.041352 | 0.29868 | 0.19802 | 0.16017 | 0.037305 | 0.29991 | 0.50371
237.1374 | -0.0206 | -0.0106 | -0.0106 | -0.012454 | -0.021054 | -0.026054 | -0.041595 | 0.041352 | 0.29868 | 0.19802 | 0.16017 | 0.037305 | 0.29991 | 0.50371
251.1886 | -0.01477 | 0.07523 | -0.02477 | -0.0082928 | -0.012894 | -0.021894 | -0.036251 | 0.041577 | 0.29871 | 0.19806 | 0.16309 | 0.036545 | 0.30168 | 0.50314
266.0725 | -0.013209 | 0.066791 | -0.033209 | -0.0081936 | -0.012393 | -0.020393 | -0.034404 | 0.030703 | 0.2989 | 0.19835 | 0.15909 | 0.02674 | 0.29904 | 0.50256
281.8383 | -0.01353 | 0.06647 | -0.03353 | -0.0057064 | -0.013207 | -0.014207 | -0.029779 | 0.030703 | 0.2989 | 0.19835 | 0.15909 | 0.02674 | 0.29904 | 0.50256
298.5383 | -0.017216 | 0.072784 | -0.027216 | -0.011465 | -0.016181 | -0.021181 | -0.034679 | 0.033194 | 0.29917 | 0.19875 | 0.16188 | 0.028081 | 0.29952 | 0.50262
316.2278 | -0.015382 | 0.084618 | -0.015382 | -0.0082197 | -0.0142 | -0.0162 | -0.031786 | 0.033194 | 0.29917 | 0.19875 | 0.16187 | 0.028087 | 0.29952 | 0.50262
334.9654 | -0.014725 | 0.0052746 | 0.0052746 | -0.004099 | -0.015113 | -0.015113 | -0.030528 | 0.033194 | 0.29917 | 0.19875 | 0.16187 | 0.028087 | 0.29952 | 0.50262
354.8134 | -0.013931 | 0.036069 | -0.063931 | 0.0027677 | -0.015871 | -0.010871 | -0.025672 | 0.033194 | 0.29917 | 0.19875 | 0.16187 | 0.028087 | 0.29952 | 0.50262
375.8374 | -0.015824 | 0.074176 | -0.025824 | 0.00016319 | -0.016111 | -0.013111 | -0.028865 | 0.033194 | 0.29917 | 0.19875 | 0.16187 | 0.028087 | 0.29952 | 0.50262
398.1072 | -0.01194 | 0.01806 | 0.01806 | -0.0020941 | -0.011918 | -0.0099175 | -0.026133 | 0.033194 | 0.29917 | 0.19875 | 0.16188 | 0.028142 | 0.29932 | 0.50262
421.6965 | -0.012814 | 0.077186 | -0.022814 | -0.0072746 | -0.012161 | -0.0021611 | -0.017292 | 0.033194 | 0.29917 | 0.19875 | 0.16188 | 0.028142 | 0.29932 | 0.50262
446.6836 | -0.0059924 | 0.044008 | -0.055992 | -0.023056 | -0.0043909 | 0.0046091 | -0.011646 | 0.033194 | 0.29917 | 0.19875 | 0.16188 | 0.028142 | 0.29932 | 0.50262
473.1513 | -0.0080528 | 0.021947 | 0.021947 | -0.038562 | -0.0038023 | 0.0031977 | -0.012939 | 0.033194 | 0.29917 | 0.19875 | 0.16188 | 0.028142 | 0.29932 | 0.50262
501.1872 | -0.0080502 | 0.01195 | 0.01195 | 0.014344 | -0.011357 | 0.0056429 | -0.012561 | 0.033197 | 0.29917 | 0.19875 | 0.16276 | 0.026138 | 0.29728 | 0.50301
530.8844 | -0.0048666 | 0.015133 | 0.015133 | 0.013866 | -0.0077191 | 0.0032809 | -0.014037 | 0.022221 | 0.29932 | 0.19898 | 0.15752 | 0.017585 | 0.29546 | 0.50267
562.3413 | -0.0078831 | 0.032117 | 0.032117 | 0.013986 | -0.010881 | 0.003119 | -0.014414 | 0.022221 | 0.29932 | 0.19898 | 0.15752 | 0.017585 | 0.29546 | 0.50267
595.6621 | -0.0063006 | 0.063699 | -0.036301 | 0.01511 | -0.0097575 | 0.0062425 | -0.013517 | 0.022221 | 0.29932 | 0.19898 | 0.15752 | 0.017585 | 0.29546 | 0.50267
630.9573 | -0.0040599 | 0.0059401 | 0.0059401 | 0.022159 | -0.0086367 | 0.011363 | -0.0067348 | 0.022221 | 0.29932 | 0.19898 | 0.15753 | 0.017566 | 0.29547 | 0.50268
668.3439 | -0.0042159 | 0.065784 | -0.034216 | 0.021559 | -0.010254 | 0.012746 | -0.0076531 | 0.022221 | 0.29932 | 0.19898 | 0.1541 | 0.02322 | 0.29365 | 0.50331
707.9458 | -0.0067584 | 0.043242 | 0.043242 | 0.023712 | -0.013315 | 0.014685 | -0.0048496 | 0.022221 | 0.29932 | 0.19898 | 0.1541 | 0.02322 | 0.29365 | 0.50331
749.8942 | -0.0037757 | 0.036224 | 0.036224 | 0.025222 | -0.010264 | 0.017736 | -0.0049718 | 0.022221 | 0.29932 | 0.19898 | 0.1541 | 0.02322 | 0.29365 | 0.50331
794.3282 | -0.0017129 | 0.048287 | -0.051713 | 0.034927 | -0.011242 | 0.026758 | 0.0037906 | 0.022221 | 0.29932 | 0.19898 | 0.1541 | 0.023225 | 0.29365 | 0.50331
841.3951 | -0.0012899 | 0.07871 | -0.02129 | 0.035109 | -0.012867 | 0.027133 | 0.0039652 | 0.022221 | 0.29932 | 0.19898 | 0.15051 | 0.030776 | 0.29179 | 0.50395
891.2509 | -0.0025208 | 0.037479 | 0.037479 | 0.035392 | -0.014318 | 0.026682 | 0.003543 | 0.022221 | 0.29932 | 0.19898 | 0.15051 | 0.030776 | 0.29179 | 0.50395
944.0609 | 0.0015747 | 0.011575 | 0.011575 | 0.039079 | -0.010992 | 0.032008 | 0.0071742 | 0.022221 | 0.29932 | 0.19898 | 0.15051 | 0.030776 | 0.29179 | 0.50395
1000 | 0.0039563 | 0.013956 | 0.013956 | 0.047504 | -0.010129 | 0.037871 | 0.011709 | 0.043917 | 0.29905 | 0.19857 | 0.1646 | 0.036484 | 0.29524 | 0.50425
1059.2537 | 0.0038817 | 0.043882 | -0.056118 | 0.04598 | -0.014851 | 0.035149 | 0.0082163 | 0.043908 | 0.29905 | 0.19856 | 0.16072 | 0.042114 | 0.2931 | 0.50472
1122.0185 | 0.0001162 | 0.0001162 | 0.0001162 | 0.047586 | -0.020351 | 0.037649 | 0.0075293 | 0.043908 | 0.29905 | 0.19856 | 0.16072 | 0.042114 | 0.2931 | 0.50472
1188.5022 | 0.0047336 | 0.084734 | -0.015266 | 0.053243 | -0.013811 | 0.042189 | 0.010711 | 0.043908 | 0.29905 | 0.19856 | 0.16072 | 0.042114 | 0.2931 | 0.50472
1258.9254 | 0.0027664 | 0.0027664 | 0.0027664 | 0.049163 | -0.016933 | 0.036067 | 0.0035404 | 0.043908 | 0.29905 | 0.19856 | 0.1607 | 0.042149 | 0.2931 | 0.50471
1333.5214 | 0.0059759 | 0.055976 | -0.044024 | 0.056504 | -0.026272 | 0.043728 | 0.0091646 | 0.043908 | 0.29905 | 0.19856 | 0.15305 | 0.058126 | 0.28898 | 0.50561
1412.5375 | 0.0055709 | 0.045571 | 0.045571 | 0.059209 | -0.025404 | 0.047596 | 0.010987 | 0.043908 | 0.29905 | 0.19856 | 0.15305 | 0.058126 | 0.28898 | 0.50561
1496.2357 | 0.0053396 | 0.07534 | -0.02466 | 0.060695 | -0.028238 | 0.046762 | 0.0086325 | 0.041323 | 0.29868 | 0.19801 | 0.14975 | 0.059673 | 0.28882 | 0.5058
1584.8932 | 0.008559 | 0.048559 | -0.051441 | 0.064557 | -0.028322 | 0.049678 | 0.011064 | 0.052178 | 0.29853 | 0.19779 | 0.15371 | 0.064513 | 0.2917 | 0.50629
1678.804 | 0.0079447 | 0.067945 | -0.032055 | 0.058173 | -0.035256 | 0.044744 | 0.0038872 | 0.052178 | 0.29853 | 0.19779 | 0.14663 | 0.082676 | 0.28803 | 0.50711
1778.2794 | 0.0090662 | 0.039066 | 0.039066 | 0.067967 | -0.039414 | 0.054586 | 0.012922 | 0.05204 | 0.29851 | 0.19776 | 0.14647 | 0.082724 | 0.28801 | 0.50712
1883.6491 | 0.016046 | 0.066046 | -0.033954 | 0.069587 | -0.031723 | 0.056277 | 0.01142 | 0.05204 | 0.29851 | 0.19776 | 0.14647 | 0.082724 | 0.28801 | 0.50712
1995.2623 | 0.016618 | 0.056618 | -0.043382 | 0.086218 | -0.025162 | 0.071838 | 0.023035 | 0.073483 | 0.29833 | 0.19748 | 0.23232 | 0.055169 | 0.28605 | 0.50664
2113.489 | 0.017149 | 0.10715 | 0.0071495 | 0.10233 | -0.020624 | 0.069376 | 0.033842 | 0.073462 | 0.29832 | 0.19748 | 0.23228 | 0.05519 | 0.28605 | 0.50665
2238.7211 | 0.012528 | 0.13253 | 0.032528 | 0.11253 | -0.027194 | 0.073806 | 0.036612 | 0.07142 | 0.29783 | 0.19673 | 0.22887 | 0.057114 | 0.28579 | 0.50689
2371.3737 | 0.013915 | 0.043915 | 0.043915 | 0.11084 | -0.031306 | 0.068694 | 0.030313 | 0.07142 | 0.29783 | 0.19673 | 0.22887 | 0.057114 | 0.28579 | 0.50689
2511.8864 | 0.0086508 | 0.13865 | 0.038651 | 0.10648 | -0.027979 | 0.069021 | 0.032401 | 0.071425 | 0.29783 | 0.19673 | 0.22874 | 0.056492 | 0.28669 | 0.50682
2660.7251 | 0.015453 | 0.13545 | 0.035453 | 0.11348 | -0.026898 | 0.073102 | 0.031333 | 0.081961 | 0.29769 | 0.19651 | 0.22726 | 0.071851 | 0.28823 | 0.50793
2818.3829 | 0.023267 | 0.043267 | 0.043267 | 0.11721 | -0.029464 | 0.082536 | 0.034544 | 0.092581 | 0.29761 | 0.1964 | 0.22928 | 0.079138 | 0.29312 | 0.50889
2985.3826 | 0.024832 | 0.074832 | -0.025168 | 0.11373 | -0.029895 | 0.076105 | 0.026859 | 0.090908 | 0.29709 | 0.19561 | 0.22688 | 0.079671 | 0.29255 | 0.50901
3162.2777 | 0.021246 | 0.14125 | 0.041246 | 0.12675 | -0.023032 | 0.086968 | 0.038057 | 0.09077 | 0.29705 | 0.19555 | 0.22637 | 0.080172 | 0.29248 | 0.50887
3349.6544 | 0.019916 | 0.16992 | -0.030084 | 0.11594 | -0.025734 | 0.063266 | 0.026331 | 0.10138 | 0.29695 | 0.19539 | 0.22462 | 0.094268 | 0.29472 | 0.51012
3548.1339 | 0.021383 | 0.17138 | -0.028617 | 0.11585 | -0.01755 | 0.05245 | 0.026538 | 0.11182 | 0.29682 | 0.1952 | 0.22601 | 0.10068 | 0.29979 | 0.51118
3758.374 | 0.022573 | 0.17257 | -0.027427 | 0.1223 | -0.013912 | 0.043088 | 0.027378 | 0.1223 | 0.29674 | 0.19508 | 0.2273 | 0.10696 | 0.30493 | 0.51228
3981.0717 | 0.024801 | 0.1048 | 0.0048007 | 0.10736 | 0.0091468 | 0.027147 | 0.014861 | 0.13272 | 0.29667 | 0.19497 | 0.31182 | 0.079743 | 0.29948 | 0.51008
4216.965 | 0.026214 | 0.18621 | -0.013786 | 0.10388 | 0.029892 | 0.016892 | 0.012443 | 0.13269 | 0.29666 | 0.19495 | 0.31177 | 0.079752 | 0.29947 | 0.51008
4466.8359 | 0.014597 | 0.1646 | 0.064597 | 0.072052 | 0.047529 | -0.0044712 | -0.015159 | 0.13267 | 0.29665 | 0.19494 | 0.31173 | 0.079758 | 0.29946 | 0.51008
4731.5126 | 0.017525 | 0.17752 | -0.022475 | 0.026487 | 0.059191 | -0.044809 | -0.062392 | 0.13133 | 0.29605 | 0.19403 | 0.30935 | 0.080169 | 0.29876 | 0.51019
5011.8723 | 0.023047 | 0.17305 | -0.026953 | -0.013445 | 0.089696 | -0.093304 | -0.095119 | 0.12767 | 0.29445 | 0.19157 | 0.30285 | 0.081469 | 0.29673 | 0.60875
5308.8444 | 0.023191 | 0.18319 | -0.016809 | -0.0066583 | 0.21736 | -0.10664 | -0.084476 | 0.12705 | 0.29418 | 0.19116 | 0.26734 | 0.19359 | 0.2613 | 0.60998
5623.4133 | -0.0014509 | 0.21855 | 0.11855 | 0.053086 | 0.42538 | -0.027621 | -0.020464 | 0.12703 | 0.29417 | 0.19115 | 0.23202 | 0.284 | 0.20809 | 0.60684
5956.6214 | -0.010842 | 0.15916 | 0.059158 | 0.047618 | 0.52914 | -0.060865 | 0.0030244 | 0.13767 | 0.29403 | 0.19093 | 0.23338 | 0.28542 | 0.21641 | 0.70646
6309.5734 | 0.054457 | -0.51554 | -0.015543 | 0.081783 | 0.65895 | -0.019052 | 0.055412 | 0.13663 | 0.30632 | 0.19018 | 0.23225 | 0.28735 | 0.21655 | 0.80602
6683.4392 | 0.031081 | -0.54892 | 0.051081 | 0.078248 | 0.82853 | -0.040469 | -0.098949 | 0.14521 | 0.39432 | 0.18838 | 0.24232 | 0.6442 | 0.31376 | 1.0153
7079.4578 | -0.0013789 | -0.55138 | 0.048621 | 0.059893 | 0.96911 | -0.050886 | -0.084812 | 0.16642 | 0.39408 | 0.18787 | 0.28727 | 0.65645 | 0.31033 | 1.121
7498.9421 | -0.012951 | -0.45295 | 0.047049 | 0.061919 | 1.1415 | -0.053467 | -0.075293 | 0.17663 | 0.39382 | 0.18734 | 0.333 | 0.6664 | 0.30139 | 1.1269
7943.2823 | -0.035832 | -0.51583 | 0.084168 | 0.080812 | 1.2741 | -0.035854 | -0.040306 | 0.24375 | 0.38963 | 0.17836 | 0.44531 | 0.64218 | 0.24035 | 1.1137
8413.9514 | 0.00073827 | -0.25926 | 0.14074 | 0.10667 | 1.3449 | -9.2011e-05 | -0.0052885 | 0.26369 | 0.59257 | 0.28484 | 0.42532 | 1.123 | 0.53247 | 1.1521
8912.5094 | -0.046444 | -0.27644 | 0.12356 | 0.087058 | 1.3295 | -0.063501 | 0.065649 | 0.28863 | 0.58924 | 0.27787 | 0.54512 | 1.1463 | 0.4974 | 1.175
9440.6088 | 0.14353 | -0.12647 | 0.17353 | 0.22765 | -1.0624 | 0.091569 | 0.55681 | 0.40702 | 0.58375 | 0.37519 | 0.58376 | 1.1696 | 0.53132 | 1.1981
10000 | -0.054995 | -0.054995 | 0.14501 | 0.09581 | -1.3745 | 0.0024623 | -0.42933 | 0.70218 | 0.5785 | 0.5785 | 0.80584 | 1.1662 | 0.56294 | 1.1958

## 2652765 Phase-Revised
Results highlighted in blue were corrected during the circulation of APMP.AUV.A-K5 Draft A.
4160. 2652765 | Sensitivity Phase (4160 2652765) degrees |  |  |  | Expanded Uncertainties (degrees) |  |  |  | Estimated reference sensitivity phase | Std Unc (ERSP) | Phase difference relative to estimated reference sensitivity phase (degrees)
Frequency (Hz) | NIM | NMIJ | KRISS | CMS/ITRI | NIM | NMIJ | KRISS | CMS/ITRI | (degrees) | (degrees) | NIM | NMIJ | KRISS | CMS/ITRI
1.9953 | 176.3269 | 176.8 | 176.531 | 175.5648 | 1 | 2.5 | 2.5 | 1.3 | 176.1477 | 0.36162 | 0.17917 | 0.65232 | 0.38332 | -0.58283
2.5119 | 176.3712 | 176.69 | 176.437 | 175.704 | 0.9 | 2.43 | 2 | 1.1 | 176.1773 | 0.31748 | 0.19388 | 0.51272 | 0.25972 | -0.47325
3.1623 | 176.4582 | 176.06 | 176.504 | 175.9194 | 0.8 | 2.22 | 1.5 | 0.9 | 176.2475 | 0.26941 | 0.21067 | -0.18753 | 0.25647 | -0.32814
3.9811 | 176.6084 | 176.33 | 176.672 | 176.1862 | 0.7 | 1.44 | 1.5 | 0.8 | 176.4394 | 0.23492 | 0.16902 | -0.10938 | 0.23262 | -0.25317
5.0119 | 176.8252 | 176.96 | 176.886 | 176.4785 | 0.6 | 1.18 | 1.2 | 0.7 | 176.734 | 0.2003 | 0.091253 | 0.22603 | 0.15203 | -0.25548
6.3096 | 177.1457 | 177.22 | 177.235 | 176.791 | 0.6 | 1.01 | 1.2 | 0.6 | 177.0281 | 0.18595 | 0.11761 | 0.1919 | 0.2069 | -0.23705
7.9433 | 177.3736 | 177.37 | 177.36 | 177.1107 | 0.5 | 0.76 | 1 | 0.6 | 177.2947 | 0.16214 | 0.078853 | 0.075289 | 0.065289 | -0.18398
10 | 177.655 | 177.68 | 177.616 | 177.4178 | 0.5 | 0.6 | 0.8 | 0.5 | 177.5788 | 0.14233 | 0.076205 | 0.10121 | 0.037212 | -0.16103
12.5893 | 177.8633 | 177.89 | 177.856 | 177.6979 | 0.4 | 0.46 | 0.6 | 0.5 | 177.832 | 0.11867 | 0.031277 | 0.058014 | 0.024014 | -0.13409
15.8489 | 178.0732 | 178.1 | 178.058 | 177.941 | 0.35 | 0.37 | 0.5 | 0.5 | 178.0564 | 0.10321 | 0.016786 | 0.043575 | 0.001575 | -0.11541
19.9526 | 178.2578 | 178.29 | 178.254 | 178.1512 | 0.3 | 0.29 | 0.5 | 0.5 | 178.2559 | 0.0898 | 0.001912 | 0.034078 | -0.0019224 | -0.10469
21.1349 | 178.2997 | 178.33 | 178.294 | 178.1979 | 0.3 | 0.29 | 0.5 | 0.5 | 178.2975 | 0.0898 | 0.0022511 | 0.032549 | -0.0034514 | -0.099557
22.3872 | 178.339 | 178.38 | 178.336 | 178.2425 | 0.3 | 0.29 | 0.5 | 0.5 | 178.3419 | 0.0898 | -0.0028938 | 0.038116 | -0.0058842 | -0.099382
23.7137 | 178.3759 | 178.4 | 178.367 | 178.2851 | 0.3 | 0.29 | 0.5 | 0.5 | 178.3723 | 0.0898 | 0.0036108 | 0.027719 | -0.0052807 | -0.087149
25.1189 | 178.4106 | 178.43 | 178.408 | 178.3251 | 0.25 | 0.25 | 0.4 | 0.5 | 178.4095 | 0.076923 | 0.0011445 | 0.020525 | -0.0014753 | -0.084372
26.6073 | 178.4434 | 178.46 | 178.439 | 178.3619 | 0.25 | 0.25 | 0.4 | 0.5 | 178.4413 | 0.076923 | 0.0020764 | 0.018685 | -0.0023145 | -0.079431
28.1838 | 178.4744 | 178.49 | 178.47 | 178.3991 | 0.25 | 0.25 | 0.4 | 0.5 | 178.4725 | 0.076923 | 0.001868 | 0.017469 | -0.0025313 | -0.073392
29.8538 | 178.5038 | 178.52 | 178.509 | 178.4312 | 0.25 | 0.25 | 0.4 | 0.5 | 178.5038 | 0.076923 | -3.3415e-05 | 0.016173 | 0.0051735 | -0.072644
31.6228 | 178.5317 | 178.55 | 178.525 | 178.4633 | 0.2 | 0.19 | 0.4 | 0.5 | 178.5347 | 0.063019 | -0.0030719 | 0.015278 | -0.0097223 | -0.071411
33.4965 | 178.558 | 178.57 | 178.553 | 178.4929 | 0.2 | 0.19 | 0.4 | 0.5 | 178.5586 | 0.063019 | -0.00066979 | 0.011372 | -0.0056281 | -0.065772
35.4813 | 178.5826 | 178.6 | 178.579 | 178.5183 | 0.2 | 0.19 | 0.4 | 0.5 | 178.5858 | 0.063019 | -0.0032108 | 0.014187 | -0.0068129 | -0.067536
37.5837 | 178.6054 | 178.62 | 178.599 | 178.5437 | 0.2 | 0.19 | 0.4 | 0.5 | 178.6073 | 0.063019 | -0.001887 | 0.012745 | -0.008255 | -0.063569
39.8107 | 178.626 | 178.63 | 178.62 | 178.5667 | 0.2 | 0.16 | 0.4 | 0.5 | 178.6244 | 0.058001 | 0.0015702 | 0.0056041 | -0.0043959 | -0.057672
42.1697 | 178.6441 | 178.65 | 178.635 | 178.5852 | 0.2 | 0.16 | 0.4 | 0.5 | 178.6433 | 0.058001 | 0.00081764 | 0.0067398 | -0.0082602 | -0.058022
44.6684 | 178.6594 | 178.66 | 178.652 | 178.6027 | 0.2 | 0.16 | 0.4 | 0.5 | 178.6561 | 0.058001 | 0.0033786 | 0.0039482 | -0.0040518 | -0.053342
47.3151 | 178.6719 | 178.67 | 178.661 | 178.6193 | 0.2 | 0.16 | 0.4 | 0.5 | 178.6672 | 0.058001 | 0.0047402 | 0.0028492 | -0.0061508 | -0.047839
50.1187 | 178.6816 | 178.68 | 178.675 | 178.6326 | 0.2 | 0.14 | 0.4 | 0.5 | 178.6779 | 0.053832 | 0.0036759 | 0.0021073 | -0.0028927 | -0.045333
53.0884 | 178.6889 | 178.7 | 178.686 | 178.6449 | 0.15 | 0.14 | 0.4 | 0.5 | 178.6924 | 0.04863 | -0.0035245 | 0.0075848 | -0.0064152 | -0.04756
56.2341 | 178.6946 | 178.7 | 178.683 | 178.6527 | 0.15 | 0.14 | 0.4 | 0.5 | 178.6949 | 0.04863 | -0.00033253 | 0.0050604 | -0.01194 | -0.042196
59.5662 | 178.6996 | 178.7 | 178.694 | 178.659 | 0.15 | 0.14 | 0.4 | 0.5 | 178.6979 | 0.04863 | 0.0016999 | 0.0020535 | -0.0039465 | -0.038913
63.0957 | 178.7047 | 178.71 | 178.702 | 178.6645 | 0.15 | 0.12 | 0.4 | 0.5 | 178.7063 | 0.044876 | -0.0015034 | 0.0037487 | -0.0042513 | -0.041734
66.8344 | 178.7098 | 178.71 | 178.699 | 178.6659 | 0.15 | 0.12 | 0.4 | 0.5 | 178.708 | 0.044876 | 0.0018441 | 0.0020473 | -0.0089527 | -0.042044
70.7946 | 178.7129 | 178.7 | 178.695 | 178.6658 | 0.15 | 0.12 | 0.4 | 0.5 | 178.7033 | 0.044876 | 0.009631 | -0.0032601 | -0.0082601 | -0.037507
74.9894 | 178.7094 | 178.7 | 178.691 | 178.662 | 0.15 | 0.12 | 0.4 | 0.5 | 178.7017 | 0.044876 | 0.0077164 | -0.0016898 | -0.01069 | -0.039699
79.4328 | 178.6919 | 178.69 | 178.684 | 178.6569 | 0.15 | 0.1 | 0.4 | 0.5 | 178.6894 | 0.040201 | 0.0024336 | 0.00056022 | -0.0054398 | -0.032546
84.1395 | 178.6526 | 178.68 | 178.672 | 178.6483 | 0.15 | 0.1 | 0.4 | 0.5 | 178.671 | 0.040201 | -0.018377 | 0.0090127 | 0.0010127 | -0.022711
89.1251 | 178.5932 | 178.69 | 178.662 | 178.6378 | 0.15 | 0.1 | 0.4 | 0.5 | 178.6597 | 0.040201 | -0.066511 | 0.030293 | 0.0022927 | -0.021884
94.4061 | 178.5489 | 178.65 | 178.648 | 178.6241 | 0.15 | 0.1 | 0.4 | 0.5 | 178.6202 | 0.040201 | -0.071324 | 0.029806 | 0.027806 | 0.0038756
100 | 178.6431 | 178.62 | 178.63 | 178.6078 | 0.15 | 0.09 | 0.4 | 0.5 | 178.6258 | 0.037461 | 0.017243 | -0.0058337 | 0.0041663 | -0.018043
105.9254 | 178.6224 | 178.62 | 178.611 | 178.5886 | 0.15 | 0.09 | 0.4 | 0.5 | 178.6196 | 0.037461 | 0.0027872 | 0.00043463 | -0.0085654 | -0.031
112.2018 | 178.5985 | 178.6 | 178.588 | 178.5659 | 0.15 | 0.09 | 0.4 | 0.5 | 178.5984 | 0.037461 | 8.3453e-05 | 0.0015527 | -0.010447 | -0.032526
118.8502 | 178.5716 | 178.56 | 178.56 | 178.5416 | 0.15 | 0.09 | 0.4 | 0.5 | 178.5625 | 0.037461 | 0.0091117 | -0.0024779 | -0.0024779 | -0.020892
125.8925 | 178.5415 | 178.52 | 178.532 | 178.5121 | 0.15 | 0.07 | 0.4 | 0.5 | 178.5239 | 0.031082 | 0.017643 | -0.0038608 | 0.0081392 | -0.011767
133.3521 | 178.5082 | 178.51 | 178.498 | 178.4799 | 0.15 | 0.07 | 0.4 | 0.5 | 178.5089 | 0.031082 | -0.00070997 | 0.0010591 | -0.010941 | -0.029052
141.2538 | 178.4717 | 178.48 | 178.462 | 178.4443 | 0.15 | 0.07 | 0.4 | 0.5 | 178.4776 | 0.031082 | -0.0058855 | 0.0024124 | -0.015588 | -0.033334
149.6236 | 178.4318 | 178.42 | 178.423 | 178.4048 | 0.15 | 0.07 | 0.4 | 0.5 | 178.4219 | 0.031082 | 0.0099397 | -0.0018657 | 0.0011343 | -0.017026
158.4893 | 178.3884 | 178.38 | 178.379 | 178.3612 | 0.15 | 0.07 | 0.3 | 0.5 | 178.3811 | 0.030794 | 0.0072898 | -0.0010849 | -0.0020849 | -0.019857
167.8804 | 178.3412 | 178.33 | 178.33 | 178.3133 | 0.15 | 0.07 | 0.3 | 0.5 | 178.3316 | 0.030794 | 0.0095489 | -0.0016305 | -0.0016305 | -0.01838
177.8279 | 178.2899 | 178.28 | 178.279 | 178.2619 | 0.15 | 0.07 | 0.3 | 0.5 | 178.2814 | 0.030794 | 0.0085637 | -0.0013556 | -0.0023556 | -0.019448
188.3649 | 178.2342 | 178.22 | 178.22 | 178.2045 | 0.15 | 0.07 | 0.3 | 0.5 | 178.2222 | 0.030794 | 0.012063 | -0.0021638 | -0.0021638 | -0.017626
199.5262 | 178.1737 | 178.16 | 178.159 | 178.144 | 0.15 | 0.06 | 0.3 | 0.5 | 178.1616 | 0.027223 | 0.012102 | -0.0015799 | -0.0025799 | -0.017589
211.3489 | 178.1078 | 178.1 | 178.093 | 178.0773 | 0.15 | 0.06 | 0.3 | 0.5 | 178.1005 | 0.027223 | 0.0073054 | -0.00053354 | -0.0075335 | -0.023193
223.8721 | 178.0363 | 178.02 | 178.021 | 178.0071 | 0.15 | 0.06 | 0.3 | 0.5 | 178.022 | 0.027223 | 0.014243 | -0.0020225 | -0.0010225 | -0.014964
237.1374 | 177.9586 | 177.94 | 177.945 | 177.9295 | 0.15 | 0.06 | 0.3 | 0.5 | 177.9425 | 0.027223 | 0.016109 | -0.0024902 | 0.0025098 | -0.013032
251.1886 | 177.8746 | 177.86 | 177.861 | 177.8466 | 0.15 | 0.05 | 0.3 | 0.5 | 177.8613 | 0.023324 | 0.013281 | -0.00132 | -0.00032002 | -0.014677
266.0725 | 177.7842 | 177.78 | 177.772 | 177.758 | 0.15 | 0.05 | 0.3 | 0.5 | 177.78 | 0.023324 | 0.0041787 | -2.1169e-05 | -0.0080212 | -0.022032
281.8383 | 177.6875 | 177.68 | 177.679 | 177.6634 | 0.15 | 0.05 | 0.3 | 0.5 | 177.6806 | 0.023324 | 0.0069435 | -0.00055696 | -0.001557 | -0.017129
298.5383 | 177.5847 | 177.58 | 177.575 | 177.5615 | 0.15 | 0.05 | 0.3 | 0.5 | 177.5802 | 0.023324 | 0.0045413 | -0.00017414 | -0.0051741 | -0.018672
316.2278 | 177.476 | 177.46 | 177.468 | 177.4524 | 0.15 | 0.05 | 0.3 | 0.5 | 177.4617 | 0.023324 | 0.014307 | -0.0016729 | 0.0063271 | -0.0092589
334.9654 | 177.361 | 177.35 | 177.35 | 177.3346 | 0.15 | 0.05 | 0.3 | 0.5 | 177.3509 | 0.023324 | 0.010083 | -0.00093104 | -0.00093104 | -0.016345
354.8134 | 177.2386 | 177.22 | 177.225 | 177.2102 | 0.15 | 0.05 | 0.3 | 0.5 | 177.2218 | 0.023324 | 0.0168 | -0.0018382 | 0.0031618 | -0.011639
375.8374 | 177.1063 | 177.09 | 177.093 | 177.0772 | 0.15 | 0.05 | 0.3 | 0.5 | 177.0915 | 0.023324 | 0.014739 | -0.0015354 | 0.0014646 | -0.014289
398.1072 | 176.9598 | 176.95 | 176.952 | 176.9358 | 0.15 | 0.05 | 0.3 | 0.5 | 176.9509 | 0.023324 | 0.0089488 | -0.00087467 | 0.0011253 | -0.01509
421.6965 | 176.7949 | 176.79 | 176.8 | 176.7849 | 0.15 | 0.05 | 0.3 | 0.5 | 176.7907 | 0.023324 | 0.0042168 | -0.0006697 | 0.0093303 | -0.0058003
446.6836 | 176.6113 | 176.63 | 176.639 | 176.6227 | 0.15 | 0.05 | 0.3 | 0.5 | 176.6283 | 0.023324 | -0.017015 | 0.0016507 | 0.010651 | -0.0056045
473.1513 | 176.4252 | 176.46 | 176.467 | 176.4509 | 0.15 | 0.05 | 0.3 | 0.5 | 176.4567 | 0.023324 | -0.031488 | 0.0032719 | 0.010272 | -0.0058644
501.1872 | 176.2957 | 176.28 | 176.287 | 176.2688 | 0.15 | 0.06 | 0.3 | 0.5 | 176.2822 | 0.027223 | 0.013535 | -0.0021664 | 0.0048336 | -0.013371
530.8844 | 176.1016 | 176.09 | 176.091 | 176.0737 | 0.15 | 0.06 | 0.3 | 0.5 | 176.0914 | 0.027223 | 0.010219 | -0.0013658 | -0.00036575 | -0.017684
562.3413 | 175.8949 | 175.87 | 175.884 | 175.8665 | 0.15 | 0.06 | 0.3 | 0.5 | 175.8737 | 0.027223 | 0.021171 | -0.0036955 | 0.010304 | -0.0072282
595.6621 | 175.6749 | 175.66 | 175.666 | 175.6462 | 0.15 | 0.06 | 0.3 | 0.5 | 175.662 | 0.027223 | 0.012875 | -0.0019934 | 0.0040066 | -0.015753
630.9573 | 175.4408 | 175.42 | 175.43 | 175.4119 | 0.15 | 0.06 | 0.3 | 0.5 | 175.423 | 0.027223 | 0.017823 | -0.0029733 | 0.0070267 | -0.011071
668.3439 | 175.1918 | 175.17 | 175.183 | 175.1626 | 0.15 | 0.07 | 0.3 | 0.5 | 175.1741 | 0.030794 | 0.0177 | -0.0041129 | 0.0088871 | -0.011512
707.9458 | 174.927 | 174.91 | 174.918 | 174.8985 | 0.15 | 0.07 | 0.3 | 0.5 | 174.913 | 0.030794 | 0.013994 | -0.0030326 | 0.0049674 | -0.014568
749.8942 | 174.6455 | 174.62 | 174.638 | 174.6153 | 0.15 | 0.07 | 0.3 | 0.5 | 174.625 | 0.030794 | 0.020503 | -0.0049837 | 0.013016 | -0.0096914
794.3282 | 174.3462 | 174.32 | 174.338 | 174.315 | 0.15 | 0.07 | 0.3 | 0.5 | 174.3251 | 0.030794 | 0.021075 | -0.005095 | 0.012905 | -0.010062
841.3951 | 174.028 | 174 | 174.02 | 173.9968 | 0.15 | 0.08 | 0.3 | 0.5 | 174.0067 | 0.034036 | 0.021243 | -0.0067325 | 0.013267 | -0.0099002
891.2509 | 173.6897 | 173.66 | 173.681 | 173.6579 | 0.15 | 0.08 | 0.3 | 0.5 | 173.6672 | 0.034036 | 0.02255 | -0.0071603 | 0.01384 | -0.0092989
944.0609 | 173.3301 | 173.3 | 173.323 | 173.2982 | 0.15 | 0.08 | 0.3 | 0.5 | 173.3073 | 0.034036 | 0.022728 | -0.0073432 | 0.015657 | -0.0091774
1000 | 172.9476 | 172.92 | 172.938 | 172.9118 | 0.15 | 0.08 | 0.3 | 0.5 | 172.9265 | 0.034036 | 0.021167 | -0.0064665 | 0.011534 | -0.014629
1059.2537 | 172.5408 | 172.51 | 172.53 | 172.5031 | 0.15 | 0.09 | 0.3 | 0.5 | 172.5185 | 0.03696 | 0.02228 | -0.0085498 | 0.01145 | -0.015483
1122.0185 | 172.1079 | 172.07 | 172.098 | 172.0679 | 0.15 | 0.09 | 0.3 | 0.5 | 172.0809 | 0.03696 | 0.027071 | -0.010867 | 0.017133 | -0.012986
1188.5022 | 171.6471 | 171.61 | 171.636 | 171.6045 | 0.15 | 0.09 | 0.3 | 0.5 | 171.6205 | 0.03696 | 0.026597 | -0.010458 | 0.015542 | -0.015935
1258.9254 | 171.1561 | 171.12 | 171.143 | 171.1105 | 0.15 | 0.09 | 0.3 | 0.5 | 171.13 | 0.03696 | 0.026142 | -0.009954 | 0.013046 | -0.01948
1333.5214 | 170.6328 | 170.59 | 170.62 | 170.5854 | 0.15 | 0.11 | 0.3 | 0.5 | 170.6056 | 0.04193 | 0.027191 | -0.015585 | 0.014415 | -0.020148
1412.5375 | 170.0746 | 170.03 | 170.063 | 170.0264 | 0.15 | 0.11 | 0.3 | 0.5 | 170.0464 | 0.04193 | 0.028193 | -0.016421 | 0.016579 | -0.020029
1496.2357 | 169.4789 | 169.43 | 169.465 | 169.4269 | 0.15 | 0.11 | 0.3 | 0.5 | 169.4479 | 0.04193 | 0.030993 | -0.017941 | 0.017059 | -0.02107
1584.8932 | 168.8429 | 168.8 | 168.828 | 168.7894 | 0.15 | 0.11 | 0.3 | 0.5 | 168.8153 | 0.04193 | 0.027588 | -0.015291 | 0.012709 | -0.025905
1678.804 | 168.1634 | 168.12 | 168.15 | 168.1091 | 0.15 | 0.13 | 0.3 | 0.5 | 168.1387 | 0.045888 | 0.02473 | -0.018699 | 0.011301 | -0.029556
1778.2794 | 167.4374 | 167.39 | 167.424 | 167.3823 | 0.15 | 0.13 | 0.3 | 0.5 | 167.4107 | 0.045888 | 0.026721 | -0.02066 | 0.01334 | -0.028324
1883.6491 | 166.6613 | 166.62 | 166.648 | 166.6031 | 0.15 | 0.13 | 0.3 | 0.5 | 166.6375 | 0.045888 | 0.023793 | -0.017516 | 0.010484 | -0.034373
1995.2623 | 165.8314 | 165.78 | 165.817 | 165.7682 | 0.2 | 0.13 | 0.3 | 0.5 | 165.7966 | 0.05018 | 0.034777 | -0.016603 | 0.020397 | -0.028406
2113.489 | 164.943 | 164.88 | 164.91 | 164.8745 | 0.2 | 0.13 | 0.3 | 0.5 | 164.899 | 0.05018 | 0.043965 | -0.018986 | 0.011014 | -0.02452
2238.7211 | 163.9897 | 163.92 | 163.951 | 163.9138 | 0.2 | 0.13 | 0.3 | 0.5 | 163.9408 | 0.05018 | 0.048947 | -0.020777 | 0.010223 | -0.02697
2371.3737 | 162.9621 | 162.89 | 162.92 | 162.8816 | 0.2 | 0.13 | 0.3 | 0.5 | 162.9112 | 0.05018 | 0.050956 | -0.021186 | 0.0088145 | -0.029566
2511.8864 | 161.8445 | 161.78 | 161.807 | 161.7704 | 0.2 | 0.13 | 0.3 | 0.5 | 161.7989 | 0.05018 | 0.045594 | -0.018865 | 0.0081347 | -0.028486
2660.7251 | 160.6504 | 160.58 | 160.61 | 160.5682 | 0.2 | 0.14 | 0.3 | 0.5 | 160.6024 | 0.052376 | 0.047929 | -0.022447 | 0.0075526 | -0.034216
2818.3829 | 159.3567 | 159.29 | 159.322 | 159.274 | 0.2 | 0.14 | 0.3 | 0.5 | 159.3115 | 0.052376 | 0.045187 | -0.021492 | 0.010508 | -0.037483
2985.3826 | 157.9536 | 157.88 | 157.916 | 157.8668 | 0.2 | 0.14 | 0.3 | 0.5 | 157.904 | 0.052376 | 0.049621 | -0.024006 | 0.011994 | -0.037252
3162.2777 | 156.4298 | 156.36 | 156.39 | 156.3411 | 0.2 | 0.14 | 0.3 | 0.5 | 156.382 | 0.052376 | 0.047814 | -0.021972 | 0.0080281 | -0.040883
3349.6544 | 154.7717 | 154.7 | 154.719 | 154.6821 | 0.2 | 0.15 | 0.3 | 0.5 | 154.7228 | 0.054375 | 0.048837 | -0.022841 | -0.0038409 | -0.040776
3548.1339 | 152.9634 | 152.89 | 152.9 | 152.8741 | 0.2 | 0.15 | 0.3 | 0.5 | 152.9123 | 0.054375 | 0.051135 | -0.022262 | -0.012262 | -0.038175
3758.374 | 150.9862 | 150.91 | 150.907 | 150.8913 | 0.2 | 0.15 | 0.3 | 0.5 | 150.9313 | 0.054375 | 0.054961 | -0.021255 | -0.024255 | -0.039965
3981.0717 | 148.8182 | 148.75 | 148.738 | 148.7257 | 0.25 | 0.15 | 0.3 | 0.5 | 148.7614 | 0.057522 | 0.056819 | -0.011395 | -0.023395 | -0.035681
4216.965 | 146.434 | 146.37 | 146.347 | 146.3426 | 0.25 | 0.15 | 0.3 | 0.5 | 146.3787 | 0.057522 | 0.055276 | -0.0087156 | -0.031716 | -0.036165
4466.8359 | 143.8045 | 143.75 | 143.728 | 143.7173 | 0.25 | 0.15 | 0.3 | 0.5 | 143.7566 | 0.057522 | 0.047943 | -0.0065802 | -0.02858 | -0.039268
4731.5126 | 140.8973 | 140.86 | 140.826 | 140.8084 | 0.25 | 0.15 | 0.3 | 0.5 | 140.8602 | 0.057522 | 0.037129 | -0.00016706 | -0.034167 | -0.05175
5011.8723 | 137.6769 | 137.66 | 137.597 | 137.5952 | 0.25 | 0.15 | 0.3 | 0.6 | 137.6518 | 0.057993 | 0.025069 | 0.0082105 | -0.05479 | -0.056605
5308.8444 | 134.106 | 134.13 | 134.006 | 134.0282 | 0.25 | 0.27 | 0.3 | 0.6 | 134.0831 | 0.075717 | 0.022879 | 0.046894 | -0.077106 | -0.054938
5623.4133 | 130.1477 | 130.23 | 130.067 | 130.0742 | 0.25 | 0.27 | 0.3 | 0.6 | 130.1483 | 0.091457 | -0.0006376 | 0.081656 | -0.081344 | -0.074187
5956.6214 | 125.7685 | 125.85 | 125.66 | 125.7239 | 0.25 | 0.27 | 0.3 | 0.7 | 125.7643 | 0.092605 | 0.0041559 | 0.085673 | -0.10433 | -0.040438
6309.5734 | 120.9428 | 121.02 | 120.842 | 120.9165 | 0.25 | 0.27 | 0.3 | 0.8 | 120.9404 | 0.093375 | 0.0024317 | 0.079597 | -0.098403 | -0.023939
6683.4392 | 115.6597 | 115.8 | 115.541 | 115.4825 | 0.3 | 0.62 | 0.4 | 1 | 115.6333 | 0.11669 | 0.02644 | 0.16672 | -0.092278 | -0.15076
7079.4578 | 109.9308 | 110.13 | 109.82 | 109.7861 | 0.35 | 0.62 | 0.4 | 1.1 | 109.9143 | 0.12808 | 0.016462 | 0.21568 | -0.094317 | -0.12824
7498.9421 | 103.8004 | 104.1 | 103.685 | 103.6632 | 0.4 | 0.62 | 0.4 | 1.1 | 103.7969 | 0.13697 | 0.0034612 | 0.30307 | -0.11193 | -0.13375
7943.2823 | 97.3567 | 97.74 | 97.24 | 97.2355 | 0.45 | 0.62 | 0.4 | 1.1 | 97.3682 | 0.14425 | -0.011492 | 0.37184 | -0.12816 | -0.13261
8413.9514 | 90.7418 | 91.16 | 90.635 | 90.6298 | 0.55 | 1.07 | 0.7 | 1.1 | 90.7496 | 0.20124 | -0.0077924 | 0.41045 | -0.11455 | -0.11975
8912.5094 | 84.1576 | 84.59 | 84.007 | 84.1361 | 0.65 | 1.07 | 0.7 | 1.1 | 84.1662 | 0.21855 | -0.0086372 | 0.4238 | -0.1592 | -0.030047
9440.6088 | 77.8601 | 78.28 | 77.724 | 78.1892 | 0.7 | 1.07 | 0.7 | 1.1 | 77.9225 | 0.22569 | -0.062457 | 0.35746 | -0.19854 | 0.2667
10000 | 72.1303 | 72.77 | 72.037 | 71.6052 | 0.72 | 1.07 | 0.7 | 1.1 | 72.1187 | 0.22831 | 0.011607 | 0.65126 | -0.08174 | -0.51354

## 2652765 DoE Phase-Revised
Recalculated DoEs with the corrected results of NMIJ during the circulation of APMP.AUV.A-K5 Draft A.
4160. 2652765 | Linking DoEs from CCAUV.A-K5 |  |  | 4160 2652765 DoE (degrees) |  |  |  | Linking U(DoE)s from CCAUV.A-K5 |  |  | U(DoE) (degrees)
Frequency (Hz) | NIM | NMIJ | KRISS | NIM | NMIJ | KRISS | CMS-ITRI | NIM | NMIJ | KRISS | NIM | NMIJ | KRISS | CMS-ITRI
1.9953 | 0.12068 | 0.48068 |  | 0.17785 | 0.65099 | 0.38199 | -0.58416 | 1.9012 | 2.7229 |  | 1.52 | 2.8379 | 2.9462 | 2.0298
2.5119 | 0.0041898 | 0.45419 |  | 0.14797 | 0.46682 | 0.21382 | -0.51916 | 1.8942 | 2.3297 |  | 1.4156 | 2.8066 | 2.4622 | 1.809
3.1623 | 0.20273 | 0.31273 |  | 0.2897 | -0.1085 | 0.3355 | -0.24911 | 1.7076 | 2.2444 |  | 1.3152 | 2.5791 | 1.9692 | 1.5613
3.9811 | -0.012347 | 0.25765 |  | 0.15837 | -0.12004 | 0.22196 | -0.26382 | 1.5277 | 1.7543 |  | 1.1722 | 1.6687 | 1.8993 | 1.4133
5.0119 | -0.014873 | 0.15513 |  | 0.011446 | 0.14623 | 0.072227 | -0.33529 | 1.1072 | 1.6644 |  | 0.96384 | 1.2985 | 1.4838 | 1.1187
6.3096 | -0.15021 | -0.030212 |  | -0.12276 | -0.048469 | -0.033469 | -0.47742 | 0.92822 | 1.1747 |  | 0.80928 | 1.0433 | 1.3852 | 0.91586
7.9433 | -0.12801 | -0.018008 |  | -0.082706 | -0.086269 | -0.096269 | -0.34554 | 0.79655 | 0.98086 |  | 0.68905 | 0.82832 | 1.1864 | 0.87616
10 | -0.064805 | -0.0048051 |  | -0.049887 | -0.02488 | -0.08888 | -0.28712 | 0.6608 | 0.78374 |  | 0.62583 | 0.62306 | 0.95695 | 0.72509
12.5893 | -0.080102 | -0.0001022 |  | -0.055813 | -0.029076 | -0.063076 | -0.22118 | 0.52414 | 0.58577 |  | 0.51499 | 0.44253 | 0.70408 | 0.62107
15.8489 | -0.053307 | -0.023307 |  | -0.056409 | -0.029619 | -0.071619 | -0.1886 | 0.39127 | 0.58736 |  | 0.43283 | 0.37203 | 0.6044 | 0.6044
19.9526 | -0.031477 | -0.031477 |  | -0.048105 | -0.015939 | -0.051939 | -0.15471 | 0.37348 | 0.48723 |  | 0.39572 | 0.31034 | 0.57605 | 0.57605
21.1349 | -0.10436 | -0.034358 | 0.46564 | -0.093542 | -0.063245 | -0.099245 | -0.19535 | 0.37581 | 0.48902 | 0.28131 | 0.39867 | 0.31225 | 0.56493 | 0.57675
22.3872 | -0.091011 | -0.071011 | 0.42899 | -0.10513 | -0.064124 | -0.10812 | -0.20162 | 0.35566 | 0.48948 | 0.28212 | 0.39405 | 0.29999 | 0.55982 | 0.57126
23.7137 | -0.092555 | -0.012555 | 0.38745 | -0.079722 | -0.055613 | -0.088613 | -0.17048 | 0.32457 | 0.48964 | 0.28239 | 0.38567 | 0.27962 | 0.55112 | 0.56257
25.1189 | -0.099613 | -0.069613 | 0.43039 | -0.099817 | -0.080437 | -0.10244 | -0.18533 | 0.32494 | 0.48988 | 0.28281 | 0.36721 | 0.25636 | 0.45086 | 0.5526
26.6073 | -0.090545 | -0.10055 | 0.39945 | -0.10176 | -0.085146 | -0.10615 | -0.18326 | 0.30413 | 0.48999 | 0.28301 | 0.36063 | 0.24067 | 0.44334 | 0.54636
28.1838 | -0.079244 | -0.019244 | 0.38076 | -0.070384 | -0.054784 | -0.074784 | -0.14564 | 0.29371 | 0.49007 | 0.28314 | 0.35725 | 0.23247 | 0.43958 | 0.54323
29.8538 | -0.081785 | -0.051785 | 0.34821 | -0.081559 | -0.065352 | -0.076352 | -0.15417 | 0.2941 | 0.4903 | 0.28353 | 0.35776 | 0.23265 | 0.43995 | 0.54328
31.6228 | -0.071152 | 0.018848 | 0.31885 | -0.048397 | -0.030047 | -0.055047 | -0.11674 | 0.28439 | 0.38843 | 0.28439 | 0.2739 | 0.24621 | 0.45004 | 0.55073
33.4965 | -0.071151 | -0.0011506 | 0.29885 | -0.057817 | -0.045775 | -0.062775 | -0.12292 | 0.23461 | 0.39057 | 0.2873 | 0.25749 | 0.21814 | 0.43994 | 0.54069
35.4813 | -0.061411 | -0.021411 | 0.27859 | -0.059967 | -0.042569 | -0.063569 | -0.12429 | 0.22409 | 0.39066 | 0.28743 | 0.25333 | 0.21146 | 0.43731 | 0.53847
37.5837 | -0.058085 | -0.048085 | 0.35191 | -0.063292 | -0.04866 | -0.06966 | -0.12497 | 0.21357 | 0.39079 | 0.2876 | 0.2491 | 0.20463 | 0.43473 | 0.53627
39.8107 | -0.069055 | 0.020945 | 0.32094 | -0.050613 | -0.046579 | -0.056579 | -0.10986 | 0.20397 | 0.39141 | 0.28845 | 0.25757 | 0.18149 | 0.43218 | 0.53353
42.1697 | -0.06004 | 0.0099602 | 0.20996 | -0.0058563 | 6.5867e-05 | -0.014934 | -0.064695 | 0.18279 | 0.39168 | 0.28881 | 0.23275 | 0.15498 | 0.4232 | 0.52372
44.6684 | -0.054422 | -0.0044225 | 0.19558 | -0.0074376 | -0.006868 | -0.014868 | -0.064158 | 0.17213 | 0.39183 | 0.28901 | 0.23015 | 0.1494 | 0.41913 | 0.5223
47.3151 | -0.043572 | -0.013572 | 0.18643 | -0.0011316 | -0.0030226 | -0.012023 | -0.053711 | 0.17219 | 0.39185 | 0.28905 | 0.23022 | 0.14944 | 0.41917 | 0.52231
50.1187 | -0.036237 | 0.083763 | 0.18376 | 0.014238 | 0.01267 | 0.0076699 | -0.034771 | 0.16149 | 0.39202 | 0.28927 | 0.23524 | 0.13253 | 0.41903 | 0.52037
53.0884 | -0.050777 | -0.050777 | 0.14922 | -0.029634 | -0.018524 | -0.032524 | -0.07367 | 0.15235 | 0.39283 | 0.29037 | 0.19361 | 0.12778 | 0.41604 | 0.51479
56.2341 | -0.065779 | 0.034221 | 0.13422 | -0.01731 | -0.011917 | -0.028917 | -0.059173 | 0.1435 | 0.39369 | 0.29154 | 0.18288 | 0.11635 | 0.42263 | 0.51231
59.5662 | -0.059114 | 0.030886 | 0.13089 | -0.01156 | -0.011207 | -0.017207 | -0.052173 | 0.14357 | 0.39371 | 0.29157 | 0.18296 | 0.11637 | 0.42268 | 0.51232
63.0957 | -0.063705 | 0.016295 | 0.11629 | -0.02229 | -0.017038 | -0.025038 | -0.062521 | 0.14551 | 0.39443 | 0.29253 | 0.16859 | 0.12921 | 0.43353 | 0.51759
66.8344 | -0.062945 | 0.017055 | 0.11706 | -0.022051 | -0.021848 | -0.032848 | -0.065939 | 0.13481 | 0.39456 | 0.2927 | 0.16505 | 0.12382 | 0.42873 | 0.51638
70.7946 | -0.060883 | 0.019117 | 0.11912 | -0.019835 | -0.032726 | -0.037726 | -0.066973 | 0.11525 | 0.39545 | 0.29391 | 0.15931 | 0.11398 | 0.42051 | 0.51429
74.9894 | -0.057789 | 0.022211 | 0.12221 | -0.022594 | -0.032 | -0.041 | -0.07001 | 0.10446 | 0.39562 | 0.29413 | 0.15545 | 0.10807 | 0.41567 | 0.51313
79.4328 | -0.048931 | 0.031069 | 0.13107 | -0.022363 | -0.024236 | -0.030236 | -0.057342 | 0.093628 | 0.39581 | 0.29439 | 0.15884 | 0.090229 | 0.41358 | 0.51093
84.1395 | -0.055145 | 0.034855 | 0.13485 | -0.04739 | -0.020001 | -0.028001 | -0.051725 | 0.094911 | 0.39612 | 0.2948 | 0.1599 | 0.091184 | 0.41436 | 0.51108
89.1251 | -0.041551 | -0.051551 | 0.14845 | -0.088159 | 0.0086454 | -0.019355 | -0.043531 | 0.095126 | 0.39617 | 0.29487 | 0.16008 | 0.091345 | 0.41449 | 0.51111
94.4061 | -0.046945 | -0.036945 | 0.063055 | -0.10722 | -0.0060888 | -0.0080888 | -0.03202 | 0.087471 | 0.39705 | 0.29606 | 0.15879 | 0.086621 | 0.41165 | 0.51033
100 | -0.046418 | -0.016418 | 0.083582 | -0.017387 | -0.040463 | -0.030463 | -0.052672 | 0.076796 | 0.29631 | 0.29631 | 0.15699 | 0.074092 | 0.40768 | 0.50839
105.9254 | -0.047461 | 0.0025389 | 0.10254 | -0.030132 | -0.032484 | -0.041484 | -0.063919 | 0.079342 | 0.29698 | 0.29698 | 0.15912 | 0.075873 | 0.40892 | 0.50861
112.2018 | -0.036067 | 0.033933 | 0.13393 | -0.021789 | -0.02032 | -0.03232 | -0.054399 | 0.068527 | 0.29715 | 0.29715 | 0.1556 | 0.068438 | 0.40491 | 0.50766
118.8502 | -0.03014 | 0.05986 | 0.15986 | -0.0040766 | -0.015666 | -0.015666 | -0.03408 | 0.068527 | 0.29715 | 0.29715 | 0.1556 | 0.068438 | 0.40491 | 0.50766
125.8925 | -0.029474 | -0.0094735 | 0.090526 | -0.0032996 | -0.024803 | -0.012803 | -0.032709 | 0.068536 | 0.29715 | 0.29715 | 0.16258 | 0.06179 | 0.40788 | 0.50649
133.3521 | -0.029087 | 0.020913 | 0.12091 | -0.020434 | -0.018665 | -0.030665 | -0.048776 | 0.057667 | 0.29736 | 0.29736 | 0.15897 | 0.053329 | 0.4042 | 0.50559
141.2538 | -0.02845 | 0.06155 | 0.06155 | -0.026022 | -0.017724 | -0.035724 | -0.05347 | 0.057667 | 0.29736 | 0.29736 | 0.15897 | 0.053329 | 0.4042 | 0.50559
149.6236 | -0.032149 | -0.0021493 | 0.097851 | -0.016551 | -0.028356 | -0.025356 | -0.043516 | 0.049947 | 0.29815 | 0.29815 | 0.15807 | 0.046824 | 0.40193 | 0.50498
158.4893 | -0.026561 | 0.043439 | 0.043439 | -0.014392 | -0.022767 | -0.023767 | -0.041539 | 0.049947 | 0.29815 | 0.29815 | 0.15836 | 0.047453 | 0.30066 | 0.50493
167.8804 | -0.028351 | -0.0083512 | 0.091649 | -0.0139 | -0.02508 | -0.02508 | -0.041829 | 0.049947 | 0.29815 | 0.29815 | 0.15836 | 0.047453 | 0.30066 | 0.50493
177.8279 | -0.022709 | 0.047291 | 0.047291 | -0.010761 | -0.02068 | -0.02168 | -0.038772 | 0.039233 | 0.29839 | 0.19758 | 0.15472 | 0.039436 | 0.29762 | 0.50428
188.3649 | -0.024744 | 0.005256 | 0.10526 | -0.0087419 | -0.022969 | -0.022969 | -0.038431 | 0.039233 | 0.29839 | 0.19758 | 0.15472 | 0.039436 | 0.29762 | 0.50428
199.5262 | -0.021023 | -0.031023 | -0.031023 | -0.0096579 | -0.02334 | -0.02434 | -0.039349 | 0.041352 | 0.29868 | 0.19802 | 0.16017 | 0.037305 | 0.29991 | 0.50371
211.3489 | -0.02349 | 0.03651 | -0.06349 | -0.015937 | -0.023776 | -0.030776 | -0.046435 | 0.041352 | 0.29868 | 0.19802 | 0.16017 | 0.037305 | 0.29991 | 0.50371
223.8721 | -0.02049 | 0.0095104 | 0.0095104 | -0.004847 | -0.021113 | -0.020113 | -0.034054 | 0.041352 | 0.29868 | 0.19802 | 0.16017 | 0.037305 | 0.29991 | 0.50371
237.1374 | -0.0206 | -0.0106 | -0.0106 | -0.004121 | -0.022721 | -0.017721 | -0.033262 | 0.041352 | 0.29868 | 0.19802 | 0.16017 | 0.037305 | 0.29991 | 0.50371
251.1886 | -0.01477 | 0.07523 | -0.02477 | 0.00048768 | -0.014113 | -0.013113 | -0.02747 | 0.041577 | 0.29871 | 0.19806 | 0.16309 | 0.036545 | 0.30168 | 0.50314
266.0725 | -0.013209 | 0.066791 | -0.033209 | -0.0081936 | -0.012393 | -0.020393 | -0.034404 | 0.030703 | 0.2989 | 0.19835 | 0.15909 | 0.02674 | 0.29904 | 0.50256
281.8383 | -0.01353 | 0.06647 | -0.03353 | -0.0057064 | -0.013207 | -0.014207 | -0.029779 | 0.030703 | 0.2989 | 0.19835 | 0.15909 | 0.02674 | 0.29904 | 0.50256
298.5383 | -0.017216 | 0.072784 | -0.027216 | -0.011465 | -0.016181 | -0.021181 | -0.034679 | 0.033194 | 0.29917 | 0.19875 | 0.16188 | 0.028081 | 0.29952 | 0.50262
316.2278 | -0.015382 | 0.084618 | -0.015382 | 0.00056079 | -0.015419 | -0.0074194 | -0.023005 | 0.033194 | 0.29917 | 0.19875 | 0.16187 | 0.028087 | 0.29952 | 0.50262
334.9654 | -0.014725 | 0.0052746 | 0.0052746 | -0.004099 | -0.015113 | -0.015113 | -0.030528 | 0.033194 | 0.29917 | 0.19875 | 0.16187 | 0.028087 | 0.29952 | 0.50262
354.8134 | -0.013931 | 0.036069 | -0.063931 | 0.0027677 | -0.015871 | -0.010871 | -0.025672 | 0.033194 | 0.29917 | 0.19875 | 0.16187 | 0.028087 | 0.29952 | 0.50262
375.8374 | -0.015824 | 0.074176 | -0.025824 | 0.00016319 | -0.016111 | -0.013111 | -0.028865 | 0.033194 | 0.29917 | 0.19875 | 0.16187 | 0.028087 | 0.29952 | 0.50262
398.1072 | -0.01194 | 0.01806 | 0.01806 | -0.0020941 | -0.011918 | -0.0099175 | -0.026133 | 0.033194 | 0.29917 | 0.19875 | 0.16188 | 0.028142 | 0.29932 | 0.50262
421.6965 | -0.012814 | 0.077186 | -0.022814 | -0.0072746 | -0.012161 | -0.0021611 | -0.017292 | 0.033194 | 0.29917 | 0.19875 | 0.16188 | 0.028142 | 0.29932 | 0.50262
446.6836 | -0.0059924 | 0.044008 | -0.055992 | -0.023056 | -0.0043909 | 0.0046091 | -0.011646 | 0.033194 | 0.29917 | 0.19875 | 0.16188 | 0.028142 | 0.29932 | 0.50262
473.1513 | -0.0080528 | 0.021947 | 0.021947 | -0.038562 | -0.0038023 | 0.0031977 | -0.012939 | 0.033194 | 0.29917 | 0.19875 | 0.16188 | 0.028142 | 0.29932 | 0.50262
501.1872 | -0.0080502 | 0.01195 | 0.01195 | 0.006011 | -0.0096904 | -0.0026904 | -0.020895 | 0.033197 | 0.29917 | 0.19875 | 0.16276 | 0.026138 | 0.29728 | 0.50301
530.8844 | -0.0048666 | 0.015133 | 0.015133 | 0.0055323 | -0.0060524 | -0.0050524 | -0.022371 | 0.022221 | 0.29932 | 0.19898 | 0.15752 | 0.017585 | 0.29546 | 0.50267
562.3413 | -0.0078831 | 0.032117 | 0.032117 | 0.013986 | -0.010881 | 0.003119 | -0.014414 | 0.022221 | 0.29932 | 0.19898 | 0.15752 | 0.017585 | 0.29546 | 0.50267
595.6621 | -0.0063006 | 0.063699 | -0.036301 | 0.006777 | -0.0080909 | -0.0020909 | -0.021851 | 0.022221 | 0.29932 | 0.19898 | 0.15752 | 0.017585 | 0.29546 | 0.50267
630.9573 | -0.0040599 | 0.0059401 | 0.0059401 | 0.013826 | -0.00697 | 0.00303 | -0.015068 | 0.022221 | 0.29932 | 0.19898 | 0.15753 | 0.017566 | 0.29547 | 0.50268
668.3439 | -0.0042159 | 0.065784 | -0.034216 | 0.013699 | -0.0081141 | 0.0048859 | -0.015513 | 0.022221 | 0.29932 | 0.19898 | 0.1541 | 0.02322 | 0.29365 | 0.50331
707.9458 | -0.0067584 | 0.043242 | 0.043242 | 0.007992 | -0.0090351 | -0.0010351 | -0.02057 | 0.022221 | 0.29932 | 0.19898 | 0.1541 | 0.02322 | 0.29365 | 0.50331
749.8942 | -0.0037757 | 0.036224 | 0.036224 | 0.017362 | -0.0081244 | 0.0098756 | -0.012832 | 0.022221 | 0.29932 | 0.19898 | 0.1541 | 0.02322 | 0.29365 | 0.50331
794.3282 | -0.0017129 | 0.048287 | -0.051713 | 0.019207 | -0.0069628 | 0.011037 | -0.01193 | 0.022221 | 0.29932 | 0.19898 | 0.1541 | 0.023225 | 0.29365 | 0.50331
841.3951 | -0.0012899 | 0.07871 | -0.02129 | 0.020354 | -0.0076212 | 0.012379 | -0.010789 | 0.022221 | 0.29932 | 0.19898 | 0.15051 | 0.030776 | 0.29179 | 0.50395
891.2509 | -0.0025208 | 0.037479 | 0.037479 | 0.020638 | -0.0090724 | 0.011928 | -0.011211 | 0.022221 | 0.29932 | 0.19898 | 0.15051 | 0.030776 | 0.29179 | 0.50395
944.0609 | 0.0015747 | 0.011575 | 0.011575 | 0.024325 | -0.0057458 | 0.017254 | -0.0075799 | 0.022221 | 0.29932 | 0.19898 | 0.15051 | 0.030776 | 0.29179 | 0.50395
1000 | 0.0039563 | 0.013956 | 0.013956 | 0.025373 | -0.0022601 | 0.01574 | -0.010422 | 0.043917 | 0.29905 | 0.19857 | 0.1646 | 0.036484 | 0.29524 | 0.50425
1059.2537 | 0.0038817 | 0.043882 | -0.056118 | 0.02529 | -0.0055404 | 0.01446 | -0.012473 | 0.043908 | 0.29905 | 0.19856 | 0.16072 | 0.042114 | 0.2931 | 0.50472
1122.0185 | 0.0001162 | 0.0001162 | 0.0001162 | 0.026897 | -0.011041 | 0.016959 | -0.01316 | 0.043908 | 0.29905 | 0.19856 | 0.16072 | 0.042114 | 0.2931 | 0.50472
1188.5022 | 0.0047336 | 0.084734 | -0.015266 | 0.032554 | -0.0045011 | 0.021499 | -0.0099784 | 0.043908 | 0.29905 | 0.19856 | 0.16072 | 0.042114 | 0.2931 | 0.50472
1258.9254 | 0.0027664 | 0.0027664 | 0.0027664 | 0.028473 | -0.0076229 | 0.015377 | -0.017149 | 0.043908 | 0.29905 | 0.19856 | 0.1607 | 0.042149 | 0.2931 | 0.50471
1333.5214 | 0.0059759 | 0.055976 | -0.044024 | 0.032584 | -0.010193 | 0.019807 | -0.014756 | 0.043908 | 0.29905 | 0.19856 | 0.15305 | 0.058126 | 0.28898 | 0.50561
1412.5375 | 0.0055709 | 0.045571 | 0.045571 | 0.035289 | -0.0093245 | 0.023675 | -0.012933 | 0.043908 | 0.29905 | 0.19856 | 0.15305 | 0.058126 | 0.28898 | 0.50561
1496.2357 | 0.0053396 | 0.07534 | -0.02466 | 0.036775 | -0.012158 | 0.022842 | -0.015288 | 0.041323 | 0.29868 | 0.19801 | 0.14975 | 0.059673 | 0.28882 | 0.5058
1584.8932 | 0.008559 | 0.048559 | -0.051441 | 0.034657 | -0.0082228 | 0.019777 | -0.018837 | 0.052178 | 0.29853 | 0.19779 | 0.15371 | 0.064513 | 0.2917 | 0.50629
1678.804 | 0.0079447 | 0.067945 | -0.032055 | 0.032385 | -0.011044 | 0.018956 | -0.021901 | 0.052178 | 0.29853 | 0.19779 | 0.14663 | 0.082676 | 0.28803 | 0.50711
1778.2794 | 0.0090662 | 0.039066 | 0.039066 | 0.037022 | -0.01036 | 0.02364 | -0.018023 | 0.05204 | 0.29851 | 0.19776 | 0.14647 | 0.082724 | 0.28801 | 0.50712
1883.6491 | 0.016046 | 0.066046 | -0.033954 | 0.038641 | -0.0026685 | 0.025331 | -0.019525 | 0.05204 | 0.29851 | 0.19776 | 0.14647 | 0.082724 | 0.28801 | 0.50712
1995.2623 | 0.016618 | 0.056618 | -0.043382 | 0.048957 | -0.0024231 | 0.034577 | -0.014225 | 0.073483 | 0.29833 | 0.19748 | 0.23232 | 0.055169 | 0.28605 | 0.50664
2113.489 | 0.017149 | 0.10715 | 0.0071495 | 0.065065 | 0.002115 | 0.032115 | -0.0034191 | 0.073462 | 0.29832 | 0.19748 | 0.23228 | 0.05519 | 0.28605 | 0.50665
2238.7211 | 0.012528 | 0.13253 | 0.032528 | 0.069058 | -0.0006652 | 0.030335 | -0.006859 | 0.07142 | 0.29783 | 0.19673 | 0.22887 | 0.057114 | 0.28579 | 0.50689
2371.3737 | 0.013915 | 0.043915 | 0.043915 | 0.067365 | -0.0047766 | 0.025223 | -0.013157 | 0.07142 | 0.29783 | 0.19673 | 0.22887 | 0.057114 | 0.28579 | 0.50689
2511.8864 | 0.0086508 | 0.13865 | 0.038651 | 0.06301 | -0.0014493 | 0.025551 | -0.01107 | 0.071425 | 0.29783 | 0.19673 | 0.22874 | 0.056492 | 0.28669 | 0.50682
2660.7251 | 0.015453 | 0.13545 | 0.035453 | 0.072489 | 0.0021127 | 0.032113 | -0.0096561 | 0.081961 | 0.29769 | 0.19651 | 0.22726 | 0.071851 | 0.28823 | 0.50793
2818.3829 | 0.023267 | 0.043267 | 0.043267 | 0.07037 | 0.0036913 | 0.035691 | -0.012301 | 0.092581 | 0.29761 | 0.1964 | 0.22928 | 0.079138 | 0.29312 | 0.50889
2985.3826 | 0.024832 | 0.074832 | -0.025168 | 0.072743 | -0.00088382 | 0.035116 | -0.01413 | 0.090908 | 0.29709 | 0.19561 | 0.22688 | 0.079671 | 0.29255 | 0.50901
3162.2777 | 0.021246 | 0.14125 | 0.041246 | 0.07991 | 0.010124 | 0.040124 | -0.0087871 | 0.09077 | 0.29705 | 0.19555 | 0.22637 | 0.080172 | 0.29248 | 0.50887
3349.6544 | 0.019916 | 0.16992 | -0.030084 | 0.077324 | 0.0056455 | 0.024646 | -0.01229 | 0.10138 | 0.29695 | 0.19539 | 0.22462 | 0.094268 | 0.29472 | 0.51012
3548.1339 | 0.021383 | 0.17138 | -0.028617 | 0.082744 | 0.0093469 | 0.019347 | -0.0065656 | 0.11182 | 0.29682 | 0.1952 | 0.22601 | 0.10068 | 0.29979 | 0.51118
3758.374 | 0.022573 | 0.17257 | -0.027427 | 0.0892 | 0.012985 | 0.0099846 | -0.0057251 | 0.1223 | 0.29674 | 0.19508 | 0.2273 | 0.10696 | 0.30493 | 0.51228
3981.0717 | 0.024801 | 0.1048 | 0.0048007 | 0.088728 | 0.020513 | 0.0085133 | -0.0037726 | 0.13272 | 0.29667 | 0.19497 | 0.31182 | 0.079743 | 0.29948 | 0.51008
4216.965 | 0.026214 | 0.18621 | -0.013786 | 0.097672 | 0.033681 | 0.010681 | 0.0062314 | 0.13269 | 0.29666 | 0.19495 | 0.31177 | 0.079752 | 0.29947 | 0.51008
4466.8359 | 0.014597 | 0.1646 | 0.064597 | 0.090685 | 0.036162 | 0.014162 | 0.003475 | 0.13267 | 0.29665 | 0.19494 | 0.31173 | 0.079758 | 0.29946 | 0.51008
4731.5126 | 0.017525 | 0.17752 | -0.022475 | 0.069965 | 0.032669 | -0.0013306 | -0.018914 | 0.13133 | 0.29605 | 0.19403 | 0.30935 | 0.080169 | 0.29876 | 0.51019
5011.8723 | 0.023047 | 0.17305 | -0.026953 | 0.061089 | 0.044231 | -0.018769 | -0.020585 | 0.12767 | 0.29445 | 0.19157 | 0.30285 | 0.081469 | 0.29673 | 0.60875
5308.8444 | 0.023191 | 0.18319 | -0.016809 | 0.060537 | 0.084552 | -0.039448 | -0.017281 | 0.12705 | 0.29418 | 0.19116 | 0.26734 | 0.19359 | 0.2613 | 0.60998
5623.4133 | -0.0014509 | 0.21855 | 0.11855 | 0.044443 | 0.12674 | -0.036264 | -0.029106 | 0.12703 | 0.29417 | 0.19115 | 0.26731 | 0.1936 | 0.2613 | 0.60999
5956.6214 | -0.010842 | 0.15916 | 0.059158 | 0.031292 | 0.11281 | -0.07719 | -0.013301 | 0.13767 | 0.29403 | 0.19093 | 0.26751 | 0.1968 | 0.26687 | 0.70933
6309.5734 | 0.054457 | -0.51554 | -0.015543 | 0.081783 | 0.15895 | -0.019052 | 0.055412 | 0.13663 | 0.30632 | 0.19018 | 0.23225 | 0.28735 | 0.21655 | 0.80602
6683.4392 | 0.031081 | -0.54892 | 0.051081 | 0.078248 | 0.21853 | -0.040469 | -0.098949 | 0.14521 | 0.39432 | 0.18838 | 0.24232 | 0.6442 | 0.31376 | 1.0153
7079.4578 | -0.0013789 | -0.55138 | 0.048621 | 0.059893 | 0.25911 | -0.050886 | -0.084812 | 0.16642 | 0.39408 | 0.18787 | 0.28727 | 0.65645 | 0.31033 | 1.121
7498.9421 | -0.012951 | -0.45295 | 0.047049 | 0.061919 | 0.36153 | -0.053467 | -0.075293 | 0.17663 | 0.39382 | 0.18734 | 0.333 | 0.6664 | 0.30139 | 1.1269
7943.2823 | -0.035832 | -0.51583 | 0.084168 | 0.080812 | 0.46415 | -0.035854 | -0.040306 | 0.24375 | 0.38963 | 0.17836 | 0.44531 | 0.64218 | 0.24035 | 1.1137
8413.9514 | 0.00073827 | -0.25926 | 0.14074 | 0.0083118 | 0.42655 | -0.098447 | -0.10364 | 0.26369 | 0.59257 | 0.28484 | 0.45884 | 0.96912 | 0.56954 | 1.1499
8912.5094 | -0.046444 | -0.27644 | 0.12356 | -0.037986 | 0.39446 | -0.18854 | -0.059395 | 0.28863 | 0.58924 | 0.27787 | 0.57398 | 0.9631 | 0.54863 | 1.169
9440.6088 | 0.14353 | -0.12647 | 0.17353 | 0.086586 | 0.50651 | -0.049492 | 0.41575 | 0.40702 | 0.58375 | 0.37519 | 0.60872 | 0.97201 | 0.57147 | 1.1863
10000 | -0.054995 | -0.054995 | 0.14501 | -0.070526 | 0.56913 | -0.16387 | -0.59567 | 0.70218 | 0.5785 | 0.5785 | 0.73528 | 0.9826 | 0.58935 | 1.1917