## Introduction
This Excel workbook contains all the uncertainty budgets submitted by the participating laboratories in the APMP.AUV.A-K5 key comparison.

## NIM Level
NIM | Level
Unit
| No. | Frequency | Type | Hz | 2 | 3.98 | 7.94 | 15.85 | 31.62 | 63.1 | 125.89 | 251.19 | 501.19 | 1000 | 1995.26 | 3981.07 | 5011.87 | 7943.28 | 10000
| 1 | Repeatability | A | 10-4 dB | 600 | 450 | 160 | 120 | 120 | 120 | 120 | 120 | 120 | 120 | 120 | 140 | 180 | 200 | 250
| 2 | Series impedance | B |  | 300 | 280 | 180 | 120 | 60 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 80 | 100 | 130
| 3 | Voltage ratio | B |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| 4 | Cross-talk | B |  | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
| 5 | Inherent and ambient noise | B |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 6 | Distortion | B |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 7 | Coupler length | B |  | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 12 | 15 | 20 | 40
| 8 | Coupler diameter | B |  | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 20 | 30 | 40 | 50
| 9 | Coupler surface area | B |  | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 15 | 20 | 25 | 30
| 10 | Coupler volume | B |  | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
| 11 | Unintentional coupler/microphone leakage | B |  | 500 | 300 | 200 | 80 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60
| 12 | Static pressure | B |  | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15
| 13 | Temperature | B |  | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
| 14 | Humidity | B |  | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
| 15 | Front cavity length | B |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 8 | 8 | 10
| 16 | Front cavity volume and equivalent volume | B |  | 80 | 80 | 80 | 80 | 80 | 80 | 80 | 80 | 80 | 80 | 80 | 80 | 120 | 120 | 160
| 17 | Microphone resonant frequency | B |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 2 | 2 | 10 | 10 | 20
| 18 | Microphone loss factor | B |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 50 | 50 | 100
| 19 | Polarizing voltage | B |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 20 | Heat conduction correction | B |  | 400 | 300 | 180 | 120 | 120 | 20 | 20 | 10 | 5 | 5 | 5 | 5 | 2 | 1 | 1
| 21 | Radial wave motion | B |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 30 | 30 | 50
| 22 | Rounding error | B |  | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
| 23 | Static pressure correction | B |  | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40
| 24 | Temperature correction | B |  | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
| Standard Uncertainty |  |  |  | 933.3 | 686.9 | 376 | 246.1 | 216.8 | 176 | 176 | 175.2 | 175 | 175 | 175 | 190.2 | 257 | 279.9 | 361.6
| Expanded Uncertainty (k=2) |  |  |  | 1867 | 1374 | 752 | 492 | 434 | 352 | 352 | 350 | 350 | 350 | 350 | 380 | 514 | 560 | 723
| Stated Uncertainty (k=2) |  |  |  | 2000 | 1600 | 800 | 800 | 600 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 600 | 600 | 800

## NIM Phase
NIM | Phase
Unit
| No. | Frequency | Type | Hz | 2 | 3.98 | 7.94 | 15.85 | 31.62 | 63.1 | 125.89 | 251.19 | 501.19 | 1000 | 1995.26 | 3981.07 | 7943.28 | 10000
| 1 | Repeatability | A | degrees | 0.3 | 0.2 | 0.15 | 0.1 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.1 | 0.1
| 2 | Series impedance | B |  | 0.1 | 0.08 | 0.05 | 0.03 | 0.03 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.005 | 0.005 | 0.005 | 0.005
| 3 | Voltage ratio | B |  | 0.15 | 0.1 | 0.05 | 0.05 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.01 | 0.01 | 0.01 | 0.01
| 4 | Inherent and ambient noise | B |  | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005
| 5 | Coupler length | B |  | 0.005 | 0.004 | 0.003 | 0.002 | 0.002 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | 0.005 | 0.01 | 0.015
| 6 | Coupler diameter | B |  | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.004 | 0.005
| 7 | Coupler surface area | B |  | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.004 | 0.006
| 8 | Coupler volume | B |  | 0.01 | 0.008 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.003 | 0.003 | 0.002 | 0.002 | 0.003 | 0.004 | 0.005
| 9 | Unintentional coupler/microphone leakage | B |  | 0.25 | 0.2 | 0.15 | 0.1 | 0.05 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005
| 10 | Static pressure | B |  | 0.015 | 0.012 | 0.01 | 0.008 | 0.008 | 0.006 | 0.006 | 0.004 | 0.004 | 0.002 | 0.002 | 0.003 | 0.004 | 0.006
| 11 | Temperature | B |  | 0.02 | 0.018 | 0.016 | 0.014 | 0.012 | 0.01 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.01 | 0.012
| 12 | Humidity | B |  | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005
| 13 | Microphone front cavity length | B |  | 0.005 | 0.005 | 0.004 | 0.004 | 0.004 | 0.003 | 0.003 | 0.002 | 0.002 | 0.002 | 0.002 | 0.003 | 0.005 | 0.008
| 14 | Front cavity volume and equivalent volume | B |  | 0.01 | 0.008 | 0.006 | 0.006 | 0.004 | 0.004 | 0.002 | 0.002 | 0.004 | 0.01 | 0.05 | 0.06 | 0.15 | 0.24
| 15 | Microphone resonant frequency | B |  | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.015 | 0.02 | 0.04 | 0.08 | 0.08 | 0.1
| 16 | Microphone loss factor | B |  | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.01 | 0.015 | 0.015 | 0.02 | 0.02 | 0.1
| 17 | Heat conduction correction | B |  | 0.2 | 0.15 | 0.1 | 0.05 | 0.01 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.002 | 0.002 | 0.001 | 0.001
| 18 | Radial wave motion | B |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.1 | 0.2
| 19 | Rounding error | B |  | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01
| 20 | Static pressure correction | B |  | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01
| 21 | Temperature correction | B |  | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01
| Standard Uncertainty |  |  |  | 0.476 | 0.346 | 0.247 | 0.163 | 0.084 | 0.06 | 0.06 | 0.06 | 0.062 | 0.065 | 0.086 | 0.116 | 0.224 | 0.359
| Expanded Uncertainty (k=2) |  |  |  | 0.95 | 0.69 | 0.49 | 0.33 | 0.17 | 0.12 | 0.12 | 0.12 | 0.12 | 0.13 | 0.17 | 0.23 | 0.45 | 0.72
| Stated Uncertainty (k=2) |  |  |  | 1 | 0.7 | 0.5 | 0.35 | 0.2 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.2 | 0.25 | 0.45 | 0.72

## NMIA Level
NMIA | Level
Uncertainty Components for the B&K Type 4160 Microphone with SN 2652765 as part of the APMP AUV.A-K5 intercomparison
|  | Frequency | 1.9953 | 2.5119 | 3.1623 | 3.9811 | 5.0119 | 6.3096 | 7.9433 | 10 | 12.5892 | 15.8489 | 19.9526 | 21.1349 | 22.3872 | 23.7137 | 25.1189 | 26.6073 | 28.1838 | 29.8538 | 31.6228 | 33.4965 | 35.4813 | 37.5837 | 39.8107 | 42.1696 | 44.6684 | 47.3151 | 50.1187 | 53.0885 | 56.2341 | 59.5662 | 63.0957 | 66.8344 | 70.7946 | 74.9894 | 79.4328 | 84.1395 | 89.1251 | 94.4061 | 100 | 105.9254 | 112.2018 | 118.8502 | 125.8925 | 133.3521 | 141.2538 | 149.6235 | 158.4893 | 167.8804 | 177.8279 | 188.3649 | 199.5262 | 211.3489 | 223.8721 | 237.1374 | 251.1887 | 266.0725 | 281.8383 | 298.5382 | 316.2278 | 334.9655 | 354.8134 | 375.8374 | 398.1071 | 421.6965 | 446.6837 | 473.1512 | 501.1873 | 530.8843 | 562.3413 | 595.6623 | 630.9572 | 668.344 | 707.9456 | 749.8942 | 794.3285 | 841.395 | 891.251 | 944.0607 | 1000 | 1059.254 | 1122.018 | 1188.502 | 1258.925 | 1333.521 | 1412.538 | 1496.235 | 1584.893 | 1678.804 | 1778.279 | 1883.65 | 1995.262 | 2113.489 | 2238.721 | 2371.374 | 2511.887 | 2660.725 | 2818.383 | 2985.382 | 3162.278 | 3349.655 | 3548.134 | 3758.374 | 3981.071 | 4216.965 | 4466.837 | 4731.512 | 5011.873 | 5308.843 | 5623.413 | 5956.623 | 6309.573 | 6683.44 | 7079.456 | 7498.942 | 7943.284 | 8413.95 | 8912.511 | 9440.606 | 10000
Parameters
Microphone
| Front Cavity Volume | Rect | 0.0084 | 0.0085 | 0.00865 | 0.0088 | 0.00885 | 0.00895 | 0.00905 | 0.0091 | 0.00915 | 0.00925 | 0.00925 | 0.0093 | 0.0093 | 0.0093 | 0.0093 | 0.00935 | 0.0093 | 0.00935 | 0.00935 | 0.00935 | 0.00935 | 0.0094 | 0.0094 | 0.0094 | 0.0094 | 0.0094 | 0.00945 | 0.00945 | 0.0094 | 0.00945 | 0.0095 | 0.00945 | 0.0095 | 0.0095 | 0.00945 | 0.0095 | 0.0095 | 0.00945 | 0.0095 | 0.0095 | 0.0095 | 0.0095 | 0.00955 | 0.0095 | 0.0095 | 0.00955 | 0.0095 | 0.00955 | 0.00955 | 0.0095 | 0.00955 | 0.0096 | 0.00955 | 0.00955 | 0.0096 | 0.00955 | 0.0096 | 0.0096 | 0.00955 | 0.0096 | 0.0096 | 0.00955 | 0.0096 | 0.0096 | 0.00955 | 0.00955 | 0.00955 | 0.00955 | 0.00955 | 0.0096 | 0.00955 | 0.00955 | 0.00955 | 0.00955 | 0.00955 | 0.0095 | 0.0095 | 0.0095 | 0.0095 | 0.00945 | 0.00945 | 0.0094 | 0.0094 | 0.0094 | 0.00935 | 0.0093 | 0.00925 | 0.0092 | 0.00915 | 0.00915 | 0.00905 | 0.00895 | 0.0089 | 0.0088 | 0.0087 | 0.0086 | 0.00845 | 0.0083 | 0.00815 | 0.00795 | 0.00775 | 0.0075 | 0.00725 | 0.0069 | 0.00655 | 0.0062 | 0.0057 | 0.0052 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.0053 | 0.00555 | 0.00815 | 0.01155
| U1.1 |  | 0.0048497 | 0.0049075 | 0.0049941 | 0.0050807 | 0.0051095 | 0.0051673 | 0.005225 | 0.0052539 | 0.0052828 | 0.0053405 | 0.0053405 | 0.0053694 | 0.0053694 | 0.0053694 | 0.0053694 | 0.0053982 | 0.0053694 | 0.0053982 | 0.0053982 | 0.0053982 | 0.0053982 | 0.0054271 | 0.0054271 | 0.0054271 | 0.0054271 | 0.0054271 | 0.005456 | 0.005456 | 0.0054271 | 0.005456 | 0.0054848 | 0.005456 | 0.0054848 | 0.0054848 | 0.005456 | 0.0054848 | 0.0054848 | 0.005456 | 0.0054848 | 0.0054848 | 0.0054848 | 0.0054848 | 0.0055137 | 0.0054848 | 0.0054848 | 0.0055137 | 0.0054848 | 0.0055137 | 0.0055137 | 0.0054848 | 0.0055137 | 0.0055426 | 0.0055137 | 0.0055137 | 0.0055426 | 0.0055137 | 0.0055426 | 0.0055426 | 0.0055137 | 0.0055426 | 0.0055426 | 0.0055137 | 0.0055426 | 0.0055426 | 0.0055137 | 0.0055137 | 0.0055137 | 0.0055137 | 0.0055137 | 0.0055426 | 0.0055137 | 0.0055137 | 0.0055137 | 0.0055137 | 0.0055137 | 0.0054848 | 0.0054848 | 0.0054848 | 0.0054848 | 0.005456 | 0.005456 | 0.0054271 | 0.0054271 | 0.0054271 | 0.0053982 | 0.0053694 | 0.0053405 | 0.0053116 | 0.0052828 | 0.0052828 | 0.005225 | 0.0051673 | 0.0051384 | 0.0050807 | 0.0050229 | 0.0049652 | 0.0048786 | 0.004792 | 0.0047054 | 0.0045899 | 0.0044745 | 0.0043301 | 0.0041858 | 0.0039837 | 0.0037816 | 0.0035796 | 0.0032909 | 0.0030022 | 0.00306 | 0.00306 | 0.00306 | 0.00306 | 0.00306 | 0.00306 | 0.00306 | 0.00306 | 0.0032043 | 0.0047054 | 0.0066684
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Front Cavity Depth | Rect | 0.0011 | 0.001 | 0.0009 | 0.0008 | 0.0007 | 0.0007 | 0.0006 | 0.0005 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0003 | 0.0003 | 0.0004 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0002 | 0.0003 | 0.0003 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0001 | 0.0002 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0003 | 0.0003 | 0.0004 | 0.0004 | 0.0005 | 0.0005 | 0.0005 | 0.0006 | 0.0007 | 0.0007 | 0.0008 | 0.0008 | 0.0008 | 0.0008 | 0.0009 | 0.0008 | 0.0007 | 0.0005 | 0.0004 | 0.0003 | 0.0001 | 0.0002
| U1.2 |  | 0.00063509 | 0.00057735 | 0.00051962 | 0.00046188 | 0.00040415 | 0.00040415 | 0.00034641 | 0.00028868 | 0.00023094 | 0.00023094 | 0.00023094 | 0.00023094 | 0.00023094 | 0.00017321 | 0.00017321 | 0.00023094 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00011547 | 0.00017321 | 0.00017321 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 0.00011547 | 5.7735e-05 | 5.7735e-05 | 0.00011547 | 0.00011547 | 5.7735e-05 | 0.00011547 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 0 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 0 | 0 | 0 | 5.7735e-05 | 5.7735e-05 | 0 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00017321 | 0.00017321 | 0.00023094 | 0.00023094 | 0.00028868 | 0.00028868 | 0.00028868 | 0.00034641 | 0.00040415 | 0.00040415 | 0.00046188 | 0.00046188 | 0.00046188 | 0.00046188 | 0.00051962 | 0.00046188 | 0.00040415 | 0.00028868 | 0.00023094 | 0.00017321 | 5.7735e-05 | 0.00011547
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Equivalent Volume | Rect | 0.00845 | 0.0085 | 0.0087 | 0.0088 | 0.00885 | 0.00895 | 0.00905 | 0.0091 | 0.00915 | 0.0092 | 0.00925 | 0.0093 | 0.00925 | 0.0093 | 0.0093 | 0.00935 | 0.0093 | 0.00935 | 0.00935 | 0.00935 | 0.00935 | 0.00935 | 0.00935 | 0.0094 | 0.0094 | 0.0094 | 0.0094 | 0.00945 | 0.0094 | 0.00945 | 0.00945 | 0.00945 | 0.00945 | 0.0095 | 0.00945 | 0.0095 | 0.00945 | 0.00945 | 0.0095 | 0.0095 | 0.0095 | 0.0095 | 0.0095 | 0.0095 | 0.00955 | 0.00955 | 0.0095 | 0.00955 | 0.00955 | 0.0095 | 0.00955 | 0.00955 | 0.00955 | 0.0096 | 0.0096 | 0.00955 | 0.00955 | 0.0096 | 0.00955 | 0.0096 | 0.0096 | 0.00955 | 0.00955 | 0.0096 | 0.00955 | 0.00955 | 0.0096 | 0.00955 | 0.00955 | 0.0096 | 0.00955 | 0.00955 | 0.00955 | 0.00955 | 0.00955 | 0.00955 | 0.0095 | 0.0095 | 0.0095 | 0.00945 | 0.00945 | 0.00945 | 0.0094 | 0.0094 | 0.00935 | 0.0093 | 0.0093 | 0.00925 | 0.0092 | 0.00915 | 0.0091 | 0.009 | 0.00895 | 0.00885 | 0.0087 | 0.0086 | 0.0085 | 0.0083 | 0.00815 | 0.0079 | 0.00765 | 0.0074 | 0.00705 | 0.0067 | 0.00625 | 0.0058 | 0.0053 | 0.00465 | 0.00405 | 0.0034 | 0.00375 | 0.0021 | 0.0017 | 0.00125 | 0.00115 | 0.00155 | 0.00205 | 0.003 | 0.00435
| U1.3 |  | 0.0048786 | 0.0049075 | 0.0050229 | 0.0050807 | 0.0051095 | 0.0051673 | 0.005225 | 0.0052539 | 0.0052828 | 0.0053116 | 0.0053405 | 0.0053694 | 0.0053405 | 0.0053694 | 0.0053694 | 0.0053982 | 0.0053694 | 0.0053982 | 0.0053982 | 0.0053982 | 0.0053982 | 0.0053982 | 0.0053982 | 0.0054271 | 0.0054271 | 0.0054271 | 0.0054271 | 0.005456 | 0.0054271 | 0.005456 | 0.005456 | 0.005456 | 0.005456 | 0.0054848 | 0.005456 | 0.0054848 | 0.005456 | 0.005456 | 0.0054848 | 0.0054848 | 0.0054848 | 0.0054848 | 0.0054848 | 0.0054848 | 0.0055137 | 0.0055137 | 0.0054848 | 0.0055137 | 0.0055137 | 0.0054848 | 0.0055137 | 0.0055137 | 0.0055137 | 0.0055426 | 0.0055426 | 0.0055137 | 0.0055137 | 0.0055426 | 0.0055137 | 0.0055426 | 0.0055426 | 0.0055137 | 0.0055137 | 0.0055426 | 0.0055137 | 0.0055137 | 0.0055426 | 0.0055137 | 0.0055137 | 0.0055426 | 0.0055137 | 0.0055137 | 0.0055137 | 0.0055137 | 0.0055137 | 0.0055137 | 0.0054848 | 0.0054848 | 0.0054848 | 0.005456 | 0.005456 | 0.005456 | 0.0054271 | 0.0054271 | 0.0053982 | 0.0053694 | 0.0053694 | 0.0053405 | 0.0053116 | 0.0052828 | 0.0052539 | 0.0051962 | 0.0051673 | 0.0051095 | 0.0050229 | 0.0049652 | 0.0049075 | 0.004792 | 0.0047054 | 0.0045611 | 0.0044167 | 0.0042724 | 0.0040703 | 0.0038682 | 0.0036084 | 0.0033486 | 0.00306 | 0.0026847 | 0.0023383 | 0.001963 | 0.0021651 | 0.0012124 | 0.0009815 | 0.00072169 | 0.00066395 | 0.00089489 | 0.0011836 | 0.0017321 | 0.0025115
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Resonance Frequency | Rect | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00038594 | 0.00042975 | 0.00047655 | 0.00052655 | 0.00058 | 0.00063716 | 0.00069833 | 0.0007638 | 0.00083393 | 0.00090908 | 0.00098965 | 0.0010761 | 0.0011689 | 0.0012685 | 0.0013756 | 0.0014908 | 0.0016147 | 0.0017481 | 0.0018919 | 0.002047 | 0.0022143 | 0.0023949 | 0.0025902 | 0.0028013 | 0.0030298 | 0.0032773 | 0.0035456 | 0.0038367 | 0.0041528 | 0.0044962 | 0.0048697 | 0.0052762 | 0.005719 | 0.0062018 | 0.0067286 | 0.0073038 | 0.0079325 | 0.0086202 | 0.0093731 | 0.010198 | 0.011102 | 0.012095 | 0.013185 | 0.014382 | 0.0157
| U1.4 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00022282 | 0.00024812 | 0.00027513 | 0.000304 | 0.00033486 | 0.00036787 | 0.00040318 | 0.00044098 | 0.00048147 | 0.00052486 | 0.00057138 | 0.00062128 | 0.00067484 | 0.00073238 | 0.00079421 | 0.0008607 | 0.00093224 | 0.0010093 | 0.0010923 | 0.0011818 | 0.0012784 | 0.0013827 | 0.0014954 | 0.0016173 | 0.0017493 | 0.0018922 | 0.0020471 | 0.0022151 | 0.0023976 | 0.0025959 | 0.0028115 | 0.0030462 | 0.0033019 | 0.0035806 | 0.0038847 | 0.0042169 | 0.0045798 | 0.0049769 | 0.0054115 | 0.0058878 | 0.0064099 | 0.0069829 | 0.0076122 | 0.0083037 | 0.0090644
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Loss Factor | Rect | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0 | 0.0001 | 0.0001 | 0 | 0.0001 | 0.0001 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0004 | 0.0004 | 0.0005 | 0.0005 | 0.0006 | 0.0007 | 0.0007 | 0.0008 | 0.001 | 0.0011 | 0.0012 | 0.0013 | 0.0015 | 0.0016 | 0.0018 | 0.0021 | 0.0022 | 0.0025 | 0.0028 | 0.0031 | 0.0034 | 0.0037 | 0.004 | 0.0043 | 0.0046 | 0.0048 | 0.005 | 0.0049 | 0.0047 | 0.0043 | 0.0036 | 0.0028 | 0.0021 | 0.0013 | 0.0011 | 0.0018 | 0.0029 | 0.0046 | 0.0068
| U1.5 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5.7735e-05 | 0 | 5.7735e-05 | 5.7735e-05 | 0 | 5.7735e-05 | 5.7735e-05 | 0 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00023094 | 0.00023094 | 0.00028868 | 0.00028868 | 0.00034641 | 0.00040415 | 0.00040415 | 0.00046188 | 0.00057735 | 0.00063509 | 0.00069282 | 0.00075056 | 0.00086603 | 0.00092376 | 0.0010392 | 0.0012124 | 0.0012702 | 0.0014434 | 0.0016166 | 0.0017898 | 0.001963 | 0.0021362 | 0.0023094 | 0.0024826 | 0.0026558 | 0.0027713 | 0.0028868 | 0.002829 | 0.0027135 | 0.0024826 | 0.0020785 | 0.0016166 | 0.0012124 | 0.00075056 | 0.00063509 | 0.0010392 | 0.0016743 | 0.0026558 | 0.003926
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Diaphragm Diameter | Rect | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0 | 0 | 0.0001 | 0 | 0.0001 | 0.0001 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0001 | 0.0002 | 0.0002 | 0.0002 | 0.0003 | 0.0003 | 0.0003 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0005 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0007 | 0.0007 | 0.0006 | 0.0006 | 0.0005 | 0.0004 | 0.0002 | 0 | 0.0001
| U1.6 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5.7735e-05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5.7735e-05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5.7735e-05 | 0 | 0 | 5.7735e-05 | 0 | 5.7735e-05 | 5.7735e-05 | 0 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 0.00011547 | 5.7735e-05 | 0.00011547 | 0.00011547 | 0.00011547 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00023094 | 0.00023094 | 0.00023094 | 0.00023094 | 0.00028868 | 0.00034641 | 0.00034641 | 0.00034641 | 0.00034641 | 0.00040415 | 0.00040415 | 0.00034641 | 0.00034641 | 0.00028868 | 0.00023094 | 0.00011547 | 0 | 5.7735e-05
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Microphone Parameters Combined |  | 0.0069083 | 0.0069642 | 0.0071022 | 0.0072 | 0.0072373 | 0.0073188 | 0.0073974 | 0.0074357 | 0.0074745 | 0.0075357 | 0.0075561 | 0.0075969 | 0.0075766 | 0.0075954 | 0.0075954 | 0.0076377 | 0.0075954 | 0.0076362 | 0.0076362 | 0.0076362 | 0.0076362 | 0.0076566 | 0.0076566 | 0.007677 | 0.007677 | 0.007677 | 0.0076975 | 0.0077178 | 0.0076759 | 0.0077178 | 0.0077383 | 0.0077168 | 0.0077372 | 0.0077576 | 0.0077168 | 0.0077576 | 0.0077372 | 0.0077168 | 0.0077576 | 0.0077576 | 0.0077576 | 0.0077576 | 0.007778 | 0.0077569 | 0.0077774 | 0.0077978 | 0.0077569 | 0.0077984 | 0.007798 | 0.0077569 | 0.0077984 | 0.007819 | 0.0077978 | 0.007819 | 0.0078388 | 0.0077978 | 0.0078184 | 0.0078388 | 0.0077978 | 0.0078388 | 0.0078388 | 0.007798 | 0.0078186 | 0.0078388 | 0.007798 | 0.007798 | 0.0078184 | 0.007798 | 0.007798 | 0.0078388 | 0.007798 | 0.0077986 | 0.0077986 | 0.0077988 | 0.0078018 | 0.007782 | 0.007764 | 0.0077648 | 0.0077663 | 0.0077268 | 0.0077299 | 0.0077118 | 0.007696 | 0.0076988 | 0.0076636 | 0.0076299 | 0.0076142 | 0.0075823 | 0.0075559 | 0.0075476 | 0.0075007 | 0.0074357 | 0.0074205 | 0.0073619 | 0.0072932 | 0.0072605 | 0.0071945 | 0.0071182 | 0.0070721 | 0.0069812 | 0.0069033 | 0.0068272 | 0.006731 | 0.0066236 | 0.0065156 | 0.0064213 | 0.0063018 | 0.0061138 | 0.0061334 | 0.0061233 | 0.0063026 | 0.0062169 | 0.0064354 | 0.0067377 | 0.0071736 | 0.0077532 | 0.0085123 | 0.010057 | 0.012181
| EffDoF |  | 61.0161 | 60.8275 | 60.642 | 60.4948 | 60.3748 | 60.3665 | 60.2634 | 60.181 | 60.1146 | 60.111 | 60.1121 | 60.1109 | 60.1098 | 60.0624 | 60.0624 | 60.1098 | 60.0624 | 60.0618 | 60.0618 | 60.0618 | 60.0618 | 60.0597 | 60.0597 | 60.0611 | 60.0611 | 60.0611 | 60.0591 | 60.0605 | 60.0272 | 60.0605 | 60.0585 | 60.0269 | 60.0251 | 60.0266 | 60.0269 | 60.0266 | 60.0251 | 60.0269 | 60.0266 | 60.0266 | 60.0266 | 60.0266 | 60.0248 | 60.0066 | 60.005 | 60.0066 | 60.0066 | 60.0263 | 60.0132 | 60.0066 | 60.0263 | 60.0311 | 60.0066 | 60.0311 | 60.013 | 60.0066 | 60.0115 | 60.013 | 60.0066 | 60.013 | 60.013 | 60.0132 | 60.018 | 60.013 | 60.0132 | 60.0132 | 60.0115 | 60.0132 | 60.0132 | 60.013 | 60.0132 | 60.0329 | 60.0329 | 60.0395 | 60.1309 | 60.1469 | 60.2241 | 60.2508 | 60.2967 | 60.3398 | 60.4349 | 60.5068 | 60.6552 | 60.7439 | 60.9254 | 61.1562 | 61.3052 | 61.5959 | 62.0718 | 62.4711 | 62.9567 | 63.5409 | 64.4175 | 65.2455 | 66.47 | 68.2557 | 69.6431 | 72.2648 | 75.2988 | 79.2054 | 83.7214 | 89.323 | 95.7776 | 103.2657 | 111.1802 | 117.2391 | 120.2929 | 115.8893 | 106.2469 | 92.3386 | 83.2903 | 63.092 | 54.254 | 47.9262 | 44.7294 | 43.9391 | 45.3555 | 57.8761 | 73.3322
Coupler
| Coupler Length | Rect | 0.00082909 | 0.00083377 | 0.00083798 | 0.00084179 | 0.00084522 | 0.00084828 | 0.00085104 | 0.00085351 | 0.00085571 | 0.00085769 | 0.00085946 | 0.00085986 | 0.00086027 | 0.00086065 | 0.00086102 | 0.0008614 | 0.00086175 | 0.00086209 | 0.00086243 | 0.00086275 | 0.00086307 | 0.00086338 | 0.00086368 | 0.00086396 | 0.00086425 | 0.00086452 | 0.00086478 | 0.00086504 | 0.00086528 | 0.00086552 | 0.00086576 | 0.00086598 | 0.00086619 | 0.00086641 | 0.00086661 | 0.0008668 | 0.000867 | 0.00086717 | 0.00086734 | 0.00086752 | 0.00086767 | 0.00086782 | 0.00086797 | 0.0008681 | 0.00086822 | 0.00086834 | 0.00086845 | 0.00086854 | 0.00086864 | 0.00086871 | 0.00086878 | 0.00086884 | 0.00086888 | 0.00086891 | 0.00086893 | 0.00086892 | 0.0008689 | 0.00086886 | 0.0008688 | 0.00086871 | 0.00086861 | 0.00086846 | 0.00086829 | 0.0008681 | 0.00086784 | 0.00086756 | 0.00086723 | 0.00086683 | 0.00086638 | 0.00086588 | 0.00086528 | 0.00086461 | 0.00086386 | 0.00086297 | 0.00086199 | 0.00086089 | 0.00085961 | 0.00085819 | 0.0008566 | 0.00085476 | 0.00085272 | 0.00085044 | 0.00084782 | 0.0008449 | 0.00084164 | 0.0008379 | 0.00083375 | 0.00082911 | 0.00082379 | 0.00081787 | 0.00081126 | 0.00080368 | 0.00079526 | 0.00078583 | 0.00077504 | 0.00076302 | 0.00074957 | 0.00073416 | 0.000717 | 0.00069777 | 0.00067572 | 0.00065115 | 0.00062362 | 0.00059205 | 0.00055689 | 0.00051753 | 0.00047245 | 0.00042238 | 0.00051864 | 0.00071112 | 0.00096029 | 0.0012818 | 0.0012818 | 0.0012818 | 0.0012818 | 0.0012818 | 0.0012818 | 0.0012818 | 0.0012818
| U2.1 |  | 0.00047867 | 0.00048138 | 0.00048381 | 0.00048601 | 0.00048799 | 0.00048976 | 0.00049135 | 0.00049278 | 0.00049405 | 0.00049519 | 0.00049621 | 0.00049644 | 0.00049667 | 0.0004969 | 0.00049711 | 0.00049733 | 0.00049753 | 0.00049773 | 0.00049792 | 0.00049811 | 0.00049829 | 0.00049847 | 0.00049864 | 0.00049881 | 0.00049897 | 0.00049913 | 0.00049928 | 0.00049943 | 0.00049957 | 0.00049971 | 0.00049985 | 0.00049997 | 0.0005001 | 0.00050022 | 0.00050034 | 0.00050045 | 0.00050056 | 0.00050066 | 0.00050076 | 0.00050086 | 0.00050095 | 0.00050103 | 0.00050112 | 0.0005012 | 0.00050127 | 0.00050134 | 0.0005014 | 0.00050145 | 0.00050151 | 0.00050155 | 0.00050159 | 0.00050163 | 0.00050165 | 0.00050166 | 0.00050168 | 0.00050167 | 0.00050166 | 0.00050164 | 0.0005016 | 0.00050155 | 0.00050149 | 0.00050141 | 0.00050131 | 0.0005012 | 0.00050105 | 0.00050088 | 0.0005007 | 0.00050047 | 0.00050021 | 0.00049992 | 0.00049957 | 0.00049918 | 0.00049875 | 0.00049824 | 0.00049767 | 0.00049703 | 0.0004963 | 0.00049547 | 0.00049456 | 0.0004935 | 0.00049232 | 0.000491 | 0.00048949 | 0.0004878 | 0.00048592 | 0.00048376 | 0.00048136 | 0.00047868 | 0.00047561 | 0.0004722 | 0.00046838 | 0.00046401 | 0.00045914 | 0.0004537 | 0.00044747 | 0.00044053 | 0.00043277 | 0.00042387 | 0.00041396 | 0.00040286 | 0.00039013 | 0.00037594 | 0.00036005 | 0.00034182 | 0.00032152 | 0.0002988 | 0.00027277 | 0.00024386 | 0.00029944 | 0.00041057 | 0.00055443 | 0.00074003 | 0.00074003 | 0.00074003 | 0.00074003 | 0.00074003 | 0.00074003 | 0.00074003 | 0.00074003
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Coupler Diameter | Rect | 0.00069471 | 0.00069547 | 0.00069619 | 0.00069688 | 0.00069749 | 0.00069806 | 0.00069858 | 0.00069905 | 0.00069947 | 0.00069986 | 0.0007002 | 0.00070028 | 0.00070036 | 0.00070044 | 0.00070051 | 0.00070059 | 0.00070065 | 0.00070072 | 0.0007008 | 0.00070086 | 0.00070092 | 0.00070099 | 0.00070104 | 0.0007011 | 0.00070117 | 0.00070122 | 0.00070128 | 0.00070134 | 0.00070138 | 0.00070144 | 0.00070149 | 0.00070154 | 0.00070159 | 0.00070165 | 0.00070169 | 0.00070174 | 0.00070179 | 0.00070184 | 0.00070189 | 0.00070194 | 0.00070199 | 0.00070204 | 0.00070209 | 0.00070214 | 0.00070219 | 0.00070225 | 0.00070231 | 0.00070237 | 0.00070244 | 0.0007025 | 0.00070257 | 0.00070265 | 0.00070273 | 0.00070282 | 0.00070291 | 0.00070301 | 0.00070312 | 0.00070325 | 0.00070338 | 0.00070352 | 0.00070369 | 0.00070386 | 0.00070405 | 0.00070427 | 0.00070451 | 0.00070478 | 0.00070508 | 0.0007054 | 0.00070577 | 0.00070618 | 0.00070663 | 0.00070715 | 0.00070773 | 0.00070836 | 0.00070908 | 0.00070989 | 0.00071078 | 0.00071179 | 0.00071293 | 0.00071419 | 0.00071563 | 0.00071724 | 0.00071902 | 0.00072107 | 0.00072336 | 0.0007259 | 0.00072882 | 0.00073209 | 0.00073574 | 0.00073994 | 0.00074464 | 0.00074991 | 0.00075599 | 0.00076283 | 0.00077051 | 0.00077941 | 0.00078947 | 0.00080083 | 0.00081407 | 0.00082912 | 0.00084622 | 0.00086632 | 0.00088935 | 0.00091573 | 0.00094706 | 0.00098331 | 0.0010253 | 0.0010757 | 0.0011349 | 0.0012045 | 0.0012903 | 0.0013946 | 0.0013946 | 0.0013946 | 0.0013946 | 0.0013946 | 0.0013946 | 0.0013946 | 0.0013946
| U2.2 |  | 0.00040109 | 0.00040153 | 0.00040195 | 0.00040234 | 0.0004027 | 0.00040302 | 0.00040333 | 0.00040359 | 0.00040384 | 0.00040406 | 0.00040426 | 0.00040431 | 0.00040436 | 0.0004044 | 0.00040444 | 0.00040448 | 0.00040452 | 0.00040456 | 0.0004046 | 0.00040464 | 0.00040468 | 0.00040472 | 0.00040475 | 0.00040478 | 0.00040482 | 0.00040485 | 0.00040488 | 0.00040492 | 0.00040494 | 0.00040497 | 0.00040501 | 0.00040503 | 0.00040506 | 0.00040509 | 0.00040512 | 0.00040515 | 0.00040518 | 0.00040521 | 0.00040523 | 0.00040527 | 0.00040529 | 0.00040532 | 0.00040535 | 0.00040538 | 0.00040541 | 0.00040545 | 0.00040548 | 0.00040551 | 0.00040555 | 0.00040559 | 0.00040563 | 0.00040568 | 0.00040572 | 0.00040577 | 0.00040583 | 0.00040588 | 0.00040595 | 0.00040602 | 0.00040609 | 0.00040618 | 0.00040627 | 0.00040637 | 0.00040649 | 0.00040661 | 0.00040675 | 0.0004069 | 0.00040708 | 0.00040726 | 0.00040748 | 0.00040772 | 0.00040798 | 0.00040827 | 0.00040861 | 0.00040897 | 0.00040939 | 0.00040985 | 0.00041037 | 0.00041095 | 0.00041161 | 0.00041234 | 0.00041317 | 0.0004141 | 0.00041513 | 0.00041631 | 0.00041763 | 0.0004191 | 0.00042079 | 0.00042267 | 0.00042478 | 0.0004272 | 0.00042992 | 0.00043296 | 0.00043647 | 0.00044042 | 0.00044485 | 0.00044999 | 0.0004558 | 0.00046236 | 0.00047 | 0.00047869 | 0.00048856 | 0.00050017 | 0.00051347 | 0.0005287 | 0.00054679 | 0.00056771 | 0.00059194 | 0.00062108 | 0.00065524 | 0.00069542 | 0.00074498 | 0.0008052 | 0.0008052 | 0.0008052 | 0.0008052 | 0.0008052 | 0.0008052 | 0.0008052 | 0.0008052
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Leakage | Rect | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| U2.3 |  | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| DoF | Rect | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Coupler Parameters Combined |  | 0.0058072 | 0.0058074 | 0.0058077 | 0.0058079 | 0.0058081 | 0.0058082 | 0.0058084 | 0.0058085 | 0.0058087 | 0.0058088 | 0.0058089 | 0.0058089 | 0.0058089 | 0.0058089 | 0.005809 | 0.005809 | 0.005809 | 0.005809 | 0.00086311 | 0.00086324 | 0.00086336 | 0.00086348 | 0.0008636 | 0.00086371 | 0.00086382 | 0.00086393 | 0.00086403 | 0.00086413 | 0.00086423 | 0.00086432 | 0.00086441 | 0.0008645 | 0.00086459 | 0.00086467 | 0.00086475 | 0.00086483 | 0.00086491 | 0.00086498 | 0.00086505 | 0.00086512 | 0.00086519 | 0.00086525 | 0.00086531 | 0.00086537 | 0.00086543 | 0.00086548 | 0.00086553 | 0.00086558 | 0.00086563 | 0.00086567 | 0.00086571 | 0.00064514 | 0.00064518 | 0.00064523 | 0.00064527 | 0.0006453 | 0.00064533 | 0.00064536 | 0.00064538 | 0.00064539 | 0.00064541 | 0.00064541 | 0.0006454 | 0.00064539 | 0.00064536 | 0.00064533 | 0.0006453 | 0.00064524 | 0.00064517 | 0.0006451 | 0.00064499 | 0.00064488 | 0.00064475 | 0.00064459 | 0.00064442 | 0.00064422 | 0.00064398 | 0.00064372 | 0.00064344 | 0.00064309 | 0.00064272 | 0.00064231 | 0.00064182 | 0.0006413 | 0.00064073 | 0.00064005 | 0.00063935 | 0.00063859 | 0.00063769 | 0.00063677 | 0.00063578 | 0.00063463 | 0.0006335 | 0.00063231 | 0.00063097 | 0.00062973 | 0.00062853 | 0.00062725 | 0.00062631 | 0.00062565 | 0.00062522 | 0.0006257 | 0.00062712 | 0.00062957 | 0.00063431 | 0.00064154 | 0.00065177 | 0.00066724 | 0.00072042 | 0.00080757 | 0.00092865 | 0.0010936 | 0.0010936 | 0.0010936 | 0.0010936 | 0.0010936 | 0.0010936 | 0.0010936 | 0.0010936
| EffDoF |  | 10.2351 | 10.2369 | 10.2385 | 10.24 | 10.2414 | 10.2426 | 10.2437 | 10.2446 | 10.2455 | 10.2463 | 10.247 | 10.2472 | 10.2474 | 10.2475 | 10.2477 | 10.2478 | 10.248 | 10.2481 | 83.5056 | 83.5112 | 83.5167 | 83.5224 | 83.5274 | 83.5324 | 83.5377 | 83.5423 | 83.5469 | 83.5517 | 83.5559 | 83.5601 | 83.5645 | 83.5684 | 83.5723 | 83.5764 | 83.58 | 83.5836 | 83.5875 | 83.5909 | 83.5943 | 83.5979 | 83.6011 | 83.6044 | 83.6079 | 83.611 | 83.6142 | 83.6177 | 83.6208 | 83.624 | 83.6276 | 83.6308 | 83.6342 | 57.485 | 57.4865 | 57.4885 | 57.491 | 57.4943 | 57.4984 | 57.5031 | 57.5088 | 57.5155 | 57.5232 | 57.5322 | 57.5426 | 57.5543 | 57.5679 | 57.5834 | 57.6007 | 57.6205 | 57.6429 | 57.668 | 57.6964 | 57.7283 | 57.7639 | 57.804 | 57.849 | 57.8987 | 57.9546 | 58.0169 | 58.0855 | 58.162 | 58.2467 | 58.3393 | 58.4416 | 58.5536 | 58.6744 | 58.8056 | 58.9463 | 59.0942 | 59.2493 | 59.4082 | 59.5649 | 59.7145 | 59.8468 | 59.9471 | 59.9979 | 59.9729 | 59.8393 | 59.5524 | 59.0579 | 58.2994 | 57.1993 | 55.6955 | 53.7578 | 51.3499 | 48.5304 | 45.436 | 42.191 | 39.0352 | 42.0068 | 48.6478 | 55.4304 | 59.5776 | 59.5776 | 59.5776 | 59.5776 | 59.5776 | 59.5776 | 59.5776 | 59.5776
Ambient Conditions
| Static Pressure | Rect | 0.0019044 | 0.0018917 | 0.0018797 | 0.0018691 | 0.0018595 | 0.0018506 | 0.0018427 | 0.0018355 | 0.0018287 | 0.0018227 | 0.0018171 | 0.0018155 | 0.0018144 | 0.0018131 | 0.0018116 | 0.0018105 | 0.0018093 | 0.0018078 | 0.0018068 | 0.0018056 | 0.0018042 | 0.0018032 | 0.001802 | 0.0018006 | 0.0017997 | 0.0017985 | 0.0017972 | 0.0017962 | 0.001795 | 0.0017937 | 0.0017928 | 0.0017916 | 0.0017903 | 0.0017894 | 0.0017882 | 0.0017868 | 0.0017859 | 0.0017848 | 0.0017834 | 0.0017825 | 0.0017814 | 0.00178 | 0.0017792 | 0.001778 | 0.0017767 | 0.0017759 | 0.0017748 | 0.0017736 | 0.0017728 | 0.0017718 | 0.0017707 | 0.0017701 | 0.0017692 | 0.0017683 | 0.0017679 | 0.0017673 | 0.0017666 | 0.0017665 | 0.0017662 | 0.001766 | 0.0017663 | 0.0017665 | 0.0017668 | 0.0017677 | 0.0017686 | 0.0017695 | 0.0017712 | 0.0017729 | 0.0017747 | 0.0017773 | 0.00178 | 0.0017828 | 0.0017866 | 0.0017904 | 0.0017944 | 0.0017994 | 0.0018045 | 0.0018098 | 0.0018162 | 0.0018227 | 0.0018295 | 0.0018375 | 0.0018456 | 0.0018543 | 0.0018643 | 0.0018747 | 0.001886 | 0.0018991 | 0.0019132 | 0.0019289 | 0.0019473 | 0.0019677 | 0.0019908 | 0.002018 | 0.0020484 | 0.002083 | 0.0021229 | 0.002167 | 0.0022154 | 0.0022689 | 0.0023244 | 0.0023808 | 0.0024367 | 0.0024859 | 0.0025241 | 0.0025469 | 0.0025438 | 0.002507 | 0.0024304 | 0.0023002 | 0.0021093 | 0.0018552 | 0.0014658 | 0.0011328 | 0.00077564 | 0.00041662 | 8.6737e-05 | 0.00018554 | 0.0003804
| U3.1 |  | 0.0010995 | 0.0010921 | 0.0010853 | 0.0010791 | 0.0010736 | 0.0010684 | 0.0010639 | 0.0010597 | 0.0010558 | 0.0010524 | 0.0010491 | 0.0010482 | 0.0010475 | 0.0010468 | 0.0010459 | 0.0010453 | 0.0010446 | 0.0010438 | 0.0010432 | 0.0010425 | 0.0010417 | 0.0010411 | 0.0010404 | 0.0010396 | 0.0010391 | 0.0010384 | 0.0010376 | 0.001037 | 0.0010364 | 0.0010356 | 0.0010351 | 0.0010344 | 0.0010336 | 0.0010331 | 0.0010324 | 0.0010316 | 0.0010311 | 0.0010304 | 0.0010297 | 0.0010291 | 0.0010285 | 0.0010277 | 0.0010272 | 0.0010265 | 0.0010258 | 0.0010253 | 0.0010247 | 0.001024 | 0.0010235 | 0.001023 | 0.0010223 | 0.001022 | 0.0010215 | 0.0010209 | 0.0010207 | 0.0010203 | 0.0010199 | 0.0010199 | 0.0010197 | 0.0010196 | 0.0010198 | 0.0010199 | 0.00102 | 0.0010206 | 0.0010211 | 0.0010216 | 0.0010226 | 0.0010236 | 0.0010246 | 0.0010261 | 0.0010277 | 0.0010293 | 0.0010315 | 0.0010337 | 0.001036 | 0.0010389 | 0.0010418 | 0.0010449 | 0.0010486 | 0.0010523 | 0.0010563 | 0.0010609 | 0.0010656 | 0.0010706 | 0.0010764 | 0.0010824 | 0.0010889 | 0.0010965 | 0.0011046 | 0.0011137 | 0.0011243 | 0.001136 | 0.0011494 | 0.0011651 | 0.0011827 | 0.0012026 | 0.0012257 | 0.0012511 | 0.0012791 | 0.0013099 | 0.001342 | 0.0013745 | 0.0014068 | 0.0014352 | 0.0014573 | 0.0014705 | 0.0014686 | 0.0014474 | 0.0014032 | 0.001328 | 0.0012178 | 0.0010711 | 0.0008463 | 0.00065403 | 0.00044782 | 0.00024053 | 5.0078e-05 | 0.00010712 | 0.00021962
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Drift in Pressure | Rect | 0.00076164 | 0.00075652 | 0.00075176 | 0.00074752 | 0.00074368 | 0.00074009 | 0.00073695 | 0.00073406 | 0.00073133 | 0.00072896 | 0.0007267 | 0.00072609 | 0.00072563 | 0.00072511 | 0.00072451 | 0.00072409 | 0.00072358 | 0.00072301 | 0.0007226 | 0.00072211 | 0.00072155 | 0.00072116 | 0.00072068 | 0.00072013 | 0.00071975 | 0.00071927 | 0.00071873 | 0.00071836 | 0.00071789 | 0.00071735 | 0.00071698 | 0.00071652 | 0.00071598 | 0.00071561 | 0.00071515 | 0.00071461 | 0.00071425 | 0.00071378 | 0.00071324 | 0.00071288 | 0.00071242 | 0.00071189 | 0.00071154 | 0.00071108 | 0.00071056 | 0.00071023 | 0.00070979 | 0.0007093 | 0.000709 | 0.0007086 | 0.00070815 | 0.00070791 | 0.00070756 | 0.00070719 | 0.00070703 | 0.00070678 | 0.00070651 | 0.00070648 | 0.00070637 | 0.00070626 | 0.0007064 | 0.00070648 | 0.00070658 | 0.00070696 | 0.0007073 | 0.00070768 | 0.00070836 | 0.00070902 | 0.00070975 | 0.00071081 | 0.00071187 | 0.00071301 | 0.0007145 | 0.00071602 | 0.00071764 | 0.00071963 | 0.00072166 | 0.00072381 | 0.00072635 | 0.00072895 | 0.00073169 | 0.00073487 | 0.00073813 | 0.0007416 | 0.00074559 | 0.00074977 | 0.00075429 | 0.00075953 | 0.00076517 | 0.00077144 | 0.0007788 | 0.00078694 | 0.00079618 | 0.00080706 | 0.00081924 | 0.00083306 | 0.00084905 | 0.00086667 | 0.00088603 | 0.00090742 | 0.00092964 | 0.00095218 | 0.00097454 | 0.00099421 | 0.0010095 | 0.0010186 | 0.0010174 | 0.0010027 | 0.00097201 | 0.00091993 | 0.00084361 | 0.00074193 | 0.00058621 | 0.000453 | 0.00031013 | 0.00016652 | 3.4566e-05 | 7.4086e-05 | 0.00015203
| U3.2 |  | 0.00043973 | 0.00043678 | 0.00043403 | 0.00043158 | 0.00042936 | 0.00042729 | 0.00042548 | 0.00042381 | 0.00042223 | 0.00042087 | 0.00041956 | 0.00041921 | 0.00041894 | 0.00041864 | 0.0004183 | 0.00041805 | 0.00041776 | 0.00041743 | 0.00041719 | 0.00041691 | 0.00041659 | 0.00041636 | 0.00041608 | 0.00041577 | 0.00041555 | 0.00041527 | 0.00041496 | 0.00041475 | 0.00041447 | 0.00041416 | 0.00041395 | 0.00041368 | 0.00041337 | 0.00041316 | 0.00041289 | 0.00041258 | 0.00041237 | 0.0004121 | 0.00041179 | 0.00041158 | 0.00041131 | 0.00041101 | 0.00041081 | 0.00041054 | 0.00041024 | 0.00041005 | 0.0004098 | 0.00040951 | 0.00040934 | 0.00040911 | 0.00040885 | 0.00040871 | 0.00040851 | 0.00040829 | 0.0004082 | 0.00040806 | 0.00040791 | 0.00040789 | 0.00040782 | 0.00040776 | 0.00040784 | 0.00040789 | 0.00040795 | 0.00040816 | 0.00040836 | 0.00040858 | 0.00040897 | 0.00040936 | 0.00040978 | 0.00041038 | 0.000411 | 0.00041166 | 0.00041252 | 0.0004134 | 0.00041433 | 0.00041548 | 0.00041665 | 0.00041789 | 0.00041936 | 0.00042086 | 0.00042244 | 0.00042428 | 0.00042616 | 0.00042816 | 0.00043047 | 0.00043288 | 0.00043549 | 0.00043851 | 0.00044177 | 0.00044539 | 0.00044964 | 0.00045434 | 0.00045967 | 0.00046595 | 0.00047299 | 0.00048097 | 0.0004902 | 0.00050037 | 0.00051155 | 0.0005239 | 0.00053673 | 0.00054974 | 0.00056265 | 0.00057401 | 0.00058284 | 0.00058812 | 0.00058738 | 0.0005789 | 0.00056119 | 0.00053112 | 0.00048706 | 0.00042835 | 0.00033845 | 0.00026154 | 0.00017905 | 9.6139e-05 | 1.9956e-05 | 4.2774e-05 | 8.7775e-05
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Temperature | Rect | 0.0015203 | 0.0014377 | 0.0013638 | 0.001303 | 0.0012387 | 0.0011831 | 0.0011412 | 0.0010913 | 0.0010498 | 0.001022 | 0.00098275 | 0.00097345 | 0.00097467 | 0.00096032 | 0.00095153 | 0.00095347 | 0.00093946 | 0.00093113 | 0.00093372 | 0.00092003 | 0.00091212 | 0.00091531 | 0.00090191 | 0.00089439 | 0.00089813 | 0.000885 | 0.00087785 | 0.00088211 | 0.00086925 | 0.00086245 | 0.00086721 | 0.00085462 | 0.00084819 | 0.00085345 | 0.00084115 | 0.0008351 | 0.00084088 | 0.00082893 | 0.00082333 | 0.00082969 | 0.00081817 | 0.00081311 | 0.00082015 | 0.0008092 | 0.00080482 | 0.00081271 | 0.00080252 | 0.00079903 | 0.000808 | 0.00079883 | 0.00079654 | 0.00080691 | 0.00079915 | 0.00079845 | 0.00081066 | 0.00080768 | 0.0008062 | 0.0008213 | 0.00082431 | 0.00082163 | 0.00084456 | 0.00085128 | 0.00085155 | 0.00088015 | 0.00089167 | 0.00089708 | 0.00093158 | 0.00094916 | 0.00096107 | 0.0010029 | 0.0010281 | 0.001048 | 0.001099 | 0.0011334 | 0.0011633 | 0.0012255 | 0.0012714 | 0.0013134 | 0.0013898 | 0.0014501 | 0.0015078 | 0.0016028 | 0.0016824 | 0.0017616 | 0.0018828 | 0.0019906 | 0.002102 | 0.0022633 | 0.0024152 | 0.0025782 | 0.0028039 | 0.0030275 | 0.0032745 | 0.0036038 | 0.0039423 | 0.0043214 | 0.0048093 | 0.0053189 | 0.0058885 | 0.0065949 | 0.0073282 | 0.0081304 | 0.0090889 | 0.010052 | 0.011067 | 0.012236 | 0.013347 | 0.014464 | 0.015739 | 0.016894 | 0.018093 | 0.019679 | 0.019679 | 0.019679 | 0.019679 | 0.019679 | 0.019679 | 0.019679 | 0.019679
| U3.3 |  | 0.00087774 | 0.00083008 | 0.00078737 | 0.00075231 | 0.00071515 | 0.00068306 | 0.00065887 | 0.00063005 | 0.00060613 | 0.00059007 | 0.00056739 | 0.00056202 | 0.00056273 | 0.00055444 | 0.00054937 | 0.00055048 | 0.0005424 | 0.00053759 | 0.00053908 | 0.00053118 | 0.00052661 | 0.00052845 | 0.00052072 | 0.00051638 | 0.00051854 | 0.00051096 | 0.00050683 | 0.00050929 | 0.00050186 | 0.00049794 | 0.00050069 | 0.00049341 | 0.0004897 | 0.00049274 | 0.00048564 | 0.00048215 | 0.00048548 | 0.00047858 | 0.00047535 | 0.00047902 | 0.00047237 | 0.00046945 | 0.00047352 | 0.00046719 | 0.00046466 | 0.00046922 | 0.00046333 | 0.00046132 | 0.0004665 | 0.00046121 | 0.00045988 | 0.00046587 | 0.00046139 | 0.00046099 | 0.00046804 | 0.00046632 | 0.00046546 | 0.00047418 | 0.00047592 | 0.00047437 | 0.0004876 | 0.00049149 | 0.00049164 | 0.00050816 | 0.0005148 | 0.00051793 | 0.00053785 | 0.000548 | 0.00055487 | 0.00057905 | 0.00059356 | 0.00060507 | 0.00063452 | 0.00065439 | 0.0006716 | 0.00070755 | 0.00073402 | 0.00075829 | 0.0008024 | 0.0008372 | 0.00087052 | 0.00092535 | 0.00097131 | 0.0010171 | 0.0010871 | 0.0011492 | 0.0012136 | 0.0013067 | 0.0013944 | 0.0014885 | 0.0016188 | 0.0017479 | 0.0018905 | 0.0020806 | 0.0022761 | 0.0024949 | 0.0027767 | 0.0030709 | 0.0033997 | 0.0038076 | 0.004231 | 0.0046941 | 0.0052475 | 0.0058037 | 0.0063895 | 0.0070647 | 0.0077058 | 0.0083506 | 0.009087 | 0.0097535 | 0.010446 | 0.011362 | 0.011362 | 0.011362 | 0.011362 | 0.011362 | 0.011362 | 0.011362 | 0.011362
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Drift in Temperature | Rect | 0.00060802 | 0.00057497 | 0.00054536 | 0.00052103 | 0.00049527 | 0.00047303 | 0.00045623 | 0.00043626 | 0.00041968 | 0.00040851 | 0.00039281 | 0.00038909 | 0.00038955 | 0.00038383 | 0.00038032 | 0.00038106 | 0.00037548 | 0.00037215 | 0.00037316 | 0.0003677 | 0.00036454 | 0.00036578 | 0.00036045 | 0.00035745 | 0.00035891 | 0.00035368 | 0.00035083 | 0.0003525 | 0.00034737 | 0.00034466 | 0.00034653 | 0.00034152 | 0.00033895 | 0.00034102 | 0.00033612 | 0.00033371 | 0.00033599 | 0.00033123 | 0.000329 | 0.00033151 | 0.00032693 | 0.00032491 | 0.00032769 | 0.00032334 | 0.00032159 | 0.00032471 | 0.00032066 | 0.00031927 | 0.00032282 | 0.00031919 | 0.00031827 | 0.00032239 | 0.00031931 | 0.00031904 | 0.00032388 | 0.0003227 | 0.00032213 | 0.00032813 | 0.00032935 | 0.0003283 | 0.00033743 | 0.00034014 | 0.00034028 | 0.00035167 | 0.00035629 | 0.0003585 | 0.00037224 | 0.00037929 | 0.00038409 | 0.00040079 | 0.00041086 | 0.00041888 | 0.00043922 | 0.00045301 | 0.00046498 | 0.00048983 | 0.00050819 | 0.00052504 | 0.00055555 | 0.00057968 | 0.00060281 | 0.00064075 | 0.00067261 | 0.00070436 | 0.0007528 | 0.00079591 | 0.00084056 | 0.00090503 | 0.00096579 | 0.0010311 | 0.0011213 | 0.0012108 | 0.0013097 | 0.0014413 | 0.0015768 | 0.0017285 | 0.0019236 | 0.0021275 | 0.0023553 | 0.0026379 | 0.0029311 | 0.0032518 | 0.0036352 | 0.0040204 | 0.004426 | 0.0048937 | 0.0053376 | 0.0057839 | 0.0062938 | 0.0067549 | 0.0072336 | 0.0078671 | 0.0037285 | 0.0033397 | 0.002888 | 0.0023749 | 0.0019004 | 0.00157 | 0.0014177
| U3.4 |  | 0.00035104 | 0.00033196 | 0.00031486 | 0.00030081 | 0.00028594 | 0.0002731 | 0.0002634 | 0.00025187 | 0.0002423 | 0.00023585 | 0.00022679 | 0.00022464 | 0.00022491 | 0.00022161 | 0.00021958 | 0.00022001 | 0.00021678 | 0.00021486 | 0.00021544 | 0.00021229 | 0.00021047 | 0.00021119 | 0.0002081 | 0.00020637 | 0.00020722 | 0.0002042 | 0.00020255 | 0.00020351 | 0.00020056 | 0.00019899 | 0.00020007 | 0.00019718 | 0.00019569 | 0.00019689 | 0.00019406 | 0.00019267 | 0.00019398 | 0.00019124 | 0.00018995 | 0.0001914 | 0.00018875 | 0.00018759 | 0.00018919 | 0.00018668 | 0.00018567 | 0.00018747 | 0.00018513 | 0.00018433 | 0.00018638 | 0.00018428 | 0.00018376 | 0.00018613 | 0.00018435 | 0.0001842 | 0.00018699 | 0.00018631 | 0.00018598 | 0.00018945 | 0.00019015 | 0.00018955 | 0.00019482 | 0.00019638 | 0.00019646 | 0.00020304 | 0.00020571 | 0.00020698 | 0.00021491 | 0.00021898 | 0.00022176 | 0.0002314 | 0.00023721 | 0.00024184 | 0.00025359 | 0.00026154 | 0.00026846 | 0.0002828 | 0.0002934 | 0.00030313 | 0.00032075 | 0.00033468 | 0.00034803 | 0.00036994 | 0.00038833 | 0.00040666 | 0.00043463 | 0.00045952 | 0.0004853 | 0.00052252 | 0.0005576 | 0.00059529 | 0.00064737 | 0.00069904 | 0.00075613 | 0.00083214 | 0.00091034 | 0.00099794 | 0.0011106 | 0.0012283 | 0.0013598 | 0.001523 | 0.0016923 | 0.0018774 | 0.0020988 | 0.0023212 | 0.0025554 | 0.0028254 | 0.0030817 | 0.0033393 | 0.0036337 | 0.0038999 | 0.0041763 | 0.0045421 | 0.0021526 | 0.0019282 | 0.0016674 | 0.0013711 | 0.0010972 | 0.00090643 | 0.00081852
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Humidity | Rect | 0.00017315 | 0.00019649 | 0.00021118 | 0.00022633 | 0.00024548 | 0.00025449 | 0.00026453 | 0.00028059 | 0.00028546 | 0.0002918 | 0.00030561 | 0.00030268 | 0.00030233 | 0.00031065 | 0.00030752 | 0.00030696 | 0.00031528 | 0.00031196 | 0.00031122 | 0.00031953 | 0.00031604 | 0.00031513 | 0.00032344 | 0.00031979 | 0.00031873 | 0.00032704 | 0.00032325 | 0.00032204 | 0.00033035 | 0.00032643 | 0.0003251 | 0.00033342 | 0.00032938 | 0.00032793 | 0.00033626 | 0.00033212 | 0.00033057 | 0.0003389 | 0.00033467 | 0.00033303 | 0.00034139 | 0.00033708 | 0.00033537 | 0.00034375 | 0.00033938 | 0.00033761 | 0.00034602 | 0.00034161 | 0.0003398 | 0.00034828 | 0.00034384 | 0.00034201 | 0.00035057 | 0.00034614 | 0.00034432 | 0.00035299 | 0.0003486 | 0.00034683 | 0.00035567 | 0.00035136 | 0.00034969 | 0.00035876 | 0.0003546 | 0.00035309 | 0.0003625 | 0.00035857 | 0.00035733 | 0.0003672 | 0.00036363 | 0.00036278 | 0.00037332 | 0.00037028 | 0.00037001 | 0.00038149 | 0.00037922 | 0.00037979 | 0.0003926 | 0.00039145 | 0.00039323 | 0.00040796 | 0.00040841 | 0.00041195 | 0.0004294 | 0.00043217 | 0.00043826 | 0.00045963 | 0.00046576 | 0.00047555 | 0.00050257 | 0.00051363 | 0.00052884 | 0.00056415 | 0.00058248 | 0.00060577 | 0.00065338 | 0.00068269 | 0.00071827 | 0.00078453 | 0.00083084 | 0.00088568 | 0.00098113 | 0.0010548 | 0.0011412 | 0.0012845 | 0.0014049 | 0.0015464 | 0.0017736 | 0.0019806 | 0.0022285 | 0.0026197 | 0.0030127 | 0.0035089 | 0.0035089 | 0.0035089 | 0.0035089 | 0.0035089 | 0.0035089 | 0.0035089 | 0.0035089
| U3.5 |  | 9.997e-05 | 0.00011344 | 0.00012192 | 0.00013067 | 0.00014173 | 0.00014693 | 0.00015273 | 0.000162 | 0.00016481 | 0.00016847 | 0.00017645 | 0.00017475 | 0.00017455 | 0.00017935 | 0.00017755 | 0.00017723 | 0.00018203 | 0.00018011 | 0.00017968 | 0.00018448 | 0.00018246 | 0.00018194 | 0.00018674 | 0.00018463 | 0.00018402 | 0.00018882 | 0.00018663 | 0.00018593 | 0.00019073 | 0.00018847 | 0.0001877 | 0.0001925 | 0.00019017 | 0.00018933 | 0.00019414 | 0.00019175 | 0.00019085 | 0.00019566 | 0.00019322 | 0.00019228 | 0.0001971 | 0.00019461 | 0.00019362 | 0.00019846 | 0.00019594 | 0.00019492 | 0.00019978 | 0.00019723 | 0.00019618 | 0.00020108 | 0.00019852 | 0.00019746 | 0.0002024 | 0.00019984 | 0.00019879 | 0.0002038 | 0.00020126 | 0.00020024 | 0.00020534 | 0.00020286 | 0.00020189 | 0.00020713 | 0.00020473 | 0.00020386 | 0.00020929 | 0.00020702 | 0.0002063 | 0.000212 | 0.00020994 | 0.00020945 | 0.00021554 | 0.00021378 | 0.00021362 | 0.00022025 | 0.00021894 | 0.00021927 | 0.00022667 | 0.000226 | 0.00022703 | 0.00023554 | 0.0002358 | 0.00023784 | 0.00024792 | 0.00024952 | 0.00025303 | 0.00026537 | 0.00026891 | 0.00027456 | 0.00029016 | 0.00029654 | 0.00030533 | 0.00032571 | 0.0003363 | 0.00034974 | 0.00037723 | 0.00039415 | 0.00041469 | 0.00045295 | 0.00047969 | 0.00051135 | 0.00056646 | 0.00060901 | 0.00065889 | 0.0007416 | 0.0008111 | 0.0008928 | 0.001024 | 0.0011435 | 0.0012866 | 0.0015125 | 0.0017394 | 0.0020259 | 0.0020259 | 0.0020259 | 0.0020259 | 0.0020259 | 0.0020259 | 0.0020259 | 0.0020259
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Drift in Humidity | Rect | 6.913e-05 | 7.8456e-05 | 8.4337e-05 | 9.0404e-05 | 9.8056e-05 | 0.00010166 | 0.00010569 | 0.0001121 | 0.00011405 | 0.00011659 | 0.00012211 | 0.00012094 | 0.00012081 | 0.00012413 | 0.00012288 | 0.00012266 | 0.00012598 | 0.00012465 | 0.00012436 | 0.00012768 | 0.00012629 | 0.00012593 | 0.00012924 | 0.00012779 | 0.00012737 | 0.00013068 | 0.00012917 | 0.00012869 | 0.00013201 | 0.00013045 | 0.00012992 | 0.00013324 | 0.00013162 | 0.00013105 | 0.00013437 | 0.00013272 | 0.0001321 | 0.00013543 | 0.00013374 | 0.00013309 | 0.00013642 | 0.0001347 | 0.00013402 | 0.00013737 | 0.00013562 | 0.00013492 | 0.00013828 | 0.00013652 | 0.0001358 | 0.00013918 | 0.00013741 | 0.00013668 | 0.0001401 | 0.00013833 | 0.0001376 | 0.00014107 | 0.00013931 | 0.00013861 | 0.00014214 | 0.00014042 | 0.00013975 | 0.00014337 | 0.00014171 | 0.00014111 | 0.00014487 | 0.0001433 | 0.00014281 | 0.00014675 | 0.00014533 | 0.00014499 | 0.0001492 | 0.00014799 | 0.00014788 | 0.00015247 | 0.00015156 | 0.00015179 | 0.00015691 | 0.00015646 | 0.00015717 | 0.00016306 | 0.00016324 | 0.00016466 | 0.00017164 | 0.00017275 | 0.00017519 | 0.00018373 | 0.00018619 | 0.00019011 | 0.00020091 | 0.00020534 | 0.00021143 | 0.00022555 | 0.00023289 | 0.00024221 | 0.00026125 | 0.00027298 | 0.00028722 | 0.00031372 | 0.00033226 | 0.0003542 | 0.00039239 | 0.00042188 | 0.00045645 | 0.00051376 | 0.00056193 | 0.00061855 | 0.00070939 | 0.00079218 | 0.0008913 | 0.0010477 | 0.0012049 | 0.0014032 | 0.0005008 | 0.00054712 | 0.00060004 | 0.0006506 | 0.00072372 | 0.00081402 | 0.0009191
| U3.6 |  | 3.9912e-05 | 4.5297e-05 | 4.8692e-05 | 5.2195e-05 | 5.6612e-05 | 5.8695e-05 | 6.1018e-05 | 6.4721e-05 | 6.5848e-05 | 6.7316e-05 | 7.0502e-05 | 6.9827e-05 | 6.9748e-05 | 7.1665e-05 | 7.0944e-05 | 7.0818e-05 | 7.2734e-05 | 7.1969e-05 | 7.1801e-05 | 7.3716e-05 | 7.2912e-05 | 7.2704e-05 | 7.4619e-05 | 7.3779e-05 | 7.3535e-05 | 7.545e-05 | 7.4577e-05 | 7.4301e-05 | 7.6216e-05 | 7.5313e-05 | 7.5007e-05 | 7.6923e-05 | 7.5994e-05 | 7.5661e-05 | 7.7579e-05 | 7.6625e-05 | 7.627e-05 | 7.819e-05 | 7.7215e-05 | 7.6839e-05 | 7.8764e-05 | 7.7772e-05 | 7.7378e-05 | 7.9309e-05 | 7.8302e-05 | 7.7895e-05 | 7.9836e-05 | 7.8819e-05 | 7.8402e-05 | 8.0356e-05 | 7.9333e-05 | 7.8913e-05 | 8.0885e-05 | 7.9864e-05 | 7.9446e-05 | 8.1445e-05 | 8.0432e-05 | 8.0025e-05 | 8.2062e-05 | 8.107e-05 | 8.0686e-05 | 8.2777e-05 | 8.1818e-05 | 8.1473e-05 | 8.364e-05 | 8.2736e-05 | 8.2451e-05 | 8.4727e-05 | 8.3905e-05 | 8.371e-05 | 8.614e-05 | 8.544e-05 | 8.5379e-05 | 8.8026e-05 | 8.7505e-05 | 8.7638e-05 | 9.0595e-05 | 9.033e-05 | 9.0745e-05 | 9.4142e-05 | 9.4248e-05 | 9.5069e-05 | 9.9095e-05 | 9.9738e-05 | 0.00010115 | 0.00010608 | 0.0001075 | 0.00010976 | 0.000116 | 0.00011855 | 0.00012207 | 0.00013022 | 0.00013446 | 0.00013984 | 0.00015083 | 0.00015761 | 0.00016583 | 0.00018113 | 0.00019183 | 0.0002045 | 0.00022654 | 0.00024357 | 0.00026353 | 0.00029662 | 0.00032443 | 0.00035712 | 0.00040957 | 0.00045737 | 0.00051459 | 0.00060491 | 0.00069563 | 0.00081017 | 0.00028914 | 0.00031588 | 0.00034643 | 0.00037563 | 0.00041784 | 0.00046998 | 0.00053064
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Ambient Conditions Combined |  | 0.0015191 | 0.0014825 | 0.00145 | 0.0014238 | 0.0013977 | 0.0013749 | 0.0013577 | 0.0013392 | 0.0013231 | 0.001312 | 0.0012985 | 0.0012947 | 0.0012944 | 0.0012903 | 0.0012867 | 0.0012866 | 0.0012827 | 0.0012792 | 0.0012793 | 0.0012756 | 0.0012723 | 0.0012726 | 0.001269 | 0.0012659 | 0.0012663 | 0.0012628 | 0.0012598 | 0.0012603 | 0.001257 | 0.0012541 | 0.0012547 | 0.0012516 | 0.0012487 | 0.0012494 | 0.0012464 | 0.0012437 | 0.0012445 | 0.0012416 | 0.001239 | 0.00124 | 0.0012372 | 0.0012347 | 0.0012359 | 0.0012333 | 0.001231 | 0.0012324 | 0.0012301 | 0.0012281 | 0.0012297 | 0.0012278 | 0.0012261 | 0.0012281 | 0.0012266 | 0.0012255 | 0.0012281 | 0.001228 | 0.0012267 | 0.0012304 | 0.001232 | 0.0012306 | 0.0012366 | 0.0012395 | 0.0012393 | 0.0012473 | 0.001252 | 0.0012536 | 0.001264 | 0.0012711 | 0.0012751 | 0.0012888 | 0.001299 | 0.0013064 | 0.0013244 | 0.0013388 | 0.0013504 | 0.0013741 | 0.0013941 | 0.0014116 | 0.0014429 | 0.0014703 | 0.0014959 | 0.0015376 | 0.0015756 | 0.0016129 | 0.0016699 | 0.0017241 | 0.0017798 | 0.0018608 | 0.0019412 | 0.0020275 | 0.002148 | 0.0022724 | 0.0024103 | 0.0025958 | 0.0027923 | 0.003013 | 0.0032993 | 0.0036045 | 0.0039461 | 0.0043716 | 0.0048193 | 0.0053086 | 0.0058941 | 0.0064883 | 0.0071121 | 0.0078311 | 0.0085202 | 0.0092104 | 0.0099991 | 0.010726 | 0.01148 | 0.012482 | 0.011779 | 0.011726 | 0.011676 | 0.011631 | 0.011601 | 0.011587 | 0.011585
| EffDoF |  | 75.7874 | 74.4513 | 72.9822 | 71.6822 | 70.1826 | 68.7093 | 67.6178 | 66.2869 | 65.0416 | 64.2722 | 63.1958 | 62.8393 | 62.9033 | 62.5484 | 62.2037 | 62.2887 | 61.9465 | 61.613 | 61.7173 | 61.3875 | 61.0646 | 61.1866 | 60.8688 | 60.5561 | 60.6944 | 60.3885 | 60.0858 | 60.2394 | 59.9454 | 59.6528 | 59.8212 | 59.5396 | 59.2576 | 59.4411 | 59.173 | 58.9027 | 59.1022 | 58.8494 | 58.593 | 58.8109 | 58.5763 | 58.3372 | 58.5771 | 58.3651 | 58.1483 | 58.416 | 58.233 | 58.0458 | 58.3496 | 58.2047 | 58.057 | 58.4088 | 58.3141 | 58.2198 | 58.6356 | 58.7131 | 58.585 | 59.1036 | 59.3991 | 59.2192 | 60.0251 | 60.4567 | 60.376 | 61.3941 | 62.0033 | 62.104 | 63.3395 | 64.1722 | 64.502 | 65.9913 | 67.086 | 67.6839 | 69.4279 | 70.7888 | 71.6539 | 73.5652 | 75.1087 | 76.1423 | 77.9671 | 79.4471 | 80.3748 | 81.6268 | 82.5827 | 82.9161 | 82.9427 | 82.7919 | 81.9743 | 80.3121 | 78.7219 | 76.5708 | 73.4893 | 70.8392 | 67.9316 | 64.4713 | 61.7107 | 58.9974 | 56.1496 | 54.012 | 52.0315 | 50.1106 | 48.7339 | 47.4846 | 46.3089 | 45.5111 | 44.7701 | 44.0693 | 43.6493 | 43.2357 | 42.8442 | 42.7377 | 42.6301 | 42.5706 | 34.5735 | 33.9773 | 33.4086 | 32.9069 | 32.568 | 32.4121 | 32.3891
Electrical
| Cross Talk, Noise and Distortion | Rect | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458 | 0.0033458
| U4.1 |  | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317 | 0.0019317
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Voltage Ratio | Rect | 0.0092184 | 0.0092185 | 0.0092186 | 0.0092187 | 0.013544 | 0.013545 | 0.013545 | 0.0067498 | 0.0067503 | 0.0067509 | 0.0067517 | 0.0067519 | 0.0067521 | 0.0067523 | 0.0067526 | 0.0067528 | 0.0067531 | 0.0067534 | 0.0069992 | 0.0069995 | 0.0069998 | 0.0070001 | 0.0070005 | 0.0070009 | 0.0070012 | 0.0070016 | 0.007002 | 0.0070025 | 0.0070029 | 0.0067825 | 0.006783 | 0.0067835 | 0.006784 | 0.0067845 | 0.0067851 | 0.0067856 | 0.0067862 | 0.0067868 | 0.0067873 | 0.0067879 | 0.0067885 | 0.006789 | 0.0067896 | 0.0067901 | 0.0067906 | 0.006791 | 0.0067914 | 0.0067918 | 0.0067921 | 0.0067923 | 0.0067924 | 0.0067924 | 0.0067922 | 0.0067919 | 0.0067913 | 0.0067906 | 0.0067896 | 0.0067883 | 0.0067867 | 0.0067847 | 0.0067823 | 0.0067794 | 0.0067681 | 0.006764 | 0.0067593 | 0.0067538 | 0.0067475 | 0.0067402 | 0.006732 | 0.0067226 | 0.006712 | 0.0067002 | 0.0066869 | 0.0066721 | 0.0066558 | 0.0066377 | 0.0066179 | 0.0065963 | 0.0066484 | 0.0066233 | 0.0065963 | 0.0065673 | 0.0065365 | 0.0065038 | 0.0064692 | 0.0064329 | 0.0063948 | 0.0063549 | 0.0063131 | 0.0062691 | 0.0062227 | 0.0061734 | 0.0061202 | 0.0060624 | 0.0059984 | 0.0059269 | 0.0058463 | 0.005755 | 0.0056521 | 0.0055374 | 0.005412 | 0.0052784 | 0.0051408 | 0.0050042 | 0.0048742 | 0.0047556 | 0.0049452 | 0.0048649 | 0.0048027 | 0.0047587 | 0.0047325 | 0.004724 | 0.0047328 | 0.0047588 | 0.0048018 | 0.0048631 | 0.0050191 | 0.0051407 | 0.0053313
| U4.2 |  | 0.0053223 | 0.0053223 | 0.0053224 | 0.0053224 | 0.0078199 | 0.00782 | 0.0078201 | 0.003897 | 0.0038973 | 0.0038976 | 0.0038981 | 0.0038982 | 0.0038983 | 0.0038985 | 0.0038986 | 0.0038988 | 0.0038989 | 0.0038991 | 0.004041 | 0.0040412 | 0.0040413 | 0.0040415 | 0.0040417 | 0.0040419 | 0.0040422 | 0.0040424 | 0.0040426 | 0.0040429 | 0.0040431 | 0.0039159 | 0.0039162 | 0.0039165 | 0.0039168 | 0.0039171 | 0.0039174 | 0.0039177 | 0.003918 | 0.0039183 | 0.0039187 | 0.003919 | 0.0039193 | 0.0039196 | 0.00392 | 0.0039203 | 0.0039205 | 0.0039208 | 0.003921 | 0.0039212 | 0.0039214 | 0.0039215 | 0.0039216 | 0.0039216 | 0.0039215 | 0.0039213 | 0.003921 | 0.0039206 | 0.00392 | 0.0039192 | 0.0039183 | 0.0039171 | 0.0039157 | 0.0039141 | 0.0039076 | 0.0039052 | 0.0039025 | 0.0038993 | 0.0038957 | 0.0038915 | 0.0038867 | 0.0038813 | 0.0038752 | 0.0038683 | 0.0038607 | 0.0038521 | 0.0038427 | 0.0038323 | 0.0038209 | 0.0038084 | 0.0038385 | 0.003824 | 0.0038083 | 0.0037916 | 0.0037738 | 0.003755 | 0.003735 | 0.003714 | 0.003692 | 0.003669 | 0.0036448 | 0.0036195 | 0.0035927 | 0.0035642 | 0.0035335 | 0.0035001 | 0.0034632 | 0.0034219 | 0.0033754 | 0.0033227 | 0.0032633 | 0.003197 | 0.0031246 | 0.0030475 | 0.002968 | 0.0028892 | 0.0028141 | 0.0027456 | 0.0028551 | 0.0028088 | 0.0027728 | 0.0027474 | 0.0027323 | 0.0027274 | 0.0027325 | 0.0027475 | 0.0027723 | 0.0028077 | 0.0028978 | 0.002968 | 0.003078
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Polarisation Voltage | Rect | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006 | 0.0010006
| U4.3 |  | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769 | 0.00057769
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Capacitance of Transmitter | Standard | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266 | 0.00093266
| U4.4 |  | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633 | 0.00046633
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Frequency | Rect | 0.00065 | 0.00031 | 0.00017 | 0.0001 | 5e-05 | 3e-05 | 2e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | 1e-05
| U4.5 |  | 0.00037528 | 0.00017898 | 9.815e-05 | 5.7735e-05 | 2.8868e-05 | 1.7321e-05 | 1.1547e-05 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06 | 5.7735e-06
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Electrical Combined |  | 0.0057227 | 0.0057133 | 0.0057114 | 0.0057109 | 0.0080892 | 0.0080892 | 0.0080893 | 0.0044124 | 0.0044127 | 0.004413 | 0.0044133 | 0.0044135 | 0.0044136 | 0.0044137 | 0.0044138 | 0.0044139 | 0.0044141 | 0.0044142 | 0.0045401 | 0.0045402 | 0.0045404 | 0.0045406 | 0.0045407 | 0.0045409 | 0.0045411 | 0.0045413 | 0.0045415 | 0.0045417 | 0.004542 | 0.0044291 | 0.0044293 | 0.0044296 | 0.0044299 | 0.0044301 | 0.0044304 | 0.0044307 | 0.004431 | 0.0044313 | 0.0044315 | 0.0044318 | 0.0044321 | 0.0044324 | 0.0044327 | 0.004433 | 0.0044332 | 0.0044334 | 0.0044336 | 0.0044338 | 0.004434 | 0.0044341 | 0.0044341 | 0.0044341 | 0.004434 | 0.0044339 | 0.0044336 | 0.0044332 | 0.0044327 | 0.0044321 | 0.0044312 | 0.0044302 | 0.004429 | 0.0044275 | 0.0044217 | 0.0044197 | 0.0044172 | 0.0044144 | 0.0044112 | 0.0044075 | 0.0044033 | 0.0043985 | 0.0043931 | 0.0043871 | 0.0043803 | 0.0043728 | 0.0043645 | 0.0043554 | 0.0043453 | 0.0043343 | 0.0043608 | 0.004348 | 0.0043343 | 0.0043196 | 0.004304 | 0.0042875 | 0.00427 | 0.0042517 | 0.0042325 | 0.0042124 | 0.0041914 | 0.0041693 | 0.0041461 | 0.0041214 | 0.0040949 | 0.0040661 | 0.0040344 | 0.003999 | 0.0039593 | 0.0039144 | 0.0038641 | 0.0038084 | 0.0037478 | 0.0036837 | 0.0036183 | 0.0035539 | 0.0034931 | 0.0034382 | 0.0035262 | 0.0034888 | 0.00346 | 0.0034396 | 0.0034276 | 0.0034237 | 0.0034277 | 0.0034397 | 0.0034596 | 0.003488 | 0.0035609 | 0.0036183 | 0.003709
| EffDoF |  | 39.4086 | 39.1479 | 39.0941 | 39.0788 | 34.2212 | 34.2205 | 34.2203 | 46.4685 | 46.4663 | 46.4636 | 46.4603 | 46.4593 | 46.4583 | 46.4573 | 46.4562 | 46.455 | 46.4538 | 46.4525 | 45.4016 | 45.4004 | 45.3991 | 45.3977 | 45.3963 | 45.3948 | 45.3933 | 45.3917 | 45.39 | 45.3883 | 45.3865 | 46.3231 | 46.321 | 46.3187 | 46.3165 | 46.3141 | 46.3117 | 46.3093 | 46.3068 | 46.3043 | 46.3018 | 46.2993 | 46.2968 | 46.2944 | 46.292 | 46.2897 | 46.2875 | 46.2855 | 46.2837 | 46.2821 | 46.2809 | 46.28 | 46.2796 | 46.2797 | 46.2804 | 46.2819 | 46.2842 | 46.2874 | 46.2918 | 46.2975 | 46.3046 | 46.3134 | 46.3242 | 46.337 | 46.387 | 46.4051 | 46.4262 | 46.4507 | 46.479 | 46.5115 | 46.5486 | 46.5909 | 46.6387 | 46.6926 | 46.7532 | 46.8209 | 46.8964 | 46.9801 | 47.0726 | 47.1744 | 46.9305 | 47.0475 | 47.1746 | 47.3122 | 47.4602 | 47.6191 | 47.7888 | 47.9697 | 48.162 | 48.3662 | 48.5834 | 48.8149 | 49.0634 | 49.3323 | 49.6268 | 49.954 | 50.3231 | 50.7456 | 51.2346 | 51.8041 | 52.4669 | 53.2318 | 54.0993 | 55.0581 | 56.0825 | 57.1331 | 58.1616 | 59.119 | 57.5967 | 58.2352 | 58.7363 | 59.0937 | 59.3066 | 59.3761 | 59.3041 | 59.0927 | 58.7434 | 58.2497 | 57.0169 | 56.083 | 54.6744
Error in Theory
| Heat Conduction and Thermal Diffusion | Rect | 0.10137 | 0.091692 | 0.082967 | 0.07494 | 0.067584 | 0.060839 | 0.054593 | 0.04893 | 0.043831 | 0.039183 | 0.035011 | 0.034057 | 0.033092 | 0.032165 | 0.031285 | 0.030395 | 0.02954 | 0.028729 | 0.027909 | 0.027122 | 0.026374 | 0.02562 | 0.024895 | 0.024207 | 0.023512 | 0.022846 | 0.022213 | 0.021574 | 0.02096 | 0.020378 | 0.019791 | 0.019227 | 0.018691 | 0.018151 | 0.017632 | 0.01714 | 0.016643 | 0.016166 | 0.015714 | 0.015257 | 0.014818 | 0.014402 | 0.013982 | 0.013578 | 0.013195 | 0.012808 | 0.012436 | 0.012083 | 0.011727 | 0.011384 | 0.011058 | 0.010729 | 0.010412 | 0.010111 | 0.0098061 | 0.0095126 | 0.0092329 | 0.00895 | 0.008677 | 0.0084161 | 0.0081521 | 0.0078967 | 0.0076519 | 0.0074038 | 0.007163 | 0.0069313 | 0.0066959 | 0.0064664 | 0.0062445 | 0.0060185 | 0.0057968 | 0.0055808 | 0.0053602 | 0.0051422 | 0.004928 | 0.0047083 | 0.0044892 | 0.0042716 | 0.0040473 | 0.0038212 | 0.0035939 | 0.0033584 | 0.0031181 | 0.0028733 | 0.0026181 | 0.0023545 | 0.0020821 | 0.0017964 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015
| U5.1 |  | 0.058525 | 0.052938 | 0.047901 | 0.043267 | 0.03902 | 0.035126 | 0.03152 | 0.02825 | 0.025306 | 0.022622 | 0.020214 | 0.019663 | 0.019106 | 0.01857 | 0.018062 | 0.017549 | 0.017055 | 0.016587 | 0.016113 | 0.015659 | 0.015227 | 0.014792 | 0.014373 | 0.013976 | 0.013575 | 0.01319 | 0.012824 | 0.012456 | 0.012101 | 0.011765 | 0.011426 | 0.0111 | 0.010791 | 0.010479 | 0.01018 | 0.0098958 | 0.0096091 | 0.0093337 | 0.0090724 | 0.0088087 | 0.0085554 | 0.008315 | 0.0080724 | 0.0078394 | 0.0076181 | 0.0073948 | 0.0071802 | 0.0069763 | 0.0067704 | 0.0065725 | 0.0063843 | 0.0061943 | 0.0060115 | 0.0058374 | 0.0056616 | 0.0054921 | 0.0053306 | 0.0051673 | 0.0050096 | 0.0048591 | 0.0047066 | 0.0045591 | 0.0044178 | 0.0042746 | 0.0041355 | 0.0040018 | 0.0038659 | 0.0037334 | 0.0036052 | 0.0034748 | 0.0033468 | 0.0032221 | 0.0030947 | 0.0029689 | 0.0028452 | 0.0027183 | 0.0025918 | 0.0024662 | 0.0023367 | 0.0022062 | 0.002075 | 0.001939 | 0.0018003 | 0.0016589 | 0.0015116 | 0.0013594 | 0.0012021 | 0.0010372 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603 | 0.00086603
| DoF |  | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
| Viscosity | Rect | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.5585e-05 | 0.00011213 | 0.00019507 | 0.00028628 | 0.00039448 | 0.00051213 | 0.00064266 | 0.00079539 | 0.00096241 | 0.0011495 | 0.0013673 | 0.0018398 | 0.0025133 | 0.0032535 | 0.0040552 | 0.004941 | 0.0059253 | 0.0069986 | 0.0081957 | 0.0095409 | 0.011019 | 0.012687 | 0.014591 | 0.016712 | 0.016712 | 0.016712 | 0.016712 | 0.016712 | 0.016712 | 0.016712 | 0.016712
| U5.2 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2.0545e-05 | 6.4737e-05 | 0.00011263 | 0.00016528 | 0.00022775 | 0.00029568 | 0.00037104 | 0.00045922 | 0.00055565 | 0.00066364 | 0.00078943 | 0.0010622 | 0.001451 | 0.0018784 | 0.0023413 | 0.0028527 | 0.003421 | 0.0040407 | 0.0047318 | 0.0055084 | 0.006362 | 0.0073249 | 0.0084241 | 0.009649 | 0.009649 | 0.009649 | 0.009649 | 0.009649 | 0.009649 | 0.009649 | 0.009649
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
| Radial Wave Motion | Rect | 9.9476e-14 | 3.0198e-13 | 4.0146e-13 | 6.9988e-13 | 1.1013e-12 | 1.7017e-12 | 2.7995e-12 | 4.2988e-12 | 6.9988e-12 | 1.1301e-11 | 1.7501e-11 | 1.9899e-11 | 2.2602e-11 | 2.4702e-11 | 2.8201e-11 | 3.2003e-11 | 3.4902e-11 | 3.9797e-11 | 4.54e-11 | 4.95e-11 | 5.63e-11 | 6.4201e-11 | 7.0301e-11 | 7.9901e-11 | 9.1099e-11 | 1.013e-10 | 1.133e-10 | 1.293e-10 | 1.468e-10 | 1.637e-10 | 1.881e-10 | 2.15e-10 | 2.407e-10 | 2.776e-10 | 3.187e-10 | 3.59e-10 | 4.16e-10 | 4.802e-10 | 5.448e-10 | 6.35e-10 | 7.381e-10 | 8.449e-10 | 9.917e-10 | 1.1619e-09 | 1.3436e-09 | 1.5898e-09 | 1.8794e-09 | 2.1979e-09 | 2.623e-09 | 3.1301e-09 | 3.7029e-09 | 4.4577e-09 | 5.3697e-09 | 6.4246e-09 | 7.7988e-09 | 9.4772e-09 | 1.1458e-08 | 1.4013e-08 | 1.7163e-08 | 2.0939e-08 | 2.5776e-08 | 3.178e-08 | 3.9067e-08 | 4.8347e-08 | 5.9931e-08 | 7.4122e-08 | 9.2119e-08 | 1.1468e-07 | 1.4251e-07 | 1.7769e-07 | 2.2192e-07 | 2.7677e-07 | 3.4596e-07 | 4.3315e-07 | 5.4169e-07 | 6.7836e-07 | 8.5089e-07 | 1.0663e-06 | 1.3372e-06 | 1.6796e-06 | 2.108e-06 | 2.6464e-06 | 3.3277e-06 | 4.1813e-06 | 5.2538e-06 | 6.612e-06 | 8.3158e-06 | 1.0456e-05 | 1.3169e-05 | 1.6576e-05 | 2.0856e-05 | 2.6285e-05 | 3.3109e-05 | 4.1685e-05 | 5.2569e-05 | 6.6264e-05 | 8.3481e-05 | 0.00010535 | 0.00013288 | 0.0001675 | 0.00021149 | 0.00026688 | 0.00033652 | 0.00042492 | 0.0005361 | 0.00067553 | 0.0008519 | 0.0010726 | 0.0013475 | 0.0016923 | 0.0021194 | 0.0026458 | 0.0033001 | 0.0041074 | 0.0051074 | 0.0063686 | 0.0079649 | 0.010004 | 0.012643
| U5.3 |  | 5.7432e-14 | 1.7435e-13 | 2.3178e-13 | 4.0408e-13 | 6.3586e-13 | 9.8251e-13 | 1.6163e-12 | 2.4819e-12 | 4.0408e-12 | 6.5247e-12 | 1.0104e-11 | 1.1489e-11 | 1.3049e-11 | 1.4262e-11 | 1.6282e-11 | 1.8477e-11 | 2.0151e-11 | 2.2977e-11 | 2.6212e-11 | 2.8579e-11 | 3.2505e-11 | 3.7067e-11 | 4.0588e-11 | 4.6131e-11 | 5.2596e-11 | 5.8485e-11 | 6.5416e-11 | 7.465e-11 | 8.4754e-11 | 9.4513e-11 | 1.086e-10 | 1.2413e-10 | 1.3897e-10 | 1.6027e-10 | 1.84e-10 | 2.0727e-10 | 2.4018e-10 | 2.7724e-10 | 3.1454e-10 | 3.6662e-10 | 4.2614e-10 | 4.878e-10 | 5.7256e-10 | 6.7082e-10 | 7.7573e-10 | 9.1787e-10 | 1.0851e-09 | 1.269e-09 | 1.5144e-09 | 1.8072e-09 | 2.1379e-09 | 2.5737e-09 | 3.1002e-09 | 3.7092e-09 | 4.5026e-09 | 5.4717e-09 | 6.615e-09 | 8.0906e-09 | 9.9091e-09 | 1.2089e-08 | 1.4882e-08 | 1.8348e-08 | 2.2555e-08 | 2.7913e-08 | 3.4601e-08 | 4.2794e-08 | 5.3185e-08 | 6.6208e-08 | 8.2276e-08 | 1.0259e-07 | 1.2813e-07 | 1.598e-07 | 1.9974e-07 | 2.5008e-07 | 3.1274e-07 | 3.9165e-07 | 4.9126e-07 | 6.1561e-07 | 7.7202e-07 | 9.6973e-07 | 1.217e-06 | 1.5279e-06 | 1.9212e-06 | 2.4141e-06 | 3.0333e-06 | 3.8174e-06 | 4.8011e-06 | 6.0369e-06 | 7.6032e-06 | 9.5701e-06 | 1.2041e-05 | 1.5176e-05 | 1.9115e-05 | 2.4067e-05 | 3.0351e-05 | 3.8257e-05 | 4.8198e-05 | 6.0822e-05 | 7.6718e-05 | 9.6708e-05 | 0.0001221 | 0.00015409 | 0.00019429 | 0.00024533 | 0.00030952 | 0.00039002 | 0.00049184 | 0.00061927 | 0.00077799 | 0.00097704 | 0.0012236 | 0.0015276 | 0.0019053 | 0.0023714 | 0.0029488 | 0.0036769 | 0.0045985 | 0.005776 | 0.0072995
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
| Corrections Combined |  | 0.058525 | 0.052938 | 0.047901 | 0.043267 | 0.03902 | 0.035126 | 0.03152 | 0.02825 | 0.025306 | 0.022622 | 0.020214 | 0.019663 | 0.019106 | 0.01857 | 0.018062 | 0.017549 | 0.017055 | 0.016587 | 0.016113 | 0.015659 | 0.015227 | 0.014792 | 0.014373 | 0.013976 | 0.013575 | 0.01319 | 0.012824 | 0.012456 | 0.012101 | 0.011765 | 0.011426 | 0.0111 | 0.010791 | 0.010479 | 0.01018 | 0.0098958 | 0.0096091 | 0.0093337 | 0.0090724 | 0.0088087 | 0.0085554 | 0.008315 | 0.0080724 | 0.0078394 | 0.0076181 | 0.0073948 | 0.0071802 | 0.0069763 | 0.0067704 | 0.0065725 | 0.0063843 | 0.0061943 | 0.0060115 | 0.0058374 | 0.0056616 | 0.0054921 | 0.0053306 | 0.0051673 | 0.0050096 | 0.0048591 | 0.0047066 | 0.0045591 | 0.0044178 | 0.0042746 | 0.0041355 | 0.0040018 | 0.0038659 | 0.0037334 | 0.0036052 | 0.0034748 | 0.0033468 | 0.0032221 | 0.0030947 | 0.0029689 | 0.0028452 | 0.0027183 | 0.0025918 | 0.0024662 | 0.0023367 | 0.0022062 | 0.002075 | 0.001939 | 0.0018003 | 0.0016589 | 0.0015116 | 0.0013594 | 0.0012021 | 0.0010372 | 0.0008663 | 0.00086849 | 0.0008734 | 0.00088179 | 0.00089568 | 0.00091543 | 0.00094265 | 0.00098099 | 0.0010301 | 0.0010928 | 0.0011743 | 0.0013739 | 0.0016942 | 0.0020742 | 0.0025039 | 0.0029913 | 0.0035425 | 0.0041508 | 0.0048355 | 0.0056104 | 0.0064677 | 0.0074403 | 0.0085565 | 0.0098074 | 0.0098733 | 0.0099738 | 0.010127 | 0.010362 | 0.010724 | 0.011279 | 0.01213
| EffDoF |  | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10.0001 | 10.0002 | 10.0003 | 10.0007 | 10.0128 | 10.1142 | 10.3421 | 10.7339 | 11.387 | 12.3171 | 13.5796 | 15.2579 | 17.1148 | 18.8494 | 20.0005 | 19.4114 | 16.4922 | 14.2233 | 12.8397 | 11.9878 | 11.4499 | 11.1112 | 10.8922 | 10.7528 | 10.6748 | 10.6401 | 10.6374 | 10.6659 | 10.9457 | 11.3738 | 12.0262 | 13.0248 | 14.5075 | 16.5449 | 18.8124
Physical Properties of Humid Air
| Density | Standard | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001
| U6.1 |  | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 5e-05
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Carbon Dioxide Mole Ratio | Standard | 1.6616e-05 | 1.9758e-05 | 2.1329e-05 | 2.29e-05 | 2.519e-05 | 2.6335e-05 | 2.748e-05 | 2.9127e-05 | 2.9951e-05 | 3.0774e-05 | 3.1412e-05 | 3.1731e-05 | 3.205e-05 | 3.233e-05 | 3.247e-05 | 3.261e-05 | 3.2867e-05 | 3.2996e-05 | 3.3125e-05 | 3.3361e-05 | 3.3479e-05 | 3.3597e-05 | 3.3813e-05 | 3.3921e-05 | 3.4029e-05 | 3.4227e-05 | 3.4326e-05 | 3.4424e-05 | 3.4605e-05 | 3.4696e-05 | 3.4786e-05 | 3.495e-05 | 3.5033e-05 | 3.5115e-05 | 3.5264e-05 | 3.5339e-05 | 3.5414e-05 | 3.5548e-05 | 3.5616e-05 | 3.5683e-05 | 3.5803e-05 | 3.5863e-05 | 3.5923e-05 | 3.6028e-05 | 3.6081e-05 | 3.6134e-05 | 3.6223e-05 | 3.6268e-05 | 3.6313e-05 | 3.6386e-05 | 3.6422e-05 | 3.6458e-05 | 3.6511e-05 | 3.6538e-05 | 3.6564e-05 | 3.6594e-05 | 3.6608e-05 | 3.6623e-05 | 3.6622e-05 | 3.6622e-05 | 3.6622e-05 | 3.6642e-05 | 3.6652e-05 | 3.6662e-05 | 3.668e-05 | 3.6688e-05 | 3.6697e-05 | 3.6684e-05 | 3.6677e-05 | 3.6671e-05 | 3.6617e-05 | 3.659e-05 | 3.6563e-05 | 3.6455e-05 | 3.6401e-05 | 3.6348e-05 | 3.6165e-05 | 3.6074e-05 | 3.5983e-05 | 3.5699e-05 | 3.5556e-05 | 3.5414e-05 | 3.4986e-05 | 3.4772e-05 | 3.4558e-05 | 3.3927e-05 | 3.3612e-05 | 3.3297e-05 | 3.238e-05 | 3.1922e-05 | 3.1463e-05 | 3.0138e-05 | 2.9475e-05 | 2.8812e-05 | 2.6901e-05 | 2.5945e-05 | 2.4989e-05 | 2.2233e-05 | 2.0855e-05 | 1.9477e-05 | 2.9427e-05 | 3.4402e-05 | 3.9377e-05 | 5.9417e-05 | 6.9438e-05 | 7.9458e-05 | 0.00011338 | 0.00013034 | 0.00014731 | 0.00021173 | 0.00024395 | 0.00027616 | 0.00017349 | 0.00012215 | 7.0819e-05 | 9.7209e-05 | 0.0001104 | 0.0001236 | 0.0001236
| U6.2 |  | 8.3079e-06 | 9.8788e-06 | 1.0664e-05 | 1.145e-05 | 1.2595e-05 | 1.3167e-05 | 1.374e-05 | 1.4563e-05 | 1.4975e-05 | 1.5387e-05 | 1.5706e-05 | 1.5865e-05 | 1.6025e-05 | 1.6165e-05 | 1.6235e-05 | 1.6305e-05 | 1.6434e-05 | 1.6498e-05 | 1.6562e-05 | 1.668e-05 | 1.6739e-05 | 1.6798e-05 | 1.6906e-05 | 1.696e-05 | 1.7014e-05 | 1.7113e-05 | 1.7163e-05 | 1.7212e-05 | 1.7303e-05 | 1.7348e-05 | 1.7393e-05 | 1.7475e-05 | 1.7516e-05 | 1.7558e-05 | 1.7632e-05 | 1.767e-05 | 1.7707e-05 | 1.7774e-05 | 1.7808e-05 | 1.7841e-05 | 1.7901e-05 | 1.7932e-05 | 1.7962e-05 | 1.8014e-05 | 1.804e-05 | 1.8067e-05 | 1.8112e-05 | 1.8134e-05 | 1.8157e-05 | 1.8193e-05 | 1.8211e-05 | 1.8229e-05 | 1.8256e-05 | 1.8269e-05 | 1.8282e-05 | 1.8297e-05 | 1.8304e-05 | 1.8311e-05 | 1.8311e-05 | 1.8311e-05 | 1.8311e-05 | 1.8321e-05 | 1.8326e-05 | 1.8331e-05 | 1.834e-05 | 1.8344e-05 | 1.8349e-05 | 1.8342e-05 | 1.8339e-05 | 1.8335e-05 | 1.8309e-05 | 1.8295e-05 | 1.8282e-05 | 1.8228e-05 | 1.8201e-05 | 1.8174e-05 | 1.8083e-05 | 1.8037e-05 | 1.7992e-05 | 1.7849e-05 | 1.7778e-05 | 1.7707e-05 | 1.7493e-05 | 1.7386e-05 | 1.7279e-05 | 1.6964e-05 | 1.6806e-05 | 1.6648e-05 | 1.619e-05 | 1.5961e-05 | 1.5732e-05 | 1.5069e-05 | 1.4738e-05 | 1.4406e-05 | 1.345e-05 | 1.2973e-05 | 1.2495e-05 | 1.1116e-05 | 1.0427e-05 | 9.7383e-06 | 1.4713e-05 | 1.7201e-05 | 1.9688e-05 | 2.9709e-05 | 3.4719e-05 | 3.9729e-05 | 5.6691e-05 | 6.5172e-05 | 7.3653e-05 | 0.00010587 | 0.00012197 | 0.00013808 | 8.6745e-05 | 6.1077e-05 | 3.5409e-05 | 4.8604e-05 | 5.5202e-05 | 6.18e-05 | 6.18e-05
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Speed of Sound | Standard | 0.0024889 | 0.0024869 | 0.0024859 | 0.0024849 | 0.0024835 | 0.0024827 | 0.002482 | 0.0024809 | 0.0024804 | 0.0024799 | 0.0024795 | 0.0024793 | 0.0024791 | 0.0024789 | 0.0024788 | 0.0024787 | 0.0024785 | 0.0024784 | 0.0024783 | 0.0024782 | 0.0024781 | 0.002478 | 0.0024779 | 0.0024778 | 0.0024777 | 0.0024776 | 0.0024775 | 0.0024774 | 0.0024773 | 0.0024772 | 0.0024771 | 0.002477 | 0.0024769 | 0.0024768 | 0.0024766 | 0.0024766 | 0.0024765 | 0.0024763 | 0.0024762 | 0.0024761 | 0.0024759 | 0.0024758 | 0.0024757 | 0.0024755 | 0.0024753 | 0.0024752 | 0.0024749 | 0.0024747 | 0.0024746 | 0.0024742 | 0.002474 | 0.0024738 | 0.0024732 | 0.0024729 | 0.0024727 | 0.0024719 | 0.0024716 | 0.0024712 | 0.0024702 | 0.0024697 | 0.0024692 | 0.0024678 | 0.002467 | 0.0024663 | 0.0024644 | 0.0024634 | 0.0024624 | 0.0024597 | 0.0024583 | 0.0024569 | 0.002453 | 0.0024511 | 0.0024491 | 0.0024437 | 0.0024409 | 0.0024382 | 0.0024305 | 0.0024266 | 0.0024227 | 0.0024118 | 0.0024063 | 0.0024008 | 0.0023852 | 0.0023774 | 0.0023696 | 0.0023558 | 0.0023489 | 0.002342 | 0.0023309 | 0.0023253 | 0.0023198 | 0.002304 | 0.0022962 | 0.0022883 | 0.002266 | 0.0022549 | 0.0022437 | 0.0022121 | 0.0021963 | 0.0021805 | 0.0021355 | 0.002113 | 0.0020905 | 0.0020265 | 0.0019946 | 0.0019626 | 0.0018713 | 0.0018256 | 0.00178 | 0.0016747 | 0.0016221 | 0.0015695 | 0.0013361 | 0.0012194 | 0.0011028 | 0.00072842 | 0.00054125 | 0.00035407 | 0.00035407
| U6.3 |  | 0.0012445 | 0.0012435 | 0.001243 | 0.0012425 | 0.0012417 | 0.0012414 | 0.001241 | 0.0012405 | 0.0012402 | 0.0012399 | 0.0012397 | 0.0012396 | 0.0012395 | 0.0012394 | 0.0012394 | 0.0012393 | 0.0012393 | 0.0012392 | 0.0012392 | 0.0012391 | 0.001239 | 0.001239 | 0.0012389 | 0.0012389 | 0.0012389 | 0.0012388 | 0.0012387 | 0.0012387 | 0.0012386 | 0.0012386 | 0.0012386 | 0.0012385 | 0.0012384 | 0.0012384 | 0.0012383 | 0.0012383 | 0.0012382 | 0.0012382 | 0.0012381 | 0.0012381 | 0.001238 | 0.0012379 | 0.0012379 | 0.0012377 | 0.0012377 | 0.0012376 | 0.0012374 | 0.0012374 | 0.0012373 | 0.0012371 | 0.001237 | 0.0012369 | 0.0012366 | 0.0012365 | 0.0012363 | 0.001236 | 0.0012358 | 0.0012356 | 0.0012351 | 0.0012348 | 0.0012346 | 0.0012339 | 0.0012335 | 0.0012332 | 0.0012322 | 0.0012317 | 0.0012312 | 0.0012298 | 0.0012291 | 0.0012284 | 0.0012265 | 0.0012255 | 0.0012246 | 0.0012218 | 0.0012205 | 0.0012191 | 0.0012152 | 0.0012133 | 0.0012114 | 0.0012059 | 0.0012031 | 0.0012004 | 0.0011926 | 0.0011887 | 0.0011848 | 0.0011779 | 0.0011745 | 0.001171 | 0.0011654 | 0.0011627 | 0.0011599 | 0.001152 | 0.0011481 | 0.0011441 | 0.001133 | 0.0011274 | 0.0011219 | 0.001106 | 0.0010981 | 0.0010902 | 0.0010677 | 0.0010565 | 0.0010453 | 0.0010133 | 0.00099728 | 0.00098128 | 0.00093564 | 0.00091282 | 0.00089 | 0.00083737 | 0.00081105 | 0.00078474 | 0.00066806 | 0.00060972 | 0.00055139 | 0.00036421 | 0.00027062 | 0.00017704 | 0.00017704
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Ratio of Specifc Heat | Standard | 0.00062312 | 0.00053515 | 0.00049116 | 0.00044717 | 0.00038305 | 0.00035099 | 0.00031893 | 0.00027278 | 0.00024971 | 0.00022664 | 0.00020877 | 0.00019983 | 0.00019089 | 0.00018303 | 0.0001791 | 0.00017517 | 0.00016794 | 0.00016433 | 0.00016072 | 0.00015409 | 0.00015077 | 0.00014746 | 0.00014137 | 0.00013833 | 0.00013528 | 0.00012969 | 0.0001269 | 0.0001241 | 0.00011897 | 0.00011641 | 0.00011384 | 0.00010913 | 0.00010678 | 0.00010442 | 0.0001001 | 9.7942e-05 | 9.5781e-05 | 9.1816e-05 | 8.9833e-05 | 8.7851e-05 | 8.4213e-05 | 8.2394e-05 | 8.0575e-05 | 7.7236e-05 | 7.5567e-05 | 7.3898e-05 | 7.0836e-05 | 6.9304e-05 | 6.7773e-05 | 6.4962e-05 | 6.3557e-05 | 6.2152e-05 | 5.9574e-05 | 5.8284e-05 | 5.6995e-05 | 5.4629e-05 | 5.3445e-05 | 5.2262e-05 | 5.009e-05 | 4.9004e-05 | 4.7918e-05 | 4.5924e-05 | 4.4927e-05 | 4.393e-05 | 4.2098e-05 | 4.1182e-05 | 4.0266e-05 | 3.8583e-05 | 3.7741e-05 | 3.6899e-05 | 3.535e-05 | 3.4576e-05 | 3.3801e-05 | 3.2374e-05 | 3.166e-05 | 3.0947e-05 | 2.9629e-05 | 2.8971e-05 | 2.8312e-05 | 2.7092e-05 | 2.6482e-05 | 2.5873e-05 | 2.4739e-05 | 2.4172e-05 | 2.3605e-05 | 2.2546e-05 | 2.2017e-05 | 2.1487e-05 | 2.0489e-05 | 1.9991e-05 | 1.9492e-05 | 1.8543e-05 | 1.8069e-05 | 1.7595e-05 | 1.668e-05 | 1.6223e-05 | 1.5766e-05 | 1.4868e-05 | 1.442e-05 | 1.3971e-05 | 1.3071e-05 | 1.2621e-05 | 1.2171e-05 | 1.1243e-05 | 1.0778e-05 | 1.0314e-05 | 9.3221e-06 | 8.826e-06 | 8.33e-06 | 8.2972e-06 | 8.2808e-06 | 8.2644e-06 | 5.7751e-06 | 4.5305e-06 | 3.2859e-06 | 2.0346e-06 | 1.409e-06 | 7.8337e-07 | 7.8337e-07
| U6.4 |  | 0.00031156 | 0.00026757 | 0.00024558 | 0.00022358 | 0.00019152 | 0.00017549 | 0.00015946 | 0.00013639 | 0.00012486 | 0.00011332 | 0.00010438 | 9.9915e-05 | 9.5446e-05 | 9.1515e-05 | 8.9549e-05 | 8.7583e-05 | 8.3972e-05 | 8.2167e-05 | 8.0361e-05 | 7.7045e-05 | 7.5387e-05 | 7.3729e-05 | 7.0685e-05 | 6.9163e-05 | 6.764e-05 | 6.4846e-05 | 6.3448e-05 | 6.2051e-05 | 5.9485e-05 | 5.8203e-05 | 5.692e-05 | 5.4566e-05 | 5.3389e-05 | 5.2212e-05 | 5.0051e-05 | 4.8971e-05 | 4.7891e-05 | 4.5908e-05 | 4.4917e-05 | 4.3925e-05 | 4.2106e-05 | 4.1197e-05 | 4.0287e-05 | 3.8618e-05 | 3.7784e-05 | 3.6949e-05 | 3.5418e-05 | 3.4652e-05 | 3.3886e-05 | 3.2481e-05 | 3.1779e-05 | 3.1076e-05 | 2.9787e-05 | 2.9142e-05 | 2.8497e-05 | 2.7314e-05 | 2.6723e-05 | 2.6131e-05 | 2.5045e-05 | 2.4502e-05 | 2.3959e-05 | 2.2962e-05 | 2.2464e-05 | 2.1965e-05 | 2.1049e-05 | 2.0591e-05 | 2.0133e-05 | 1.9291e-05 | 1.887e-05 | 1.845e-05 | 1.7675e-05 | 1.7288e-05 | 1.6901e-05 | 1.6187e-05 | 1.583e-05 | 1.5473e-05 | 1.4815e-05 | 1.4485e-05 | 1.4156e-05 | 1.3546e-05 | 1.3241e-05 | 1.2936e-05 | 1.237e-05 | 1.2086e-05 | 1.1803e-05 | 1.1273e-05 | 1.1008e-05 | 1.0743e-05 | 1.0245e-05 | 9.9954e-06 | 9.746e-06 | 9.2717e-06 | 9.0346e-06 | 8.7974e-06 | 8.3401e-06 | 8.1115e-06 | 7.8828e-06 | 7.4342e-06 | 7.2099e-06 | 6.9855e-06 | 6.5355e-06 | 6.3105e-06 | 6.0854e-06 | 5.6213e-06 | 5.3892e-06 | 5.1571e-06 | 4.661e-06 | 4.413e-06 | 4.165e-06 | 4.1486e-06 | 4.1404e-06 | 4.1322e-06 | 2.8876e-06 | 2.2653e-06 | 1.6429e-06 | 1.0173e-06 | 7.045e-07 | 3.9169e-07 | 3.9169e-07
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Physical properties of air combined |  | 0.0012839 | 0.0012729 | 0.001268 | 0.0012635 | 0.0012575 | 0.0012548 | 0.0012523 | 0.001249 | 0.0012476 | 0.0012462 | 0.0012452 | 0.0012448 | 0.0012443 | 0.0012439 | 0.0012437 | 0.0012435 | 0.0012432 | 0.0012431 | 0.0012429 | 0.0012426 | 0.0012425 | 0.0012423 | 0.0012421 | 0.0012419 | 0.0012418 | 0.0012416 | 0.0012415 | 0.0012414 | 0.0012412 | 0.0012411 | 0.001241 | 0.0012408 | 0.0012407 | 0.0012406 | 0.0012405 | 0.0012404 | 0.0012403 | 0.0012401 | 0.0012401 | 0.00124 | 0.0012398 | 0.0012397 | 0.0012397 | 0.0012395 | 0.0012394 | 0.0012393 | 0.0012391 | 0.001239 | 0.0012389 | 0.0012387 | 0.0012385 | 0.0012384 | 0.0012381 | 0.001238 | 0.0012378 | 0.0012374 | 0.0012372 | 0.001237 | 0.0012365 | 0.0012362 | 0.001236 | 0.0012352 | 0.0012349 | 0.0012345 | 0.0012335 | 0.001233 | 0.0012325 | 0.0012311 | 0.0012304 | 0.0012297 | 0.0012278 | 0.0012268 | 0.0012258 | 0.0012231 | 0.0012217 | 0.0012204 | 0.0012165 | 0.0012146 | 0.0012126 | 0.0012071 | 0.0012044 | 0.0012016 | 0.0011939 | 0.00119 | 0.0011861 | 0.0011791 | 0.0011757 | 0.0011722 | 0.0011667 | 0.0011639 | 0.0011611 | 0.0011532 | 0.0011493 | 0.0011454 | 0.0011342 | 0.0011286 | 0.0011231 | 0.0011073 | 0.0010993 | 0.0010914 | 0.001069 | 0.0010578 | 0.0010467 | 0.001015 | 0.00099915 | 0.00098337 | 0.0009387 | 0.00091652 | 0.00089445 | 0.00084552 | 0.00082171 | 0.00079837 | 0.00067553 | 0.00061482 | 0.00055478 | 0.00037083 | 0.00028069 | 0.00019406 | 0.00019406
| EffDoF |  | 33.8514 | 32.8775 | 32.4439 | 32.0464 | 31.5324 | 31.3048 | 31.0969 | 30.8323 | 30.7154 | 30.6088 | 30.5333 | 30.4979 | 30.4641 | 30.4355 | 30.4217 | 30.4082 | 30.3842 | 30.3726 | 30.3612 | 30.341 | 30.3312 | 30.3216 | 30.3046 | 30.2963 | 30.2883 | 30.2739 | 30.267 | 30.2602 | 30.2481 | 30.2423 | 30.2366 | 30.2264 | 30.2215 | 30.2167 | 30.2082 | 30.2041 | 30.2 | 30.1929 | 30.1894 | 30.186 | 30.18 | 30.1771 | 30.1742 | 30.1692 | 30.1667 | 30.1643 | 30.1601 | 30.158 | 30.156 | 30.1525 | 30.1507 | 30.149 | 30.1461 | 30.1446 | 30.1432 | 30.1407 | 30.1395 | 30.1383 | 30.1363 | 30.1353 | 30.1343 | 30.1326 | 30.1318 | 30.131 | 30.1297 | 30.129 | 30.1284 | 30.1273 | 30.1268 | 30.1263 | 30.1256 | 30.1252 | 30.1249 | 30.1244 | 30.1242 | 30.124 | 30.1238 | 30.1237 | 30.1237 | 30.1239 | 30.124 | 30.1242 | 30.1249 | 30.1252 | 30.1256 | 30.1261 | 30.1263 | 30.1266 | 30.1267 | 30.1267 | 30.1268 | 30.1272 | 30.1274 | 30.1277 | 30.1286 | 30.1291 | 30.1296 | 30.1314 | 30.1324 | 30.1335 | 30.1452 | 30.1525 | 30.1607 | 30.1997 | 30.2255 | 30.256 | 30.3937 | 30.4882 | 30.6029 | 31.1776 | 31.591 | 32.1085 | 31.3541 | 31.0103 | 30.7433 | 32.2185 | 34.6169 | 42.4174 | 42.4174
| Error in Adding Volumes | Rect | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| U7.1 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Type A | Standard | 0.042181 | 0.029675 | 0.021269 | 0.015618 | 0.01182 | 0.0092669 | 0.0075506 | 0.0063969 | 0.0056213 | 0.0051 | 0.0047496 | 0.0046817 | 0.0046202 | 0.0045645 | 0.004514 | 0.0044684 | 0.004427 | 0.0043896 | 0.0043557 | 0.004325 | 0.0042972 | 0.0042721 | 0.0042493 | 0.0042286 | 0.00421 | 0.004193 | 0.0041777 | 0.0041639 | 0.0041513 | 0.0041399 | 0.0041296 | 0.0041203 | 0.0041119 | 0.0041042 | 0.0040973 | 0.004091 | 0.0040854 | 0.0040802 | 0.0040756 | 0.0040714 | 0.0040675 | 0.0040641 | 0.004061 | 0.0040581 | 0.0040556 | 0.0040532 | 0.0040511 | 0.0040492 | 0.0040475 | 0.0040459 | 0.0040445 | 0.0040433 | 0.0040421 | 0.004041 | 0.0040401 | 0.0040392 | 0.0040385 | 0.0040378 | 0.0040371 | 0.0040365 | 0.004036 | 0.0040355 | 0.0040351 | 0.0040347 | 0.0040344 | 0.0040341 | 0.0040338 | 0.0040335 | 0.0040333 | 0.0040331 | 0.0040329 | 0.0040327 | 0.0040325 | 0.0040324 | 0.0040322 | 0.0040321 | 0.004032 | 0.0040319 | 0.0040318 | 0.0040318 | 0.0040317 | 0.0040316 | 0.0040316 | 0.0040315 | 0.0040315 | 0.0040314 | 0.0040314 | 0.0040313 | 0.0040313 | 0.0040313 | 0.0040313 | 0.0040312 | 0.0040312 | 0.0040312 | 0.0040312 | 0.0040312 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.004 | 0.0048782 | 0.0066349 | 0.0089619 | 0.011573 | 0.014168 | 0.01647 | 0.018281 | 0.019573 | 0.020616 | 0.022158 | 0.025691 | 0.0338
| U8.1 |  | 0.042181 | 0.029675 | 0.021269 | 0.015618 | 0.01182 | 0.0092669 | 0.0075506 | 0.0063969 | 0.0056213 | 0.0051 | 0.0047496 | 0.0046817 | 0.0046202 | 0.0045645 | 0.004514 | 0.0044684 | 0.004427 | 0.0043896 | 0.0043557 | 0.004325 | 0.0042972 | 0.0042721 | 0.0042493 | 0.0042286 | 0.00421 | 0.004193 | 0.0041777 | 0.0041639 | 0.0041513 | 0.0041399 | 0.0041296 | 0.0041203 | 0.0041119 | 0.0041042 | 0.0040973 | 0.004091 | 0.0040854 | 0.0040802 | 0.0040756 | 0.0040714 | 0.0040675 | 0.0040641 | 0.004061 | 0.0040581 | 0.0040556 | 0.0040532 | 0.0040511 | 0.0040492 | 0.0040475 | 0.0040459 | 0.0040445 | 0.0040433 | 0.0040421 | 0.004041 | 0.0040401 | 0.0040392 | 0.0040385 | 0.0040378 | 0.0040371 | 0.0040365 | 0.004036 | 0.0040355 | 0.0040351 | 0.0040347 | 0.0040344 | 0.0040341 | 0.0040338 | 0.0040335 | 0.0040333 | 0.0040331 | 0.0040329 | 0.0040327 | 0.0040325 | 0.0040324 | 0.0040322 | 0.0040321 | 0.004032 | 0.0040319 | 0.0040318 | 0.0040318 | 0.0040317 | 0.0040316 | 0.0040316 | 0.0040315 | 0.0040315 | 0.0040314 | 0.0040314 | 0.0040313 | 0.0040313 | 0.0040313 | 0.0040313 | 0.0040312 | 0.0040312 | 0.0040312 | 0.0040312 | 0.0040312 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.0040311 | 0.004 | 0.0048782 | 0.0066349 | 0.0089619 | 0.011573 | 0.014168 | 0.01647 | 0.018281 | 0.019573 | 0.020616 | 0.022158 | 0.025691 | 0.0338
| DoF |  | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14
| Rounding | Rect | 0.1 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01
| U9.1 |  | 0.057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735
| DoF |  | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| Frequency |  | 1.9953 | 2.5119 | 3.1623 | 3.9811 | 5.0119 | 6.3096 | 7.9433 | 10 | 12.5892 | 15.8489 | 19.9526 | 21.1349 | 22.3872 | 23.7137 | 25.1189 | 26.6073 | 28.1838 | 29.8538 | 31.6228 | 33.4965 | 35.4813 | 37.5837 | 39.8107 | 42.1696 | 44.6684 | 47.3151 | 50.1187 | 53.0885 | 56.2341 | 59.5662 | 63.0957 | 66.8344 | 70.7946 | 74.9894 | 79.4328 | 84.1395 | 89.1251 | 94.4061 | 100 | 105.9254 | 112.2018 | 118.8502 | 125.8925 | 133.3521 | 141.2538 | 149.6235 | 158.4893 | 167.8804 | 177.8279 | 188.3649 | 199.5262 | 211.3489 | 223.8721 | 237.1374 | 251.1887 | 266.0725 | 281.8383 | 298.5382 | 316.2278 | 334.9655 | 354.8134 | 375.8374 | 398.1071 | 421.6965 | 446.6837 | 473.1512 | 501.1873 | 530.8843 | 562.3413 | 595.6623 | 630.9572 | 668.344 | 707.9456 | 749.8942 | 794.3285 | 841.395 | 891.251 | 944.0607 | 1000 | 1059.254 | 1122.018 | 1188.502 | 1258.925 | 1333.521 | 1412.538 | 1496.235 | 1584.893 | 1678.804 | 1778.279 | 1883.65 | 1995.262 | 2113.489 | 2238.721 | 2371.374 | 2511.887 | 2660.725 | 2818.383 | 2985.382 | 3162.278 | 3349.655 | 3548.134 | 3758.374 | 3981.071 | 4216.965 | 4466.837 | 4731.512 | 5011.873 | 5308.843 | 5623.413 | 5956.623 | 6309.573 | 6683.44 | 7079.456 | 7498.942 | 7943.284 | 8413.95 | 8912.511 | 9440.606 | 10000
| Ucombined |  | 0.093037 | 0.061928 | 0.053858 | 0.047656 | 0.043019 | 0.038849 | 0.03523 | 0.031371 | 0.028596 | 0.026161 | 0.024043 | 0.02358 | 0.023099 | 0.022653 | 0.022228 | 0.021818 | 0.021399 | 0.021034 | 0.019869 | 0.019496 | 0.019144 | 0.018802 | 0.018469 | 0.018166 | 0.017855 | 0.017559 | 0.017292 | 0.017026 | 0.016746 | 0.016491 | 0.016257 | 0.016017 | 0.015812 | 0.015609 | 0.015387 | 0.01522 | 0.015023 | 0.014836 | 0.014693 | 0.014531 | 0.014378 | 0.014235 | 0.014105 | 0.01396 | 0.013848 | 0.013737 | 0.013599 | 0.013515 | 0.01341 | 0.013286 | 0.013218 | 0.013127 | 0.013028 | 0.012961 | 0.012895 | 0.012796 | 0.012739 | 0.012684 | 0.012595 | 0.012561 | 0.012502 | 0.012422 | 0.012381 | 0.012343 | 0.012269 | 0.012224 | 0.012193 | 0.012138 | 0.012098 | 0.012085 | 0.012022 | 0.011986 | 0.011952 | 0.011919 | 0.011888 | 0.011845 | 0.011803 | 0.011774 | 0.011762 | 0.011709 | 0.011685 | 0.011649 | 0.011615 | 0.011594 | 0.011552 | 0.011511 | 0.011484 | 0.011452 | 0.011425 | 0.011426 | 0.011409 | 0.011381 | 0.011391 | 0.011384 | 0.011376 | 0.011402 | 0.011429 | 0.011462 | 0.011536 | 0.011636 | 0.011786 | 0.01199 | 0.012267 | 0.012592 | 0.012991 | 0.013512 | 0.014098 | 0.014976 | 0.016434 | 0.018296 | 0.020622 | 0.023201 | 0.024405 | 0.025755 | 0.026837 | 0.027832 | 0.029331 | 0.032715 | 0.040255
| EffDoF |  | 42.3339 | 17.489 | 15.544 | 14.5327 | 14.6664 | 14.8797 | 15.5196 | 15.1302 | 16.1941 | 17.7091 | 19.7218 | 20.3485 | 20.9814 | 21.6984 | 22.426 | 23.2999 | 24.1016 | 25.0696 | 22.6871 | 23.5264 | 24.4056 | 25.4308 | 26.4768 | 27.6206 | 28.8487 | 30.1505 | 31.5727 | 33.1617 | 34.6403 | 36.2212 | 38.1639 | 40.0496 | 42.1896 | 44.5652 | 46.7691 | 49.463 | 52.1006 | 54.8489 | 58.044 | 61.2806 | 64.6315 | 68.0492 | 71.8936 | 75.523 | 79.4851 | 83.7076 | 87.4895 | 91.8705 | 96.116 | 100.0032 | 104.4032 | 108.3233 | 112.1875 | 116.1611 | 120.0981 | 123.4674 | 126.9228 | 130.2814 | 133.1052 | 135.9469 | 138.5072 | 140.6917 | 142.6838 | 144.536 | 146.0693 | 147.3709 | 148.5243 | 149.4581 | 150.157 | 150.6715 | 151.1249 | 151.356 | 151.4842 | 151.4559 | 151.3924 | 151.2612 | 151.0675 | 150.672 | 150.6337 | 150.2655 | 149.7818 | 149.3739 | 148.9604 | 148.4167 | 148.1878 | 147.9663 | 147.6512 | 147.7443 | 148.0728 | 149.0284 | 150.4555 | 152.0488 | 154.0993 | 156.6671 | 159.7813 | 163.8188 | 168.2727 | 173.9826 | 180.3987 | 188.3155 | 196.3382 | 203.3364 | 206.6999 | 204.4161 | 195.6171 | 180.3164 | 163.2003 | 143.4233 | 123.9483 | 100.666 | 80.3477 | 66.1648 | 52.5473 | 46.3812 | 43.1832 | 41.4797 | 39.2797 | 34.8536 | 27.5618
| k |  | 2.0204 | 2.1109 | 2.131 | 2.1437 | 2.1419 | 2.1391 | 2.1313 | 2.136 | 2.1237 | 2.1089 | 2.093 | 2.0887 | 2.0846 | 2.0803 | 2.0763 | 2.0717 | 2.0679 | 2.0635 | 2.0749 | 2.0706 | 2.0665 | 2.062 | 2.0578 | 2.0536 | 2.0495 | 2.0455 | 2.0416 | 2.0375 | 2.0341 | 2.0308 | 2.0271 | 2.0239 | 2.0206 | 2.0173 | 2.0145 | 2.0115 | 2.0088 | 2.0064 | 2.0038 | 2.0014 | 1.9992 | 1.9972 | 1.9952 | 1.9935 | 1.9918 | 1.9902 | 1.9889 | 1.9875 | 1.9863 | 1.9852 | 1.9842 | 1.9833 | 1.9825 | 1.9817 | 1.981 | 1.9804 | 1.9798 | 1.9793 | 1.9789 | 1.9785 | 1.9782 | 1.9779 | 1.9776 | 1.9774 | 1.9772 | 1.9771 | 1.9769 | 1.9768 | 1.9768 | 1.9767 | 1.9766 | 1.9766 | 1.9766 | 1.9766 | 1.9766 | 1.9766 | 1.9767 | 1.9767 | 1.9767 | 1.9767 | 1.9768 | 1.9768 | 1.9769 | 1.9769 | 1.977 | 1.977 | 1.977 | 1.977 | 1.977 | 1.9769 | 1.9767 | 1.9765 | 1.9763 | 1.9761 | 1.9757 | 1.9753 | 1.9749 | 1.9744 | 1.9739 | 1.9733 | 1.9728 | 1.9724 | 1.9721 | 1.9723 | 1.9728 | 1.9739 | 1.9754 | 1.9775 | 1.9803 | 1.9851 | 1.9915 | 1.9983 | 2.0084 | 2.015 | 2.0192 | 2.0216 | 2.0252 | 2.0337 | 2.0538
| Uexp |  | 0.18797 | 0.13073 | 0.11477 | 0.10216 | 0.092144 | 0.083103 | 0.075085 | 0.067008 | 0.060729 | 0.055172 | 0.050321 | 0.049252 | 0.048153 | 0.047126 | 0.04615 | 0.0452 | 0.04425 | 0.043405 | 0.041227 | 0.040368 | 0.039561 | 0.03877 | 0.038007 | 0.037306 | 0.036593 | 0.035918 | 0.035302 | 0.034691 | 0.034064 | 0.033489 | 0.032956 | 0.032417 | 0.03195 | 0.031488 | 0.030998 | 0.030615 | 0.030179 | 0.029767 | 0.029442 | 0.029082 | 0.028744 | 0.02843 | 0.028142 | 0.02783 | 0.027582 | 0.02734 | 0.027046 | 0.026862 | 0.026635 | 0.026376 | 0.026227 | 0.026034 | 0.025828 | 0.025685 | 0.025544 | 0.02534 | 0.025222 | 0.025106 | 0.024924 | 0.024851 | 0.024732 | 0.024568 | 0.024485 | 0.024408 | 0.024259 | 0.024168 | 0.024105 | 0.023994 | 0.023914 | 0.023889 | 0.023762 | 0.023692 | 0.023624 | 0.023559 | 0.023499 | 0.023413 | 0.02333 | 0.023274 | 0.02325 | 0.023145 | 0.023098 | 0.023028 | 0.022961 | 0.022921 | 0.022837 | 0.022757 | 0.022704 | 0.02264 | 0.022587 | 0.022588 | 0.022552 | 0.022496 | 0.022512 | 0.022496 | 0.022477 | 0.022523 | 0.022572 | 0.022631 | 0.022771 | 0.022963 | 0.023252 | 0.023648 | 0.024192 | 0.024835 | 0.025629 | 0.026671 | 0.02785 | 0.029615 | 0.032545 | 0.036319 | 0.041069 | 0.046362 | 0.049016 | 0.051895 | 0.054189 | 0.056266 | 0.059401 | 0.066532 | 0.082678
| Rounding to 1 or 2 dp |  | 0.18797 | 0.13073 | 0.11477 | 0.10216 | 0.092144 | 0.083103 | 0.075085 | 0.067008 | 0.060729 | 0.055172 | 0.050321 | 0.049252 | 0.048153 | 0.047126 | 0.04615 | 0.0452 | 0.04425 | 0.043405 | 0.041227 | 0.040368 | 0.039561 | 0.03877 | 0.038007 | 0.037306 | 0.036593 | 0.035918 | 0.035302 | 0.034691 | 0.034064 | 0.033489 | 0.032956 | 0.032417 | 0.03195 | 0.031488 | 0.030998 | 0.030615 | 0.030179 | 0.029767 | 0.029442 | 0.029082 | 0.028744 | 0.02843 | 0.028142 | 0.02783 | 0.027582 | 0.02734 | 0.027046 | 0.026862 | 0.026635 | 0.026376 | 0.026227 | 0.026034 | 0.025828 | 0.025685 | 0.025544 | 0.02534 | 0.025222 | 0.025106 | 0.024924 | 0.024851 | 0.024732 | 0.024568 | 0.024485 | 0.024408 | 0.024259 | 0.024168 | 0.024105 | 0.023994 | 0.023914 | 0.023889 | 0.023762 | 0.023692 | 0.023624 | 0.023559 | 0.023499 | 0.023413 | 0.02333 | 0.023274 | 0.02325 | 0.023145 | 0.023098 | 0.023028 | 0.022961 | 0.022921 | 0.022837 | 0.022757 | 0.022704 | 0.02264 | 0.022587 | 0.022588 | 0.022552 | 0.022496 | 0.022512 | 0.022496 | 0.022477 | 0.022523 | 0.022572 | 0.022631 | 0.022771 | 0.022963 | 0.023252 | 0.023648 | 0.024192 | 0.024835 | 0.025629 | 0.026671 | 0.02785 | 0.029615 | 0.032545 | 0.036319 | 0.041069 | 0.046362 | 0.049016 | 0.051895 | 0.054189 | 0.056266 | 0.059401 | 0.066532 | 0.082678
| Final U95 |  | 0.2 | 0.14 | 0.12 | 0.11 | 0.1 | 0.09 | 0.08 | 0.07 | 0.07 | 0.06 | 0.06 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.04 | 0.05 | 0.05 | 0.05 | 0.06 | 0.06 | 0.06 | 0.06 | 0.07 | 0.09

## SCL Level
SCL | Level
Brüel & Kjær - Microphone Reciprocity Calibration System Type 9699
| Uncertainty Budget: |  | Parameter Uncertainty |  | Uncertainty of Microphone Sensitivity and Contributions to this [dB]
Frequency (Hz)
| Type 4160 - LS1 |  | Value | Unit | 19.95 | 25.12 | 31.62 | 39.81 | 50.12 | 63.1 | 79.43 | 100 | 125.89 | 158.49 | 199.53 | 251.19 | 316.23 | 398.11 | 501.19 | 630.96 | 794.33 | 1000 | 1258.9 | 1584.9 | 1995.3 | 2511.9 | 3162.3 | 3981.1 | 5011.9 | 6309.6 | 7943.3 | 10000
| Mp Expanded Uncertainty at reference environmental conditions |  | S | dB | 0.06 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 | 0.06 | 0.09
| Electrical Transfer Impedance |  | S | dB | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006
| 1 | Series Capacitor | 2.7 | pF | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027 | 0.0027
| 2 | Voltage Ratios | 1 | - | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004
| 3 | Cross Talk | 1 | - | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013 | 0.0013
| 4 | Inherent and Ambient Noise | 1 | - | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001
| 5 | Distortion | 1 | - | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002
| 6 | Frequency | 1 | - | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 7a | Receiver Ground Shielding | 1 | - | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035 | 0.0035
| 7b | Receiver Ground Shielding (guard length) | 0.09 | mm | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001
| 8a | Transmitter Ground Shielding | 1 | - | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002
| 8b | Transmitter Ground Shielding (guard length) | 0.09 | mm | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001
| Acoustic Transfer Impedance (coupler and gas) |  | S | dB | 0.012 | 0.011 | 0.011 | 0.011 | 0.011 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.011 | 0.011 | 0.011 | 0.011 | 0.012 | 0.014
| 9 | Coupler Length | 0.006 | mm | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0003 | 0.0004 | 0.0006 | 0.001 | 0.0015 | 0.0024 | 0.0044
(normalized at 500 Hz)
| 10 | Coupler Diameter (normalized at 500 Hz) | 0.006 | mm | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001
| 11a | Coupler Volume (short | 2.89 | mm3 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104
7.5 mm) + volume fitting
| 11b | Coupler Volume (long | 4.04 | mm3 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104 | 0.0104
15 mm) + volume fitting
| 12a | Coupler Surface Area | 0.115 | % | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
(short 7.5 mm)
| 12b | Coupler Surface Area | 0.115 | % | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | - | - | - | - | - | - | -
(long 15.0 mm)
| 13 | Coupler Leakage | 0.58 | - | 0.0063 | 0.0041 | 0.0025 | 0.002 | 0.0014 | 0.0006 | 0.0005 | 0.0003 | 0.0002 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 14 | Static Pressure of air in Coupler | 0.015 | kPa | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006
| 15 | Temperature of air in Coupler | 0.58 | °C | 0.0002 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0003 | 0.0005 | 0.0012 | 0.0012 | 0.0019 | 0.003 | 0.0048 | 0.0084
| 16 | Relative Humidity of air in Coupler | 0.35 | % | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0002
| Acoustic Transfer Impedance (microphones) |  | S | dB | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.004 | 0.005 | 0.007 | 0.01 | 0.012 | 0.012 | 0.011 | 0.021
| 17 | Microphone Front Cavity Depth | 0.01 | mm | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0003 | 0.0005 | 0.0009 | 0.0013 | 0.0021 | 0.0034 | 0.0054 | 0.0086 | 0.0149 | 0.0253
| 18 | Microphone Front Cavity Volume | 1.73 | mm3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 19 | Microphone Equivalent Volume | 2.29 | mm3 | 0.0002 | 0.0002 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0003 | 0.0007 | 0.0011 | 0.0017 | 0.0028 | 0.0044 | 0.0052 | 0.0033
| 20 | Microphone Diaphragm Resonance Frequency | 318 | Hz | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0002 | 0.0003 | 0.001 | 0.0017 | 0.0031 | 0.0061 | 0.0069 | 0.0023 | 0.0073
| 21 | Microphone Diaphragm Loss Factor | 5.77 | % | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0004 | 0.0006 | 0.0011 | 0.0017 | 0.0025 | 0.004 | 0.0058 | 0.0082 | 0.0096 | 0.0065 | 0.0022 | 0.0126
| 22 | Microphone Diaphragm Compliance | Follow 21,22,23 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 23 | Microphone Diaphragm Mass | Follow 21,22,23 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 24 | Microphone Diaphragm Resistance | Follow 21,22,23 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 25a | Mic. Cavity Thread - Short Coup. (heat cond.) | 0 | % | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 25b | Mic. Cavity Thread - Long Coup. (heat cond.) | 0 | % | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 26 | Microphone Polarization Voltage | 0.05 | V | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026
| Acoustic Transfer Imp. (imperfection of theory) |  | S | dB | 0.009 | 0.008 | 0.007 | 0.007 | 0.006 | 0.005 | 0.005 | 0.004 | 0.004 | 0.003 | 0.003 | 0.003 | 0.002 | 0.002 | 0.002 | 0.002 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | 0.002 | 0.004 | 0.005 | 0.007 | 0.011
| 27 | Heat Conduction | 58 | % | 0.0089 | 0.008 | 0.0071 | 0.0066 | 0.0059 | 0.0051 | 0.0047 | 0.0042 | 0.0036 | 0.0034 | 0.003 | 0.0026 | 0.0024 | 0.0021 | 0.0018 | 0.0016 | 0.0015 | 0.0011 | 0.001 | 0.0008 | 0.0004 | 0.0002 | 0.0002 | 0.0008 | 0.0014 | 0.0023 | 0.0037 | 0.0064
| 28 | Adding of Excess Volume | 0 | mm3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 29 | Viscosity Losses | 1 | - | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 30 | Radial Wave-motion | 58 | % | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0003 | 0.0004 | 0.0006 | 0.001 | 0.0015 | 0.0022 | 0.0034 | 0.0049 | 0.0063 | 0.0086
| Processing of Results |  | S | dB | 0.009 | 0.01 | 0.007 | 0.007 | 0.006 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.006 | 0.006 | 0.006 | 0.006 | 0.005 | 0.005 | 0.005 | 0.007 | 0.012 | 0.016 | 0.024 | 0.035
| 31 | Rounding to number of decimals: | 3 | - | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003
| 32 | Repeatability of measurements | 0.58 | - | 0.0087 | 0.0093 | 0.00638 | 0.00638 | 0.00522 | 0.00464 | 0.00464 | 0.00464 | 0.00464 | 0.00464 | 0.00464 | 0.00406 | 0.00464 | 0.00464 | 0.00464 | 0.00464 | 0.00522 | 0.00522 | 0.00522 | 0.00522 | 0.00464 | 0.00406 | 0.00464 | 0.0064 | 0.011 | 0.0162 | 0.0232 | 0.0336
| 33a | Static Pressure Correction UDP x Coef(250 Hz) | 0.015 | kPa | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002
| 33b | Static Pressure Correction DP x Ucoef(250 Hz) | 1.4 | kPa | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022 | 0.0022
| 33c | Static Pressure Correction UDP x normCoef(f) | 0.015 | kPa | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0001 | 0.0002
| 33d | Static Pressure Correction DP x UnormCoef(f) | 1.4 | kPa | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0003 | 0.0003 | 0.0004 | 0.0007 | 0.0011 | 0.0015 | 0.0017 | 0.0003 | 0.0056 | 0.0092
| 34a | Temperature Correction UDP x Coef(250 Hz) | 0.1 | °C | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002
| 34b | Temperature Correction | 1 | °C | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001
DT x Ucoef(250 Hz)
| 34c | Temperature Correction UDT x normCoef(f) | 0.1 | °C | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0002 | 0.0003 | 0.0005 | 0.0007 | 0.0011 | 0.0014 | 0.0011 | 0.0003
| 34d | Temperature Correction | 1 | °C | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0004 | 0.0006 | 0.0009 | 0.0011 | 0.0003 | 0.0022 | 0.0051
DT x UnormCoef(f)

## NMIJ Level
NMIJ | Level
Freq. [Hz]
| Uncertainty components | Type | Dist. | 1.9953 | 2.5119 | 3.1623 | 3.9811 | 5.0119 | 6.3096 | 7.9433 | 10 | 12.5893 | 15.8489 | 19.9526 | 25.1189 | 31.6228 | 39.8107 | 50.1187 | 63.0957 | 79.4328 | 100 | 125.8925 | 158.4893 | 199.5262 | 251.1886 | 316.2278 | 398.1072 | 501.1872 | 630.9573 | 794.3282 | 1000 | 1258.9254 | 1584.8932 | 1995.2623 | 2511.8864 | 3162.2777 | 3981.0717 | 5011.8723 | 6309.5734 | 7943.2823 | 10000
| Electrical impedance of transmitter |  | Combined | 0.0010466 | 0.00079549 | 0.00069781 | 0.00065962 | 0.00078172 | 0.00089237 | 0.001038 | 0.0011496 | 0.0013558 | 0.0014066 | 0.0028632 | 0.0027244 | 0.0026387 | 0.0026257 | 0.0026132 | 0.001803 | 0.0017843 | 0.0018677 | 0.0021419 | 0.0019037 | 0.0019425 | 0.0018518 | 0.0019711 | 0.0017393 | 0.0018875 | 0.00096019 | 0.00088645 | 0.0014768 | 0.00075435 | 0.00076418 | 0.00053946 | 0.00057189 | 0.00060005 | 0.00042286 | 0.00045306 | 0.00035643 | 0.00048533 | 0.00046701
| Reference capacitance | B | Rect. | 0.00060814 | 0.00058642 | 0.00059945 | 0.00053573 | 0.00054297 | 0.00047056 | 0.00036051 | 0.00028232 | 0.00021282 | 0.00011437 | 0.00065708 | 0.00075837 | 0.00076705 | 0.00072943 | 0.0006976 | 0.0002606 | 0.00034748 | 0.00039237 | 0.00043436 | 0.00046477 | 0.00071531 | 0.00069938 | 0.00069504 | 0.00069214 | 0.00068055 | 0.00028666 | 0.00023019 | 0.00019979 | 0.00015056 | 0.00013174 | 7.8171e-05 | 9.5542e-05 | 0.00012015 | 0.00013897 | 0.00015779 | 0.00022003 | 0.00023305 | 0.00025476
| Stray capacitance | B | Rect. | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Repeatability | A | Normal | 0.00085175 | 0.00053751 | 0.00035721 | 0.00038483 | 0.00056238 | 0.00075822 | 0.00097337 | 0.0011144 | 0.001339 | 0.001402 | 0.0027868 | 0.0026168 | 0.0025248 | 0.0025224 | 0.0025183 | 0.0017841 | 0.0017501 | 0.0018261 | 0.0020974 | 0.0018461 | 0.001806 | 0.0017146 | 0.0018445 | 0.0015956 | 0.0017605 | 0.0009164 | 0.00085604 | 0.0014633 | 0.00073917 | 0.00075274 | 0.00053377 | 0.00056386 | 0.0005879 | 0.00039937 | 0.0004247 | 0.00028041 | 0.00042571 | 0.0003914
| Voltage transfer function between transmitter and receiver |  | Combined | 0.10949 | 0.10556 | 0.10546 | 0.093342 | 0.041341 | 0.034865 | 0.02047 | 0.015755 | 0.013357 | 0.014588 | 0.012821 | 0.0096769 | 0.010506 | 0.010033 | 0.010942 | 0.00981 | 0.0092606 | 0.011783 | 0.010873 | 0.011541 | 0.010209 | 0.0095454 | 0.011286 | 0.012278 | 0.012001 | 0.011028 | 0.011986 | 0.010378 | 0.011574 | 0.0098036 | 0.0096078 | 0.0095579 | 0.0091727 | 0.009344 | 0.010256 | 0.010825 | 0.011086 | 0.038467
| Attenuator | B | Rect. | 0.0011547 | 0.0013279 | 0.00092376 | 0.0012124 | 0.0013279 | 0.00092376 | 0.0012124 | 0.0014434 | 0.0009815 | 0.0012124 | 0.0012124 | 0.0012702 | 0.00075056 | 0.0009815 | 0.0013279 | 0.0009815 | 0.0012124 | 0.0017898 | 0.00046188 | 0.0015588 | 0.00086603 | 0.001097 | 0.0015588 | 0.0009815 | 0.0019053 | 0.00023094 | 0.00011547 | 0.00028868 | 0.00034641 | 0.00075056 | 0.00086603 | 0.00069282 | 0.0013856 | 0.00080829 | 0.0012124 | 0.00063509 | 0.0027135 | 0.001963
| Crosstalk | B | Rect. | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301 | 0.00043301
| Linearity of FFT analyzer | B | Rect. | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 0.00025
| Static pressure | B | Rect. | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269 | 0.0064269
| Aoustic coupler difference | B | Rect. | 0.06 | 0.06 | 0.059836 | 0.033341 | 0.03 | 0.025449 | 0.01 | 0.0087137 | 0.0089716 | 0.0089413 | 0.0063822 | 0.0031272 | 0.0022945 | 0.0042208 | 0.0020733 | 0.0044273 | 0.0015604 | 0.0053302 | 0.0014858 | 0.0020261 | 0.00035755 | 0.0010828 | 0.0010628 | 0.00048514 | 0.0015514 | 0.0039888 | 0.003144 | 0.0024829 | 0.0047671 | 0.00099879 | 0.0017003 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Polarizing voltage | B | Rect. | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708 | 0.00030708
| Repeatability | A | Normal | 0.091347 | 0.086603 | 0.086603 | 0.086937 | 0.027673 | 0.022924 | 0.016613 | 0.011342 | 0.0074436 | 0.0094781 | 0.008979 | 0.0063792 | 0.0079369 | 0.0063505 | 0.0084918 | 0.005841 | 0.0063481 | 0.0081039 | 0.0086162 | 0.0092256 | 0.0078611 | 0.006869 | 0.0090702 | 0.010392 | 0.0098202 | 0.0080058 | 0.0096021 | 0.0077391 | 0.0083395 | 0.0072797 | 0.0068639 | 0.0070227 | 0.0063768 | 0.0067158 | 0.0078845 | 0.0086729 | 0.0086011 | 0.037872
| Coupler volume | A | Normal | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461 | 0.00461
| Specific heat ratio | B | Rect. | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014
| Coupler correction |  | Combined | 0.009 | 0.009 | 0.009 | 0.009 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.009 | 0.009 | 0.009 | 0.008 | 0.007 | 0.007 | 0.009
| Environmental corrections of microphone sensitivities |  | Combined | 0.0055906 | 0.0055906 | 0.0055906 | 0.0055906 | 0.0055906 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055907 | 0.0055911 | 0.0055916 | 0.0055925 | 0.0055939 | 0.0055972 | 0.0056039 | 0.0056125 | 0.0056064 | 0.0055905
| Temperature coeff. | B | Rect. | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014445 | 0.0014446 | 0.0014455 | 0.0014469 | 0.0014492 | 0.001453 | 0.0014614 | 0.0014783 | 0.0015001 | 0.0014846 | 0.0014436
| Pressure coeff. | B | Rect. | 0.0043301 | 0.0043301 | 0.0043301 | 0.0043301 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043302 | 0.0043301 | 0.0043301 | 0.0043301 | 0.0043301 | 0.0043303
| Discrepancy by frequency drift | B | Rect. | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Rounding error | B | Rect. | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868
| Combined standard uncertainty |  |  | 0.11015 | 0.10624 | 0.10615 | 0.094112 | 0.043273 | 0.037137 | 0.024143 | 0.020306 | 0.018522 | 0.019432 | 0.018313 | 0.016244 | 0.016737 | 0.016443 | 0.017011 | 0.016195 | 0.015866 | 0.017468 | 0.0169 | 0.01731 | 0.016456 | 0.016042 | 0.017148 | 0.017793 | 0.017618 | 0.016892 | 0.017528 | 0.016513 | 0.017243 | 0.016109 | 0.015982 | 0.015346 | 0.015111 | 0.015211 | 0.015243 | 0.015144 | 0.015333 | 0.040295
| Expanded uncertainty (k=2) |  |  | 0.22029 | 0.21249 | 0.21229 | 0.18822 | 0.086545 | 0.074274 | 0.048286 | 0.040612 | 0.037043 | 0.038863 | 0.036627 | 0.032488 | 0.033475 | 0.032885 | 0.034021 | 0.032391 | 0.031733 | 0.034936 | 0.0338 | 0.03462 | 0.032912 | 0.032084 | 0.034297 | 0.035585 | 0.035236 | 0.033784 | 0.035057 | 0.033026 | 0.034486 | 0.032218 | 0.031963 | 0.030692 | 0.030222 | 0.030421 | 0.030486 | 0.030289 | 0.030667 | 0.080591
| Expanded uncertainty (Stated) |  |  | 0.23 | 0.22 | 0.22 | 0.19 | 0.09 | 0.08 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.09

## NMIJ Phase
NMIJ | Phase
Freq. [Hz]
| Uncertainty components | Type | Dist. | 1.9953 | 2.5119 | 3.1623 | 3.9811 | 5.0119 | 6.3096 | 7.9433 | 10 | 12.5893 | 15.8489 | 19.9526 | 25.1189 | 31.6228 | 39.8107 | 50.1187 | 63.0957 | 79.4328 | 100 | 125.8925 | 158.4893 | 199.5262 | 251.1886 | 316.2278 | 398.1072 | 501.1872 | 630.9573 | 794.3282 | 1000 | 1258.9254 | 1584.8932 | 1995.2623 | 2511.8864 | 3162.2777 | 3981.0717 | 5011.8723 | 6309.5734 | 7943.2823 | 10000
| Electrical impedance of transmitter |  | Combined | 0.005878 | 0.0061311 | 0.0073791 | 0.008182 | 0.0083515 | 0.0088495 | 0.0060305 | 0.0055212 | 0.0049498 | 0.0091469 | 0.0081777 | 0.0072634 | 0.0056859 | 0.0044879 | 0.004348 | 0.0019774 | 0.0017463 | 0.0015969 | 0.001457 | 0.0012985 | 0.0011958 | 0.0011405 | 0.0010837 | 0.0010004 | 0.0009658 | 0.00087369 | 0.00094266 | 0.00078793 | 0.00091984 | 0.0007 | 0.00079443 | 0.0008531 | 0.00097781 | 0.0017243 | 0.0013189 | 0.0021275 | 0.0010214 | 0.0022127
| Reference capacitance | B | Rect. | 0.00016667 | 0.0007 | 0.00043333 | 0.00145 | 0.00135 | 0.00093333 | 0.0014333 | 0.00185 | 0.0021 | 0.00205 | 0.0045667 | 0.0040167 | 0.0033833 | 0.0028667 | 0.00265 | 0.0013667 | 0.0013167 | 0.00135 | 0.0012833 | 0.0011833 | 0.00115 | 0.0011 | 0.0010167 | 0.0009 | 0.00088333 | 0.00085 | 0.00076667 | 0.0007 | 0.00071667 | 0.0006 | 0.00046667 | 0.00051667 | 0.00041667 | 0.0005 | 0.00041667 | 0.00058333 | 0.00055 | 0.00061667
| Stray capacitance | B | Rect. | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Repeatability | A | Normal | 0.0058757 | 0.006091 | 0.0073664 | 0.0080525 | 0.0082417 | 0.0088002 | 0.0058577 | 0.005202 | 0.0044823 | 0.0089142 | 0.0067839 | 0.0060517 | 0.0045697 | 0.003453 | 0.0034471 | 0.0014292 | 0.0011471 | 0.00085294 | 0.00068981 | 0.00053463 | 0.00032787 | 0.00030139 | 0.00037528 | 0.00043684 | 0.00039051 | 0.00020207 | 0.00054848 | 0.00036171 | 0.00057663 | 0.00036056 | 0.00064291 | 0.00067885 | 0.00088459 | 0.0016503 | 0.0012513 | 0.0020459 | 0.00086072 | 0.002125
| Voltage transfer function between transmitter and receiver |  | Combined | 1.249 | 1.2147 | 1.1093 | 0.71611 | 0.58846 | 0.50018 | 0.37586 | 0.29668 | 0.22625 | 0.18227 | 0.14154 | 0.11974 | 0.093284 | 0.076572 | 0.065887 | 0.057572 | 0.046022 | 0.038095 | 0.030611 | 0.026828 | 0.025515 | 0.020574 | 0.020287 | 0.020545 | 0.023866 | 0.023052 | 0.028722 | 0.033038 | 0.036133 | 0.045095 | 0.058539 | 0.05558 | 0.059367 | 0.062327 | 0.063315 | 0.12804 | 0.30137 | 0.52574
| Attenuator | B | Rect. | 0.00011547 | 5.7735e-05 | 0 | 0 | 5.7735e-05 | 5.7735e-05 | 0 | 0.00011547 | 0 | 5.7735e-05 | 5.7735e-05 | 0.00028868 | 0.00023094 | 0 | 0.00046188 | 0.00023094 | 0.00040415 | 0.001963 | 0.00017321 | 0.001963 | 0.00080829 | 0.00075056 | 0.0050229 | 0.0048497 | 0.011201 | 0.0040415 | 0.0040415 | 0.0021362 | 0.001963 | 0.010335 | 0.0092953 | 0.010796 | 0.01403 | 0.013337 | 0.017378 | 0.016743 | 0.029041 | 0.030369
| Crosstalk | B | Rect. | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981 | 0.0025981
| Linearity of FFT analyzer | B | Rect. | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005
| Static pressure | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | -----
| Aoustic coupler difference | B | Rect. | 0.69282 | 0.69282 | 0.68532 | 0.21752 | 0.21651 | 0.20844 | 0.1299 | 0.12425 | 0.096741 | 0.082375 | 0.079821 | 0.066754 | 0.059648 | 0.048482 | 0.03954 | 0.036817 | 0.029757 | 0.028162 | 0.021449 | 0.017321 | 0.017321 | 0.017321 | 0.017321 | 0.017321 | 0.017321 | 0.017321 | 0.017321 | 0.017321 | 0.017321 | 0.017321 | 0.017321 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Polarizing voltage | B | Rect. | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Repeatability | A | Normal | 1.0392 | 0.99775 | 0.87225 | 0.68225 | 0.54715 | 0.45465 | 0.35265 | 0.26935 | 0.20445 | 0.1625 | 0.11675 | 0.09925 | 0.0715 | 0.059 | 0.0524 | 0.0439 | 0.03465 | 0.02495 | 0.0211 | 0.0196 | 0.01785 | 0.009538 | 0.0073888 | 0.0081752 | 0.0106 | 0.013539 | 0.021837 | 0.027481 | 0.031144 | 0.039937 | 0.054851 | 0.054229 | 0.05741 | 0.060622 | 0.060622 | 0.12682 | 0.29991 | 0.52484
| Coupler volume | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | -----
| Specific heat ratio | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | ----- | -----
| Coupler correction |  | Combined | 0.0071474 | 0.0065819 | 0.0060393 | 0.0055211 | 0.005031 | 0.0045694 | 0.0041357 | 0.00373 | 0.0033503 | 0.0029959 | 0.0026628 | 0.0023491 | 0.0020529 | 0.0017698 | 0.0014996 | 0.0012434 | 0.0010131 | 0.00084253 | 0.00090719 | 0.0012186 | 0.0016918 | 0.0023164 | 0.0031039 | 0.0040876 | 0.0053129 | 0.00684 | 0.0087454 | 0.011122 | 0.01408 | 0.017736 | 0.022174 | 0.027295 | 0.032327 | 0.034426 | 0.033547 | 0.02067 | 0.066818 | 0.086042
| Environmental corrections of microphone sensitivities |  | Combined | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.01301 | 0.013011 | 0.013011 | 0.013012 | 0.013013 | 0.013015 | 0.013017 | 0.013021 | 0.013026 | 0.013031 | 0.01303 | 0.013016 | 0.013026 | 0.013271 | 0.013452
| Temperature coeff. | B | Rect. | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0072171 | 0.007218 | 0.0072201 | 0.0072224 | 0.0072255 | 0.0072295 | 0.0072355 | 0.0072445 | 0.0072533 | 0.007253 | 0.007227 | 0.0072438 | 0.0076721 | 0.0079786
| Pressure coeff. | B | Rect. | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010825 | 0.010826 | 0.010829 | 0.010831
| Discrepancy by frequency drift | B | Rect. | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Rounding error | B | Rect. | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868
| Combined standard uncertainty |  |  | 1.2491 | 1.2148 | 1.1094 | 0.7163 | 0.58869 | 0.50046 | 0.37616 | 0.29705 | 0.22672 | 0.18301 | 0.14243 | 0.12072 | 0.094425 | 0.077872 | 0.067379 | 0.05914 | 0.047955 | 0.040399 | 0.03343 | 0.030009 | 0.02886 | 0.024649 | 0.024495 | 0.024848 | 0.027863 | 0.027506 | 0.032863 | 0.03733 | 0.041017 | 0.050263 | 0.064008 | 0.063347 | 0.06891 | 0.072463 | 0.072895 | 0.1304 | 0.30899 | 0.53292
| Expanded uncertainty (k=2) |  |  | 2.4982 | 2.4296 | 2.2188 | 1.4326 | 1.1774 | 1.0009 | 0.75233 | 0.59411 | 0.45345 | 0.36603 | 0.28485 | 0.24145 | 0.18885 | 0.15574 | 0.13476 | 0.11828 | 0.09591 | 0.080798 | 0.066861 | 0.060018 | 0.057721 | 0.049297 | 0.048989 | 0.049696 | 0.055726 | 0.055011 | 0.065725 | 0.07466 | 0.082034 | 0.10053 | 0.12802 | 0.12669 | 0.13782 | 0.14493 | 0.14579 | 0.26081 | 0.61798 | 1.0658
| Expanded uncertainty (Stated) |  |  | 2.5 | 2.43 | 2.22 | 1.44 | 1.18 | 1.01 | 0.76 | 0.6 | 0.46 | 0.37 | 0.29 | 0.25 | 0.19 | 0.16 | 0.14 | 0.12 | 0.1 | 0.09 | 0.07 | 0.07 | 0.06 | 0.05 | 0.05 | 0.05 | 0.06 | 0.06 | 0.07 | 0.08 | 0.09 | 0.11 | 0.13 | 0.13 | 0.14 | 0.15 | 0.15 | 0.27 | 0.62 | 1.07

## KRISS Level
KRISS | Level
| Uncertainty Budget for LS1P (Modulus) | S/N | 2652762
| MEASURED QUANTITY | Type | Unc. | Unit
| Frequency |  |  | Hz | 1.9953 | 2.5119 | 3.1623 | 3.9811 | 5.0119 | 6.3096 | 7.9433 | 10 | 12.5893 | 15.8489 | 19.9526 | 25.1189 | 31.6228 | 39.8107 | 50.1187 | 63.0957 | 79.4328 | 100 | 125.8925 | 158.4893 | 199.5262 | 251.1886 | 316.2278 | 398.1072 | 501.1872 | 630.9573 | 794.3282 | 1000 | 1258.9254 | 1584.8932 | 1995.2623 | 2511.8864 | 3162.2777 | 3981.0717 | 5011.8723 | 6309.5734 | 7943.2823 | 10000
| Electrical Transfer Impedance |  |  |  | 0.026412 | 0.026351 | 0.026293 | 0.026237 | 0.026182 | 0.02613 | 0.02608 | 0.026031 | 0.025985 | 0.025942 | 0.020913 | 0.020725 | 0.020585 | 0.021969 | 0.013571 | 0.0068915 | 0.0063462 | 0.0062616 | 0.0062026 | 0.0059188 | 0.0059087 | 0.0058999 | 0.005889 | 0.0057808 | 0.0057916 | 0.0057677 | 0.005768 | 0.0057611 | 0.0057267 | 0.0057511 | 0.0057553 | 0.0057331 | 0.0057487 | 0.0057547 | 0.0058772 | 0.0062145 | 0.0068645 | 0.0090671
| Series Impedance | B | Table | nF | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.017635 | 0.0040515 | 0.0040515 | 0.0032893 | 0.0032893 | 0.0032893 | 0.002902 | 0.002902 | 0.002902 | 0.002902 | 0.0026309 | 0.0026309 | 0.0026309 | 0.0026307 | 0.0026307 | 0.0026307 | 0.0026307 | 0.0027059 | 0.0027059 | 0.0027059 | 0.0026117 | 0.0026117 | 0.0026117 | 0.0029013 | 0.0029013
| Voltage Ratio | B | Table |  | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.0075225 | 0.0075 | 0.0075 | 0.01183 | 0.0118 | 0.0022 | 0.0021931 | 0.0022 | 0.0022 | 0.0020655 | 0.0021 | 0.0021 | 0.0020655 | 0.0021 | 0.0021 | 0.0020071 | 0.002 | 0.002 | 0.0019528 | 0.0021 | 0.0021 | 0.0021 | 0.0021287 | 0.0021 | 0.0020071 | 0.0021 | 0.0021 | 0.0052
| Cross-talk | B | Meas. |  | 0.0045136 | 0.0044861 | 0.0044648 | 0.0044489 | 0.0044378 | 0.0044308 | 0.0044274 | 0.004427 | 0.0044291 | 0.0044333 | 0.0044393 | 0.0044733 | 0.0045116 | 0.0045438 | 0.0045651 | 0.0045747 | 0.0045746 | 0.0045679 | 0.0045585 | 0.0045498 | 0.0045443 | 0.0045435 | 0.0045471 | 0.0045538 | 0.0045608 | 0.0045649 | 0.0045627 | 0.0045518 | 0.0045315 | 0.0045043 | 0.0044772 | 0.0044637 | 0.0044852 | 0.0045736 | 0.004773 | 0.0051427 | 0.0057596 | 0.0067213
| Inherent Noise | B | Meas. |  | 0.0074478 | 0.0072491 | 0.0070479 | 0.006845 | 0.0066409 | 0.0064362 | 0.0062315 | 0.0060273 | 0.0058241 | 0.0056223 | 0.0054224 | 0.0046491 | 0.0039333 | 0.0032903 | 0.0027296 | 0.0022561 | 0.0018703 | 0.001569 | 0.0013461 | 0.0011933 | 0.0011001 | 0.0010554 | 0.0010469 | 0.0010627 | 0.0010911 | 0.0011215 | 0.0011447 | 0.0011535 | 0.0011431 | 0.0011118 | 0.001061 | 0.0009958 | 0.00092567 | 0.00086447 | 0.000831 | 0.0008493 | 0.00094903 | 0.0011658
| Distortion | B | Meas. |  | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002
| Frequency | B | Table | Hz | 0.000136 | 0.00012 | 0.000104 | 8.8e-05 | 7.2e-05 | 6e-05 | 5.2e-05 | 4.8e-05 | 4.4e-05 | 4e-05 | 3.4e-05 | 3e-05 | 2.6e-05 | 2.2e-05 | 1.8e-05 | 1.5e-05 | 1.3e-05 | 1.2e-05 | 1.1e-05 | 1e-05 | 9e-06 | 8e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 6e-06 | 6e-06 | 5e-06 | 5e-06 | 5e-06
| Reveiver Ground Shield | B | B&K |  | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003
| Transmitter Ground Shield | B | B&K |  | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003
| Coupler Properties |  |  |  | 0.00053751 | 0.000536 | 0.00053682 | 0.00053799 | 0.00053799 | 0.00053649 | 0.00053768 | 0.00053737 | 0.00053737 | 0.00053588 | 0.00053708 | 0.00053708 | 0.00053708 | 0.00053828 | 0.0005356 | 0.0005368 | 0.0005368 | 0.0005368 | 0.0005368 | 0.0005368 | 0.0005368 | 0.0005368 | 0.0005368 | 0.0005368 | 0.00053801 | 0.00053828 | 0.00053708 | 0.00053737 | 0.00053649 | 0.000536 | 0.00053639 | 0.00053538 | 0.00053744 | 0.00054762 | 0.0005688 | 0.00062153 | 0.0007288 | 0.0010081
| Coupler Length | B | 0.001 | mm | 0.00032616 | 0.00032616 | 0.00032997 | 0.00033188 | 0.00033188 | 0.00033188 | 0.00033379 | 0.00033569 | 0.00033569 | 0.00033569 | 0.0003376 | 0.0003376 | 0.0003376 | 0.00033951 | 0.0003376 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00034142 | 0.00033951 | 0.0003376 | 0.00033569 | 0.00033188 | 0.00032616 | 0.00031662 | 0.00030136 | 0.00027847 | 0.00024414 | 0.00018501 | 9.5367e-05 | 5.5313e-05 | 0.00038719
| Coupler Diameter | B | 0.0015 | mm | 0.00042725 | 0.00042534 | 0.00042343 | 0.00042343 | 0.00042343 | 0.00042152 | 0.00042152 | 0.00041962 | 0.00041962 | 0.00041771 | 0.00041771 | 0.00041771 | 0.00041771 | 0.00041771 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.00041771 | 0.00041771 | 0.00041962 | 0.00042152 | 0.00042534 | 0.00043297 | 0.0004425 | 0.00045967 | 0.00049019 | 0.00053787 | 0.00061417 | 0.0007267 | 0.00093079
| Coupler Volume | B | 0.4266 | mm3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Coupler Surface Area | B | 0.1362 | mm2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Environmental Conditions |  |  |  | 0.002219 | 0.0022138 | 0.0022098 | 0.0022079 | 0.0022069 | 0.0022058 | 0.0022068 | 0.0022074 | 0.0022092 | 0.0022109 | 0.002213 | 0.0022185 | 0.0022252 | 0.0022303 | 0.0022322 | 0.0022318 | 0.0022306 | 0.0022289 | 0.0022277 | 0.0022277 | 0.0022303 | 0.0022327 | 0.0022363 | 0.0022392 | 0.0022411 | 0.0022395 | 0.0022343 | 0.0022226 | 0.0022039 | 0.0021761 | 0.0021475 | 0.0021488 | 0.002298 | 0.0028679 | 0.0040811 | 0.005542 | 0.0056331 | 0.0065715
| Static Pressure | B | 2.26 | kPa | 0.0021368 | 0.002131 | 0.0021265 | 0.0021231 | 0.0021206 | 0.0021189 | 0.0021179 | 0.0021176 | 0.0021177 | 0.0021183 | 0.0021192 | 0.0021236 | 0.0021285 | 0.002132 | 0.0021335 | 0.0021332 | 0.0021317 | 0.0021298 | 0.0021282 | 0.0021275 | 0.0021278 | 0.0021288 | 0.0021296 | 0.0021286 | 0.0021238 | 0.0021125 | 0.0020916 | 0.0020576 | 0.0020062 | 0.001931 | 0.0018192 | 0.0016435 | 0.0013526 | 0.00087749 | 0.00022097 | 0.0002191 | 0.00099098 | 0.0059804
| Temperature | B | 0.7 | K | 0.00057945 | 0.00057368 | 0.00056904 | 0.00056537 | 0.00056251 | 0.00056033 | 0.00055868 | 0.00055746 | 0.00055653 | 0.00055579 | 0.00055514 | 0.00055735 | 0.00055985 | 0.00056083 | 0.00055952 | 0.00055598 | 0.00055075 | 0.00054475 | 0.00053903 | 0.00053474 | 0.00053309 | 0.00053547 | 0.00054349 | 0.00055907 | 0.00058437 | 0.00062143 | 0.00067155 | 0.00073476 | 0.00081089 | 0.00090501 | 0.001041 | 0.0012839 | 0.0017592 | 0.0026306 | 0.0039663 | 0.0053973 | 0.0052825 | 0.00093868
| Relative Humidity | B | 5.85 | % | 0.00014925 | 0.00017413 | 0.00019403 | 0.0002189 | 0.0002388 | 0.00024876 | 0.00026866 | 0.00027861 | 0.00029353 | 0.00030348 | 0.00031343 | 0.00031841 | 0.00032836 | 0.00033831 | 0.00034328 | 0.00034826 | 0.00035821 | 0.00036816 | 0.00037811 | 0.00038806 | 0.00040298 | 0.00040796 | 0.00041293 | 0.00041293 | 0.00041293 | 0.00040796 | 0.00040796 | 0.00040796 | 0.00041791 | 0.00043283 | 0.00046766 | 0.00051741 | 0.00059701 | 0.00073134 | 0.00093532 | 0.0012388 | 0.0016866 | 0.0025572
| Microphone Parameters |  |  |  | 0.010657 | 0.010751 | 0.010889 | 0.011019 | 0.011142 | 0.011202 | 0.01131 | 0.011357 | 0.011401 | 0.011493 | 0.011583 | 0.011585 | 0.011665 | 0.011688 | 0.011709 | 0.011726 | 0.011797 | 0.01181 | 0.011822 | 0.011836 | 0.011843 | 0.011847 | 0.011847 | 0.011846 | 0.011838 | 0.011822 | 0.011788 | 0.011737 | 0.011597 | 0.011413 | 0.013085 | 0.012972 | 0.012563 | 0.011972 | 0.010732 | 0.0069508 | 0.0031819 | 0.011963
| Front Cavity Depth | B | 0.0025 | mm | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0 | 0.0001
| Front Cavity Volume | B | 2.59 | mm3 | 0.0058 | 0.0058 | 0.0059 | 0.006 | 0.0061 | 0.0061 | 0.0062 | 0.0062 | 0.0062 | 0.0063 | 0.0063 | 0.0063 | 0.0064 | 0.0064 | 0.0064 | 0.0064 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0064 | 0.0063 | 0.0062 | 0.0059 | 0.0055 | 0.0049 | 0.0039 | 0.0021 | 0.0011 | 0.0078
| Equivalent Volume | B | 6.9 | mm3 | 0.0089306 | 0.009043 | 0.0091428 | 0.0092327 | 0.0093146 | 0.009386 | 0.0094499 | 0.0095065 | 0.0095587 | 0.0096036 | 0.0097122 | 0.0097142 | 0.0097448 | 0.0097722 | 0.009797 | 0.0098176 | 0.009837 | 0.0098527 | 0.0098662 | 0.0098837 | 0.0098914 | 0.0098963 | 0.0098968 | 0.0098954 | 0.0098851 | 0.0098644 | 0.0098244 | 0.0097569 | 0.0096445 | 0.0094553 | 0.011403 | 0.011289 | 0.010648 | 0.0095483 | 0.0076719 | 0.0047501 | 0.0024524 | 0.005233
| Resonance Frequency | B | 150 | Hz | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0003 | 0.0006 | 0.0014 | 0.0027 | 0.0031 | 0.0007 | 0.0012
| Loss Factor | B | 0.18 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0004 | 0.0006 | 0.001 | 0.0016 | 0.0024 | 0.0037 | 0.0051 | 0.0058 | 0.0034 | 0.0015 | 0.0073
Additional Heat Conduction
| Caused by Front Cavity Thread | N.A. | Table
| Polarizing Voltage | B |  | V | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004
| Imperfection of Theory |  |  |  | 0.014871 | 3.1729e-05 | 0.0062044 | 0.0064707 | 0.0044762 | 0.002484 | 0.0014048 | 0.0011021 | 0.0010359 | 0.00095605 | 0.00085384 | 0.00075648 | 0.00067051 | 0.00059334 | 0.00052959 | 0.00047099 | 0.00041738 | 0.0003729 | 0.00033317 | 0.00029409 | 0.00026264 | 0.00023442 | 0.00020814 | 0.00018367 | 0.00016285 | 0.00014286 | 0.00012324 | 0.0001051 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Heat Conduction Theory | N.A. |  |  | 0.014871 | 3.1729e-05 | 0.0062044 | 0.0064707 | 0.0044762 | 0.002484 | 0.0014048 | 0.0011021 | 0.0010359 | 0.00095605 | 0.00085384 | 0.00075648 | 0.00067051 | 0.00059334 | 0.00052959 | 0.00047099 | 0.00041738 | 0.0003729 | 0.00033317 | 0.00029409 | 0.00026264 | 0.00023442 | 0.00020814 | 0.00018367 | 0.00016285 | 0.00014286 | 0.00012324 | 0.0001051 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Adding of Excess Volume | N.A.
| Radial Wave Motion | N.A.
| Processing of Results |  |  |  | 0.02117 | 0.0076023 | 0.0032049 | 0.011973 | 0.0078377 | 0.0090018 | 0.0070208 | 0.0062377 | 0.0074699 | 0.006933 | 0.0086629 | 0.010354 | 0.011735 | 0.011735 | 0.010255 | 0.010553 | 0.010553 | 0.0096941 | 0.010879 | 0.010879 | 0.0039511 | 0.0039512 | 0.003826 | 0.0040048 | 0.0044386 | 0.0041952 | 0.004513 | 0.0045627 | 0.0045627 | 0.0037536 | 0.0045685 | 0.0051539 | 0.0050955 | 0.0052692 | 0.0065403 | 0.0088751 | 0.0095165 | 0.0082346
| Rounding Error | B | 0.005 | dB | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868
| Repeatability of Measurements | A | Table | dB | 0.020972 | 0.0070329 | 0.0013921 | 0.011619 | 0.0072868 | 0.0085264 | 0.0063999 | 0.0055295 | 0.0068896 | 0.0063035 | 0.0081678 | 0.009943 | 0.011374 | 0.011374 | 0.00984 | 0.010151 | 0.010151 | 0.0092543 | 0.010489 | 0.010489 | 0.0026977 | 0.002698 | 0.002511 | 0.0027758 | 0.0033716 | 0.003044 | 0.003469 | 0.0035334 | 0.0035334 | 0.0023993 | 0.0035409 | 0.0042696 | 0.0041989 | 0.004408 | 0.0058688 | 0.0083925 | 0.0090681 | 0.0077121
| Combined Uncertainty(Sum) |  |  |  | 0.038544 | 0.029546 | 0.029391 | 0.031625 | 0.029938 | 0.03001 | 0.029402 | 0.029187 | 0.029449 | 0.029312 | 0.025544 | 0.026014 | 0.026518 | 0.027615 | 0.020784 | 0.017374 | 0.017212 | 0.016675 | 0.017376 | 0.017286 | 0.014004 | 0.014003 | 0.013964 | 0.013969 | 0.014097 | 0.013998 | 0.014068 | 0.014036 | 0.013902 | 0.013507 | 0.015169 | 0.015251 | 0.014913 | 0.014585 | 0.014473 | 0.014029 | 0.013419 | 0.018367
| Expanded Uncertainty (k=2) |  |  |  | 0.077089 | 0.059092 | 0.058782 | 0.06325 | 0.059876 | 0.06002 | 0.058804 | 0.058375 | 0.058899 | 0.058625 | 0.051088 | 0.052028 | 0.053037 | 0.055229 | 0.041568 | 0.034748 | 0.034425 | 0.03335 | 0.034753 | 0.034573 | 0.028007 | 0.028007 | 0.027929 | 0.027938 | 0.028193 | 0.027996 | 0.028135 | 0.028072 | 0.027803 | 0.027014 | 0.030338 | 0.030502 | 0.029827 | 0.02917 | 0.028946 | 0.028057 | 0.026838 | 0.036734
| Stated Uncertainty |  |  |  | 0.09 | 0.08 | 0.07 | 0.07 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.07
| Uncertainty Budget for LS1P (Modulus) | S/N | 2652765
| MEASURED QUANTITY | Type | Unc. | Unit
| Frequency |  |  | Hz | 1.9953 | 2.5119 | 3.1623 | 3.9811 | 5.0119 | 6.3096 | 7.9433 | 10 | 12.5893 | 15.8489 | 19.9526 | 25.1189 | 31.6228 | 39.8107 | 50.1187 | 63.0957 | 79.4328 | 100 | 125.8925 | 158.4893 | 199.5262 | 251.1886 | 316.2278 | 398.1072 | 501.1872 | 630.9573 | 794.3282 | 1000 | 1258.9254 | 1584.8932 | 1995.2623 | 2511.8864 | 3162.2777 | 3981.0717 | 5011.8723 | 6309.5734 | 7943.2823 | 10000
| Electrical Transfer Impedance |  |  |  | 0.026412 | 0.026351 | 0.026293 | 0.026237 | 0.026182 | 0.02613 | 0.02608 | 0.026031 | 0.025985 | 0.025942 | 0.020913 | 0.020725 | 0.020585 | 0.021969 | 0.013571 | 0.0068915 | 0.0063462 | 0.0062616 | 0.0062026 | 0.0059188 | 0.0059087 | 0.0058999 | 0.005889 | 0.0057808 | 0.0057916 | 0.0057677 | 0.005768 | 0.0057611 | 0.0057267 | 0.0057511 | 0.0057553 | 0.0057331 | 0.0057487 | 0.0057547 | 0.0058772 | 0.0062145 | 0.0068645 | 0.0090671
| Series Impedance | B | Table | nF | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.018206 | 0.017635 | 0.0040515 | 0.0040515 | 0.0032893 | 0.0032893 | 0.0032893 | 0.002902 | 0.002902 | 0.002902 | 0.002902 | 0.0026309 | 0.0026309 | 0.0026309 | 0.0026307 | 0.0026307 | 0.0026307 | 0.0026307 | 0.0027059 | 0.0027059 | 0.0027059 | 0.0026117 | 0.0026117 | 0.0026117 | 0.0029013 | 0.0029013
| Voltage Ratio | B | Table |  | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.017031 | 0.0075225 | 0.0075 | 0.0075 | 0.01183 | 0.0118 | 0.0022 | 0.0021931 | 0.0022 | 0.0022 | 0.0020655 | 0.0021 | 0.0021 | 0.0020655 | 0.0021 | 0.0021 | 0.0020071 | 0.002 | 0.002 | 0.0019528 | 0.0021 | 0.0021 | 0.0021 | 0.0021287 | 0.0021 | 0.0020071 | 0.0021 | 0.0021 | 0.0052
| Cross-talk | B | Meas. |  | 0.0045136 | 0.0044861 | 0.0044648 | 0.0044489 | 0.0044378 | 0.0044308 | 0.0044274 | 0.004427 | 0.0044291 | 0.0044333 | 0.0044393 | 0.0044733 | 0.0045116 | 0.0045438 | 0.0045651 | 0.0045747 | 0.0045746 | 0.0045679 | 0.0045585 | 0.0045498 | 0.0045443 | 0.0045435 | 0.0045471 | 0.0045538 | 0.0045608 | 0.0045649 | 0.0045627 | 0.0045518 | 0.0045315 | 0.0045043 | 0.0044772 | 0.0044637 | 0.0044852 | 0.0045736 | 0.004773 | 0.0051427 | 0.0057596 | 0.0067213
| Inherent Noise | B | Meas. |  | 0.0074478 | 0.0072491 | 0.0070479 | 0.006845 | 0.0066409 | 0.0064362 | 0.0062315 | 0.0060273 | 0.0058241 | 0.0056223 | 0.0054224 | 0.0046491 | 0.0039333 | 0.0032903 | 0.0027296 | 0.0022561 | 0.0018703 | 0.001569 | 0.0013461 | 0.0011933 | 0.0011001 | 0.0010554 | 0.0010469 | 0.0010627 | 0.0010911 | 0.0011215 | 0.0011447 | 0.0011535 | 0.0011431 | 0.0011118 | 0.001061 | 0.0009958 | 0.00092567 | 0.00086447 | 0.000831 | 0.0008493 | 0.00094903 | 0.0011658
| Distortion | B | Meas. |  | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002
| Frequency | B | Table | Hz | 0.000136 | 0.00012 | 0.000104 | 8.8e-05 | 7.2e-05 | 6e-05 | 5.2e-05 | 4.8e-05 | 4.4e-05 | 4e-05 | 3.4e-05 | 3e-05 | 2.6e-05 | 2.2e-05 | 1.8e-05 | 1.5e-05 | 1.3e-05 | 1.2e-05 | 1.1e-05 | 1e-05 | 9e-06 | 8e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 7e-06 | 6e-06 | 6e-06 | 5e-06 | 5e-06 | 5e-06
| Reveiver Ground Shield | B | B&K |  | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003
| Transmitter Ground Shield | B | B&K |  | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003 | 0.0003
| Coupler Properties |  |  |  | 0.00053751 | 0.000536 | 0.00053682 | 0.00053799 | 0.00053799 | 0.00053649 | 0.00053768 | 0.00053737 | 0.00053737 | 0.00053588 | 0.00053708 | 0.00053708 | 0.00053708 | 0.00053828 | 0.0005356 | 0.0005368 | 0.0005368 | 0.0005368 | 0.0005368 | 0.0005368 | 0.0005368 | 0.0005368 | 0.0005368 | 0.0005368 | 0.00053801 | 0.00053828 | 0.00053708 | 0.00053737 | 0.00053649 | 0.000536 | 0.00053639 | 0.00053538 | 0.00053744 | 0.00054762 | 0.0005688 | 0.00062153 | 0.0007288 | 0.0010081
| Coupler Length | B | 0.001 | mm | 0.00032616 | 0.00032616 | 0.00032997 | 0.00033188 | 0.00033188 | 0.00033188 | 0.00033379 | 0.00033569 | 0.00033569 | 0.00033569 | 0.0003376 | 0.0003376 | 0.0003376 | 0.00033951 | 0.0003376 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00033951 | 0.00034142 | 0.00033951 | 0.0003376 | 0.00033569 | 0.00033188 | 0.00032616 | 0.00031662 | 0.00030136 | 0.00027847 | 0.00024414 | 0.00018501 | 9.5367e-05 | 5.5313e-05 | 0.00038719
| Coupler Diameter | B | 0.0015 | mm | 0.00042725 | 0.00042534 | 0.00042343 | 0.00042343 | 0.00042343 | 0.00042152 | 0.00042152 | 0.00041962 | 0.00041962 | 0.00041771 | 0.00041771 | 0.00041771 | 0.00041771 | 0.00041771 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.0004158 | 0.00041771 | 0.00041771 | 0.00041962 | 0.00042152 | 0.00042534 | 0.00043297 | 0.0004425 | 0.00045967 | 0.00049019 | 0.00053787 | 0.00061417 | 0.0007267 | 0.00093079
| Coupler Volume | B | 0.4266 | mm3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Coupler Surface Area | B | 0.1362 | mm2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Environmental Conditions |  |  |  | 0.002219 | 0.0022138 | 0.0022098 | 0.0022079 | 0.0022069 | 0.0022058 | 0.0022068 | 0.0022074 | 0.0022092 | 0.0022109 | 0.002213 | 0.0022185 | 0.0022252 | 0.0022303 | 0.0022322 | 0.0022318 | 0.0022306 | 0.0022289 | 0.0022277 | 0.0022277 | 0.0022303 | 0.0022327 | 0.0022363 | 0.0022392 | 0.0022411 | 0.0022395 | 0.0022343 | 0.0022226 | 0.0022039 | 0.0021761 | 0.0021475 | 0.0021488 | 0.002298 | 0.0028679 | 0.0040811 | 0.005542 | 0.0056331 | 0.0065715
| Static Pressure | B | 2.26 | kPa | 0.0021368 | 0.002131 | 0.0021265 | 0.0021231 | 0.0021206 | 0.0021189 | 0.0021179 | 0.0021176 | 0.0021177 | 0.0021183 | 0.0021192 | 0.0021236 | 0.0021285 | 0.002132 | 0.0021335 | 0.0021332 | 0.0021317 | 0.0021298 | 0.0021282 | 0.0021275 | 0.0021278 | 0.0021288 | 0.0021296 | 0.0021286 | 0.0021238 | 0.0021125 | 0.0020916 | 0.0020576 | 0.0020062 | 0.001931 | 0.0018192 | 0.0016435 | 0.0013526 | 0.00087749 | 0.00022097 | 0.0002191 | 0.00099098 | 0.0059804
| Temperature | B | 0.7 | K | 0.00057945 | 0.00057368 | 0.00056904 | 0.00056537 | 0.00056251 | 0.00056033 | 0.00055868 | 0.00055746 | 0.00055653 | 0.00055579 | 0.00055514 | 0.00055735 | 0.00055985 | 0.00056083 | 0.00055952 | 0.00055598 | 0.00055075 | 0.00054475 | 0.00053903 | 0.00053474 | 0.00053309 | 0.00053547 | 0.00054349 | 0.00055907 | 0.00058437 | 0.00062143 | 0.00067155 | 0.00073476 | 0.00081089 | 0.00090501 | 0.001041 | 0.0012839 | 0.0017592 | 0.0026306 | 0.0039663 | 0.0053973 | 0.0052825 | 0.00093868
| Relative Humidity | B | 5.85 | % | 0.00014925 | 0.00017413 | 0.00019403 | 0.0002189 | 0.0002388 | 0.00024876 | 0.00026866 | 0.00027861 | 0.00029353 | 0.00030348 | 0.00031343 | 0.00031841 | 0.00032836 | 0.00033831 | 0.00034328 | 0.00034826 | 0.00035821 | 0.00036816 | 0.00037811 | 0.00038806 | 0.00040298 | 0.00040796 | 0.00041293 | 0.00041293 | 0.00041293 | 0.00040796 | 0.00040796 | 0.00040796 | 0.00041791 | 0.00043283 | 0.00046766 | 0.00051741 | 0.00059701 | 0.00073134 | 0.00093532 | 0.0012388 | 0.0016866 | 0.0025572
| Microphone Parameters |  |  |  | 0.0066593 | 0.0066788 | 0.0067836 | 0.0068865 | 0.0069881 | 0.0070006 | 0.0070994 | 0.0071094 | 0.0071181 | 0.0072134 | 0.0072218 | 0.0072324 | 0.0073249 | 0.0073297 | 0.0073338 | 0.0073375 | 0.0074284 | 0.0074318 | 0.0074338 | 0.0074366 | 0.0074379 | 0.0074386 | 0.0074386 | 0.0074388 | 0.0074373 | 0.0074356 | 0.0074287 | 0.0074247 | 0.0073317 | 0.0072603 | 0.0076429 | 0.0075966 | 0.007711 | 0.008011 | 0.0080009 | 0.0053539 | 0.002233 | 0.010933
| Front Cavity Depth | B | 0.0025 | mm | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0 | 0.0001
| Front Cavity Volume | B | 2.59 | mm3 | 0.0058 | 0.0058 | 0.0059 | 0.006 | 0.0061 | 0.0061 | 0.0062 | 0.0062 | 0.0062 | 0.0063 | 0.0063 | 0.0063 | 0.0064 | 0.0064 | 0.0064 | 0.0064 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0065 | 0.0064 | 0.0063 | 0.0062 | 0.0059 | 0.0055 | 0.0049 | 0.0039 | 0.0021 | 0.0011 | 0.0078
| Equivalent Volume | B | 6.9 | mm3 | 0.0032459 | 0.0032859 | 0.0033223 | 0.0033547 | 0.0033843 | 0.00341 | 0.0034338 | 0.0034545 | 0.0034725 | 0.0034891 | 0.0035077 | 0.0035295 | 0.0035403 | 0.0035502 | 0.0035588 | 0.0035664 | 0.0035736 | 0.0035808 | 0.0035848 | 0.0035907 | 0.0035934 | 0.0035947 | 0.0035947 | 0.0035938 | 0.0035907 | 0.003583 | 0.0035687 | 0.0035435 | 0.0035035 | 0.0034441 | 0.0041526 | 0.0041095 | 0.003873 | 0.003468 | 0.0027756 | 0.001707 | 0.000936 | 0.0019466
| Resonance Frequency | B | 150 | Hz | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0003 | 0.0006 | 0.0014 | 0.0027 | 0.0031 | 0.0007 | 0.0012
| Loss Factor | B | 0.18 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0004 | 0.0006 | 0.001 | 0.0016 | 0.0024 | 0.0037 | 0.0051 | 0.0058 | 0.0034 | 0.0015 | 0.0073
Additional Heat Conduction
| Caused by Front Cavity Thread | N.A. | Table
| Polarizing Voltage | B |  | V | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004
| Imperfection of Theory |  |  |  | 0.014871 | 3.1729e-05 | 0.0062044 | 0.0064707 | 0.0044762 | 0.002484 | 0.0014048 | 0.0011021 | 0.0010359 | 0.00095605 | 0.00085384 | 0.00075648 | 0.00067051 | 0.00059334 | 0.00052959 | 0.00047099 | 0.00041738 | 0.0003729 | 0.00033317 | 0.00029409 | 0.00026264 | 0.00023442 | 0.00020814 | 0.00018367 | 0.00016285 | 0.00014286 | 0.00012324 | 0.0001051 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Heat Conduction Theory | N.A. |  |  | 0.014871 | 3.1729e-05 | 0.0062044 | 0.0064707 | 0.0044762 | 0.002484 | 0.0014048 | 0.0011021 | 0.0010359 | 0.00095605 | 0.00085384 | 0.00075648 | 0.00067051 | 0.00059334 | 0.00052959 | 0.00047099 | 0.00041738 | 0.0003729 | 0.00033317 | 0.00029409 | 0.00026264 | 0.00023442 | 0.00020814 | 0.00018367 | 0.00016285 | 0.00014286 | 0.00012324 | 0.0001051 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Adding of Excess Volume | N.A.
| Radial Wave Motion | N.A.
| Processing of Results |  |  |  | 0.013814 | 0.042396 | 0.010196 | 0.0055832 | 0.0058482 | 0.010241 | 0.0060157 | 0.0047383 | 0.0058205 | 0.0030966 | 0.0041147 | 0.0040472 | 0.0033887 | 0.004108 | 0.004108 | 0.0042421 | 0.0042421 | 0.0040996 | 0.0042776 | 0.0039454 | 0.0036877 | 0.0039976 | 0.0039976 | 0.0039606 | 0.0039413 | 0.0037081 | 0.0037042 | 0.0036974 | 0.0036376 | 0.0032413 | 0.0058132 | 0.0058132 | 0.0064795 | 0.0075414 | 0.0092303 | 0.010209 | 0.011772 | 0.011772
| Rounding Error | B | 0.005 | dB | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868
| Repeatability of Measurements | A | Table | dB | 0.013508 | 0.042297 | 0.0097788 | 0.004779 | 0.0050861 | 0.0098254 | 0.0052779 | 0.0037574 | 0.0050542 | 0.0011205 | 0.0029321 | 0.0028366 | 0.0017748 | 0.0029227 | 0.0029227 | 0.0031084 | 0.0031084 | 0.0029109 | 0.0031566 | 0.0026894 | 0.0022947 | 0.0027654 | 0.0027654 | 0.0027116 | 0.0026834 | 0.0023274 | 0.0023211 | 0.0023103 | 0.0022133 | 0.001474 | 0.0050457 | 0.0050457 | 0.005801 | 0.006967 | 0.0087672 | 0.0097922 | 0.011413 | 0.011413
| Combined Uncertainty(Sum) |  |  |  | 0.034045 | 0.050414 | 0.029748 | 0.028531 | 0.028173 | 0.02912 | 0.027818 | 0.027514 | 0.027677 | 0.027216 | 0.022635 | 0.02245 | 0.022239 | 0.02364 | 0.016136 | 0.011172 | 0.010904 | 0.0108 | 0.010835 | 0.010547 | 0.010448 | 0.010557 | 0.010551 | 0.010477 | 0.010475 | 0.010374 | 0.010367 | 0.010355 | 0.010243 | 0.010066 | 0.011412 | 0.01137 | 0.011835 | 0.012755 | 0.014168 | 0.014234 | 0.014932 | 0.019609
| Expanded Uncertainty (k=2) |  |  |  | 0.068091 | 0.10083 | 0.059497 | 0.057061 | 0.056347 | 0.05824 | 0.055637 | 0.055028 | 0.055355 | 0.054431 | 0.045271 | 0.0449 | 0.044478 | 0.04728 | 0.032272 | 0.022344 | 0.021807 | 0.0216 | 0.02167 | 0.021094 | 0.020896 | 0.021114 | 0.021102 | 0.020955 | 0.02095 | 0.020749 | 0.020734 | 0.02071 | 0.020487 | 0.020131 | 0.022824 | 0.02274 | 0.02367 | 0.02551 | 0.028336 | 0.028468 | 0.029863 | 0.039219
| Stated Uncertainty |  |  |  | 0.1 | 0.1 | 0.07 | 0.07 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.05 | 0.05 | 0.05 | 0.05 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.07

## KRISS Phase
KRISS | Phase
| Uncertainty Budget for LS1P (phase) | S/N | 2652762
| MEASURED QUANTITY | Type | Unc. | Unit
| Frequency |  |  | Hz | 1.9953 | 2.5119 | 3.1623 | 3.9811 | 5.0119 | 6.3096 | 7.9433 | 10 | 12.5893 | 15.8489 | 19.9526 | 25.1189 | 31.6228 | 39.8107 | 50.1187 | 63.0957 | 79.4328 | 100 | 125.8925 | 158.4893 | 199.5262 | 251.1886 | 316.2278 | 398.1072 | 501.1872 | 630.9573 | 794.3282 | 1000 | 1258.9254 | 1584.8932 | 1995.2623 | 2511.8864 | 3162.2777 | 3981.0717 | 5011.8723 | 6309.5734 | 7943.2823 | 10000
| Electrical Transfer Impedance |  |  |  | 0.30138 | 0.28346 | 0.26935 | 0.2582 | 0.24938 | 0.24238 | 0.23682 | 0.23238 | 0.15914 | 0.1568 | 0.15491 | 0.1526 | 0.15078 | 0.14936 | 0.14828 | 0.14745 | 0.14683 | 0.14636 | 0.146 | 0.12264 | 0.1225 | 0.12241 | 0.12235 | 0.11681 | 0.11681 | 0.11681 | 0.11681 | 0.11678 | 0.11673 | 0.11668 | 0.11662 | 0.12206 | 0.1221 | 0.12229 | 0.12264 | 0.1234 | 0.12482 | 0.14931
| Series Impedance | B | Table | nF | 0.29583 | 0.27781 | 0.26364 | 0.2525 | 0.24373 | 0.23682 | 0.23136 | 0.22706 | 0.15161 | 0.14948 | 0.14782 | 0.14651 | 0.14548 | 0.14467 | 0.14403 | 0.14353 | 0.14313 | 0.14281 | 0.14256 | 0.11861 | 0.11851 | 0.11844 | 0.11838 | 0.11261 | 0.11259 | 0.11257 | 0.11256 | 0.11254 | 0.11253 | 0.11253 | 0.11252 | 0.11818 | 0.11817 | 0.11816 | 0.11816 | 0.11816 | 0.11816 | 0.14161
| Voltage Ratio | B | Table |  | 0.00024506 | 0.0003745 | 6.7484e-05 | 0.00025661 | 0.00038931 | 0.00014589 | 0.000176 | 0.00033079 | 0.00026821 | 0.00033141 | 4.4154e-05 | 0.00017569 | 0.0003132 | 7.6094e-05 | 0.00045119 | 0.0004285 | 0.0001371 | 0.00012185 | 6.3396e-05 | 1.9117e-05 | 0.00052703 | 0.0005668 | 0.00012995 | 0.00054842 | 0.00071228 | 0.0010305 | 0.0016339 | 0.00095917 | 0.0015664 | 0.0023482 | 0.0027298 | 0.0034776 | 0.0046905 | 0.00618 | 0.0068933 | 0.0085996 | 0.011257 | 0.014288
| Cross-talk | B | Meas. |  | 0.029781 | 0.0296 | 0.029459 | 0.029354 | 0.029281 | 0.029235 | 0.029212 | 0.02921 | 0.029224 | 0.029252 | 0.029291 | 0.029515 | 0.029768 | 0.02998 | 0.030121 | 0.030185 | 0.030184 | 0.03014 | 0.030078 | 0.03002 | 0.029984 | 0.029979 | 0.030003 | 0.030046 | 0.030093 | 0.03012 | 0.030106 | 0.030034 | 0.0299 | 0.02972 | 0.029541 | 0.029452 | 0.029594 | 0.030177 | 0.031493 | 0.033934 | 0.038006 | 0.044353
| Inherent Noise | B | Meas. |  | 0.04915 | 0.047838 | 0.04651 | 0.04517 | 0.043823 | 0.042472 | 0.04112 | 0.039772 | 0.038431 | 0.037099 | 0.035779 | 0.030676 | 0.025952 | 0.021708 | 0.018009 | 0.014884 | 0.012339 | 0.010351 | 0.0088803 | 0.0078717 | 0.0072574 | 0.0069621 | 0.0069064 | 0.0070105 | 0.0071979 | 0.0073983 | 0.0075512 | 0.0076092 | 0.0075411 | 0.0073345 | 0.006999 | 0.0065691 | 0.0061065 | 0.0057027 | 0.0054819 | 0.0056026 | 0.0062605 | 0.0076903
| Distortion | B | Meas. |  | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193
| Frequency | B | Table | Hz | 0.00089712 | 0.00079158 | 0.00068603 | 0.00058049 | 0.00047494 | 0.00039579 | 0.00034301 | 0.00031663 | 0.00029024 | 0.00026386 | 0.00022428 | 0.00019789 | 0.00017151 | 0.00014512 | 0.00011874 | 9.8946e-05 | 8.5754e-05 | 7.9157e-05 | 7.2561e-05 | 6.5964e-05 | 5.9368e-05 | 5.2771e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 3.9579e-05 | 3.9579e-05 | 3.2982e-05 | 3.2982e-05 | 3.2982e-05
| Reveiver Ground Shield | B | B&K |  | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979
| Transmitter Ground Shield | B | B&K |  | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979
| Coupler Properties |  |  |  | 0.00011917 | 0.00010843 | 9.187e-05 | 8.6988e-05 | 7.6294e-05 | 7.1976e-05 | 6.5631e-05 | 5.9587e-05 | 5.5016e-05 | 4.8852e-05 | 4.4487e-05 | 3.8147e-05 | 3.2369e-05 | 2.7508e-05 | 2.7508e-05 | 2.7508e-05 | 1.706e-05 | 2.4126e-05 | 1.706e-05 | 1.706e-05 | 2.2888e-05 | 1.5259e-05 | 1.5259e-05 | 2.4126e-05 | 2.7508e-05 | 3.8147e-05 | 3.2369e-05 | 4.8852e-05 | 5.9587e-05 | 7.034e-05 | 9.187e-05 | 0.00011917 | 0.00014814 | 0.00020984 | 0.00030537 | 0.00048093 | 0.00074756 | 0.0011007
| Coupler Length | B | 0.001 | mm | 9.1553e-05 | 8.3923e-05 | 6.8665e-05 | 6.8665e-05 | 6.1035e-05 | 6.1035e-05 | 5.3406e-05 | 4.5776e-05 | 4.5776e-05 | 3.8147e-05 | 3.8147e-05 | 3.0518e-05 | 2.2888e-05 | 2.2888e-05 | 2.2888e-05 | 2.2888e-05 | 1.5259e-05 | 2.2888e-05 | 1.5259e-05 | 1.5259e-05 | 2.2888e-05 | 1.5259e-05 | 1.5259e-05 | 2.2888e-05 | 2.2888e-05 | 3.0518e-05 | 2.2888e-05 | 3.8147e-05 | 4.5776e-05 | 5.3406e-05 | 6.8665e-05 | 9.1553e-05 | 0.00012207 | 0.00018692 | 0.00029755 | 0.00047874 | 0.00068736 | 0.00089073
| Coupler Diameter | B | 0.0015 | mm | 7.6294e-05 | 6.8665e-05 | 6.1035e-05 | 5.3406e-05 | 4.5776e-05 | 3.8147e-05 | 3.8147e-05 | 3.8147e-05 | 3.0518e-05 | 3.0518e-05 | 2.2888e-05 | 2.2888e-05 | 2.2888e-05 | 1.5259e-05 | 1.5259e-05 | 1.5259e-05 | 7.6294e-06 | 7.6294e-06 | 7.6294e-06 | 7.6294e-06 | 0 | 0 | 0 | 7.6294e-06 | 1.5259e-05 | 2.2888e-05 | 2.2888e-05 | 3.0518e-05 | 3.8147e-05 | 4.5776e-05 | 6.1035e-05 | 7.6294e-05 | 8.3923e-05 | 9.5367e-05 | 6.8665e-05 | 4.5776e-05 | 0.00029391 | 0.00064659
| Coupler Volume | B | 0.4266 | mm3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Coupler Surface Area | B | 0.1362 | mm2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Environmental Conditions |  |  |  | 0.0013333 | 0.001194 | 0.0010747 | 0.00099507 | 0.00089561 | 0.00081607 | 0.00073659 | 0.0006572 | 0.00059786 | 0.00053881 | 0.00048028 | 0.0004806 | 0.00046121 | 0.00044192 | 0.00044242 | 0.00042333 | 0.00040439 | 0.00040517 | 0.00038655 | 0.00038757 | 0.00038872 | 0.00037074 | 0.00037225 | 0.00035492 | 0.0003381 | 0.00034044 | 0.00034304 | 0.0003277 | 0.0003311 | 0.00033487 | 0.00032166 | 0.00032653 | 0.00033192 | 0.00032164 | 0.0003285 | 0.00032096 | 0.0003296 | 0.00032532
| Static Pressure | B | 2.26 | kPa | 6.0966e-07 | 7.6791e-07 | 9.6859e-07 | 1.2236e-06 | 1.5484e-06 | 1.9631e-06 | 2.4947e-06 | 3.1786e-06 | 4.0627e-06 | 5.2119e-06 | 6.7154e-06 | 7.1623e-06 | 7.6419e-06 | 8.1569e-06 | 8.7098e-06 | 9.3037e-06 | 9.9417e-06 | 1.0627e-05 | 1.1365e-05 | 1.2157e-05 | 1.3011e-05 | 1.3929e-05 | 1.4918e-05 | 1.5984e-05 | 1.7133e-05 | 1.8373e-05 | 1.9711e-05 | 2.1157e-05 | 2.2719e-05 | 2.4408e-05 | 2.6236e-05 | 2.8216e-05 | 3.036e-05 | 3.2686e-05 | 3.5208e-05 | 3.7946e-05 | 4.0919e-05 | 4.415e-05
| Temperature | B | 0.7 | K | 5.041e-06 | 6.3354e-06 | 7.9649e-06 | 1.0017e-05 | 1.26e-05 | 1.5854e-05 | 1.995e-05 | 2.511e-05 | 3.1607e-05 | 3.9788e-05 | 5.0089e-05 | 5.3066e-05 | 5.6222e-05 | 5.9568e-05 | 6.3114e-05 | 6.6871e-05 | 7.0852e-05 | 7.5069e-05 | 7.9535e-05 | 8.4264e-05 | 8.9271e-05 | 9.4572e-05 | 0.00010018 | 0.00010612 | 0.00011241 | 0.00011906 | 0.0001261 | 0.00013355 | 0.00014144 | 0.00014978 | 0.00015861 | 0.00016795 | 0.00017783 | 0.00018829 | 0.00019936 | 0.00021107 | 0.00022346 | 0.00023657
| Relative Humidity | B | 5.85 | % | 0.0013333 | 0.001194 | 0.0010746 | 0.00099502 | 0.00089552 | 0.00081592 | 0.00073632 | 0.00065671 | 0.00059701 | 0.00053731 | 0.00047761 | 0.00047761 | 0.00045771 | 0.00043781 | 0.00043781 | 0.00041791 | 0.00039801 | 0.00039801 | 0.00037811 | 0.00037811 | 0.00037811 | 0.00035821 | 0.00035821 | 0.00033831 | 0.00031841 | 0.00031841 | 0.00031841 | 0.00029851 | 0.00029851 | 0.00029851 | 0.00027861 | 0.00027861 | 0.00027861 | 0.00025871 | 0.00025871 | 0.0002388 | 0.0002388 | 0.0002189
| Microphone Parameters |  |  |  | 0.010008 | 0.0092667 | 0.0085237 | 0.0078891 | 0.0072531 | 0.0066506 | 0.006132 | 0.0056092 | 0.0051503 | 0.0047383 | 0.0043597 | 0.0041747 | 0.003847 | 0.0035812 | 0.0033539 | 0.0031525 | 0.0030147 | 0.0029398 | 0.0029789 | 0.0032441 | 0.0036846 | 0.0043569 | 0.0053445 | 0.0066162 | 0.0083176 | 0.010398 | 0.013064 | 0.016359 | 0.020507 | 0.025356 | 0.036055 | 0.043962 | 0.051955 | 0.056931 | 0.05573 | 0.049413 | 0.05019 | 0.11021
| Front Cavity Depth | B | 0.0025 | mm | 0.0007 | 0.0006 | 0.0006 | 0.0005 | 0.0005 | 0.0004 | 0.0004 | 0.0004 | 0.0003 | 0.0003 | 0.0003 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0003 | 0.0006 | 0.0012 | 0.0017 | 0.0018
| Front Cavity Volume | B | 2.59 | mm3 | 0.0042 | 0.0039 | 0.0035 | 0.0033 | 0.003 | 0.0027 | 0.0025 | 0.0022 | 0.002 | 0.0018 | 0.0016 | 0.0015 | 0.0013 | 0.0012 | 0.0011 | 0.001 | 0.0009 | 0.0008 | 0.0007 | 0.0007 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0007 | 0.0007 | 0.0008 | 0.001 | 0.0012 | 0.0014 | 0.0018 | 0.0023 | 0.003 | 0.0041 | 0.0059 | 0.01
| Equivalent Volume | B | 6.9 | mm3 | 0.0087051 | 0.0080032 | 0.0073344 | 0.0066961 | 0.0060908 | 0.0055237 | 0.0049922 | 0.0044937 | 0.0040207 | 0.0035807 | 0.0031713 | 0.0029678 | 0.0025864 | 0.0022125 | 0.0018463 | 0.0014826 | 0.0011037 | 0.0007019 | 0.00027212 | 0.00075785 | 0.0013987 | 0.0021617 | 0.0030746 | 0.0041885 | 0.0055491 | 0.007225 | 0.0092901 | 0.011826 | 0.014918 | 0.018603 | 0.029485 | 0.036641 | 0.044398 | 0.050201 | 0.050598 | 0.045904 | 0.045794 | 0.078331
| Resonance Frequency | B | 150 | Hz | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0002 | 0.0003 | 0.0003 | 0.0004 | 0.0005 | 0.0006 | 0.0008 | 0.001 | 0.0012 | 0.0015 | 0.0019 | 0.0024 | 0.0031 | 0.0039 | 0.005 | 0.0065 | 0.0084 | 0.0108 | 0.0142 | 0.0185 | 0.0229 | 0.023 | 0.0123 | 0.0161 | 0.0718
| Loss Factor | B | 0.18 |  | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0002 | 0.0003 | 0.0004 | 0.0005 | 0.0006 | 0.0008 | 0.001 | 0.0013 | 0.0016 | 0.002 | 0.0025 | 0.0032 | 0.004 | 0.0051 | 0.0063 | 0.0079 | 0.0098 | 0.0122 | 0.0148 | 0.0175 | 0.0195 | 0.0194 | 0.0136 | 0.001 | 0.0126 | 0.0109 | 0.0273
Additional Heat Conduction
| Caused by Front Cavity Thread | N.A. | Table
| Polarizing Voltage | B | 0.0063 | V | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025
| Imperfection of Theory |  |  |  | 0.10426 | 0.049932 | 0.0081625 | 0.010478 | 0.014226 | 0.011112 | 0.0080039 | 0.0064342 | 0.0057346 | 0.0052565 | 0.0047533 | 0.0042646 | 0.0038184 | 0.0034491 | 0.0031103 | 0.002804 | 0.002556 | 0.0023433 | 0.0021482 | 0.0020079 | 0.0019025 | 0.0018318 | 0.0018032 | 0.0018216 | 0.0018926 | 0.0020306 | 0.0022291 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Heat Conduction Theory | N.A. |  |  | 0.10426 | 0.049932 | 0.0081625 | 0.010478 | 0.014226 | 0.011112 | 0.0080039 | 0.0064342 | 0.0057346 | 0.0052565 | 0.0047533 | 0.0042646 | 0.0038184 | 0.0034491 | 0.0031103 | 0.002804 | 0.002556 | 0.0023433 | 0.0021482 | 0.0020079 | 0.0019025 | 0.0018318 | 0.0018032 | 0.0018216 | 0.0018926 | 0.0020306 | 0.0022291 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Adding of Excess Volume | N.A.
| Radial Wave Motion | N.A.
| Processing of Results |  |  |  | 0.64016 | 0.61509 | 0.45201 | 0.36734 | 0.28054 | 0.30714 | 0.18325 | 0.13967 | 0.11975 | 0.095465 | 0.074534 | 0.068391 | 0.060111 | 0.046923 | 0.041649 | 0.037428 | 0.033437 | 0.032023 | 0.030293 | 0.029853 | 0.029689 | 0.029283 | 0.029464 | 0.029464 | 0.02976 | 0.030175 | 0.030958 | 0.031954 | 0.033943 | 0.037027 | 0.041708 | 0.047603 | 0.053511 | 0.060204 | 0.060641 | 0.056248 | 0.056989 | 0.086631
| Rounding Error | B | 0.05 | degree | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868
| Repeatability of Measurements | A | Table | degree | 0.6395 | 0.61441 | 0.45109 | 0.3662 | 0.27905 | 0.30578 | 0.18096 | 0.13665 | 0.11622 | 0.090996 | 0.068717 | 0.062 | 0.052726 | 0.036993 | 0.030021 | 0.023822 | 0.016872 | 0.013862 | 0.0091848 | 0.007609 | 0.0069362 | 0.0049162 | 0.0058965 | 0.0058965 | 0.0072324 | 0.0087854 | 0.011183 | 0.013702 | 0.017855 | 0.023188 | 0.030104 | 0.037851 | 0.045057 | 0.052832 | 0.05333 | 0.048276 | 0.049137 | 0.081679
| Combined Uncertainty(Sum) |  |  |  | 0.71526 | 0.67916 | 0.52631 | 0.4492 | 0.3757 | 0.39147 | 0.29961 | 0.27126 | 0.19932 | 0.18371 | 0.17203 | 0.16733 | 0.16241 | 0.15664 | 0.15409 | 0.15219 | 0.15064 | 0.14987 | 0.14916 | 0.12628 | 0.12612 | 0.12595 | 0.12598 | 0.12066 | 0.12084 | 0.12111 | 0.12157 | 0.12217 | 0.12329 | 0.12501 | 0.12899 | 0.13819 | 0.14308 | 0.14772 | 0.14773 | 0.14434 | 0.14611 | 0.20481
| Expanded Uncertainty (k=2) |  |  |  | 1.4305 | 1.3583 | 1.0526 | 0.89839 | 0.7514 | 0.78295 | 0.59922 | 0.54251 | 0.39863 | 0.36743 | 0.34406 | 0.33466 | 0.32482 | 0.31328 | 0.30817 | 0.30438 | 0.30128 | 0.29973 | 0.29832 | 0.25256 | 0.25223 | 0.25191 | 0.25196 | 0.24132 | 0.24168 | 0.24222 | 0.24313 | 0.24434 | 0.24657 | 0.25003 | 0.25799 | 0.27638 | 0.28616 | 0.29543 | 0.29546 | 0.28868 | 0.29222 | 0.40961
| Stated Uncertainty |  |  |  | 1.8 | 1.5 | 1.1 | 1 | 0.8 | 0.8 | 0.7 | 0.6 | 0.5 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.4 | 0.7
| Uncertainty Budget for LS1P (phase) | S/N | 2652765
| MEASURED QUANTITY | Type | Unc. | Unit
| Frequency |  |  | Hz | 1.9953 | 2.5119 | 3.1623 | 3.9811 | 5.0119 | 6.3096 | 7.9433 | 10 | 12.5893 | 15.8489 | 19.9526 | 25.1189 | 31.6228 | 39.8107 | 50.1187 | 63.0957 | 79.4328 | 100 | 125.8925 | 158.4893 | 199.5262 | 251.1886 | 316.2278 | 398.1072 | 501.1872 | 630.9573 | 794.3282 | 1000 | 1258.9254 | 1584.8932 | 1995.2623 | 2511.8864 | 3162.2777 | 3981.0717 | 5011.8723 | 6309.5734 | 7943.2823 | 10000
| Electrical Transfer Impedance |  |  |  | 0.30138 | 0.28346 | 0.26935 | 0.2582 | 0.24938 | 0.24238 | 0.23682 | 0.23238 | 0.15914 | 0.1568 | 0.15491 | 0.1526 | 0.15078 | 0.14936 | 0.14828 | 0.14745 | 0.14683 | 0.14636 | 0.146 | 0.12264 | 0.1225 | 0.12241 | 0.12235 | 0.11681 | 0.11681 | 0.11681 | 0.11681 | 0.11678 | 0.11673 | 0.11668 | 0.11662 | 0.12206 | 0.1221 | 0.12229 | 0.12264 | 0.1234 | 0.12482 | 0.14931
| Series Impedance | B | Table | nF | 0.29583 | 0.27781 | 0.26364 | 0.2525 | 0.24373 | 0.23682 | 0.23136 | 0.22706 | 0.15161 | 0.14948 | 0.14782 | 0.14651 | 0.14548 | 0.14467 | 0.14403 | 0.14353 | 0.14313 | 0.14281 | 0.14256 | 0.11861 | 0.11851 | 0.11844 | 0.11838 | 0.11261 | 0.11259 | 0.11257 | 0.11256 | 0.11254 | 0.11253 | 0.11253 | 0.11252 | 0.11818 | 0.11817 | 0.11816 | 0.11816 | 0.11816 | 0.11816 | 0.14161
| Voltage Ratio | B | Table |  | 0.00024506 | 0.0003745 | 6.7484e-05 | 0.00025661 | 0.00038931 | 0.00014589 | 0.000176 | 0.00033079 | 0.00026821 | 0.00033141 | 4.4154e-05 | 0.00017569 | 0.0003132 | 7.6094e-05 | 0.00045119 | 0.0004285 | 0.0001371 | 0.00012185 | 6.3396e-05 | 1.9117e-05 | 0.00052703 | 0.0005668 | 0.00012995 | 0.00054842 | 0.00071228 | 0.0010305 | 0.0016339 | 0.00095917 | 0.0015664 | 0.0023482 | 0.0027298 | 0.0034776 | 0.0046905 | 0.00618 | 0.0068933 | 0.0085996 | 0.011257 | 0.014288
| Cross-talk | B | Meas. |  | 0.029781 | 0.0296 | 0.029459 | 0.029354 | 0.029281 | 0.029235 | 0.029212 | 0.02921 | 0.029224 | 0.029252 | 0.029291 | 0.029515 | 0.029768 | 0.02998 | 0.030121 | 0.030185 | 0.030184 | 0.03014 | 0.030078 | 0.03002 | 0.029984 | 0.029979 | 0.030003 | 0.030046 | 0.030093 | 0.03012 | 0.030106 | 0.030034 | 0.0299 | 0.02972 | 0.029541 | 0.029452 | 0.029594 | 0.030177 | 0.031493 | 0.033934 | 0.038006 | 0.044353
| Inherent Noise | B | Meas. |  | 0.04915 | 0.047838 | 0.04651 | 0.04517 | 0.043823 | 0.042472 | 0.04112 | 0.039772 | 0.038431 | 0.037099 | 0.035779 | 0.030676 | 0.025952 | 0.021708 | 0.018009 | 0.014884 | 0.012339 | 0.010351 | 0.0088803 | 0.0078717 | 0.0072574 | 0.0069621 | 0.0069064 | 0.0070105 | 0.0071979 | 0.0073983 | 0.0075512 | 0.0076092 | 0.0075411 | 0.0073345 | 0.006999 | 0.0065691 | 0.0061065 | 0.0057027 | 0.0054819 | 0.0056026 | 0.0062605 | 0.0076903
| Distortion | B | Meas. |  | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193 | 0.0013193
| Frequency | B | Table | Hz | 0.00089712 | 0.00079158 | 0.00068603 | 0.00058049 | 0.00047494 | 0.00039579 | 0.00034301 | 0.00031663 | 0.00029024 | 0.00026386 | 0.00022428 | 0.00019789 | 0.00017151 | 0.00014512 | 0.00011874 | 9.8946e-05 | 8.5754e-05 | 7.9157e-05 | 7.2561e-05 | 6.5964e-05 | 5.9368e-05 | 5.2771e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 4.6175e-05 | 3.9579e-05 | 3.9579e-05 | 3.2982e-05 | 3.2982e-05 | 3.2982e-05
| Reveiver Ground Shield | B | B&K |  | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979
| Transmitter Ground Shield | B | B&K |  | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979 | 0.001979
| Coupler Properties |  |  |  | 0.00011917 | 0.00010843 | 9.187e-05 | 8.6988e-05 | 7.6294e-05 | 7.1976e-05 | 6.5631e-05 | 5.9587e-05 | 5.5016e-05 | 4.8852e-05 | 4.4487e-05 | 3.8147e-05 | 3.2369e-05 | 2.7508e-05 | 2.7508e-05 | 2.7508e-05 | 1.706e-05 | 2.4126e-05 | 1.706e-05 | 1.706e-05 | 2.2888e-05 | 1.5259e-05 | 1.5259e-05 | 2.4126e-05 | 2.7508e-05 | 3.8147e-05 | 3.2369e-05 | 4.8852e-05 | 5.9587e-05 | 7.034e-05 | 9.187e-05 | 0.00011917 | 0.00014814 | 0.00020984 | 0.00030537 | 0.00048093 | 0.00074756 | 0.0011007
| Coupler Length | B | 0.001 | mm | 9.1553e-05 | 8.3923e-05 | 6.8665e-05 | 6.8665e-05 | 6.1035e-05 | 6.1035e-05 | 5.3406e-05 | 4.5776e-05 | 4.5776e-05 | 3.8147e-05 | 3.8147e-05 | 3.0518e-05 | 2.2888e-05 | 2.2888e-05 | 2.2888e-05 | 2.2888e-05 | 1.5259e-05 | 2.2888e-05 | 1.5259e-05 | 1.5259e-05 | 2.2888e-05 | 1.5259e-05 | 1.5259e-05 | 2.2888e-05 | 2.2888e-05 | 3.0518e-05 | 2.2888e-05 | 3.8147e-05 | 4.5776e-05 | 5.3406e-05 | 6.8665e-05 | 9.1553e-05 | 0.00012207 | 0.00018692 | 0.00029755 | 0.00047874 | 0.00068736 | 0.00089073
| Coupler Diameter | B | 0.0015 | mm | 7.6294e-05 | 6.8665e-05 | 6.1035e-05 | 5.3406e-05 | 4.5776e-05 | 3.8147e-05 | 3.8147e-05 | 3.8147e-05 | 3.0518e-05 | 3.0518e-05 | 2.2888e-05 | 2.2888e-05 | 2.2888e-05 | 1.5259e-05 | 1.5259e-05 | 1.5259e-05 | 7.6294e-06 | 7.6294e-06 | 7.6294e-06 | 7.6294e-06 | 0 | 0 | 0 | 7.6294e-06 | 1.5259e-05 | 2.2888e-05 | 2.2888e-05 | 3.0518e-05 | 3.8147e-05 | 4.5776e-05 | 6.1035e-05 | 7.6294e-05 | 8.3923e-05 | 9.5367e-05 | 6.8665e-05 | 4.5776e-05 | 0.00029391 | 0.00064659
| Coupler Volume | B | 0.4266 | mm3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Coupler Surface Area | B | 0.1362 | mm2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Environmental Conditions |  |  |  | 0.0013333 | 0.001194 | 0.0010747 | 0.00099507 | 0.00089561 | 0.00081607 | 0.00073659 | 0.0006572 | 0.00059786 | 0.00053881 | 0.00048028 | 0.0004806 | 0.00046121 | 0.00044192 | 0.00044242 | 0.00042333 | 0.00040439 | 0.00040517 | 0.00038655 | 0.00038757 | 0.00038872 | 0.00037074 | 0.00037225 | 0.00035492 | 0.0003381 | 0.00034044 | 0.00034304 | 0.0003277 | 0.0003311 | 0.00033487 | 0.00032166 | 0.00032653 | 0.00033192 | 0.00032164 | 0.0003285 | 0.00032096 | 0.0003296 | 0.00032532
| Static Pressure | B | 2.26 | kPa | 6.0966e-07 | 7.6791e-07 | 9.6859e-07 | 1.2236e-06 | 1.5484e-06 | 1.9631e-06 | 2.4947e-06 | 3.1786e-06 | 4.0627e-06 | 5.2119e-06 | 6.7154e-06 | 7.1623e-06 | 7.6419e-06 | 8.1569e-06 | 8.7098e-06 | 9.3037e-06 | 9.9417e-06 | 1.0627e-05 | 1.1365e-05 | 1.2157e-05 | 1.3011e-05 | 1.3929e-05 | 1.4918e-05 | 1.5984e-05 | 1.7133e-05 | 1.8373e-05 | 1.9711e-05 | 2.1157e-05 | 2.2719e-05 | 2.4408e-05 | 2.6236e-05 | 2.8216e-05 | 3.036e-05 | 3.2686e-05 | 3.5208e-05 | 3.7946e-05 | 4.0919e-05 | 4.415e-05
| Temperature | B | 0.7 | K | 5.041e-06 | 6.3354e-06 | 7.9649e-06 | 1.0017e-05 | 1.26e-05 | 1.5854e-05 | 1.995e-05 | 2.511e-05 | 3.1607e-05 | 3.9788e-05 | 5.0089e-05 | 5.3066e-05 | 5.6222e-05 | 5.9568e-05 | 6.3114e-05 | 6.6871e-05 | 7.0852e-05 | 7.5069e-05 | 7.9535e-05 | 8.4264e-05 | 8.9271e-05 | 9.4572e-05 | 0.00010018 | 0.00010612 | 0.00011241 | 0.00011906 | 0.0001261 | 0.00013355 | 0.00014144 | 0.00014978 | 0.00015861 | 0.00016795 | 0.00017783 | 0.00018829 | 0.00019936 | 0.00021107 | 0.00022346 | 0.00023657
| Relative Humidity | B | 5.85 | % | 0.0013333 | 0.001194 | 0.0010746 | 0.00099502 | 0.00089552 | 0.00081592 | 0.00073632 | 0.00065671 | 0.00059701 | 0.00053731 | 0.00047761 | 0.00047761 | 0.00045771 | 0.00043781 | 0.00043781 | 0.00041791 | 0.00039801 | 0.00039801 | 0.00037811 | 0.00037811 | 0.00037811 | 0.00035821 | 0.00035821 | 0.00033831 | 0.00031841 | 0.00031841 | 0.00031841 | 0.00029851 | 0.00029851 | 0.00029851 | 0.00027861 | 0.00027861 | 0.00027861 | 0.00025871 | 0.00025871 | 0.0002388 | 0.0002388 | 0.0002189
| Microphone Parameters |  |  |  | 0.0058546 | 0.0054937 | 0.0050888 | 0.004823 | 0.0045122 | 0.0042078 | 0.0039925 | 0.0037319 | 0.0035332 | 0.0033634 | 0.0032031 | 0.0031269 | 0.0029976 | 0.0029271 | 0.0028794 | 0.0028329 | 0.0028333 | 0.0028664 | 0.0029681 | 0.0031667 | 0.0034462 | 0.0038636 | 0.0045112 | 0.0053404 | 0.0065112 | 0.0079203 | 0.009779 | 0.012084 | 0.015069 | 0.018495 | 0.023329 | 0.02767 | 0.031392 | 0.032406 | 0.029664 | 0.024678 | 0.026485 | 0.082572
| Front Cavity Depth | B | 0.0025 | mm | 0.0007 | 0.0006 | 0.0006 | 0.0005 | 0.0005 | 0.0004 | 0.0004 | 0.0004 | 0.0003 | 0.0003 | 0.0003 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0003 | 0.0006 | 0.0012 | 0.0017 | 0.0018
| Front Cavity Volume | B | 2.59 | mm3 | 0.0042 | 0.0039 | 0.0035 | 0.0033 | 0.003 | 0.0027 | 0.0025 | 0.0022 | 0.002 | 0.0018 | 0.0016 | 0.0015 | 0.0013 | 0.0012 | 0.0011 | 0.001 | 0.0009 | 0.0008 | 0.0007 | 0.0007 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 0.0007 | 0.0007 | 0.0008 | 0.001 | 0.0012 | 0.0014 | 0.0018 | 0.0023 | 0.003 | 0.0041 | 0.0059 | 0.01
| Equivalent Volume | B | 6.9 | mm3 | 0.0031459 | 0.0028915 | 0.0026525 | 0.0024211 | 0.0022024 | 0.0019964 | 0.0018056 | 0.0016301 | 0.0014572 | 0.001297 | 0.0011444 | 0.0010757 | 0.00093587 | 0.00079854 | 0.00067139 | 0.00053406 | 0.00039673 | 0.00025686 | 9.9182e-05 | 0.00027974 | 0.00050608 | 0.00078583 | 0.0011139 | 0.0015132 | 0.0020014 | 0.0026118 | 0.0033569 | 0.0042725 | 0.005394 | 0.006724 | 0.010661 | 0.01325 | 0.016042 | 0.018143 | 0.018285 | 0.016566 | 0.016719 | 0.028427
| Resonance Frequency | B | 150 | Hz | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0002 | 0.0003 | 0.0003 | 0.0004 | 0.0005 | 0.0006 | 0.0008 | 0.001 | 0.0012 | 0.0015 | 0.0019 | 0.0024 | 0.0031 | 0.0039 | 0.005 | 0.0065 | 0.0084 | 0.0108 | 0.0142 | 0.0185 | 0.0229 | 0.023 | 0.0123 | 0.0161 | 0.0718
| Loss Factor | B | 0.18 |  | 0 | 0 | 0 | 0 | 0 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0002 | 0.0003 | 0.0004 | 0.0005 | 0.0006 | 0.0008 | 0.001 | 0.0013 | 0.0016 | 0.002 | 0.0025 | 0.0032 | 0.004 | 0.0051 | 0.0063 | 0.0079 | 0.0098 | 0.0122 | 0.0148 | 0.0175 | 0.0195 | 0.0194 | 0.0136 | 0.001 | 0.0126 | 0.0109 | 0.0273
Additional Heat Conduction
| Caused by Front Cavity Thread | N.A. | Table
| Polarizing Voltage | B | 0.0063 | V | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025
| Imperfection of Theory |  |  |  | 0.10426 | 0.049932 | 0.0081625 | 0.010478 | 0.014226 | 0.011112 | 0.0080039 | 0.0064342 | 0.0057346 | 0.0052565 | 0.0047533 | 0.0042646 | 0.0038184 | 0.0034491 | 0.0031103 | 0.002804 | 0.002556 | 0.0023433 | 0.0021482 | 0.0020079 | 0.0019025 | 0.0018318 | 0.0018032 | 0.0018216 | 0.0018926 | 0.0020306 | 0.0022291 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Heat Conduction Theory | N.A. |  |  | 0.10426 | 0.049932 | 0.0081625 | 0.010478 | 0.014226 | 0.011112 | 0.0080039 | 0.0064342 | 0.0057346 | 0.0052565 | 0.0047533 | 0.0042646 | 0.0038184 | 0.0034491 | 0.0031103 | 0.002804 | 0.002556 | 0.0023433 | 0.0021482 | 0.0020079 | 0.0019025 | 0.0018318 | 0.0018032 | 0.0018216 | 0.0018926 | 0.0020306 | 0.0022291 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Adding of Excess Volume | N.A.
| Radial Wave Motion | N.A.
| Processing of Results |  |  |  | 1.1018 | 0.92268 | 0.54817 | 0.58938 | 0.48318 | 0.45173 | 0.29627 | 0.23934 | 0.19215 | 0.15394 | 0.12388 | 0.11272 | 0.087222 | 0.074655 | 0.054992 | 0.050651 | 0.042963 | 0.037899 | 0.035841 | 0.033125 | 0.032085 | 0.030587 | 0.030154 | 0.029804 | 0.029837 | 0.030292 | 0.030292 | 0.030422 | 0.031136 | 0.032274 | 0.032717 | 0.033114 | 0.04691 | 0.055105 | 0.043351 | 0.03721 | 0.035288 | 0.05006
| Rounding Error | B | 0.05 | degree | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868 | 0.028868
| Repeatability of Measurements | A | Table | degree | 1.1014 | 0.92223 | 0.54741 | 0.58867 | 0.48232 | 0.45081 | 0.29486 | 0.23759 | 0.18997 | 0.15121 | 0.12047 | 0.10896 | 0.082306 | 0.068847 | 0.046806 | 0.04162 | 0.03182 | 0.024556 | 0.021242 | 0.016245 | 0.014004 | 0.010112 | 0.0087155 | 0.0074129 | 0.0075436 | 0.0091799 | 0.0091799 | 0.0095998 | 0.011668 | 0.014432 | 0.015398 | 0.016223 | 0.036976 | 0.046938 | 0.032342 | 0.023479 | 0.020296 | 0.040898
| Combined Uncertainty(Sum) |  |  |  | 1.1471 | 0.96655 | 0.61084 | 0.64356 | 0.54395 | 0.51279 | 0.37939 | 0.33368 | 0.24959 | 0.21982 | 0.19843 | 0.18979 | 0.17426 | 0.16704 | 0.15821 | 0.15596 | 0.15303 | 0.15123 | 0.15038 | 0.12709 | 0.1267 | 0.12625 | 0.12611 | 0.12068 | 0.12075 | 0.12095 | 0.12109 | 0.12128 | 0.12175 | 0.12247 | 0.12335 | 0.12946 | 0.13452 | 0.13799 | 0.13342 | 0.13123 | 0.13239 | 0.17782
| Expanded Uncertainty (k=2) |  |  |  | 2.2941 | 1.9331 | 1.2217 | 1.2871 | 1.0879 | 1.0256 | 0.75878 | 0.66735 | 0.49918 | 0.43965 | 0.39687 | 0.37958 | 0.34851 | 0.33409 | 0.31641 | 0.31192 | 0.30607 | 0.30246 | 0.30077 | 0.25418 | 0.25339 | 0.25249 | 0.25222 | 0.24136 | 0.2415 | 0.2419 | 0.24218 | 0.24255 | 0.2435 | 0.24493 | 0.24669 | 0.25892 | 0.26903 | 0.27598 | 0.26684 | 0.26246 | 0.26479 | 0.35564
| Stated Uncertainty |  |  |  | 2.5 | 2 | 1.5 | 1.5 | 1.2 | 1.2 | 1 | 0.8 | 0.6 | 0.5 | 0.5 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.4 | 0.7

## NIMT Level
NIMT | Level
Uncertainty Components for the B&K Type 4160 Microphone with SN 2652762 as part of the APMP AUV.A-K5 intercomparison
| Symbol | Type | Distrubution | Component |  | DOF | [5 empty] | Uncertainty of Microphone Sensitivity for 4160 (LS1P) [dB]
[7 empty] | 1.9953 | 2.5119 | 3.1623 | 3.9811 | 5.0119 | 6.3096 | 7.9433 | 10 | 12.5892 | 15.8489 | 19.9528
| uElec |  |  | Electrical Transfer Impedance |  |  | 0.0063069 | 0.006283 | 0.00626 | 0.0062378 | 0.0062164 | 0.0061959 | 0.0061762 | 0.0061573 | 0.0061394 | 0.0061223 | 0.0059805
|  | B | Rectangualar | 1 | Series Capacitor | ¥ | 0.0027566 | 0.0027016 | 0.0026476 | 0.0025946 | 0.0025428 | 0.0024922 | 0.0024428 | 0.0023948 | 0.0023482 | 0.0023032 | 0.0018944
|  | B | Rectangualar | 2 | Voltage Ratios | ¥ | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992
|  | B | Rectangualar | 3 | Cross Talk | ¥ | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564
|  | B | Rectangualar | 4 | Inherent and Ambient Noise | ¥ | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05
|  | B | Rectangualar | 5 | Distortion | ¥ | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417
|  | B | Rectangualar | 6 | Frequency | ¥ | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05
|  | B | Rectangualar | 7 | Receiver Ground Shielding | ¥ | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641
|  | B | Rectangualar | 8 | Receiver Ground Shielding (guard length) | ¥ | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376
|  | B | Rectangualar | 9 | Transmitter Ground Shielding | ¥ | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509
|  | B | Rectangualar | 10 | Transmitter Ground Shielding (guard length) | ¥ | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376
| uCoup |  |  | Acoustic Transfer Impedance (coupler and gas) |  |  | 0.098096 | 0.098096 | 0.098096 | 0.098096 | 0.098096 | 0.062536 | 0.039948 | 0.025764 | 0.01703 | 0.011855 | 0.0089925
|  | B | Rectangular | 11 | Coupler Length | ¥ | 6.2769e-05 | 5.506e-05 | 4.9554e-05 | 4.2947e-05 | 3.634e-05 | 3.1935e-05 | 2.753e-05 | 2.4227e-05 | 1.9822e-05 | 1.7619e-05 | 7.7085e-06
|  | B | Rectangular | 12 | Coupler Diameter | ¥ | 7.7085e-06 | 6.6072e-06 | 4.4048e-06 | 3.3036e-06 | 3.3036e-06 | 2.2024e-06 | 2.2024e-06 | 2.2024e-06 | 1.1012e-06 | 1.1012e-06 | 4.4048e-06
|  | B | Rectangular | 13 | Coupler Volume (short 7. 5 mm) + volume fitting | ¥ | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776
|  | B | Rectangular | 11b | Coupler Volume (long 15 mm) + volume fitting | ¥ | 0.0106 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776
|  | B | Rectangular | 14 | Coupler Surface Area (short 7.5 mm) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.9282e-05
|  | B | Rectangular | 15 | Coupler Surface Area (long 15.0 mm) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8.0829e-05
|  | B | Rectangular | 13 | Capillary Tube Length | ¥ | - | - | - | - | - | - | - | - | - | -
|  | B | Rectangular | 14 | Capillary Tube Diameter | ¥ | - | - | - | - | - | - | - | - | - | -
|  | B | Rectangular | 16 | Coupler Leakage | ¥ | 0.097886 | 0.097886 | 0.097886 | 0.097886 | 0.097886 | 0.062205 | 0.039428 | 0.024949 | 0.015771 | 0.0099623 | 0.0062904
|  | B | Rectangular | 17 | Static Pressure of air in Coupler | ¥ | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733
|  | B | Rectangular | 18 | Temperature of air in Coupler | ¥ | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321
|  | B | Rectangular | 19 | Relative Humidity of air in Coupler | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| uMic |  |  | Acoustic Transfer Impedance (microphones) |  |  | 0.0025384 | 0.0025247 | 0.0025135 | 0.0025046 | 0.0024977 | 0.0024921 | 0.0024877 | 0.0024842 | 0.0024817 | 0.0024796 | 0.0024779
|  | B | Rectangular | 20 | Microphone Front Cavity Depth | ¥ | 0.00022795 | 0.00020262 | 0.0001784 | 0.00015747 | 0.00013985 | 0.00012113 | 0.00010682 | 9.2501e-05 | 8.0388e-05 | 6.9376e-05 | 6.8275e-05
|  | B | Rectangular | 21 | Microphone Front Cavity Volume | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 22 | Microphone Equivalent Volume | ¥ | 0.00052087 | 0.00046141 | 0.00040745 | 0.00035789 | 0.00031495 | 0.0002753 | 0.00023896 | 0.00020703 | 0.0001795 | 0.00015307 | 0.00012334
|  | B | Rectangular | 23 | Microphone Diaphragm Resonance Frequency | ¥ | 0 | 0 | 1.1012e-06 | 1.1012e-06 | 0 | 0 | 0 | 0 | 0 | 0 | 2.6665e-14
|  | B | Rectangular | 24 | Microphone Diaphragm Loss Factor | ¥ | 0 | 1.1012e-06 | 1.1012e-06 | 0 | 2.2024e-06 | 1.1012e-06 | 1.1012e-06 | 1.1012e-06 | 1.1012e-06 | 1.1012e-06 | 1.1012e-06
|  | B | Rectangular | 25 | Microphone Diaphragm Compliance | ¥ | - | - | - | - | - | - | - | - | - | -
|  | B | Rectangular | 26 | Microphone Diaphragm Mass | ¥ | - | - | - | - | - | - | - | - | - | -
|  | B | Rectangular | 27 | Microphone Diaphragm Resistance | ¥ | - | - | - | - | - | - | - | - | - | -
|  | B | Rectangular | 25 | Mic. Cavity Thread - Short Coup. (heat cond.) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 27b | Mic. Cavity Thread - Long Coup. (heat cond.) | ¥ | [8 empty] | - | -
|  | B | Rectangular | 26 | Microphone Polarization Voltage | ¥ | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739
| uThe |  |  | Acoustic Transfer Imp. (imperfection of theory) |  |  | 0.0065643 | 0.0066568 | 0.0067405 | 0.0068143 | 0.0068848 | 0.0069442 | 0.0069993 | 0.0070488 | 0.0070929 | 0.0071314 | 0.017725
|  | B | Rectangular | 27 | Heat Conduction | ¥ | 0.0065643 | 0.0066568 | 0.0067405 | 0.0068143 | 0.0068848 | 0.0069442 | 0.0069993 | 0.0070488 | 0.0070929 | 0.0071314 | 0.017725
|  |  |  | 28 | Adding of Excess Volume | ¥ | - | - | - | - | - | - | - | - | - | -
|  |  |  | 28 | Viscosity Losses | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 28 | Radial Wave-motion | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Processing of Results
| uRoud | B | Rectangular | 29 | Rounding to number of decimals: | ¥ | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868
| uRepeat | A | Normal | 30 | Repeatability of measurements | 3 | 0.0016969 | 0.0016964 | 0.0017025 | 0.0018628 | 0.0019058 | 0.0017954 | 0.0018187 | 0.0020749 | 0.001823 | 0.0020161 | 0.0021426
| uPcorr |  |  | Pressure correction |  |  | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573
|  | B | Rectangular | 31 | Static Pressure Correction UDP x Coef(500 Hz) | ¥ | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014
|  | B | Rectangular | 32 | Static Pressure Correction DP x Ucoef(500 Hz) | ¥ | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156
|  | B | Rectangular | 33 | Static Pressure Correction UDP x normCoef(f) | ¥ | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05
|  | B | Rectangular | 34 | Static Pressure Correction DP x UnormCoef(f) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| uTcorr |  |  | Temperature correction |  |  | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645
|  | B | Rectangular | 35 | Temperature Correction UDT x Coef(500 Hz) | ¥ | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033
|  | B | Rectangular | 36 | Temperature Correction DT x Ucoef(500 Hz) | ¥ | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403
|  | B | Rectangular | 37 | Temperature Correction UDT x normCoef(f) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 38 | Temperature Correction DT x UnormCoef(f) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| uC | Combined standard uncertainty |  |  |  |  | 0.098621 | 0.098625 | 0.098629 | 0.098635 | 0.09864 | 0.063386 | 0.041273 | 0.0278 | 0.019962 | 0.015813 | 0.021272
| Uk | Expanded uncertainty (k » 2) |  |  |  |  | 0.19724 | 0.19725 | 0.19726 | 0.19727 | 0.19728 | 0.12677 | 0.082546 | 0.055599 | 0.039923 | 0.031627 | 0.042544
| Reported uncertainty (dB) | [5 empty] | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.13 | 0.09 | 0.06 | 0.04 | 0.04 | 0.05
| Confidence uncertainty (dB) | [5 empty] | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.13 | 0.09 | 0.06 | 0.05 | 0.05 | 0.05
Uncertainty Components for the B&K Type 4160 Microphone with SN 2652762 as part of the APMP AUV.A-K5 intercomparison
| Symbol | Type | Distrubution | Component |  | DOF | Uncertainty of Microphone Sensitivity for 4160 (LS1P) [dB]
[7 empty] | 19.9528 | 25.119 | 31.623 | 39.811 | 50.119 | 63.096 | 79.433 | 100 | 125.893 | 158.49 | 199.527 | 251.189 | 316.228 | 398.108 | 501.19 | 630.96 | 794.33 | 1000 | 1258.9 | 1584.9 | 1995.3 | 2511.9 | 3162.3 | 3981.1 | 5011.9 | 6309.6 | 7943.3 | 10000
| uElec |  |  | Electrical Transfer Impedance |  |  | 0.0059805 | 0.0059474 | 0.0059109 | 0.0059109 | 0.0058313 | 0.0058143 | 0.0058143 | 0.00578 | 0.005763 | 0.005763 | 0.0057439 | 0.0057097 | 0.0057042 | 0.0057042 | 0.0056961 | 0.0056936 | 0.0056936 | 0.0056916 | 0.0056921 | 0.0056921 | 0.0056937 | 0.0056995 | 0.005704 | 0.005704 | 0.0057159 | 0.0057234 | 0.0057234 | 0.0057412
|  | B | Rectangualar | 1 | Series Capacitor | ¥ | 0.0018944 | 0.0017869 | 0.0016617 | 0.0016617 | 0.0013515 | 0.0012762 | 0.0012762 | 0.0011094 | 0.0010172 | 0.0010172 | 0.00090268 | 0.00064981 | 0.00059956 | 0.00059956 | 0.00051788 | 0.00048956 | 0.00048956 | 0.00046567 | 0.00047135 | 0.00047135 | 0.00049037 | 0.0005537 | 0.00059796 | 0.00059796 | 0.00070264 | 0.00076086 | 0.00076086 | 0.00088516
|  | B | Rectangualar | 2 | Voltage Ratios | ¥ | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992
|  | B | Rectangualar | 3 | Cross Talk | ¥ | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564
|  | B | Rectangualar | 4 | Inherent and Ambient Noise | ¥ | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05
|  | B | Rectangualar | 5 | Distortion | ¥ | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417
|  | B | Rectangualar | 6 | Frequency | ¥ | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05
|  | B | Rectangualar | 7 | Receiver Ground Shielding | ¥ | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641
|  | B | Rectangualar | 8 | Receiver Ground Shielding (guard length) | ¥ | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376
|  | B | Rectangualar | 9 | Transmitter Ground Shielding | ¥ | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509
|  | B | Rectangualar | 10 | Transmitter Ground Shielding (guard length) | ¥ | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376
| uCoup |  |  | Acoustic Transfer Impedance (coupler and gas) |  |  | 0.0089925 | 0.0075537 | 0.0068971 | 0.0066174 | 0.0065026 | 0.0064562 | 0.0064377 | 0.0064303 | 0.0064273 | 0.0064261 | 0.0064256 | 0.0064254 | 0.0064254 | 0.0064253 | 0.0064253 | 0.0064253 | 0.0064253 | 0.0064254 | 0.0064255 | 0.0064263 | 0.0064282 | 0.006428 | 0.0064254 | 0.0064288 | 0.0064545 | 0.0065466 | 0.0068239 | 0.0078078
|  | B | Rectangular | 11 | Coupler Length | ¥ | 7.7085e-06 | 5.506e-06 | 3.3036e-06 | 1.1012e-06 | 1.1012e-06 | 2.2024e-06 | 4.4048e-06 | 4.4048e-06 | 4.4048e-06 | 6.6072e-06 | 6.6072e-06 | 6.6072e-06 | 5.506e-06 | 4.4048e-06 | 0 | 7.7085e-06 | 1.7619e-05 | 3.4137e-05 | 5.1757e-05 | 0.00010241 | 0.00017069 | 0.00016848 | 3.5239e-05 | 0.0001806 | 0.00052968 | 0.0010924 | 0.0020372 | 0.0040282
|  | B | Rectangular | 12 | Coupler Diameter | ¥ | 4.4048e-06 | 4.4048e-06 | 4.4048e-06 | 4.4048e-06 | 4.4048e-06 | 4.4048e-06 | 5.506e-06 | 5.506e-06 | 5.506e-06 | 5.506e-06 | 4.4048e-06 | 3.3036e-06 | 3.3036e-06 | 1.1012e-06 | 0 | 3.3036e-06 | 8.8097e-06 | 1.7619e-05 | 2.753e-05 | 5.506e-05 | 9.14e-05 | 8.4793e-05 | 9.9109e-06 | 0.00011232 | 0.00030944 | 0.00061668 | 0.0010638 | 0.0018577
|  | B | Rectangular | 13 | Coupler Volume (short 7. 5 mm) + volume fitting | ¥ | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776
|  | B | Rectangular | 11b | Coupler Volume (long 15 mm) + volume fitting | ¥
|  | B | Rectangular | 14 | Coupler Surface Area (short 7.5 mm) | ¥ | 6.9282e-05 | 5.7735e-05 | 4.6188e-05 | 4.6188e-05 | 4.6188e-05 | 3.4641e-05 | 3.4641e-05 | 3.4641e-05 | 2.3094e-05 | 2.3094e-05 | 2.3094e-05 | 2.3094e-05 | 2.3094e-05 | 2.3094e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 15 | Coupler Surface Area (long 15.0 mm) | ¥ | 8.0829e-05 | 5.7735e-05 | 5.7735e-05 | 4.6188e-05 | 3.4641e-05 | 3.4641e-05 | 2.3094e-05 | 1.1547e-05 | 1.1547e-05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 13 | Capillary Tube Length | ¥
|  | B | Rectangular | 14 | Capillary Tube Diameter | ¥
|  | B | Rectangular | 16 | Coupler Leakage | ¥ | 0.0062904 | 0.0039708 | 0.0025061 | 0.0015816 | 0.00099802 | 0.00062975 | 0.00039736 | 0.00025073 | 0.0001582 | 9.9819e-05 | 6.2982e-05 | 3.9739e-05 | 2.5074e-05 | 1.5821e-05 | 9.9821e-06 | 6.2983e-06 | 3.974e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06
|  | B | Rectangular | 17 | Static Pressure of air in Coupler | ¥ | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733
|  | B | Rectangular | 18 | Temperature of air in Coupler | ¥ | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321
|  | B | Rectangular | 19 | Relative Humidity of air in Coupler | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| uMic |  |  | Acoustic Transfer Impedance (microphones) |  |  | 0.0024779 | 0.0024767 | 0.0024758 | 0.0024752 | 0.0024748 | 0.0024744 | 0.0024742 | 0.0024741 | 0.002474 | 0.0024739 | 0.0024739 | 0.002474 | 0.0024741 | 0.0024746 | 0.0024758 | 0.0024788 | 0.0024867 | 0.0025068 | 0.00254 | 0.0026778 | 0.0029509 | 0.0040401 | 0.005479 | 0.0077231 | 0.010101 | 0.0094179 | 0.0054524 | 0.01482
|  | B | Rectangular | 20 | Microphone Front Cavity Depth | ¥ | 6.8275e-05 | 5.9465e-05 | 5.0656e-05 | 4.515e-05 | 3.8542e-05 | 3.3036e-05 | 2.753e-05 | 2.3125e-05 | 1.8721e-05 | 1.5417e-05 | 1.2113e-05 | 9.9109e-06 | 5.506e-06 | 4.4048e-06 | 0 | 2.2024e-06 | 6.6072e-06 | 1.1012e-05 | 1.5417e-05 | 2.6429e-05 | 3.9643e-05 | 6.0566e-05 | 9.3603e-05 | 0.00014095 | 0.00020262 | 0.00024337 | 0.00017289 | 3.8542e-05
|  | B | Rectangular | 21 | Microphone Front Cavity Volume | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 22 | Microphone Equivalent Volume | ¥ | 0.00012334 | 0.00010351 | 8.4793e-05 | 6.9376e-05 | 5.3959e-05 | 4.0745e-05 | 2.9733e-05 | 1.9822e-05 | 1.1012e-05 | 4.4048e-06 | 1.1012e-06 | 5.506e-06 | 7.7085e-06 | 5.506e-06 | 0 | 1.1012e-05 | 3.5239e-05 | 7.4882e-05 | 0.00012113 | 0.00024997 | 0.00043057 | 0.00059796 | 0.00022355 | 0.00042176 | 0.0015065 | 0.0030283 | 0.0039379 | 0.0021264
|  | B | Rectangular | 23 | Microphone Diaphragm Resonance Frequency | ¥ | 2.6665e-14 | 0 | 1.1012e-06 | 0 | 0 | 0 | 0 | 1.1012e-06 | 2.2024e-06 | 2.2024e-06 | 3.3036e-06 | 2.2024e-06 | 4.4048e-06 | 3.3036e-06 | 1.8721e-05 | 2.3125e-05 | 2.753e-05 | 3.4137e-05 | 3.5239e-05 | 2.3125e-05 | 8.8097e-06 | 0.00013655 | 0.00043828 | 0.0013336 | 0.0033047 | 0.0048563 | 0.0012245 | 0.0067978
|  | B | Rectangular | 24 | Microphone Diaphragm Loss Factor | ¥ | 1.1012e-06 | 1.1012e-06 | 1.1012e-06 | 1.1012e-06 | 1.1012e-06 | 0 | 0 | 0 | 2.2024e-06 | 6.6072e-06 | 1.1012e-05 | 2.0923e-05 | 3.5239e-05 | 5.9465e-05 | 9.5805e-05 | 0.00015417 | 0.00024777 | 0.00039643 | 0.00056162 | 0.00099329 | 0.0015494 | 0.003134 | 0.0048629 | 0.0071799 | 0.0090927 | 0.0070543 | 0.0025636 | 0.012759
|  | B | Rectangular | 25 | Microphone Diaphragm Compliance | ¥
|  | B | Rectangular | 26 | Microphone Diaphragm Mass | ¥
|  | B | Rectangular | 27 | Microphone Diaphragm Resistance | ¥
|  | B | Rectangular | 25 | Mic. Cavity Thread - Short Coup. (heat cond.) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 27b | Mic. Cavity Thread - Long Coup. (heat cond.) | ¥
|  | B | Rectangular | 26 | Microphone Polarization Voltage | ¥ | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739
| uThe |  |  | Acoustic Transfer Imp. (imperfection of theory) |  |  | 0.017725 | 0.015877 | 0.014203 | 0.014203 | 0.014203 | 0.010161 | 0.010161 | 0.010161 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0050807 | 0.005081 | 0.005081 | 0.0035223 | 0.0035336 | 0.0035336 | 0.0022701 | 0.0025047 | 0.0025047 | 0.0013976 | 0.0046375 | 0.0046375 | 0.0046375 | 0.0073768 | 0.010795 | 0.014566 | 0.021351
|  | B | Rectangular | 27 | Heat Conduction | ¥ | 0.017725 | 0.015877 | 0.014203 | 0.014203 | 0.014203 | 0.010161 | 0.010161 | 0.010161 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0050807 | 0.0050807 | 0.0050807 | 0.0035218 | 0.0035218 | 0.0035218 | 0.0022517 | 0.0022517 | 0.0022517 | 0.00086603 | 0.0015011 | 0.0015011 | 0.0015011 | 0.002829 | 0.0046188 | 0.0073323 | 0.012644
|  |  |  | 28 | Adding of Excess Volume | ¥
|  |  |  | 28 | Viscosity Losses | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 28 | Radial Wave-motion | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 0.00028868 | 0.00028868 | 0.00028868 | 0.001097 | 0.001097 | 0.001097 | 0.0043879 | 0.0043879 | 0.0043879 | 0.0068127 | 0.0097572 | 0.012586 | 0.017205
Processing of Results
| uRoud | B | Rectangular | 29 | Rounding to number of decimals: | ¥ | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868
| uRepeat | A | Normal | 30 | Repeatability of measurements | 3 | 0.0033724 | 0.0034207 | 0.003356 | 0.0032677 | 0.0032284 | 0.0031579 | 0.0033713 | 0.0032594 | 0.0032778 | 0.003553 | 0.0032614 | 0.0033124 | 0.0032285 | 0.003326 | 0.0032403 | 0.0032138 | 0.0033579 | 0.0032536 | 0.0031203 | 0.0030896 | 0.0027161 | 0.0027941 | 0.0017581 | 0.0021896 | 0.0047131 | 0.0077832 | 0.010778 | 0.013103
| uPcorr |  |  | Pressure correction |  |  | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015731 | 0.0015738 | 0.0015753 | 0.0015785 | 0.0015843 | 0.0015945 | 0.0016341 | 0.0017103 | 0.0018425 | 0.0018931 | 0.0015966 | 0.0038068 | 0.0059298
|  | B | Rectangular | 31 | Static Pressure Correction UDP x Coef(500 Hz) | ¥ | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014
|  | B | Rectangular | 32 | Static Pressure Correction DP x Ucoef(500 Hz) | ¥ | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156
|  | B | Rectangular | 33 | Static Pressure Correction UDP x normCoef(f) | ¥ | 1.5492e-05 | 1.5492e-05 | 1.4201e-05 | 1.101e-05 | 6.9934e-06 | 1.9365e-06 | 1.9365e-06 | 1.9365e-06 | 1.9365e-06 | 1.9365e-06 | 1.9365e-06 | 1.9365e-06 | 1.4327e-06 | 7.9847e-07 | 0 | 1.3435e-06 | 3.0348e-06 | 5.164e-06 | 1.0537e-05 | 1.7303e-05 | 2.582e-05 | 4.8826e-05 | 7.779e-05 | 0.00011425 | 0.0001717 | 0.00021172 | 0.00013297 | 0.00013426
|  | B | Rectangular | 34 | Static Pressure Correction DP x UnormCoef(f) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2.253e-05 | 5.0894e-05 | 8.6603e-05 | 0.00013166 | 0.00018839 | 0.00025981 | 0.00044004 | 0.00066696 | 0.00095263 | 0.0010392 | 0.00017321 | 0.0034641 | 0.0057158
| uTcorr |  |  | Temperature correction |  |  | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00042511 | 0.00045892 | 0.00053016 | 0.00063208 | 0.00078134 | 0.00098694 | 0.0017972 | 0.0028534 | 0.0041963 | 0.0064558 | 0.0083616 | 0.0069321 | 0.0025398
|  | B | Rectangular | 35 | Temperature Correction UDT x Coef(500 Hz) | ¥ | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033
|  | B | Rectangular | 36 | Temperature Correction DT x Ucoef(500 Hz) | ¥ | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403
|  | B | Rectangular | 37 | Temperature Correction UDT x normCoef(f) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2.1463e-05 | 4.8484e-05 | 8.25e-05 | 0.00011957 | 0.00016624 | 0.000225 | 0.00043962 | 0.00070983 | 0.00105 | 0.00162 | 0.0021 | 0.00174 | 0.00063
|  | B | Rectangular | 38 | Temperature Correction DT x UnormCoef(f) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -8.2612e-05 | -0.00018661 | -0.00031754 | -0.00046022 | -0.00063987 | -0.00086603 | -0.0016921 | -0.0027321 | -0.0040415 | -0.0062354 | -0.0080829 | -0.0066973 | 0.0024249
| uC | Combined standard uncertainty |  |  |  |  | 0.021431 | 0.019322 | 0.017681 | 0.017557 | 0.01748 | 0.014346 | 0.014387 | 0.014343 | 0.012428 | 0.012503 | 0.012414 | 0.011304 | 0.011277 | 0.011305 | 0.010665 | 0.010661 | 0.010708 | 0.010334 | 0.010362 | 0.010399 | 0.01018 | 0.011562 | 0.012159 | 0.013753 | 0.01751 | 0.020548 | 0.022545 | 0.031481
| Uk | Expanded uncertainty (k » 2) |  |  |  |  | 0.042862 | 0.038643 | 0.035361 | 0.035113 | 0.034959 | 0.028693 | 0.028773 | 0.028687 | 0.024856 | 0.025005 | 0.024828 | 0.022607 | 0.022553 | 0.02261 | 0.02133 | 0.021321 | 0.021417 | 0.020669 | 0.020724 | 0.020798 | 0.02036 | 0.023123 | 0.024318 | 0.027506 | 0.035019 | 0.041096 | 0.045089 | 0.062963
| Reported uncertainty (dB) | [5 empty] | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.05 | 0.05 | 0.07
| Confidence uncertainty (dB) | [5 empty] | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 | 0.05 | 0.08
Uncertainty Components for the B&K Type 4160 Microphone with SN 2652765 as part of the APMP AUV.A-K5 intercomparison
| Symbol | Type | Distrubution | Component |  | DOF | [5 empty] | Uncertainty of Microphone Sensitivity for 4160 (LS1P) [dB]
[7 empty] | 1.9953 | 2.5119 | 3.1623 | 3.9811 | 5.0119 | 6.3096 | 7.9433 | 10 | 12.5892 | 15.8489 | 19.9528
| uElec |  |  | Electrical Transfer Impedance |  |  | 0.0063069 | 0.006283 | 0.00626 | 0.0062378 | 0.0062164 | 0.0061959 | 0.0061762 | 0.0061573 | 0.0061394 | 0.0061223 | 0.0059805
|  | B | Rectangualar | 1 | Series Capacitor | ¥ | 0.0027566 | 0.0027016 | 0.0026476 | 0.0025946 | 0.0025428 | 0.0024922 | 0.0024428 | 0.0023948 | 0.0023482 | 0.0023032 | 0.0018944
|  | B | Rectangualar | 2 | Voltage Ratios | ¥ | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992
|  | B | Rectangualar | 3 | Cross Talk | ¥ | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564
|  | B | Rectangualar | 4 | Inherent and Ambient Noise | ¥ | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05
|  | B | Rectangualar | 5 | Distortion | ¥ | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417
|  | B | Rectangualar | 6 | Frequency | ¥ | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05
|  | B | Rectangualar | 7 | Receiver Ground Shielding | ¥ | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641
|  | B | Rectangualar | 8 | Receiver Ground Shielding (guard length) | ¥ | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376
|  | B | Rectangualar | 9 | Transmitter Ground Shielding | ¥ | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509
|  | B | Rectangualar | 10 | Transmitter Ground Shielding (guard length) | ¥ | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376
| uCoup |  |  | Acoustic Transfer Impedance (coupler and gas) |  |  | 0.098096 | 0.098096 | 0.098096 | 0.098096 | 0.098096 | 0.062536 | 0.039948 | 0.025764 | 0.01703 | 0.011855 | 0.0089925
|  | B | Rectangular | 11 | Coupler Length | ¥ | 6.2769e-05 | 5.506e-05 | 4.9554e-05 | 4.2947e-05 | 3.634e-05 | 3.1935e-05 | 2.753e-05 | 2.4227e-05 | 1.9822e-05 | 1.7619e-05 | 7.7085e-06
|  | B | Rectangular | 12 | Coupler Diameter | ¥ | 7.7085e-06 | 6.6072e-06 | 4.4048e-06 | 3.3036e-06 | 3.3036e-06 | 2.2024e-06 | 2.2024e-06 | 2.2024e-06 | 1.1012e-06 | 1.1012e-06 | 4.4048e-06
|  | B | Rectangular | 13 | Coupler Volume (short 7. 5 mm) + volume fitting | ¥ | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776
|  | B | Rectangular | 11b | Coupler Volume (long 15 mm) + volume fitting | ¥ | 0.0106 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776
|  | B | Rectangular | 14 | Coupler Surface Area (short 7.5 mm) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.9282e-05
|  | B | Rectangular | 15 | Coupler Surface Area (long 15.0 mm) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8.0829e-05
|  | B | Rectangular | 13 | Capillary Tube Length | ¥ | - | - | - | - | - | - | - | - | - | -
|  | B | Rectangular | 14 | Capillary Tube Diameter | ¥ | - | - | - | - | - | - | - | - | - | -
|  | B | Rectangular | 16 | Coupler Leakage | ¥ | 0.097886 | 0.097886 | 0.097886 | 0.097886 | 0.097886 | 0.062205 | 0.039428 | 0.024949 | 0.015771 | 0.0099623 | 0.0062904
|  | B | Rectangular | 17 | Static Pressure of air in Coupler | ¥ | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733
|  | B | Rectangular | 18 | Temperature of air in Coupler | ¥ | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321
|  | B | Rectangular | 19 | Relative Humidity of air in Coupler | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| uMic |  |  | Acoustic Transfer Impedance (microphones) |  |  | 0.0025384 | 0.0025247 | 0.0025135 | 0.0025046 | 0.0024977 | 0.0024921 | 0.0024877 | 0.0024842 | 0.0024817 | 0.0024796 | 0.0024779
|  | B | Rectangular | 20 | Microphone Front Cavity Depth | ¥ | 0.00022795 | 0.00020262 | 0.0001784 | 0.00015747 | 0.00013985 | 0.00012113 | 0.00010682 | 9.2501e-05 | 8.0388e-05 | 6.9376e-05 | 6.8275e-05
|  | B | Rectangular | 21 | Microphone Front Cavity Volume | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 22 | Microphone Equivalent Volume | ¥ | 0.00052087 | 0.00046141 | 0.00040745 | 0.00035789 | 0.00031495 | 0.0002753 | 0.00023896 | 0.00020703 | 0.0001795 | 0.00015307 | 0.00012334
|  | B | Rectangular | 23 | Microphone Diaphragm Resonance Frequency | ¥ | 0 | 0 | 1.1012e-06 | 1.1012e-06 | 0 | 0 | 0 | 0 | 0 | 0 | 2.6665e-14
|  | B | Rectangular | 24 | Microphone Diaphragm Loss Factor | ¥ | 0 | 1.1012e-06 | 1.1012e-06 | 0 | 2.2024e-06 | 1.1012e-06 | 1.1012e-06 | 1.1012e-06 | 1.1012e-06 | 1.1012e-06 | 1.1012e-06
|  | B | Rectangular | 25 | Microphone Diaphragm Compliance | ¥ | - | - | - | - | - | - | - | - | - | -
|  | B | Rectangular | 26 | Microphone Diaphragm Mass | ¥ | - | - | - | - | - | - | - | - | - | -
|  | B | Rectangular | 27 | Microphone Diaphragm Resistance | ¥ | - | - | - | - | - | - | - | - | - | -
|  | B | Rectangular | 25 | Mic. Cavity Thread - Short Coup. (heat cond.) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 27b | Mic. Cavity Thread - Long Coup. (heat cond.) | ¥ | [8 empty] | - | -
|  | B | Rectangular | 26 | Microphone Polarization Voltage | ¥ | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739
| uThe |  |  | Acoustic Transfer Imp. (imperfection of theory) |  |  | 0.0065643 | 0.0066568 | 0.0067405 | 0.0068143 | 0.0068848 | 0.0069442 | 0.0069993 | 0.0070488 | 0.0070929 | 0.0071314 | 0.017725
|  | B | Rectangular | 27 | Heat Conduction | ¥ | 0.0065643 | 0.0066568 | 0.0067405 | 0.0068143 | 0.0068848 | 0.0069442 | 0.0069993 | 0.0070488 | 0.0070929 | 0.0071314 | 0.017725
|  |  |  | 28 | Adding of Excess Volume | ¥ | - | - | - | - | - | - | - | - | - | -
|  |  |  | 28 | Viscosity Losses | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 28 | Radial Wave-motion | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Processing of Results
| uRoud | B | Rectangular | 29 | Rounding to number of decimals: | ¥ | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868
| uRepeat | A | Normal | 30 | Repeatability of measurements | 3 | 0.0025063 | 0.0027231 | 0.0033118 | 0.0037928 | 0.0035451 | 0.0042796 | 0.0045438 | 0.0042064 | 0.0038995 | 0.0041642 | 0.0042793
| uPcorr |  |  | Pressure correction |  |  | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573
|  | B | Rectangular | 31 | Static Pressure Correction UDP x Coef(500 Hz) | ¥ | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014
|  | B | Rectangular | 32 | Static Pressure Correction DP x Ucoef(500 Hz) | ¥ | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156
|  | B | Rectangular | 33 | Static Pressure Correction UDP x normCoef(f) | ¥ | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05 | 1.5492e-05
|  | B | Rectangular | 34 | Static Pressure Correction DP x UnormCoef(f) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| uTcorr |  |  | Temperature correction |  |  | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645
|  | B | Rectangular | 35 | Temperature Correction UDT x Coef(500 Hz) | ¥ | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033
|  | B | Rectangular | 36 | Temperature Correction DT x Ucoef(500 Hz) | ¥ | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403
|  | B | Rectangular | 37 | Temperature Correction UDT x normCoef(f) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 38 | Temperature Correction DT x UnormCoef(f) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| uC | Combined standard uncertainty |  |  |  |  | 0.098638 | 0.098648 | 0.09867 | 0.098691 | 0.098685 | 0.063505 | 0.041483 | 0.028039 | 0.020257 | 0.016228 | 0.021592
| Uk | Expanded uncertainty (k » 2) |  |  |  |  | 0.19728 | 0.1973 | 0.19734 | 0.19738 | 0.19737 | 0.12701 | 0.082965 | 0.056079 | 0.040514 | 0.032456 | 0.043185
| Reported uncertainty (dB) | [5 empty] | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.13 | 0.09 | 0.06 | 0.05 | 0.04 | 0.05
| Confidence uncertainty (dB) | [5 empty] | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.13 | 0.09 | 0.06 | 0.05 | 0.05 | 0.05
Uncertainty Components for the B&K Type 4160 Microphone with SN 2652765 as part of the APMP AUV.A-K5 intercomparison
| Symbol | Type | Distrubution | Component |  | DOF | Uncertainty of Microphone Sensitivity for 4160 (LS1P) [dB]
[7 empty] | 19.9528 | 25.119 | 31.623 | 39.811 | 50.119 | 63.096 | 79.433 | 100 | 125.893 | 158.49 | 199.527 | 251.189 | 316.228 | 398.108 | 501.19 | 630.96 | 794.33 | 1000 | 1258.9 | 1584.9 | 1995.3 | 2511.9 | 3162.3 | 3981.1 | 5011.9 | 6309.6 | 7943.3 | 10000
| uElec |  |  | Electrical Transfer Impedance |  |  | 0.0059805 | 0.0059474 | 0.0059109 | 0.0059109 | 0.0058313 | 0.0058143 | 0.0058143 | 0.00578 | 0.005763 | 0.005763 | 0.0057439 | 0.0057097 | 0.0057042 | 0.0057042 | 0.0056961 | 0.0056936 | 0.0056936 | 0.0056916 | 0.0056921 | 0.0056921 | 0.0056937 | 0.0056995 | 0.005704 | 0.005704 | 0.0057159 | 0.0057234 | 0.0057234 | 0.0057412
|  | B | Rectangualar | 1 | Series Capacitor | ¥ | 0.0018944 | 0.0017869 | 0.0016617 | 0.0016617 | 0.0013515 | 0.0012762 | 0.0012762 | 0.0011094 | 0.0010172 | 0.0010172 | 0.00090268 | 0.00064981 | 0.00059956 | 0.00059956 | 0.00051788 | 0.00048956 | 0.00048956 | 0.00046567 | 0.00047135 | 0.00047135 | 0.00049037 | 0.0005537 | 0.00059796 | 0.00059796 | 0.00070264 | 0.00076086 | 0.00076086 | 0.00088516
|  | B | Rectangualar | 2 | Voltage Ratios | ¥ | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992 | 0.0040992
|  | B | Rectangualar | 3 | Cross Talk | ¥ | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564 | 0.0012564
|  | B | Rectangualar | 4 | Inherent and Ambient Noise | ¥ | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05 | 3.1491e-05
|  | B | Rectangualar | 5 | Distortion | ¥ | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417 | 0.00022417
|  | B | Rectangualar | 6 | Frequency | ¥ | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05 | 1.0447e-05
|  | B | Rectangualar | 7 | Receiver Ground Shielding | ¥ | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641
|  | B | Rectangualar | 8 | Receiver Ground Shielding (guard length) | ¥ | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376
|  | B | Rectangualar | 9 | Transmitter Ground Shielding | ¥ | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509 | 0.00019509
|  | B | Rectangualar | 10 | Transmitter Ground Shielding (guard length) | ¥ | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376 | 0.00092376
| uCoup |  |  | Acoustic Transfer Impedance (coupler and gas) |  |  | 0.0089925 | 0.0075537 | 0.0068971 | 0.0066174 | 0.0065026 | 0.0064562 | 0.0064377 | 0.0064303 | 0.0064273 | 0.0064261 | 0.0064256 | 0.0064254 | 0.0064254 | 0.0064253 | 0.0064253 | 0.0064253 | 0.0064253 | 0.0064254 | 0.0064255 | 0.0064263 | 0.0064282 | 0.006428 | 0.0064254 | 0.0064288 | 0.0064545 | 0.0065466 | 0.0068239 | 0.0078078
|  | B | Rectangular | 11 | Coupler Length | ¥ | 7.7085e-06 | 5.506e-06 | 3.3036e-06 | 1.1012e-06 | 1.1012e-06 | 2.2024e-06 | 4.4048e-06 | 4.4048e-06 | 4.4048e-06 | 6.6072e-06 | 6.6072e-06 | 6.6072e-06 | 5.506e-06 | 4.4048e-06 | 0 | 7.7085e-06 | 1.7619e-05 | 3.4137e-05 | 5.1757e-05 | 0.00010241 | 0.00017069 | 0.00016848 | 3.5239e-05 | 0.0001806 | 0.00052968 | 0.0010924 | 0.0020372 | 0.0040282
|  | B | Rectangular | 12 | Coupler Diameter | ¥ | 4.4048e-06 | 4.4048e-06 | 4.4048e-06 | 4.4048e-06 | 4.4048e-06 | 4.4048e-06 | 5.506e-06 | 5.506e-06 | 5.506e-06 | 5.506e-06 | 4.4048e-06 | 3.3036e-06 | 3.3036e-06 | 1.1012e-06 | 0 | 3.3036e-06 | 8.8097e-06 | 1.7619e-05 | 2.753e-05 | 5.506e-05 | 9.14e-05 | 8.4793e-05 | 9.9109e-06 | 0.00011232 | 0.00030944 | 0.00061668 | 0.0010638 | 0.0018577
|  | B | Rectangular | 13 | Coupler Volume (short 7. 5 mm) + volume fitting | ¥ | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776 | 0.0061776
|  | B | Rectangular | 11b | Coupler Volume (long 15 mm) + volume fitting | ¥
|  | B | Rectangular | 14 | Coupler Surface Area (short 7.5 mm) | ¥ | 6.9282e-05 | 5.7735e-05 | 4.6188e-05 | 4.6188e-05 | 4.6188e-05 | 3.4641e-05 | 3.4641e-05 | 3.4641e-05 | 2.3094e-05 | 2.3094e-05 | 2.3094e-05 | 2.3094e-05 | 2.3094e-05 | 2.3094e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 1.1547e-05 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 15 | Coupler Surface Area (long 15.0 mm) | ¥ | 8.0829e-05 | 5.7735e-05 | 5.7735e-05 | 4.6188e-05 | 3.4641e-05 | 3.4641e-05 | 2.3094e-05 | 1.1547e-05 | 1.1547e-05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 13 | Capillary Tube Length | ¥
|  | B | Rectangular | 14 | Capillary Tube Diameter | ¥
|  | B | Rectangular | 16 | Coupler Leakage | ¥ | 0.0062904 | 0.0039708 | 0.0025061 | 0.0015816 | 0.00099802 | 0.00062975 | 0.00039736 | 0.00025073 | 0.0001582 | 9.9819e-05 | 6.2982e-05 | 3.9739e-05 | 2.5074e-05 | 1.5821e-05 | 9.9821e-06 | 6.2983e-06 | 3.974e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06 | 2.5074e-06
|  | B | Rectangular | 17 | Static Pressure of air in Coupler | ¥ | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733 | 0.00034733
|  | B | Rectangular | 18 | Temperature of air in Coupler | ¥ | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321
|  | B | Rectangular | 19 | Relative Humidity of air in Coupler | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| uMic |  |  | Acoustic Transfer Impedance (microphones) |  |  | 0.0024779 | 0.0024767 | 0.0024758 | 0.0024752 | 0.0024748 | 0.0024744 | 0.0024742 | 0.0024741 | 0.002474 | 0.0024739 | 0.0024739 | 0.002474 | 0.0024741 | 0.0024746 | 0.0024758 | 0.0024788 | 0.0024867 | 0.0025068 | 0.00254 | 0.0026778 | 0.0029509 | 0.0040401 | 0.005479 | 0.0077231 | 0.010101 | 0.0094179 | 0.0054524 | 0.01482
|  | B | Rectangular | 20 | Microphone Front Cavity Depth | ¥ | 6.8275e-05 | 5.9465e-05 | 5.0656e-05 | 4.515e-05 | 3.8542e-05 | 3.3036e-05 | 2.753e-05 | 2.3125e-05 | 1.8721e-05 | 1.5417e-05 | 1.2113e-05 | 9.9109e-06 | 5.506e-06 | 4.4048e-06 | 0 | 2.2024e-06 | 6.6072e-06 | 1.1012e-05 | 1.5417e-05 | 2.6429e-05 | 3.9643e-05 | 6.0566e-05 | 9.3603e-05 | 0.00014095 | 0.00020262 | 0.00024337 | 0.00017289 | 3.8542e-05
|  | B | Rectangular | 21 | Microphone Front Cavity Volume | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 22 | Microphone Equivalent Volume | ¥ | 0.00012334 | 0.00010351 | 8.4793e-05 | 6.9376e-05 | 5.3959e-05 | 4.0745e-05 | 2.9733e-05 | 1.9822e-05 | 1.1012e-05 | 4.4048e-06 | 1.1012e-06 | 5.506e-06 | 7.7085e-06 | 5.506e-06 | 0 | 1.1012e-05 | 3.5239e-05 | 7.4882e-05 | 0.00012113 | 0.00024997 | 0.00043057 | 0.00059796 | 0.00022355 | 0.00042176 | 0.0015065 | 0.0030283 | 0.0039379 | 0.0021264
|  | B | Rectangular | 23 | Microphone Diaphragm Resonance Frequency | ¥ | 2.6665e-14 | 0 | 1.1012e-06 | 0 | 0 | 0 | 0 | 1.1012e-06 | 2.2024e-06 | 2.2024e-06 | 3.3036e-06 | 2.2024e-06 | 4.4048e-06 | 3.3036e-06 | 1.8721e-05 | 2.3125e-05 | 2.753e-05 | 3.4137e-05 | 3.5239e-05 | 2.3125e-05 | 8.8097e-06 | 0.00013655 | 0.00043828 | 0.0013336 | 0.0033047 | 0.0048563 | 0.0012245 | 0.0067978
|  | B | Rectangular | 24 | Microphone Diaphragm Loss Factor | ¥ | 1.1012e-06 | 1.1012e-06 | 1.1012e-06 | 1.1012e-06 | 1.1012e-06 | 0 | 0 | 0 | 2.2024e-06 | 6.6072e-06 | 1.1012e-05 | 2.0923e-05 | 3.5239e-05 | 5.9465e-05 | 9.5805e-05 | 0.00015417 | 0.00024777 | 0.00039643 | 0.00056162 | 0.00099329 | 0.0015494 | 0.003134 | 0.0048629 | 0.0071799 | 0.0090927 | 0.0070543 | 0.0025636 | 0.012759
|  | B | Rectangular | 25 | Microphone Diaphragm Compliance | ¥
|  | B | Rectangular | 26 | Microphone Diaphragm Mass | ¥
|  | B | Rectangular | 27 | Microphone Diaphragm Resistance | ¥
|  | B | Rectangular | 25 | Mic. Cavity Thread - Short Coup. (heat cond.) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 27b | Mic. Cavity Thread - Long Coup. (heat cond.) | ¥
|  | B | Rectangular | 26 | Microphone Polarization Voltage | ¥ | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739 | 0.0024739
| uThe |  |  | Acoustic Transfer Imp. (imperfection of theory) |  |  | 0.017725 | 0.015877 | 0.014203 | 0.014203 | 0.014203 | 0.010161 | 0.010161 | 0.010161 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0050807 | 0.005081 | 0.005081 | 0.0035223 | 0.0035336 | 0.0035336 | 0.0022701 | 0.0025047 | 0.0025047 | 0.0013976 | 0.0046375 | 0.0046375 | 0.0046375 | 0.0073768 | 0.010795 | 0.014566 | 0.021351
|  | B | Rectangular | 27 | Heat Conduction | ¥ | 0.017725 | 0.015877 | 0.014203 | 0.014203 | 0.014203 | 0.010161 | 0.010161 | 0.010161 | 0.0072169 | 0.0072169 | 0.0072169 | 0.0050807 | 0.0050807 | 0.0050807 | 0.0035218 | 0.0035218 | 0.0035218 | 0.0022517 | 0.0022517 | 0.0022517 | 0.00086603 | 0.0015011 | 0.0015011 | 0.0015011 | 0.002829 | 0.0046188 | 0.0073323 | 0.012644
|  |  |  | 28 | Adding of Excess Volume | ¥
|  |  |  | 28 | Viscosity Losses | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | B | Rectangular | 28 | Radial Wave-motion | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5.7735e-05 | 5.7735e-05 | 5.7735e-05 | 0.00028868 | 0.00028868 | 0.00028868 | 0.001097 | 0.001097 | 0.001097 | 0.0043879 | 0.0043879 | 0.0043879 | 0.0068127 | 0.0097572 | 0.012586 | 0.017205
Processing of Results
| uRoud | B | Rectangular | 29 | Rounding to number of decimals: | ¥ | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868
| uRepeat | A | Normal | 30 | Repeatability of measurements | 5 | 0.0019155 | 0.0017003 | 0.0018147 | 0.0017698 | 0.0018661 | 0.0017457 | 0.0017326 | 0.0017813 | 0.0017001 | 0.0018561 | 0.0017079 | 0.0016353 | 0.0016836 | 0.0016476 | 0.0016598 | 0.0016126 | 0.0015262 | 0.0015623 | 0.0016214 | 0.0014959 | 0.0013377 | 0.0013389 | 0.0010698 | 0.00099785 | 0.002232 | 0.0042171 | 0.0046041 | 0.001856
| uPcorr |  |  | Pressure correction |  |  | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.001573 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015729 | 0.0015731 | 0.0015738 | 0.0015753 | 0.0015785 | 0.0015843 | 0.0015945 | 0.0016341 | 0.0017103 | 0.0018425 | 0.0018931 | 0.0015966 | 0.0038068 | 0.0059298
|  | B | Rectangular | 31 | Static Pressure Correction UDP x Coef(500 Hz) | ¥ | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014 | 0.0002014
|  | B | Rectangular | 32 | Static Pressure Correction DP x Ucoef(500 Hz) | ¥ | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156 | 0.00156
|  | B | Rectangular | 33 | Static Pressure Correction UDP x normCoef(f) | ¥ | 1.5492e-05 | 1.5492e-05 | 1.4201e-05 | 1.101e-05 | 6.9934e-06 | 1.9365e-06 | 1.9365e-06 | 1.9365e-06 | 1.9365e-06 | 1.9365e-06 | 1.9365e-06 | 1.9365e-06 | 1.4327e-06 | 7.9847e-07 | 0 | 1.3435e-06 | 3.0348e-06 | 5.164e-06 | 1.0537e-05 | 1.7303e-05 | 2.582e-05 | 4.8826e-05 | 7.779e-05 | 0.00011425 | 0.0001717 | 0.00021172 | 0.00013297 | 0.00013426
|  | B | Rectangular | 34 | Static Pressure Correction DP x UnormCoef(f) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2.253e-05 | 5.0894e-05 | 8.6603e-05 | 0.00013166 | 0.00018839 | 0.00025981 | 0.00044004 | 0.00066696 | 0.00095263 | 0.0010392 | 0.00017321 | 0.0034641 | 0.0057158
| uTcorr |  |  | Temperature correction |  |  | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00041645 | 0.00042511 | 0.00045892 | 0.00053016 | 0.00063208 | 0.00078134 | 0.00098694 | 0.0017972 | 0.0028534 | 0.0041963 | 0.0064558 | 0.0083616 | 0.0069321 | 0.0025398
|  | B | Rectangular | 35 | Temperature Correction UDT x Coef(500 Hz) | ¥ | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033 | 0.00033
|  | B | Rectangular | 36 | Temperature Correction DT x Ucoef(500 Hz) | ¥ | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403 | 0.00025403
|  | B | Rectangular | 37 | Temperature Correction UDT x normCoef(f) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2.1463e-05 | 4.8484e-05 | 8.25e-05 | 0.00011957 | 0.00016624 | 0.000225 | 0.00043962 | 0.00070983 | 0.00105 | 0.00162 | 0.0021 | 0.00174 | 0.00063
|  | B | Rectangular | 38 | Temperature Correction DT x UnormCoef(f) | ¥ | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -8.2612e-05 | -0.00018661 | -0.00031754 | -0.00046022 | -0.00063987 | -0.00086603 | -0.0016921 | -0.0027321 | -0.0040415 | -0.0062354 | -0.0080829 | -0.0066973 | 0.0024249
| uC | Combined standard uncertainty |  |  |  |  | 0.021251 | 0.019092 | 0.017454 | 0.01734 | 0.01728 | 0.014103 | 0.014093 | 0.014081 | 0.012108 | 0.01213 | 0.012099 | 0.01093 | 0.010935 | 0.01093 | 0.010296 | 0.010292 | 0.010282 | 0.0099325 | 0.010013 | 0.010042 | 0.0099015 | 0.011299 | 0.012079 | 0.013614 | 0.01701 | 0.019479 | 0.02033 | 0.028685
| Uk | Expanded uncertainty (k » 2) |  |  |  |  | 0.042501 | 0.038185 | 0.034907 | 0.034681 | 0.03456 | 0.028206 | 0.028186 | 0.028163 | 0.024215 | 0.02426 | 0.024198 | 0.021861 | 0.02187 | 0.021859 | 0.020591 | 0.020584 | 0.020564 | 0.019865 | 0.020027 | 0.020083 | 0.019803 | 0.022597 | 0.024157 | 0.027229 | 0.034021 | 0.038958 | 0.04066 | 0.05737
| Reported uncertainty (dB) | [5 empty] | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.03 | 0.03 | 0.02 | 0.03 | 0.03 | 0.03 | 0.04 | 0.04 | 0.05 | 0.06
| Confidence uncertainty (dB) | [5 empty] | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 | 0.05 | 0.08

## NMIM Level
NMIM | Level
Uncertainty budget for LS1P microphones (B&K 4160)
| No | Input Parameter | Symbol | Std U | Unit | 31.5 | 63 | 125 | 250 | 500 | 1000 | 1250 | 1600 | 2000 | 2500 | 3150 | 4000 | 5000 | 6300 | 8000 | 10000
| 2.1.1 | Static Pressure | Ps1, Ps2, Ps3 | 0.039 | kPa | 0.0015673 | 0.0015583 | 0.0015525 | 0.0015484 | 0.001546 | 0.0015459 | 0.0015469 | 0.0015493 | 0.0015528 | 0.0015592 | 0.0015705 | 0.0015891 | 0.0016195 | 0.0016596 | 0.0016876 | 0.0016652
| 2.1.2 | Temperature | T1, T2, T3 | 3.057 | K | 0.00063 | 0.00045 | 0.00032 | 0.00025 | 0.00025 | 0.00049 | 0.0007 | 0.00107 | 0.00161 | 0.00248 | 0.00393 | 0.00638 | 0.01007 | 0.01604 | 0.0259 | 0.04436
| 2.1.3 | Relative Humidity | RH1, RH2, RH3 | 6.066 | % | 0.0003282 | 0.00032711 | 0.0003262 | 0.00032627 | 0.00032937 | 0.00034397 | 0.00035521 | 0.00037526 | 0.00040475 | 0.00045116 | 0.00052818 | 0.00066006 | 0.00085692 | 0.0011777 | 0.0016968 | 0.0026487
| 2.1 | Amb. Conditions |  |  |  | 0.0017207 | 0.0016547 | 0.0016183 | 0.001602 | 0.0016003 | 0.0016578 | 0.0017347 | 0.0019199 | 0.0022732 | 0.002964 | 0.004265 | 0.006608 | 0.010235 | 0.016169 | 0.02601 | 0.04447
| 2.2.1 | Coupler Length | Icoup | 0.00122 | mm | 0.00026502 | 0.00026599 | 0.0002664 | 0.00026698 | 0.00026868 | 0.00027525 | 0.00028018 | 0.00028889 | 0.00030178 | 0.00032191 | 0.00035531 | 0.0004135 | 0.00049953 | 0.00064123 | 0.00086659 | 0.0012722
| 2.2.2 | Coupler Diameter | dcoup | 0.00576 | mm | 2.2396e-05 | 1.535e-05 | 1.0787e-05 | 7.5928e-06 | 5.4691e-06 | 4.3143e-06 | 4.2227e-06 | 4.4625e-06 | 4.8323e-06 | 5.8989e-06 | 7.8863e-06 | 9.7457e-06 | 1.425e-05 | 1.8117e-05 | 2.5857e-05 | 4.2708e-05
| 2.2.3 | Coupler Volume | Vcoup | 1.5849 | mm3 | 0.0020345 | 0.0020357 | 0.0020361 | 0.0020361 | 0.0020352 | 0.0020311 | 0.002028 | 0.0020223 | 0.0020145 | 0.0020021 | 0.0019821 | 0.0019512 | 0.0019043 | 0.0018245 | 0.0016391 | 0.0011885
| 2.2.4 | Coup. Surf. Area | Scoup | 0.4764 | mm2 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001
| 2.2.5 | Coupler Leakage | - | see table | - | 0.0126 | 0.0114 | 0.0121 | 0.0128 | 0.0121 | 0.0129 | 0.0125 | 0.0119 | 0.0132 | 0.0132 | 0.0129 | 0.012 | 0.012 | 0.012 | 0.012 | 0.012
| 2.2 | Coupler |  |  |  | 0.012766 | 0.011584 | 0.012273 | 0.012964 | 0.012273 | 0.013062 | 0.012667 | 0.012074 | 0.013357 | 0.013355 | 0.013057 | 0.012165 | 0.012161 | 0.012155 | 0.012143 | 0.012126
| 2.3.1 | Mic. Front Length | La | 0.01 | mm | 9.7938e-05 | 6.9122e-05 | 5.0587e-05 | 4.2314e-05 | 5.6311e-05 | 0.00014572 | 0.00021658 | 0.00034438 | 0.00053113 | 0.00082723 | 0.0013188 | 0.0021472 | 0.0033946 | 0.0054093 | 0.0087317 | 0.014946
| 2.3.2 | Front Vol | Va | 2 | mm3 | 0.0051522 | 0.0051495 | 0.0051472 | 0.0051458 | 0.0051457 | 0.0051495 | 0.005153 | 0.0051593 | 0.0051691 | 0.0051852 | 0.0052134 | 0.0052658 | 0.0053457 | 0.0054597 | 0.0055139 | 0.0053553
| 2.3.3 | Eqv. Vol | Veoa | 3 | mm3 | 0.0074961 | 0.0075602 | 0.0076038 | 0.0076309 | 0.0076323 | 0.00756 | 0.0074962 | 0.0073768 | 0.0071958 | 0.0068965 | 0.0063729 | 0.0054355 | 0.0039828 | 0.0019554 | 0.00088331 | 0.003291
| 2.3.4 | Diaphragm Res. | foa | 600 | Hz | 4.4136e-06 | 4.6213e-06 | 5.5426e-06 | 5.5646e-06 | 1.538e-06 | 5.7807e-05 | 0.00012159 | 0.00027774 | 0.00059937 | 0.0013238 | 0.0030481 | 0.0070184 | 0.013277 | 0.015135 | 0.0062709 | 0.017979
| 2.3.5 | Diaph. Damp. Fac. | Da | 0.1 | - | 9.1604e-07 | 4.8884e-07 | 1.0096e-05 | 5.7914e-05 | 0.00025805 | 0.0010667 | 0.0016701 | 0.0027274 | 0.0042163 | 0.0064296 | 0.0096475 | 0.013728 | 0.016587 | 0.01169 | 0.0040407 | 0.02296
| 2.3 | Microphones |  |  |  | 0.0090965 | 0.0091476 | 0.0091823 | 0.0092041 | 0.0092087 | 0.0092106 | 0.0092519 | 0.0094164 | 0.0098447 | 0.010873 | 0.013111 | 0.017309 | 0.022525 | 0.020703 | 0.01277 | 0.033366
| 2.4.1 | Series Cap. | C | see table | pF | 0.0013704 | 0.00091313 | 0.00091283 | 0.00091269 | 0.00091271 | 0.00091307 | 0.00091336 | 0.00091398 | 0.00091464 | 0.00091597 | 0.0009182 | 0.00092027 | 0.00092494 | 0.00092898 | 0.00093694 | 0.00095404
| 2.4.2 | Volt. Ratio, DVM | VRab, VRac, VRbc | see table | - | 0.00941 | 0.00935 | 0.00932 | 0.00932 | 0.00932 | 0.00932 | 0.00932 | 0.00932 | 0.00935 | 0.00935 | 0.00935 | 0.00935 | 0.00935 | 0.00935 | 0.00935 | 0.0105
| 2.4.3 | Volt. Ratio, Cr-talk | VRab, VRac, VRbc | < -66 dB | - | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025
| 2.4.4 | Volt. Ratio, Noise | VRab, VRac, VRbc | < -46 dB | - | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001
| 2.4.5 | Volt. Ratio, Distort. | VRab, VRac, VRbc | < -46 dB | - | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001
| 2.4.6 | Frequency | f | 25 | ppm | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002 | 0.0003 | 0.0004
| 2.4.7 | Pol. Voltage | Uo | 0.023 | V | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041 | 0.0014041
| 2.4 | Electrical Parameters |  |  |  | 0.0099337 | 0.0098238 | 0.0097953 | 0.0097953 | 0.0097953 | 0.0097953 | 0.0097953 | 0.0097954 | 0.009824 | 0.0098241 | 0.0098243 | 0.0098245 | 0.0098265 | 0.0098269 | 0.0098302 | 0.010934
| 2.5 | Reproduc./Stability | Mp | see table | dB | 0.00843 | 0.00472 | 0.00434 | 0.00426 | 0.00432 | 0.00436 | 0.0045 | 0.00453 | 0.00482 | 0.00519 | 0.00583 | 0.00675 | 0.00844 | 0.01058 | 0.01376 | 0.0177
| 2.6 | Result Rounding | Mp | < 0.005 | dB | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735 | 0.0057735
|  | Uncertainty at Measurement Conditions |  |  | dB | 0.021255 | 0.019306 | 0.019639 | 0.020069 | 0.019645 | 0.020162 | 0.019965 | 0.019697 | 0.020822 | 0.021498 | 0.0229 | 0.02582 | 0.031002 | 0.032858 | 0.036144 | 0.060862
| 2.7 | Sensitivity Correction for Static Pressure |  |  | dB | 0.0032 | 0.0032 | 0.0032 | 0.0032 | 0.0032 | 0.003 | 0.0028 | 0.0026 | 0.0023 | 0.0019 | 0.001 | 0.0009 | 0.034 | 0.0053 | 0.0013 | 0.0204
| 2.7 | Sensitivity Correction for Temperature |  |  | dB | 0.0012 | 0.0012 | 0.0011 | 0.0011 | 0.0013 | 0.0019 | 0.0023 | 0.0027 | 0.0034 | 0.0044 | 0.0065 | 0.0102 | 0.016 | 0.0213 | 0.0185 | 0.0072
|  | Uncertainty at Reference Conditions |  |  | dB | 0.021528 | 0.019606 | 0.019928 | 0.020353 | 0.019947 | 0.020472 | 0.020291 | 0.02005 | 0.021223 | 0.022026 | 0.023825 | 0.027776 | 0.048715 | 0.039515 | 0.040624 | 0.064592
|  | Reported Uncertainty at Ref Conditions |  |  | dB | 0.07 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.06 | 0.07 | 0.1

## CMS-ITRI Level
CMS-ITRI | Level
| Uncertainty Budget - A01 LS1P microphone calibration system_reciprocity method | [13 empty] | Unit : dB
| Frequency | 1.9953 | 2.5119 | 3.1623 | 3.9811 | 5.0119 | 6.3096 | 7.9433 | 10 | 12.5892 | 15.8489 | 19.9526 | 21.1349 | 22.3872 | 23.7137 | 25.1189 | 26.6073 | 28.1838 | 29.8538 | 31.6228 | 33.4965 | 35.4813 | 37.5837 | 39.8107 | 42.1697 | 44.6684 | 47.3151 | 50.1187 | 53.0884 | 56.2341 | 59.5662 | 63.0957 | 66.8344 | 70.7946 | 74.9894 | 79.4328 | 84.1395 | 89.1251 | 94.4061 | 100 | 105.9254 | 112.2018 | 118.8502 | 125.8925 | 133.3521 | 141.2538 | 149.6235 | 158.4893 | 167.8804 | 177.8279 | 188.3649 | 199.5262 | 211.3489 | 223.8721 | 237.1374 | 251.1887 | 266.0725 | 281.8383 | 298.5382 | 316.2278 | 334.9655 | 354.8134 | 375.8374 | 398.1071 | 421.6965 | 446.6837 | 473.1512 | 501.1873 | 530.8843 | 562.3412 | 595.6623 | 630.9572 | 668.344 | 707.9456 | 749.8942 | 794.3285 | 841.395 | 891.251 | 944.0607 | 1000 | 1059.254 | 1122.0179 | 1188.502 | 1258.9249 | 1333.521 | 1412.5381 | 1496.235 | 1584.8929 | 1678.8041 | 1778.2789 | 1883.6499 | 1995.2621 | 2113.489 | 2238.7209 | 2371.3738 | 2511.887 | 2660.7251 | 2818.3831 | 2985.3821 | 3162.2778 | 3349.655 | 3548.134 | 3758.374 | 3981.071 | 4216.9653 | 4466.8369 | 4731.5122 | 5011.8726 | 5308.8433 | 5623.4131 | 5956.623 | 6309.5732 | 6683.4404 | 7079.4561 | 7498.9419 | 7943.2842 | 8413.9502 | 8912.5107 | 9440.6064 | 10000 | 10592.54 | 11220.1904 | 11885.0205 | 12589.25
Voltage ratio (CorR,,n)
| umeas,repeat | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496
| umeas,specification | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.0092889 | 0.018395 | 0.018395 | 0.018395 | 0.018395 | 0.018395
Coupler volume(CorCV)
| Vcoup,deviation | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641
Static pressure (CorPs)
| Ps,deviation | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026
| Ps,specification | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004
Capacitance (CorC)
| Ctraceability | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011
| Cerror | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321
Microphone properties (Sref.)
| Ps,error | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0046188 | 0.0040415 | 0.0034641 | 0.0028868 | 0.0017321 | 0.0011547 | 0.00057735 | 0.00057735 | 0.00057735
| Terror | 0.0034641 | 0.0034641 | 0.0034641 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0046188 | 0.0046188 | 0.0051962 | 0.0051962 | 0.0057735 | 0.0057735 | 0.0063509 | 0.0069282 | 0.0075056 | 0.0080829 | 0.0092376 | 0.009815 | 0.01097 | 0.012124 | 0.013279 | 0.015011 | 0.016743 | 0.018475 | 0.020207 | 0.021939 | 0.023671 | 0.025981 | 0.02829 | 0.0306 | 0.020207 | 0.01963 | 0.018475 | 0.016743 | 0.014434 | 0.012124 | 0.010392 | 0.01097 | 0.015011 | 0.026558
| Herror | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0046188 | 0.0046188 | 0.0040415 | 0.0034641 | 0.0028868 | 0.0017321 | 0.0017321 | 0.0023094 | 0.0028868 | 0.0034641
| LF,error | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321
| Equivalent volume | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0028868 | 0.0028868 | 0.0023094 | 0.0023094 | 0.0017321 | 0.0017321 | 0.0011547 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0017321 | 0.0023094 | 0.0028868 | 0.0034641 | 0.0051962 | 0.0063509 | 0.0086603
| Leakage | 0.018475 | 0.014434 | 0.012702 | 0.0092376 | 0.0086603 | 0.0069282 | 0.0069282 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0057735 | 0.0057735 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0069282 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0069282 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0063509 | 0.0063509 | 0.0063509 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0063509 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0063509 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0069282 | 0.0075056 | 0.0075056 | 0.0069282 | 0.0069282 | 0.0075056 | 0.0069282 | 0.0069282 | 0.0075056 | 0.0080829 | 0.0080829 | 0.0080829 | 0.009815 | 0.010392 | 0.01097 | 0.013279 | 0.016743 | 0.020207 | 0.025403 | 0.032332 | 0.045033 | 0.045033 | 0.045033 | 0.045033 | 0.045033 | 0.045033 | 0.045033 | 0.045033 | 0.045033 | 0.045033 | 0.045033 | 0.045033 | 0.045033
| Resonance Frequency | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0 | 0 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.0011547 | 0.0017321 | 0.0028868 | 0.0051962 | 0.0092376 | 0.015011 | 0.022517 | 0.0069282 | 0.0057735 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0051962 | 0.0046188 | 0.0046188 | 0.0063509 | 0.0086603
| Loss factor | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0028868 | 0.0028868 | 0.0034641 | 0.0034641 | 0.0040415 | 0.0040415 | 0.0046188 | 0.0046188 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0051962 | 0.0046188 | 0.0034641 | 0.0023094 | 0.00057735 | 0.0023094 | 0.0046188 | 0.0057735 | 0.0051962 | 0.0034641 | 0.0034641 | 0.0040415 | 0.0063509 | 0.009815 | 0.013856 | 0.017898 | 0.021939 | 0.025981 | 0.031177
| Rounding error | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029 | 0.0029
| MP,repeat | 0.004 | 0.003 | 0.003 | 0.003 | 0.0025 | 0.002 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.001 | 0.001 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0.002 | 0.002 | 0.0015 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.003 | 0.003 | 0.003 | 0.0035 | 0.004 | 0.004 | 0.0045 | 0.005 | 0.0055 | 0.006 | 0.007 | 0.007 | 0.008 | 0.009 | 0.009 | 0.004 | 0.0035 | 0.004 | 0.0035 | 0.003 | 0.0045 | 0.0045 | 0.006 | 0.0065 | 0.004 | 0.004 | 0.006
| u= | 0.024 | 0.021 | 0.02 | 0.018 | 0.017 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.016 | 0.017 | 0.017 | 0.017 | 0.017 | 0.017 | 0.017 | 0.017 | 0.017 | 0.017 | 0.017 | 0.018 | 0.018 | 0.018 | 0.019 | 0.02 | 0.02 | 0.021 | 0.023 | 0.024 | 0.025 | 0.027 | 0.03 | 0.033 | 0.038 | 0.044 | 0.056 | 0.058 | 0.061 | 0.053 | 0.052 | 0.051 | 0.051 | 0.051 | 0.054 | 0.055 | 0.056 | 0.059 | 0.066
| v= | 126.3037 | 173.497 | 206.8624 | 264.3989 | 236.3246 | 233.0327 | 235.3207 | 247.3504 | 247.3504 | 247.3504 | 247.3504 | 247.3504 | 257.3686 | 257.3686 | 247.3504 | 247.3504 | 242.3401 | 242.3401 | 243.0772 | 243.0772 | 243.0772 | 243.0772 | 243.0772 | 243.0772 | 243.0772 | 243.0772 | 243.0772 | 243.0772 | 243.0772 | 243.0772 | 243.0772 | 243.0772 | 243.9963 | 244.0004 | 243.0792 | 243.0772 | 243.0772 | 243.0772 | 243.0772 | 244.429 | 244.429 | 244.429 | 244.429 | 244.429 | 244.429 | 244.429 | 244.429 | 244.429 | 245.796 | 245.796 | 245.796 | 245.796 | 245.796 | 245.796 | 233.9134 | 245.796 | 245.796 | 245.796 | 245.796 | 233.9134 | 245.796 | 245.796 | 245.796 | 233.9134 | 233.9134 | 233.9134 | 245.796 | 245.796 | 244.429 | 231.9997 | 231.9997 | 231.9997 | 230.7815 | 242.3401 | 230.7815 | 230.7815 | 230.7815 | 230.7815 | 230.7815 | 230.7544 | 234.0396 | 234.0396 | 245.9354 | 231.7764 | 231.7764 | 295.6742 | 292.6912 | 292.6912 | 292.6912 | 288.49 | 288.49 | 262.5265 | 261.8516 | 270.1053 | 272.1077 | 313.6953 | 305.3948 | 286.7894 | 320.4836 | 316.1612 | 276.8167 | 285.7877 | 288.6838 | 263.3875 | 219.1911 | 192.5519 | 170.5513 | 157.817 | 146.5085 | 124.7978 | 105.2862 | 117.3763 | 131.5804 | 91.8117 | 85.4949 | 79.7349 | 80.3541 | 80.8855 | 98.7336 | 104.7516 | 109.7319 | 127.7261 | 166.4321
| k= | 1.979 | 1.9738 | 1.9715 | 1.969 | 1.9701 | 1.9702 | 1.9701 | 1.9696 | 1.9696 | 1.9696 | 1.9696 | 1.9696 | 1.9692 | 1.9692 | 1.9696 | 1.9696 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9697 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9698 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9702 | 1.9697 | 1.9697 | 1.9697 | 1.9697 | 1.9702 | 1.9697 | 1.9697 | 1.9697 | 1.9702 | 1.9702 | 1.9702 | 1.9697 | 1.9697 | 1.9697 | 1.9703 | 1.9703 | 1.9703 | 1.9703 | 1.9698 | 1.9703 | 1.9703 | 1.9703 | 1.9703 | 1.9703 | 1.9703 | 1.9702 | 1.9702 | 1.9697 | 1.9703 | 1.9703 | 1.968 | 1.9681 | 1.9681 | 1.9681 | 1.9682 | 1.9682 | 1.9691 | 1.9691 | 1.9688 | 1.9687 | 1.9676 | 1.9678 | 1.9683 | 1.9674 | 1.9675 | 1.9686 | 1.9683 | 1.9682 | 1.969 | 1.9709 | 1.9724 | 1.974 | 1.9752 | 1.9763 | 1.9793 | 1.9828 | 1.9804 | 1.9782 | 1.9864 | 1.9883 | 1.9905 | 1.9901 | 1.9901 | 1.9845 | 1.983 | 1.982 | 1.9788 | 1.9744
| U= | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.06 | 0.06 | 0.07 | 0.08 | 0.09 | 0.12 | 0.12 | 0.13 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.12 | 0.12 | 0.14
| Ureport= | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.06 | 0.06 | 0.07 | 0.08 | 0.09 | 0.12 | 0.12 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.14

## CMS-ITRI Phase
CMS-ITRI | Phase
| Uncertainty Budget - A01 LS1P microphone calibration system_reciprocity method | [11 empty] | Unit : degree
| Frequency | 1.9953 | 2.5119 | 3.1623 | 3.9811 | 5.0119 | 6.3096 | 7.9433 | 10 | 12.5892 | 15.8489 | 19.9526 | 21.1349 | 22.3872 | 23.7137 | 25.1189 | 26.6073 | 28.1838 | 29.8538 | 31.6228 | 33.4965 | 35.4813 | 37.5837 | 39.8107 | 42.1697 | 44.6684 | 47.3151 | 50.1187 | 53.0884 | 56.2341 | 59.5662 | 63.0957 | 66.8344 | 70.7946 | 74.9894 | 79.4328 | 84.1395 | 89.1251 | 94.4061 | 100 | 105.9254 | 112.2018 | 118.8502 | 125.8925 | 133.3521 | 141.2538 | 149.6235 | 158.4893 | 167.8804 | 177.8279 | 188.3649 | 199.5262 | 211.3489 | 223.8721 | 237.1374 | 251.1887 | 266.0725 | 281.8383 | 298.5382 | 316.2278 | 334.9655 | 354.8134 | 375.8374 | 398.1071 | 421.6965 | 446.6837 | 473.1512 | 501.1873 | 530.8843 | 562.3412 | 595.6623 | 630.9572 | 668.344 | 707.9456 | 749.8942 | 794.3285 | 841.395 | 891.251 | 944.0607 | 1000 | 1059.254 | 1122.0179 | 1188.502 | 1258.9249 | 1333.521 | 1412.5381 | 1496.235 | 1584.8929 | 1678.8041 | 1778.2789 | 1883.6499 | 1995.2621 | 2113.489 | 2238.7209 | 2371.3738 | 2511.887 | 2660.7251 | 2818.3831 | 2985.3821 | 3162.2778 | 3349.655 | 3548.134 | 3758.374 | 3981.071 | 4216.9653 | 4466.8369 | 4731.5122 | 5011.8726 | 5308.8433 | 5623.4131 | 5956.623 | 6309.5732 | 6683.4404 | 7079.4561 | 7498.9419 | 7943.2842 | 8413.9502 | 8912.5107 | 9440.6064 | 10000 | 10592.54 | 11220.1904 | 11885.0205 | 12589.25
Voltage ratio (CorR,,n)
| umeas,repeat | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496 | 0.0048496
| umeas,specification | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.1926 | 0.3816 | 0.3816 | 0.3816 | 0.3816 | 0.3816
Coupler volume(CorCV)
| Vcoup,deviation | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0017321 | 0.0017321 | 0.0023094 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0023094 | 0.0028868 | 0.0040415
Static pressure (CorPs)
| Ps,deviation | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026 | 0.0026
| Ps,specification | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 0.0004
Capacitance (CorC)
| Ctraceability | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011 | 0.0011
| Cerror | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0 | 0.00057735 | 0 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0 | 0 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735
Microphone properties (Sref.)
| Ps,error | 0.0023094 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0017321 | 0.0011547 | 0.00057735 | 0.0011547 | 0.0023094 | 0.0040415 | 0.0063509 | 0.0086603 | 0.011547 | 0.015011 | 0.018475 | 0.018475 | 0.021939 | 0.024826 | 0.026558 | 0.027713 | 0.027713 | 0.025981 | 0.022517 | 0.017898 | 0.013279
| Terror | 0.0080829 | 0.0075056 | 0.0075056 | 0.0069282 | 0.0063509 | 0.0063509 | 0.0057735 | 0.0051962 | 0.0051962 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0011547 | 0.0011547 | 0.0011547 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.0011547 | 0.0011547 | 0.0017321 | 0.0017321 | 0.0023094 | 0.0023094 | 0.0028868 | 0.0028868 | 0.0034641 | 0.0034641 | 0.0040415 | 0.0046188 | 0.0046188 | 0.0051962 | 0.0057735 | 0.0063509 | 0.0069282 | 0.0069282 | 0.0075056 | 0.0080829 | 0.0086603 | 0.0092376 | 0.009815 | 0.010392 | 0.01097 | 0.012124 | 0.012702 | 0.013279 | 0.013856 | 0.015011 | 0.015588 | 0.016743 | 0.017321 | 0.018475 | 0.019053 | 0.01963 | 0.020207 | 0.020785 | 0.021362 | 0.020785 | 0.020207 | 0.019053 | 0.017321 | 0.014434 | 0.009815 | 0.0040415 | 0.0040415 | 0.014434 | 0.026558 | 0.042147 | 0.060622 | 0.083716 | 0.081984 | 0.098727 | 0.11489 | 0.12759 | 0.13452 | 0.13337 | 0.1224 | 0.099882 | 0.072169 | 0.057735
| Herror | 0.0023094 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0017321 | 0.0011547 | 0.00057735 | 0.0011547 | 0.0023094 | 0.0040415 | 0.0063509 | 0.0086603 | 0.011547 | 0.015011 | 0.018475 | 0.018475 | 0.021939 | 0.024826 | 0.026558 | 0.027713 | 0.027713 | 0.025981 | 0.022517 | 0.017898 | 0.013279
| LF,error | 0.0028868 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0017321 | 0.0017321 | 0.0023094 | 0.0023094 | 0.0028868 | 0.0034641 | 0.0040415 | 0.0051962 | 0.0057735 | 0.0063509 | 0.0069282 | 0.0075056 | 0.0080829 | 0.0080829 | 0.0080829 | 0.0080829 | 0.0080829 | 0.0075056 | 0.0069282 | 0.0069282
| Equivalent volume | 0.0034641 | 0.0028868 | 0.0028868 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0046188 | 0.0046188 | 0.0051962 | 0.0051962 | 0.0057735 | 0.0057735 | 0.0063509 | 0.0063509 | 0.0069282 | 0.0075056 | 0.0075056 | 0.0080829 | 0.0080829 | 0.0086603 | 0.0092376 | 0.0092376 | 0.009815 | 0.009815 | 0.009815 | 0.010392 | 0.009815 | 0.009815 | 0.0092376 | 0.0086603 | 0.0069282 | 0.0051962 | 0.0023094 | 0.0023094 | 0.0080829 | 0.015011 | 0.024826 | 0.03695 | 0.0011547 | 0.0040415 | 0.009815 | 0.015011 | 0.020785 | 0.025981 | 0.031177 | 0.036373 | 0.042147 | 0.049075
| Leakage | 0.59756 | 0.48671 | 0.39375 | 0.3187 | 0.24942 | 0.20669 | 0.16628 | 0.13452 | 0.10796 | 0.086025 | 0.069859 | 0.066395 | 0.063509 | 0.059467 | 0.057735 | 0.053694 | 0.052539 | 0.049652 | 0.047343 | 0.044456 | 0.042147 | 0.039837 | 0.03926 | 0.03695 | 0.034641 | 0.033486 | 0.031754 | 0.0306 | 0.028868 | 0.027713 | 0.025981 | 0.025981 | 0.024249 | 0.023671 | 0.022517 | 0.022517 | 0.020785 | 0.019053 | 0.018475 | 0.017898 | 0.017321 | 0.015588 | 0.014434 | 0.015011 | 0.015011 | 0.015011 | 0.012702 | 0.011547 | 0.011547 | 0.012124 | 0.010392 | 0.01097 | 0.009815 | 0.0092376 | 0.0080829 | 0.0086603 | 0.0075056 | 0.0080829 | 0.0075056 | 0.0057735 | 0.0069282 | 0.0057735 | 0.0063509 | 0.0051962 | 0.0057735 | 0.0057735 | 0.0069282 | 0.0063509 | 0.0051962 | 0.0063509 | 0.0069282 | 0.0057735 | 0.0069282 | 0.009815 | 0.0069282 | 0.009815 | 0.0092376 | 0.011547 | 0.0092376 | 0.011547 | 0.011547 | 0.013856 | 0.015588 | 0.012124 | 0.017321 | 0.018475 | 0.017321 | 0.020785 | 0.024249 | 0.023094 | 0.027713 | 0.028868 | 0.034064 | 0.035218 | 0.03926 | 0.040415 | 0.046765 | 0.053116 | 0.057158 | 0.061199 | 0.068705 | 0.083716 | 0.091799 | 0.10508 | 0.11432 | 0.13625 | 0.15473 | 0.16281 | 0.19919 | 0.23902 | 0.2829 | 0.41685 | 0.43763 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Resonance Frequency | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0040415 | 0.0040415 | 0.0040415 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0051962 | 0.0051962 | 0.0057735 | 0.0057735 | 0.0063509 | 0.0069282 | 0.0069282 | 0.0075056 | 0.0080829 | 0.0086603 | 0.0086603 | 0.0092376 | 0.009815 | 0.010392 | 0.01097 | 0.012124 | 0.012702 | 0.013279 | 0.014434 | 0.015011 | 0.016166 | 0.017321 | 0.018475 | 0.01963 | 0.020785 | 0.022517 | 0.023671 | 0.025403 | 0.027135 | 0.02829 | 0.0306 | 0.032909 | 0.034641 | 0.03695 | 0.03926 | 0.040992 | 0.042724 | 0.043301 | 0.044456 | 0.043879 | 0.042724 | 0.041569 | 0.040415 | 0.041569 | 0.049075 | 0.069282 | 0.11316 | 0.062354 | 0.069859 | 0.084293 | 0.10681 | 0.1351 | 0.16397 | 0.19572 | 0.22979 | 0.27597 | 0.40703
| Loss factor | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.00057735 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0011547 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0017321 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0023094 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0028868 | 0.0034641 | 0.0034641 | 0.0034641 | 0.0040415 | 0.0040415 | 0.0046188 | 0.0046188 | 0.0046188 | 0.0051962 | 0.0051962 | 0.0057735 | 0.0057735 | 0.0063509 | 0.0069282 | 0.0069282 | 0.0075056 | 0.0080829 | 0.0080829 | 0.0086603 | 0.0092376 | 0.009815 | 0.010392 | 0.01097 | 0.011547 | 0.012124 | 0.012702 | 0.013279 | 0.014434 | 0.015011 | 0.015588 | 0.016743 | 0.017321 | 0.018475 | 0.019053 | 0.020207 | 0.020785 | 0.021939 | 0.022517 | 0.023671 | 0.024249 | 0.024826 | 0.025403 | 0.025981 | 0.026558 | 0.026558 | 0.026558 | 0.025981 | 0.024826 | 0.023671 | 0.021362 | 0.019053 | 0.016166 | 0.013279 | 0.009815 | 0.0075056 | 0.0075056 | 0.011547 | 0.022517 | 0.043879 | 0.079097 | 0.13048 | 0.0017321 | 0.016743 | 0.034064 | 0.049075 | 0.057735 | 0.059467 | 0.054271 | 0.045611 | 0.031177 | 0.010392
| Rounding error | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029
| MP,repeat | 0.1405 | 0.114 | 0.094 | 0.078 | 0.06 | 0.0495 | 0.0395 | 0.0325 | 0.026 | 0.02 | 0.0165 | 0.0155 | 0.015 | 0.014 | 0.014 | 0.0125 | 0.0125 | 0.0115 | 0.0105 | 0.01 | 0.0095 | 0.009 | 0.009 | 0.008 | 0.0075 | 0.0075 | 0.007 | 0.0065 | 0.006 | 0.0055 | 0.006 | 0.005 | 0.005 | 0.005 | 0.005 | 0.0045 | 0.004 | 0.0045 | 0.004 | 0.0035 | 0.003 | 0.0035 | 0.003 | 0.003 | 0.003 | 0.0025 | 0.0025 | 0.0025 | 0.003 | 0.0025 | 0.0025 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.0025 | 0.002 | 0.0015 | 0.0025 | 0.0025 | 0.002 | 0.002 | 0.0025 | 0.0025 | 0.0025 | 0.0025 | 0.002 | 0.0025 | 0.0025 | 0.0035 | 0.003 | 0.0025 | 0.0025 | 0.0035 | 0.0035 | 0.003 | 0.003 | 0.004 | 0.004 | 0.0045 | 0.0045 | 0.004 | 0.0055 | 0.0045 | 0.0045 | 0.005 | 0.0055 | 0.006 | 0.0055 | 0.006 | 0.0065 | 0.0065 | 0.0055 | 0.006 | 0.006 | 0.007 | 0.006 | 0.0075 | 0.008 | 0.01 | 0.0095 | 0.014 | 0.016 | 0.014 | 0.0165 | 0.013 | 0.015 | 0.014 | 0.023 | 0.022 | 0.0255 | 0.089 | 0.0365 | 0.0215 | 0.015
| u= | 0.65 | 0.54 | 0.45 | 0.39 | 0.33 | 0.29 | 0.26 | 0.24 | 0.23 | 0.22 | 0.21 | 0.21 | 0.21 | 0.21 | 0.21 | 0.21 | 0.21 | 0.21 | 0.21 | 0.21 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.21 | 0.21 | 0.21 | 0.21 | 0.21 | 0.21 | 0.21 | 0.21 | 0.21 | 0.22 | 0.22 | 0.22 | 0.23 | 0.23 | 0.24 | 0.25 | 0.26 | 0.26 | 0.29 | 0.32 | 0.36 | 0.48 | 0.52 | 0.23 | 0.24 | 0.25 | 0.27 | 0.29 | 0.45 | 0.47 | 0.47 | 0.49 | 0.57
| v= | 66.7314 | 71.3332 | 77.6929 | 95.1646 | 109.6443 | 107.937 | 105.2434 | 96.5799 | 92.1873 | 81.7381 | 69.4106 | 69.6387 | 69.8014 | 69.9981 | 70.0685 | 70.219 | 70.2541 | 70.3368 | 70.3939 | 70.4523 | 57.9938 | 58.0214 | 58.0274 | 58.0502 | 58.0686 | 58.0763 | 58.0868 | 58.0931 | 58.101 | 58.1056 | 58.1111 | 58.1115 | 58.1161 | 58.1174 | 58.1198 | 58.12 | 58.123 | 58.1252 | 58.126 | 58.1266 | 58.1272 | 58.1285 | 58.1292 | 58.1288 | 58.1288 | 58.1289 | 58.1299 | 58.1303 | 58.1302 | 58.1301 | 58.1305 | 58.1304 | 58.1306 | 58.1307 | 58.1308 | 58.1308 | 58.1309 | 58.1308 | 58.1309 | 58.131 | 58.1309 | 58.1309 | 58.1309 | 58.1309 | 58.1309 | 58.1309 | 58.1308 | 58.1308 | 58.1308 | 58.1307 | 58.1306 | 58.1306 | 58.1304 | 58.13 | 58.1301 | 58.1296 | 58.1295 | 58.1287 | 58.1289 | 58.1278 | 58.1273 | 58.1257 | 58.1237 | 58.1245 | 58.1199 | 58.1176 | 58.1162 | 58.1101 | 58.1001 | 58.1001 | 70.5993 | 70.5886 | 70.5495 | 70.5319 | 70.479 | 70.4558 | 70.3354 | 70.1579 | 70.0081 | 84.1006 | 83.5803 | 81.9913 | 96.4413 | 93.1671 | 106.9572 | 113.2483 | 117.0068 | 109.745 | 119.7506 | 112.8836 | 107.7353 | 83.8722 | 94.8019 | 97.364 | 110.8728 | 121.8012 | 149.1347 | 172.3206 | 92.0927 | 102.9832 | 101.1491 | 106.5992 | 108.4429
| k= | 1.9966 | 1.9939 | 1.9913 | 1.9853 | 1.982 | 1.9824 | 1.9828 | 1.985 | 1.9861 | 1.9897 | 1.9949 | 1.9949 | 1.9949 | 1.9949 | 1.9944 | 1.9944 | 1.9944 | 1.9944 | 1.9944 | 1.9944 | 2.0025 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 2.0017 | 1.9944 | 1.9944 | 1.9944 | 1.9944 | 1.9944 | 1.9944 | 1.9944 | 1.9944 | 1.9944 | 1.9886 | 1.989 | 1.9897 | 1.985 | 1.9858 | 1.9826 | 1.9812 | 1.9804 | 1.982 | 1.9801 | 1.9814 | 1.9824 | 1.989 | 1.9855 | 1.9847 | 1.9818 | 1.9798 | 1.976 | 1.9739 | 1.9861 | 1.9835 | 1.9837 | 1.9826 | 1.9822
| U= | 1.3 | 1.1 | 0.9 | 0.8 | 0.7 | 0.6 | 0.6 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.40049 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.40034 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.6 | 0.6 | 0.6 | 0.7 | 0.8 | 1 | 1.1 | 0.5 | 0.5 | 0.5 | 0.6 | 0.6 | 0.9 | 1 | 1 | 1 | 1.2
| Ureport= | 1.3 | 1.1 | 0.9 | 0.8 | 0.7 | 0.6 | 0.6 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.6 | 0.6 | 0.6 | 0.7 | 0.8 | 1 | 1.1 | 1.1 | 1.1 | 1.1 | 1.1 | 1.1 | 1.1 | 1.1 | 1.1 | 1.1 | 1.2