MEASUREMENT REPORT

## Transportation

Conditions of the transfer standards on arrival:

| Date of arrival: |  |
| --- | --- |
| Method of transportation: |  |
| Transportation box: |  |
| Other remarks: |  |

Initial inspection of transfer standards:

| Weight | 200 g | 20 g | 2 g | 200 mg |
| --- | --- | --- | --- | --- |
| Top |  |  |  |  |
| Side envelope |  |  |  | / |
| Bottom |  |  |  |  |
| Comments |  |  |  |  |

Conditions of the transfer standards on departure:

| Date of departure: |  |
| --- | --- |
| Method of transportation: |  |
| Transportation box: |  |
| Other remarks: |  |

Final inspection of transfer standards:

| Weight | 200 g | 20 g | 2 g | 200 mg |
| --- | --- | --- | --- | --- |
| Top |  |  |  |  |
| Side envelope |  |  |  | / |
| Bottom |  |  |  |  |
| Comments |  |  |  |  |

## Calibration results

| Nominal value | Mass (true) m | Standard uncertainty um (k=1) | Conventional mass mc | Standard uncertainty umc (k=1) | No. of measurement cycles *) |
| --- | --- | --- | --- | --- | --- |
| 200 g |  |  |  |  |  |
| 20 g |  |  |  |  |  |
| 2 g |  |  |  |  |  |
| 200 mg |  |  |  |  |  |

*) If the calibration was made by other procedures than one-to-one comparison (e.g. subdivision using a system of weighing equations), give the degrees of freedom of this procedure.

## Data provided by MIRS

| Weight | Volume V (cm3) | Uncertainty uV (cm3) (k = 1) | Centre of gravity z (mm) | Uncertainty uz (mm) (k = 1) | Magnetic susceptibility  |
| --- | --- | --- | --- | --- | --- |
| 200 g | 25,102 | 0,006 | 21 | 2,0 |  < 0,02 |
| 20 g | 2,5091 | 0,0016 | 9,7 | 2,0 |  < 0,02 |
| 2 g | 0,2514 | 0,0009 | 4,6 | 1,0 |  < 0,06 |
| Weight | Density ρ (kgm-3) | Relative uncertainty ρ (k = 1) |  |  |  |
| 200 mg | 8600 | 2 % |  |  |  |

The assumed volume temperature expansion coefficient for the 2 g, 20 g and 200 g weight is α = 48 ± 5   10-6 K-1

## Traceability of reference standards

Insert the standard(s) including additional weights used for the calibration and its traceability to the international prototype of the kilogram

200 g

| Identification | Mass m | Standard uncertainty um (k=1) | Volume V | Standard uncertainty uV (k=1) | Calibration date |
| --- | --- | --- | --- | --- | --- |

20 g

| Identification | Mass m | Standard uncertainty um (k=1) | Volume V | Standard uncertainty uV (k=1) | Calibration date |
| --- | --- | --- | --- | --- | --- |

2 g

| Identification | Mass m | Standard uncertainty um (k=1) | Volume V | Standard uncertainty uV (k=1) | Calibration date |
| --- | --- | --- | --- | --- | --- |

200 mg

| Identification | Mass m | Standard uncertainty um (k=1) | Volume V | Standard uncertainty uV (k=1) | Calibration date |
| --- | --- | --- | --- | --- | --- |

## Description of the measurement procedure

## Uncertainty budget

The uncertainties shall be estimated and combined following the Guide to the Expression of Uncertainty in Measurement JCGM 100:2008.

|  | (Standard) uncertainty contributions (mg) |  |  |  |
| --- | --- | --- | --- | --- |
| Uncertainty component | 200 g | 20 g | 2 g | 200 mg |
| Combined uncertainty uc (k=1) |  |  |  |  |

## Details of the balance used for the calibration

| Weight | Manufacturer | Type | Maximum load | Resolution | Standard deviation* | Manual/ automatic |
| --- | --- | --- | --- | --- | --- | --- |
| 200 g |  |  |  |  |  |  |
| 20 g |  |  |  |  |  |  |
| 2 g |  |  |  |  |  |  |
| 200 mg |  |  |  |  |  |  |

*Indicate the weighing procedure and number of measurements

## Details of the instruments used for air density determination

|  | Manufacturer | Type | Range | Resolution | Standard uncertainty (k=1) |
| --- | --- | --- | --- | --- | --- |
| Temperature |  |  |  |  |  |
| Pressure |  |  |  |  |  |
| Humidity |  |  |  |  |  |
| CO2 |  |  |  |  |  |

## Description of air buoyancy determination

## Environmental data during calibration

Insert the maximum and minimum values of the measured quantities and their standard uncertainties.

| Temperature t (oC) | Pressure p (mbar) | Relative humidity h (%) | x(CO2)  106 | Air density a (kg/m3) |
| --- | --- | --- | --- | --- |

## Description of the measuring room

| Mean temperature |  |
| --- | --- |
| Minimum, maximum temperature |  |
| Maximum change of temperature during 8 hours |  |
| Mean humidity |  |
| Minimum, maximum humidity |  |

| Laboratory: | Responsible person: |
| --- | --- |
| Date: | Signature: |