## Slide 1
Update on CCQM-K19.2018.1
pH measurement of Borate buffer
Narine Oganyan and Sergey Prokunin
Russian Metrological Institute of Technical Physics
and Radio Engineering (VNIIFTRI)
Toshiaki Asakai and Igor Maksimov
National Metrology Institute of Japan (NMIJ) / AIST
11 October 2022, CCQM-EAWG virtual meeting

## Slide 2: Scope
Sample:
Borate buffer prepared around pH 9.2
Method:
Primary Harned cell method
Secondary methods, differential potentiometric cell
or glass electrode with highest metrological level
Temperature:
15 oC, 25 oC and 37 oC, optionally 5 oC and 50 oC
How far the light shines:
Alkaline range of pH (25 oC) 8.5 to 9.5 for Harned cell

## Slide 3: History and Participants
CCQM-K19 (2006):
11 institutes
CCQM-K19.1 (2011):
8 institutes
CCQM-K19.2018:
17 institutes from 24 institutes
(BFKH, BIM, CENAM, CMI, DFM, GUM, IBMETRO, INACAL, INM, INMETRO, INTI, IPQ, LATU, LNE, NIM, NIMT, NIST, NMCC/SASO, NMIJ, PTB, SMU, TÜBİTAK UME, UkrCSM and VNIIFTRI)
CCQM-K19.2018.1:
16 institutes registered
(BSN, CENAM, CMI, DFM, IBMETRO, INACAL, INM-CO, IPQ, LATU, NIMT, NIST, NMIJ, PTB, UkCSM, TÜBİTAK UME and VNIIFTRI)

## Slide 4: Schedule
Proposal:
April 2021 in EAWG
Invitation:
June 2021
Registration deadline:
31 July 2021
Dispatch of samples:
17 - 18 January 2022
Reporting deadline:
15 July 2022 -> 31 August 2022

## Slide 5: Sample preparation
The borate buffer solution with pH around ~9.2 at 25 oC was prepared using the ultrapure water with the electrical resistivity 18 MΩ cm, a reagent of disodium tetraborate decahydrate (Na2B4O7·10H2O, 99.6 % to 100.5 %, CAS 1303-96-4) as the source material. Disodium tetraborate decahydrate had been stored at a desiccator with a saturated potassium bromide over 6 months prior to the preparation.
An approximately 50 L sample solution was prepared on 11 November 2021, by dissolving the above-mentioned reagent in the Nalgene polypropylene bottle. The composition of solution was close to that of standard borate buffer with the slightly altered molality of the solute from 0.05 mol kg−1 nominal values.
The solution was later subdivided into 1 L high density polyethylene bottles (47 bottles) on 18 November 2021. The bottles were numbered, weighted and sealed in the aluminized plastic bags.
The mass fraction of water in the solution was w(H2O) = 0.996201.

## Slide 6: Solution homogeneity - I
Homogeneity of the bottled buffer solutions was checked at 25 oC by measuring the values of pH for 6 bottles with a precise pH meter PHL-90 with a combined glass electrode, and 2 bottles by primary Harned cell method. The precise pH meter was a custom-built one (DKK-TOA Corporation, Tokyo, Japan) with the minimum resolution 0.0001.

## Slide 7: Solution homogeneity - II
Bottle No. | ΔpH | u(ΔpH) | Method
1 | 0.0003 | 0.0014 | Harned
2 | ‒0.0004 | 0.0014 | Glass
10 | 0.0007 | 0.0007 | Glass
19 | 0.0002 | 0.0007 | Glass
28 | ‒0.0003 | 0.0007 | Glass
37 | 0.0004 | 0.0001 | Glass
46 | ‒0.0006 | 0.0017 | Glass
49 | ‒0.0003 | 0.0016 | Harned

## Slide 8: Solution homogeneity - III
Results of the homogeneity assessment.
●: primary Harned cell method; ▲: glass electrode method
Each bar corresponds to the combined standard uncertainty.

## Slide 9: Solution stability - I
The stability of the sample solutions was assessed at 25 oC by Harned cell method over the duration of the key comparison from 26 November 2021 to 14 July 2022. Each point in the stability assessment represents the independent measurement result from a single bottle of the borate buffer by pa0 determination from 4 solutions with chloride molalities from 0.005 mol kg−1 to 0.020 mol kg−1.
All the points lay within the ±0.001 range from the average pa0; therefore, the sample solutions were stable enough over the period of the key comparison.

## Slide 10: Solution stability - II
Date | Bottle No. | Δpa0 | U(pa0)
26 Nov 2021 | 1 | 0.00070 | 0.00278
30 Nov 2021 | 49 | 0.00012 | 0.00315
20 May 2022 | 12 | ‒0.00082 | 0.00375
14 July 2022 | 7 | ‒0.00096 | 0.00253

## Slide 11: Solution stability - III
Results of the stability assessment.
Each bar corresponds to the expanded uncertainty.

## Slide 12: Shipping integrity
The samples were shipped by NMIJ on 17 or 18 January 2022 via EMS (Express Mail Service) or FedEx. All registered participants have received the samples by 30 March 2022.
No visible signs of any damage to the samples were reported to the coordinating laboratories. To ensure the integrity of the received samples, each participant reported: the ambient pressure, temperature, relative humidity, balance reading, and the calculated mass of each bottle after air buoyancy correction.
The reported maximum relative change in bottle mass was ‒0.007 %, the average change was ‒0.001 % (±0.003 % (SD)), and acceptable for the comparison on pH.

## Slide 13: Coordinating laboratories
Reporting:
Russian Metrological Institute of Technical Physics and Radio Engineering
141570, Mendeleevo, Moscow Region, Russia
Contact person: Narine Oganyan
Email: oganyan@vniiftri.ru; Phone: +7 495 526 63 21
Alternative contact person: Sergey Prokunin
Email: prokunin@vniiftri.ru; Phone: +7 495 526 63 91
Sample preparation, homogeneity and stability test and shipping:
National Metrology Institute of Japan / AIST
1-1-1 Umezono, Tsukuba, Ibaraki 305-8563, Japan
Contact person: Toshiaki Asakai
Email: t-asakai@aist.go.jp; Phone: +81 29 861 4162
Alternative contact person: Igor Maksimov
Email: maksimov.igor@aist.go.jp; Phone: +81 29 861 6881