The Technical Protocol for EURAMET Project No. 1590

Inter-laboratory Calibration Comparison

of the Rotary Piston Gas Meter G650

1. Introduction

The aim of the comparison is to compare the performance of calibrations of a gas meter in different gas flow laboratories in Europe. A rotary piston gas meter IRM-3-DUO G650 will be used to compare low pressure test facilities over the flow range (50 ÷ 1000) m3/h.  If laboratories are not able to cover all the flow rates, they may calibrate the meter over a part of the flow range.

2. Participants and time schedule

Each country will take 3 weeks to perform the calibration of the meter and to transfer the meter to a next laboratory. The participants and the time schedule are mentioned in     table 1. In the table 1 there the “address of the place of calibration” is the address where the meter should be sent. Please, mention also the name of responsible person of next laboratory  on the label on the wooden box.

Table 1 - Participants and the time schedule

| Country | Laboratory | Address of the place of calibration | e-mail telephone | Date of calibration | Responsible person |
| --- | --- | --- | --- | --- | --- |
| Czech Republic (PILOT LAB) | CMI Czech Metrology Institute | Czech Metrology Institute Prumyslova 455 53003 Pardubice Czech Republic | tvalenta@cmi.cz +420 466 670 728 | 11.12.2023 -22.12.2023 | Tomas Valenta |
| Germany | PTB Physikalisch-Technische Bundesanstalt | PTB Bundesallee 100 38116 Braunschweig Germany | Bodo.Mickan@ptb.de +49 531 592 1331 | 2.1.2024 - 19.1.2024 | Bodo Mickan |
| Denmark | FORCE Technology | FORCE Technology Navervej 1 6600 Vejen Denmark | JRB@forcetechnology.com +45 22697620 | 22.1.2024 - 9.2.2024 | Jesper Busk |
| Netherlands | VSL (Van Swinden Laboratorium) | VSL Thijsseweg 11 2629 JA Delft The Netherlands | twendel@vsl.nl +31631119906 | 12.2.2024 – 1.3.2024 | Thomas Wendel |
| Austria | BEV Bundesamt für Eich- und Vermessungswesen | BEV Arltgasse 35 A-1160 Wien Austria | claudia.berkmann@bev.gv.at +43 1 211 10-826539 | 4.3.2024 - 22.3.2024 | Claudia Berkmann |
| Poland | GUM Główny Urząd Miar (Central Office of Measures) | Central Office of Measures 00-950 Warszawa P-10 ul. Elektoralna 2 Poland | adam.urbanowicz@gum.gov.pl +48 22 581 9319 | 25.3.2024 - 12.4.2024 | Adam Urbanowicz |
| Lithuania | Lithuanian Energy Institute (Heat equipment research and testing laboratory) | Lithuanian Energy Institute Breslaujos st. 3 LT-44403 Kaunas Lithuania | arunas.stankevicius@lei.lt +370 68620850 | 15.4.2024 - 3.5.2024 | Arūnas Stankevičius |
| Country | Laboratory | Address of the place of calibration | e-mail telephone | Date of calibration | Responsible person |
| Czech Republic (PILOT LAB) | CMI Czech Metrology Institute | Czech Metrology Institute Prumyslova 455 53003 Pardubice Czech Republic | tvalenta@cmi.cz +420 466 670 728 | 6.5.2024 - 24.5.2024 | Tomas Valenta |
| Switzerland (ATA-CARNET) | METAS Metrology and Accreditation Switzerland | Federal Institute of Metrology METAS Laboratory for Flow and Hydrometry Lindenweg 50, 3084 Wabern-Bern Switzerland | marc.dehuu@metas.ch +41 58 387 02 67 | 27.5.2024 - 14.6.2024 | Marc de Huu |
| Bosnia and Herzegovina (ATA-CARNET) | Institut za mjeriteljstvo Bosne i Hercegovine | LABSAGAS KJKP Sarajevogas d.o.o Trg Fadile Odžaković 4 71000 Sarajevo Bosnia and Herzegovina | ibusuladzic@sarajevogas.ba | 17.6.2024 – 5.7.2024 | Ibrahim Busuladžić |
| Czech Republic (PILOT LAB) | CMI Czech Metrology Institute | Czech Metrology Institute Prumyslova 455 53003 Pardubice Czech Republic | tvalenta@cmi.cz +420 466 670 728 | 8.7.2024 – 26.7.2024 | Tomas Valenta |

3. Device

A rotary piston gas meter will be the instrument which will be tested. The description and the picture of the meter are mentioned down in table 2 and in the figure 1.

Figure 1 – Rotary piston gas meter IRM-3-DUO G650  ELSTER

Table 2 - Technical specification of the meter

| Manufacturer: ELSTER INSTROMET, Stará Turá, Slovak Republic |  |
| --- | --- |
| EC-type examination certificate: T10198 (NMi) | Pmax : 16 bar |
| Size: G650 | Inside diameter: DN150 |
| Qmin : 6 m3/h | Qmax : 1000 m3/h |
| Weight: approximately 62 kg |  |

The dimensions of the meter are mentioned in table 3 and in the figure 2.

Table 3  - Dimensions of  the meter

| Nominal diameter DN | A | B | C | E | F |
| --- | --- | --- | --- | --- | --- |
| 150 | 598 mm | 260 mm | 308 mm | 336 mm | 262 mm |

Figure 2 - Dimensions of  the meter

The HF pulse generator type REPROX will be used.  Only one of two HF  pulse generators will be used. The right one will be labelled.  This emitter is made according to EN 60947-5-6 (NAMUR). The pulse emitter is mentioned in the figure 3 and in the figure 4.

Figure 3 – HF pulse emitter REPROX

Figure 4 – HF pulse emitter determined for tests

This pulse emitter shall be used for tests.

The meter will be packed in wooden box that is mentioned in the figure 5.  The diameter of the box is (800x370x500)mm. The weight of the complete box with the meter is approximately 75 kg.

Figure 5 – The wooden box for the rotary piston gas meter  G650

500 mm

800 mm

370 mm

In the box there will be the meter, a pulse emitter connector and the copy of this protocol.

4. Advice on handling and on travelling

It is necessary to handle with care with the meter during travelling and also during packing and unpacking.  The meter is not resistant against any beats and downfalls. During packing and unpacking each laboratory will have to use its own ropes with hooks and its own crane (figure 1).

During test the air may flow through the rotary piston gas meter only in the direction marked by the arrow on the meter. It is advisable to protect the rotary piston gas meter against pressure shocks, big gas flow changes and against overload.

There will be two loops of the comparison during the project. The first loop will be among EU countries where no ATA CARNET is needed. One loop will be among countries which are out of EU and for this purpose the pilot laboratory will ensure the issue of ATA CARNET. For the transport of the gas meter among countries which are out of EU it is obligatory to use service of company DHL (http://www.dhl.com) which is able to ensure all customs affairs without problems.

The non-EU laboratory (METAS, Institut za mjeriteljstvo Bosne i Hercegovine) will be responsible to use only aircraft transport for delivery of the meter to next laboratory.

5. Actions to be taken on receipt and after sending of the meter

Each participant laboratory has to give information to the coordinator after sending of the meter to the next laboratory. Each participant laboratory has to give information to the coordinator after receipt of the meter. The best way is to send e-mail to the coordinator’s address tvalenta@cmi.cz.

6. Test procedure

The participating laboratories can use their usual calibration procedure. Only instructions mentioned down have to be fulfilled.

- The rotary piston gas meter has to be tested in horizontal position by air near barometric pressure.

- The reference pressure from the rotary piston gas meter has to be measured from the output “pr”. This pressure output “pr” is in the inlet (upstream) of the gas meter.

This pressure output shall be used for tests like reference pressure “pr”.

- The upstream straightening pipe of the length at least 5xDN shall be used.

- The reference temperature from the rotary piston gas meter should be measured in the distance (23)x DN downstream of the rotary piston  gas meter.

- It is necessary to use only a HF pulse emitter for the tests. The right HF pulse emitter is the one which is close to the centre of a gas meter. The correct one is not labelled “DO NOT USE”.

This HF pulse emitter shall be used for tests.

- The test should be performed in the laboratory where the temperature is from 19.5°C to 23.5°C.  The upstream pressure of the meter should be near atmospheric pressure.

- Before the beginning of the test the gas meter has to work 20 minutes in a flow rate Q=400 m3/h.

- The rotary piston gas meter has to be tested in 9  flow rates:

1000 m3/h, 800 m3/h, 650 m3/h, 450 m3/h, 350 m3/h, 250 m3/h, 160 m3/h, 100 m3/h, 50 m3/h.

- The test in one flow rate should be repeated at least 3 times and then the means of values in the table 3 have to be calculated.  The flow rate has to be in the interval ± 3% of the required value.

- The one single test in one flow rate has to take more than 3 minutes. Beforehand the flow rate has to be accurately stabilised.

- Each participant has to record the results in the form of table 3 mentioned down.

Table 3

| Flow rate in the meter | Absolute pressure in the meter | Temperature in the meter | Pressure loss of the meter | Error of the meter |
| --- | --- | --- | --- | --- |
| (m3/h) | (Pa) | (°C) | (Pa) | (%) |
| 1000 |  |  |  |  |
| 800 |  |  |  |  |
| 650 |  |  |  |  |
| 450 |  |  |  |  |
| 350 |  |  |  |  |
| 250 |  |  |  |  |
| 160 |  |  |  |  |
| 100 |  |  |  |  |
| 50 |  |  |  |  |

Error of the meter is value which shows the relationship in percentage terms of the difference between the volume indicated by the meter and the volume which has actually flowed through the meter, to the later value.

(%)              [1]

where     	E is the error of the meter

Vi   is the indicated volume by the meter (m3)

Vc   is the real volume which has actually flowed through the meter (m3)

7. Instruction for reporting results

During 5 weeks after the test of the meter in the laboratory all the data mentioned down have to be sent to coordinator by e-mail.

- the table of means of the measured values and the uncertainty of the error of the meter

| Flow rate in the meter | Absolute pressure in the meter | Temperature in the meter | Pressure loss of the meter | Error of the meter | Uncertainty of the error U(k=2) |
| --- | --- | --- | --- | --- | --- |
| (m3/h) | (Pa) | (°C) | (Pa) | (%) | (%) |
| 1000 |  |  |  |  |  |
| 800 |  |  |  |  |  |
| 650 |  |  |  |  |  |
| 450 |  |  |  |  |  |
| 350 |  |  |  |  |  |
| 250 |  |  |  |  |  |
| 160 |  |  |  |  |  |
| 100 |  |  |  |  |  |
| 50 |  |  |  |  |  |

- the description and CMC in KCDB of the test facility which was used for the test of the rotary piston gas meter and the address where the facility is situated

- the traceability of the standard meters and other measuring instruments which were used during the test including the recalibration interval of the meters

- the description of the test procedure

The determination of the error of the meter is usually based on the comparison of the volume (or mass) of the air indicated by standard meter and of the volume which was indicated by the meter under test after calculations of corrections concerning different temperatures and pressures in the standard meter and in the meter under test.

The volume of the air depends on the pressure and temperature. Hence it is clear that the main sources of uncertainty of measurement are the uncertainties of the standard meter (for example standard gas meter, sonic nozzles), of the temperature meters and of pressure meters. The humidity, the time measurement and the barometric pressure can also be involved in the uncertainty budget. The uncertainty of the error of the meter U(k=2) has to be calculated according to the Guide to Expression of Uncertainty in Measurement (published by ISO, Geneva, 1995).

8. Evaluation

8.1. Description of the method

The reference value will be determined in each flow rate separately. The method of determination of the reference value in each flow rate will correspond to the procedure A presented by M.G.Cox1). Only results from independent laboratories will be taken into account for the determination of the key comparison reference value (KCRV) and of the uncertainty of the key comparison reference value. Then the results from dependent laboratories will be compared with the key comparison reference value and with the uncertainty of the key comparison reference value.

8.1.1.	The determination of the Key Comparison Reference Value (KCRV) and its uncertainty

The reference value y will be calculated as weighted mean error (WME):

,                     [2]

where    x1,  x2, ….. xn  	are errors of the meter in one flow rate in different independent laboratories    1,2, …...n

ux1, ux2,…..uxn	are standard uncertainties (not expanded) of the error in different independent laboratories  1,2, …...n  including the uncertainty caused by stability of the meter

The standard uncertainties (not expanded) of the error in different laboratories ux1, ux2,…..uxn   (equation [2] ) will include the stability of the meter. These uncertainties will calculated by

[3]

where  	is the expanded uncertainty (k=2) determined by laboratory i and presented in results of laboratory i

Utm      	is estimated expanded uncertainty caused by the stability (reproducibility) of the rotary piston  gas meter (The meter will be tested three times in the pilot laboratory and from these results  Utm will be determined.)

The standard uncertainty of the reference value uy  is given by

[4]

The expanded uncertainty of the reference value U(y) is

[5]

The  chi-squared test for consistency check  will be performed using values of errors of the meter in each flow rate. At first the chi-squared value will be calculated by

[6]

The degrees of freedom  will be assigned

[7]

where  n is number of evaluated laboratories.

The consistency check will be  failing if

Pr{}<0,05                        [8]

(The function CHIINV(0,05;) in MS Excel will be used. The consistency check will be  failing if   CHIINV(0,05; )< )

If the consistency check does not fail then y will be accepted as the key comparison reference value xref and U(y)  will be  accepted as the expanded uncertainty of the key comparison  reference value U(xref).

If the consistency check  fails then the laboratory with the highest value of  will be excluded for the next round of evaluation and the new reference value y (WME), the new standard uncertainty of the reference value uy and the chi-squared value will be  calculated again without the values of excluded laboratory. The consistency check will be calculated again, too. This procedure will be  repeated till the consistency check will pass.

8.1.2.	The determination of the differences “Lab to KCRV” and “Lab to Lab” as well as their uncertainties and Degrees of Equivalence

When the KCRV will be determined, the differences between the participating laboratories and the KCRV will be  calculated according to

[9]

[10]

Based on these differences, the Degree of Equivalence (DoE) will be calculated according to:

[11]

and                           ,      respectively.          [12]

The DoE is a measure for the equivalence of the results of any laboratory with the KCRV or with any other laboratory, respectively:

- The results of a laboratory will be equivalent (passed) if  Ei or Eij ≤ 1.

- The laboratory will be determined as not equivalent (failed) if Ei or Eij >1.2.

- For values of DoE in the range 1 < Ei or Eij ≤ 1.2 the “warning level” is defined.  In this case some actions to check are recommended to the laboratory.

The reason for such “warning level” is that it is necessary to consider the confidence in the determination of the uncertainties (for the results of labs as well the KCRV). Conventionally we work at a 95% confidence level. Therefore in some comparisons a range up to E < 1.5 is used for these “warnings”2). This is a reasonable value where stochastic influences dominate the uncertainty budgets.  In the case of comparisons for gas flow, the smaller value 1.2 was chosen, which reflects the dominance of non-stochastic parts of uncertainty compared to the stochastic parts. (The reproducibility is usually much better than the total uncertainty of a laboratory). 3)

The calculation of the DoE needs the information about the uncertainty of the differences di and dij  (equations [11] and [12]). To make statements about this, it is necessary to consider first the general problem of the difference of two values x1 and x2. If we look to the pure propagation of (standard) uncertainty we find:

[13]

Simply spoken, the (standard) uncertainty of the difference is the quadratic sum of the uncertainties of the inputs (u1 and u2) subtracting twice the covariance (cov) between the two input values.

Therefore it is possible find the different cases in this comparison:

A)	Differences to the KCRV

A1) Independent laboratories with contribution to the KCRV

The covariance between the result of a laboratory (with contribution to the KCRV) and the KCRV is the variance of the KCRV itself. 1)

=> 	[14]

A2) Independent laboratories without contribution to the KCRV

There is no covariance between the result of a laboratory without contribution and the KCRV.

=> 	[15]

A3) Laboratories with traceability to a laboratory contributing to the KCRV

In this case we have covariance between the laboratory and the KCRV because the laboratory is linked to the KCRV via the source of traceability. Although we have no detailed information about it, we can determine a conservative estimation of an upper limit of this covariance. The upper limit is determined for the theoretical case if we have no additional stochastic influence in the traceability of the lab from its source (which is the lab contributing to the KCRV). Then the results of the lab considered here would be strongly correlated with the results of the laboratory contributing to the KCRV (correlation coefficient = 1) and there would be the same covariance to the KCRV as in case A1. In any case of additional uncertainty caused stochastically the correlation and consequently the covariance is smaller.

=> 	[16]

B)	Differences Lab to Lab

B1) Independent laboratories

There is no covariance between the results of two independent laboratory i   and j

=> 	[17]

B2) Dependent laboratories with common source of traceability

In the case of two labs i and j with a common source of traceability we will find again a covariance between these labs which is caused by the common source. In our case the common source is another laboratory from which the traceabilities of both labs are derived. Again we can determine a conservative upper limit of the covariance for the same reason as in A3 as cov = u2SourceLab.

=> 	[18]

The equations from [14] to [18] use the standard uncertainties (k = 1). The expanded uncertainties U(di) and U(dij)  (see equations [11],[12]) are determined by

[19]

[20]

9. Financial aspects

The participation of any laboratory in this comparison is free of charge. Of course, each laboratory is responsible for the delivery of the meter to the next laboratory. It means the expense of the delivery of the meter to the next laboratory will be paid by the previous participating laboratory.

The cost of ATA CARNET will be paid by the pilot laboratory. The meter is the property of the pilot laboratory.

1) Cox M.G., Evaluation of  key comparison data, Metrologia, 2002, 39, 589-595

2) C.Ullner et al., Special features in proficiency tests of mechanical testing laboratories, and
P. Robouch et al., The „Naji Plot“, a simple graphical tool for the evaluation of inter-laboratory comparisons,
Both in: D. Richter, W. Wöger, W. Hässelbarth (ed.) Data analysis of key comparisons, 178. PTB-Seminar/International Workshop, ISBN 3-89701-933-3.

3) 	D.Dopheide, B.Mickan, R.Kramer, H.-J.Hotze, J.-P.Vallet, M.R.Harris, Jiunn-Haur Shaw, Kyung-Am Park,  CIPM Key Comparisons for Compressed Air and Nitrogen, CCM.FF-5.b – Final Report, 07/09/2006

http://kcdb.bipm.org/appendixB/appbresults/ccm.ff-k5.b/ccm.ff-k5.b_final_report.pdf

1)  Cox M.G., Evaluation of  key comparison data, Metrologia, 2002, 39, 589-595