2025-10-09 v. 1.4

Bilateral key comparisons of gas pressure between LEI and VTT MIKES in the range 25 kPa to 200 kPa (gauge and absolute mode) and 0.7 MPa to 7 MPa (gauge mode)

Technical Protocol

Euramet Project No. 1694

# Introduction

This key comparison was initiated by the Lithuanian Energy Institute (Lithuanian DI) to demonstrate the metrological equivalence of their national standard in the range from 0.008 MPa to 7.0 MPa (gauge mode) and from 0.008 MPa to 0.35 MPa (absolute mode). The comparisons will be piloted by LEI, and VTT MIKES who took part in the previous EURAMET key comparisons (EURAMET.M.P-K8. 2012 and EURAMET.M.P-K1.c. 2020) has agreed to link the reference value. The parameter to be compared is the determined effective area of a Piston Cylinder Assembly (PCA-unit) for each selected pressure level in the range of 25 kPa to 200 kPa (gauge and absolute mode) and in the range 0.7 MPa to 7.0 MPa gauge mode. This technical protocol specifies the procedures to be followed in the comparison.

# Participants

Pilot laboratory: 	Lithuanian Energy Institute (LEI)

Contact person: 	Andrius Tonkonogovas

Invoice address: 	Lithuanian Energy Institute

Breslaujos 3, Kaunas, Lithuania

LT-44403

Phone:		+370 671 13961

E-mail: 		Andrius.Tonkonogovas@lei.lt

Reference laboratory: VTT MIKES

Contact persons: 	Antti Lakka, Richard Högström

Address: 		VTT Technical Research Centre of Finland Ltd

National Metrology Institute VTT MIKES

Tekniikantie 1, Espoo

P. O.Box 1000, FI-02044 VTT, Finland

Phone: 		+358505952231, +358 50 3039341

E-mail 		Antti.Lakka@vtt.fi, Richard.Hogstrom@vtt.fi

# Transfer standard

The comparisons will be performed using transfer standard – pressure balance Ruska 2465 with two piston-cylinder units and a mass set. It will be provided by JSC “Vilduja” Central calibration and testing laboratory, through the agreement with LEI. As a result of these comparisons, VTT will issue a calibration certificate.

Technical parameters of the transfer standard are presented in the Table.

| Parameter | Value |
| --- | --- |
| Base unit | 2465-754 |
| Manufacturer | Ruska |
| Serial number |  |

Low range piston-cylinder unit

| Parameter | Value |
| --- | --- |
| Model | 2465-725 |
| Serial number | TL1711 |
| Nominal area of the piston | 335 mm2 |
| Pressure range | 0.0014 – 0.172 MPa |
| Piston material | 440C Stainless steel |
| Cylinder material | Tungsten Carbide |
| Thermal expansion coefficient | 1.5·10-5 1/°C |
| Conventional mass ± U (k=2) | (47.0363 g ± 10·10-5) g |
| Assumed density of the piston | (8000±200) kg/m3 |

High range piston-cylinder unit

| Parameter | Value |
| --- | --- |
| Model | 2465-729 |
| Serial number | V1692 |
| Pressure range | 0.014 – 0.7 MPa |
| Nominal area of the piston | 8 mm2 |
| Piston material | Tungsten Carbide |
| Cylinder material | Tungsten Carbide |
| Thermal expansion coefficient | 9.1·10-6 1/°C |
| Conventional mass ± U (k=2) | (11.63791 ± 5·10-5) g |
| Assumed density of the piston | (8000±200) kg/m3 |

Mass set

| Parameter | Value |
| --- | --- |
| Model | 2465-799 |
| Serial number | 65990 |
| Approximate Total Mass | 6.31 kg |
| Assumed density | (8000±200) kg/m3 |
| Mass Material | 300 Series, Non-magnetic, Austenitic, Stainless Steel |

| Mass number | Conventional mass, g | Uncertainty, mg (k = 2) | Density |
| --- | --- | --- | --- |
| Mass carrier + bell (1) | 113.6121 | 0.16 | 8000 |
| 2 | 1000.005 | 1.03 | 8000 |
| 3 | 1000.041 | 1.03 | 8000 |
| 4 | 1000.018 | 1.03 | 8000 |
| 5 | 999.989 | 1.03 | 8000 |
| 6 | 1000.052 | 1.03 | 8000 |
| 7 | 500.050 | 1.05 | 8000 |
| 8 | 300.0202 | 0.39 | 8000 |
| 9 | 200.0075 | 0.25 | 8000 |
| 10 | 100.0078 | 0.16 | 8000 |
| 11 | 49.9981 | 0.08 | 8000 |
| 12 | 30.00340 | 0.12 | 8000 |
| 13 | 20.00971 | 0.07 | 8000 |
| 14 | 10.00782 | 0.05 | 8000 |

# Time schedule

The expected preliminary time schedule is:

| Dates | Action |
| --- | --- |
| 2024-11-01 .. 2025-02-01 | Measurements (gauge pressure) at LEI |
| 2025-05-21..2025-05-22 | TS transportation to VTT MIKES |
| 2025-05-26..2025-08-30 | Measurements at VTT MIKES |
| 2025-09-02..2025-09-05 | TS transportation to LEI |
| 2025-10-25..2025-12-31 | Measurement at LEI |

# Transportation of the equipment and cost

The shipping costs of the equipment are covered by LEI. Equipment will be delivered to VTT and back to LEI personally by the LEI.

The cost of the calibration performed by VTT MIKES will be covered by LEI.

# Laboratory standards

## LEI

LEI national standard is DH Instruments pressure balance with 2 piston-cylinder assemblies and mass set (traceability to SI is transferred via calibration at PTB, calibration certificates No PTB 11090/23 dated 2023-10-12 and No. PTB 30079/23 dated 2024-01-16 ) and Pfeifer vacuum pressure gauge (traceability to SI is transferred via calibration at MIKES, calibration certificate No M-25P052, dated 2025-05-27 . Instruments characteristics are presented in tables bellow.

Base

| Parameter | Value |
| --- | --- |
| Model | PG7601 |
| Serial number | 296 |
| PCU temperature sensor (integrated in PG7601 base) | PT100 |
| PCU temperature probe serial number | U323 |
| Temperature calibration uncertainty (k=2) | 0.07 °C |

Weight carrier

| Parameter | Value |
| --- | --- |
| Serial number | 396 |
| Mass | (0.2999865±5058) kg |
| Density ± U (k=2) | (5058±100) kg/m3 |

Low pressure PCU

| Parameter | Value |
| --- | --- |
| Piston cylinder assembly | PC-7100/7600-10 TC |
| Serial number | 0525 |
| Piston material | Tungsten carbide |
| Cylinder material | Tungsten carbide |
| Nominal diameter | 35 mm |
| Calibration range | (0.008-0.5) MPa |
| Mass ± U (k=2) | (0.4999790±5·10-6 )kg |
| Density ± U (k=2) | (10080±150) kg/m3 |
| A0 ± σ | (9.80528±0.012) mm2 |
| λ ± U (k=2) | (4.2±0.73)·10-12 Pa-1 |
| (α+β) ± U (k=2) | (9.0±1.0)·10-6 °C-1 |

High pressure PCU

| Parameter | Value |
| --- | --- |
| Piston cylinder assembly | PC-7100/7600-200 |
| Serial number | 0487 |
| Piston material | Tungsten carbide |
| Cylinder material | Tungsten carbide |
| Nominal diameter | 8 mm |
| Calibration range | (0.1-10) MPa |
| Mass ± U (k=2) | (0.1999948±2·10-6) kg |
| Density ± U (k=2) | (8030±100) kg/m3 |
| A0 ± σ | (49,01484±0.0085) mm2 |
| λ ± U (k=2) | (-3.1±0.65)·10-12 Pa-1 |
| (α+β) ± U (k=2) | (9.0±1.0)·10-6 °C-1 |

Mass set

| Parameter | Value |
| --- | --- |
| Model | MS-7002-35 |
| Serial number | 2143 |
| Material | Stainless steel AISI 304L |
| Nominal value | 100 g- 5 kg |

| Mass number | True mass, g | Uncertainty, mg (k = 2) | Density |
| --- | --- | --- | --- |
| 5 kg – 1 | 5000.037 | 15 | (7930±140) kg/m3 |
| 5 kg – 2 | 5000.048 | 15 |  |
| 5 kg – 3 | 5000.043 | 15 |  |
| 5 kg – 4 | 5000.052 | 15 |  |
| 5 kg – 5 | 5000-043 | 15 |  |
| 4 kg | 4000.039 | 12 |  |
| 2 kg – 1 | 2000.021 | 6 |  |
| 2 kg – 1 | 2000.021 | 6 |  |
| 1 kg | 1000.006 | 3 |  |
| 0.5 kg | 500.0025 | 3 |  |
| 0.2 kg – 1 | 200.0002 | 0.6 |  |
| 0.2 kg – 2 | 200.0009 | 0.6 |  |
| 0.1 kg | 100.0001 | 0.3 |  |

Vacuum gauge

| Parameter | Value |
| --- | --- |
| Type | Capacitance Diaphragm Gauge |
| Model | CMR 364 |
| Serial number | 47018136 |
| Measurement range | (10-4 – 1.1) hPa |
| Controller | OmniControl200 |
| Serial number | 80108137 |

## MIKES

VTT MIKES national standards include DH Instruments and Desgranges et Huot pressure balances with 5 piston-cylinder assemblies and 3 mass sets and a Pirani vacuum pressure gauge. Instruments characteristics are presented in tables below.

Bases

| Parameter | Value |
| --- | --- |
| Model | PG7607 |
| Serial number | 397 |
| PCU temperature transducer serial number | MIKES010391 |
| Temperature calibration uncertainty (k=2) | 0.5 °C |

| Parameter | Value |
| --- | --- |
| Model | PG7601 |
| Serial number | 149 |
| PCU temperature transducer serial number | MIKES010390 |
| Temperature calibration uncertainty (k=2) | 0.5 °C |

| Parameter | Value |
| --- | --- |
| Model | D&H 5203 |
| Serial number | 4401 |
| PCU temperature transducer serial number | 1300 ch101 |
| Temperature calibration uncertainty (k=2) | 0.02 °C |

| Parameter | Value |
| --- | --- |
| Model | D&H 5203 |
| Serial number | 6402 |
| PCU temperature transducer serial number | 1123 ch102 |
| Temperature calibration uncertainty (k=2) | 0.02 °C |

Weight carriers

| Parameter | Value |
| --- | --- |
| Serial number | 3105 |
| Mass | (800.007±0.004) g |
| Density ± U (k=2) | (7920±30) kg/m3 |

| Parameter | Value |
| --- | --- |
| Serial number | 2926 |
| Mass | (799.978±0.004) g |
| Density ± U (k=2) | (7920±30) kg/m3 |

| Parameter | Value |
| --- | --- |
| Serial number | 2376 |
| Mass | (492.972±0.005) g |
| Density ± U (k=2) | (4610±160) kg/m3 |

| Parameter | Value |
| --- | --- |
| Serial number | 626 |
| Mass | (299.995±0.003) g |
| Density ± U (k=2) | (5013±60) kg/m3 |

PCUs

| Parameter | Value |
| --- | --- |
| Piston cylinder assembly manufacturer | DHI |
| Serial number | 451 |
| Piston material | Tungsten carbide |
| Cylinder material | Tungsten carbide |
| Calibration range | (0.005-0.17) MPa |
| Mass ± U (k=2) | (699.9507±0.0030) g |
| Density ± U (k=2) | (970360) kg/m3 |
| A0 ± σ | (1961.046±0.044) mm2 |
| λ ± U (k=2) | (7.15±)·10-12 Pa-1 |
| (α+β) ± U (k=2) | (9.0±1.0)·10-6 °C-1 |

| Parameter | Value |
| --- | --- |
| Piston cylinder assembly manufacturer | DHI |
| Serial number | 1931 |
| Piston material | Tungsten carbide |
| Cylinder material | Tungsten carbide |
| Calibration range | (0.007-0.35) MPa |
| Mass ± U (k=2) | (400.0037±0.0008) g |
| Density ± U (k=2) | (9363100) kg/m3 |
| A0 ± σ | (980.5260±0.0216) mm2 |
| λ ± U (k=2) | 0 Pa-1 |
| (α+β) ± U (k=2) | (9.0±1.0)·10-6 °C-1 |

| Parameter | Value |
| --- | --- |
| Piston cylinder assembly manufacturer | DHI |
| Serial number | 620 |
| Piston material | Tungsten carbide |
| Cylinder material | Tungsten carbide |
| Calibration range | (0.04-1.75) MPa |
| Mass ± U (k=2) | (199.9945±0.0007) g |
| Density ± U (k=2) | (691030) kg/m3 |
| A0 ± σ | (196.1238±0.0044) mm2 |
| λ ± U (k=2) | 0 Pa-1 |
| (α+β) ± U (k=2) | (9.0±1.0)·10-6 °C-1 |

| Parameter | Value |
| --- | --- |
| Piston cylinder assembly manufacturer | D&H |
| Serial number | 4012 |
| Piston material | Tungsten carbide |
| Cylinder material | Tungsten carbide |
| Calibration range | (0.1-5) MPa |
| Mass ± U (k=2) | (199.9945±0.0008) g |
| Density ± U (k=2) | (9800300) kg/m3 |
| A0 ± σ | (98.04905±0.00216) mm2 |
| λ ± U (k=2) | 0 Pa-1 |
| (α+β) ± U (k=2) | (9.0±1.0)·10-6 °C-1 |

| Parameter | Value |
| --- | --- |
| Piston cylinder assembly manufacturer | D&H |
| Serial number | 6804 |
| Piston material | Tungsten carbide |
| Cylinder material | Tungsten carbide |
| Calibration range | (0.2-16) MPa |
| Mass ± U (k=2) | (199.9953±0.0007) g |
| Density ± U (k=2) | (8800200) kg/m3 |
| A0 ± σ | (49.02622±0.00197) mm2 |
| λ ± U (k=2) | (1.6±0.2)·10-12 Pa-1 |
| (α+β) ± U (k=2) | (9.0±1.0)·10-6 °C-1 |

Mass sets

| Parameter | Value |
| --- | --- |
| Manufacturer | DHI |
| Serial number | 2376 |
| Material | AISI 304L |
| Nominal value | 0.1 kg- 6.2 kg |

| Mass number | True mass, g | Uncertainty, g (k = 2) | Density |
| --- | --- | --- | --- |
| 6.2kg 1 | 6 247.046 | 0.010 | (7920±……) kg/m3 |
| 6.2kg 2 | 6 241.484 | 0.010 |  |
| 6.2kg 3 | 6 244.967 | 0.010 |  |
| 6.2kg 4 | 6 243.793 | 0.010 |  |
| 6.2kg 5 | 6 242.422 | 0.010 |  |
| 3.2kg | 3 209.041 | 0.004 |  |
| 1.6 kg | 1 608.994 | 0.003 |  |
| 0.8 kg | 806.6795 | 0.0015 |  |
| 0.4 kg | 402.4284 | 0.0005 |  |
| 0.2 kg | 200.56179 | 0.00020 |  |
| 0.1 kg | 100.97399 | 0.00020 |  |

| Parameter | Value |
| --- | --- |
| Manufacturer | DHI |
| Serial number | 2229 |
| Material | AISI 304L |
| Nominal value | 100 g- 5000 g |

| Mass number | True mass, g | Uncertainty, g (k = 2) | Density |
| --- | --- | --- | --- |
| 100 g | 100.0003 | 0.0003 | (8000±…….)kg/m3 |
| 200 g 1 | 199.9990 | 0.0006 |  |
| 200 g 2 | 200.0006 | 0.0006 |  |
| 500 g | 500.002 | 0.002 |  |
| 1000 g | 1000.014 | 0.003 |  |
| 2000 g 1 | 2000.020 | 0.004 |  |
| 2000 g 2 | 2000.018 | 0.004 |  |
| 4500 g | 4500.003 | 0.010 |  |
| 5000 g 1 | 5000.046 | 0.010 |  |
| 5000 g 2 | 5000.037 | 0.010 |  |
| 5000 g 3 | 5000.029 | 0.010 |  |
| 5000 g 4 | 5000.038 | 0.010 |  |
| 5000 g 5 | 5000.030 | 0.010 |  |

| Parameter | Value |
| --- | --- |
| Manufacturer | D&H |
| Serial number | 2926 |
| Material | AISI 304L |
| Nominal value | 100 g- 5 kg |

| Mass number | True mass, g | Uncertainty, g (k = 2) | Density |
| --- | --- | --- | --- |
| 100 g | 99.9995 | 0.0003 | (7920±……) kg/m3 |
| 200 g 1 | 199.9996 | 0.0006 |  |
| 200 g 2 | 200.0004 | 0.0006 |  |
| 500 g | 500.002 | 0.002 |  |
| 1 kg | 999.989 | 0.003 |  |
| 2 kg 1 | 1999.996 | 0.004 |  |
| 2 kg 2 | 1999.995 | 0.004 |  |
| 4 kg | 4000.005 | 0.008 |  |
| 5 kg 1 | 4999.984 | 0.010 |  |
| 5 kg 2 | 4999.991 | 0.010 |  |
| 5 kg 3 | 4999.998 | 0.010 |  |
| 5 kg 4 | 4999.986 | 0.010 |  |
| 5 kg 5 | 4999.985 | 0.010 |  |
| 5 kg 6 | 4999.962 | 0.010 |  |
| 5 kg 7 | 4999.984 | 0.010 |  |
| 5 kg 8 | 4999.983 | 0.010 |  |

Vacuum gauge

| Parameter | Value |
| --- | --- |
| Type | Pirani vacuum gauge |
| Model | PG7601 |
| Serial number | 149 |
| Measurement range | (0.006 – 0.01) hPa |

# Measurements procedures

The preparation of the measurements, as well as their execution and reporting, should be done according to procedures based on EURAMET cg-3 methodology (including measurement uncertainty estimation), and include the following specified steps:

## Preparation

Upon arrival at the laboratory, personnel should inspect the transfer standard for potential damage or deformation, which could impact measurements.

The piston cylinder unit should be cleaned when the laboratory personnel think it is necessary.

- Visual inspection of the instrument. The piston and cylinder shall also be inspected with respect to potential damage or deformations, which could influence the measurement results. Cleaning of the PCA-unit should be performed if deemed necessary.

- Reassembly and initial pressure balance functionality test including leak rate, and evaluation of piston fall rate and piston rotation speed.

- Alignment of reference levels between the transfer standard and the reference standard of the laboratory (if required).

As gas medium, dry nitrogen (N2) or synthetic air is to be used.

## Calibration

The comparison measurements shall be carried out in four measurement series at increasing and/or decreasing pressure in the nominal pressure points, specified below.

5-200 kPa gauge and absolute pressure

| Number | Nominal pressure level [kPa] | TS piston, used for measurements | Weights combination |
| --- | --- | --- | --- |
| 1 | 25 | 2465-725 | 1, 7, 10, 11, 12, 14 |
| 2 | 50 | 2465-725 | 1, 2, 7, 12, 14 |
| 3 | 75 | 2465-725 | 1, 2, 3, 8, 10 |
| 4 | 100 | 2465-725 | 1, 2, 3, 4, 9, 11 |
| 5 | 125 | 2465-725 | 1, 2, 3, 4, 5, 10, 14 |
| 6 | 150 | 2465-725 | 1, 2, 3, 4, 5, 7, 8, 10, 11, 14 |
| 7 | 175 | 2465-725 | 1, 2, 3, 4, 5, 6, 7, 8, 13 |
| 8 | 200 | 2465-729 | 1, 12, 14 |

0,7 – 7,0 MPa gauge pressure

| Number | Nominal pressure level [MPa] | TS piston, used for measurements | Weights combination |
| --- | --- | --- | --- |
| 1 | 0.73 | 2465-729 | 1, 7 |
| 2 | 1.08 | 2465-729 | 1, 7, 9, 11, 12, 14 |
| 3 | 1.77 | 2465-729 | 1, 2, 8, 11, 12 |
| 4 | 2.94 | 2465-729 | 1, 2, 3, 8, 11, 12 |
| 5 | 4.12 | 2465-729 | 1, 2, 3, 4, 8, 11, 12 |
| 6 | 5.28 | 2465-729 | 1, 2, 3, 4, 5, 8, 11, 12 |
| 7 | 6.45 | 2465-729 | 1, 2, 3, 4, 5, 6, 8, 11, 12 |
| 8 | 6.77 | 2465-729 | 1, 2, 3, 4, 5, 6, 7, 10, 11, 12 |

## Stability check

Prior to the comparison a calibration of the transfer standard was performed at LEI, which will serve as the stability check. The calibration followed the same procedures as those specified in this technical protocol, but was limited to gauge pressure with four repetitions at each of the pressure points. The results of these measurements will be used to verify that the transfer standard maintained its functionality during transports and that no significant drift has occurred between the time of the calibrations and the stability check. In the case that the instrument proves to have drifted, uncertainty due to this drift will be included in the uncertainty of the reference value.

# Reporting

The participants should report the data of their own pressure standards filling the table given in Annex 1, as well as giving any additional information, if useful.

It should be reported how TS was connected to LS and how their pressure equilibrium was achieved and controlled.

Results of individual measurement series and a summary of the 4-measurement series should be reported using forms given in Annexes 1, 2 and 3.

In addition, a list of main uncertainty sources and their contributions to 𝑢(𝐴𝑝) for minimum and maximum pressure must be reported. Reports with the results of the measurements should be prepared as a WORD file with the tables from Annexes 2 and 3 being additionally provided in an EXCEL file.

The reports should initially be sent to the head of the EURAMET TC-M Subcommittee Pressure as an impartial instance by e-mail to wladimir.sabuga@ptb.de within three months after finishing the measurements. Then they will be forwarded to the pilot laboratory. The pilot laboratory will evaluate the measurement data and prepare a report.

# Evaluation

The reference value of the comparison is provided by the reference laboratory including associated uncertainties, and (if required) including influence due to drift.

The results will be compared in accordance with the statistical approach presented in the standard ISO/IEC 17043:2010 “Conformity assessment – General requirements for proficiency testing”. The degree of equivalence will be described by the number 𝐸𝑛 calculated from the mean values of determined effective areas at each pressure level, according to the following equation

En=Ap,t,LEI-Ap,t,MIKESULEI2+UMIKES2

where:

Ap,t,LEI - is the value of the effective area determined by LEI (participant’s result);

Ap,t,MIKES – is the value of the effective area determined by MIKES (reference value);

ULEI– is the estimated expanded uncertainty of the LEI result;

UMIKES– is the estimated expanded uncertainty of the MIKES result.

# Link to KCRV

The results of the present KC will be linked to the reference values of EURAMET.M.P-K8 and EURAMET.M.P-K1.c using results of VTT MIKES obtained in these two comparisons. The link will be performed at gauge and absolute pressure points (25, 50, 75, 100, 125, 150, 175 and 200) kPa and (0.73; 1.08; 1.77; 2.94; 4.12; 5.28; 6.45 and 6.77) MPa. To do this, the results of the present KC will be linearly interpolated to the pressure points of the link.

# References

- EURAMET Guide on Comparisons EURAMET Guide No. 4 Version 2.0 (04/2021)

- EURAMET Calibration Guide No. 3. Guidelines on the Calibration of Pressure Balances Version 2.0, 05/2024

- JCGM 100:2008 Evaluation of measurement data — Guide to the expression of uncertainty in measurement

# Annexes

Annex 1. Results of measurements (gauge pressure)

| pnom, kPa/MPa | tLS, °C | tamb, °C | Pamb, Pa | φ, % | p, kPa/MPa | tTS, °C | Ap, cm2 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 20 |  |  |  |  |  |  |  |
| … |  |  |  |  |  |  |  |
| … |  |  |  |  |  |  |  |

pnom, – nominal pressure;

tLS – temperature of laboratory standard;

tamb – ambient temperature;

pamb – ambient pressure;

φ amb – ambient humidity;

p – applied reference pressure at the reference level of transfer standard;

tTS – temperature of transfer standard;

𝐴𝑝,t is effective area of TS pressure p and reference temperature 20 °C

Annex 2. Results of measurements (absolute pressure)

| pnom, kPa | μLS, Pa | t, °C | tamb, °C | Pamb, Pa | φ, % | p, kPa | μTS, Pa | tTS, °C | Ap, cm2 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

pnom, – nominal pressure;

μLS – absolute pressure under the bell jar of laboratory standard;

tLS – temperature of laboratory standard;

tamb – ambient temperature;

pamb – ambient pressure;

φ amb – ambient humidity;

p – applied reference pressure at the reference level of transfer standard;

μLS – absolute pressure under the bell jar of transfer standard;

tTS – temperature of transfer standard;

𝐴𝑝,t is effective area of TS pressure p and reference temperature 20 °C

Annex 3. Results of measurements

| pnom, kPa/MPa | Ap, cm2 | σ(Ap )/Ap ·106 | u(p)/p·106 | u(t), °C | u(Ap )/Ap ·106 |
| --- | --- | --- | --- | --- | --- |

Ap  – mean value of the Ap, measured at the same nominal pressure;

σ(Ap )/Ap – standard deviation of the Ap mean value;

u(p)/p –type B uncertainty of the pressure at the reference level of transfer standard,

u(t) – uncertainty of the transfer standard temperature measured values,

u(Ap )/Ap – combined uncertainty of the Ap mean value.

In addition, a list of main uncertainty sources and their contributions to u(Ap) for pressures of 20 kPa and 175 kPa (both gauge and absolute pressure), and 0.7 MPa and 7.0 MPa (gauge mode) must be presented.

All the uncertainties should be expressed as standard ones.