## DoE BIPM.EM-K12-100 ohm
100 ohm resistance measurement | Di / 10-9 | 10-9
BIPM.EM-K12
| Di | Ui only or Ui- | Ui+ (optional)
PTB | 0.4 | 4.2
|  |  |  | Ui | symmetric uncertainty around Di. This is the most common representation ofuncertainties.
|  |  |  | Ui- | "negative uncertainty" where the lower uncertainty limit is Di - Ui-
|  |  |  | Ui+ | "positive uncertainty" where the upper uncertainty limit is Di + Ui+
All uncertainty values should be given in absolute values.

## TABLE 1 BIPM.EM-K12 - 100 ohm
Comparison of 100 ohm standards
| Lab i | participating institute
| xi | measurement value submitted by institute i
| u(xi) | standard uncertainty of xi
| Labi | xi | u(xi)
|  | 10-9 | 10-9
| BIPM | -155.78 | 2.1
| PTB | -155.36 | 0.21
(The KCRV is defined by the BIPM measurement)

## TABLE 2 BIPM.EM-K12 - 100 ohm
Comparison of 100 ohm standards
| Lab i | participating institute
| xi | relative difference between the result of measurement of laboratory i and that of the BIPM
| u(xi) | standard uncertainty of xi
| Labi | xi | u(xi) | Date of Measurement
|  | 10-9 | 10-9
| PTB | 0.4 | 2.1 | 2025-05-08
(The KCRV is defined by the BIPM measurement)

## DoE BIPM.EM-K12 - K1 ratio
10 kohm / 100 ohm resistance ratio (K1) | Di / 10-9 | 10-9
BIPM.EM-K12
| Di | Ui only or Ui- | Ui+ (optional)
PTB | -0.7 | 4.2
|  |  |  | Ui | symmetric uncertainty around Di. This is the most common representation ofuncertainties.
|  |  |  | Ui- | "negative uncertainty" where the lower uncertainty limit is Di - Ui-
|  |  |  | Ui+ | "positive uncertainty" where the upper uncertainty limit is Di + Ui+
All uncertainty values should be given in absolute values.

## TABLE 1 BIPM.EM-K12 - K1 ratio
Comparison of 10 kohm / 100 ohm resistance ratio (K1)
| Lab i | participating institute
| xi | measurement value submitted by institute i
| u(xi) | standard uncertainty of xi
| Labi | xi | u(xi)
|  | 10-9 | 10-9
| BIPM | 1648.15 | 2
| PTB | 1647.47 | 0.56
(The KCRV is defined by the BIPM measurement)

## TABLE 2 BIPM.EM-K12 - K1 ratio
Comparison of 10 kohm / 100 ohm resistance ratio (K1)
| Lab i | participating institute
| xi | relative difference between the result of measurement of laboratory i and that of the BIPM
| u(xi) | standard uncertainty of xi
| Labi | xi | u(xi) | Date of Measurement
|  | 10-9 | 10-9
| PTB | -0.7 | 2.1 | 2025-05-09
(The KCRV is defined by the BIPM measurement)

## DoE BIPM.EM-K12 - K2 ratio
100 ohm / 1 ohm resistance ratio (K2) | Di / 10-9 | 10-9
BIPM.EM-K12
| Di | Ui only or Ui- | Ui+ (optional)
PTB | -0.7 | 4.8
|  |  |  | Ui | symmetric uncertainty around Di. This is the most common representation ofuncertainties.
|  |  |  | Ui- | "negative uncertainty" where the lower uncertainty limit is Di - Ui-
|  |  |  | Ui+ | "positive uncertainty" where the upper uncertainty limit is Di + Ui+
All uncertainty values should be given in absolute values.

## TABLE 1 BIPM.EM-K12 - K2 ratio
Comparison of 100 ohm / 1 ohm resistance ratio (K2)
| Lab i | participating institute
| xi | measurement value submitted by institute i
| u(xi) | standard uncertainty of xi
| Labi | xi | u(xi)
|  | 10-9 | 10-9
| BIPM | 864.53 | 2
| PTB | 863.88 | 0.94
(The KCRV is defined by the BIPM measurement)

## TABLE 2 BIPM.EM-K12 - K2 ratio
Comparison of 100 ohm / 1 ohm resistance ratio (K2)
| Lab i | participating institute
| xi | relative difference between the result of measurement of laboratory i and that of the BIPM
| u(xi) | standard uncertainty of xi
| Labi | xi | u(xi) | Date of Measurement
|  | 10-9 | 10-9
| PTB | -0.7 | 2.4 | 2025-05-12
(The KCRV is defined by the BIPM measurement)