Publication of a Final Report in Metrologia’s Technical Supplement

To submit your final report for publication in Metrologia’s Technical Supplement, please complete the following form.

| Title of REPORT | On-site comparison of Quantum Hall Effect resistance standards of the PTB and the BIPM: Ongoing key comparison BIPM.EM-K12 |
| --- | --- |
| Contact Person | Pierre Gournay |
| CONTACT PERSON e-mail | pierre.gournay@bipm.org |
| FIELD | Electricity and Magnetism |
| COMPARISON TYPE | KEY COMPARISON |
| COMPARISON identifier | BIPM.EM-K12 |
| APPROVING BODY | CCEM |
| final report URL | TO BE COMPLETED BY THE KCDB OFFICE |

| ABSTRACT | The ongoing on-site comparison BIPM.EM-K12 is part of the BIPM program implemented to verify the international coherence of the primary resistance standards. It allows National Metrology Institutes (NMIs) to validate their implementations of the Quantum Hall Effect (QHE) for dc resistance traceability by comparison to the reference maintained at the BIPM. The realization of the ohm from the QHE-based standard of the NMIs at 100 is compared with that realized by the BIPM from its own transportable quantum Hall resistance standard (QHRS). This comparison is usually completed by scaling measurements from 100 to 1 and 10 k through the measurement of the resistance ratios 100 Ω/1 Ω and 10 kΩ/100 Ω, respectively. In May 2025, a new BIPM.EM-K12 on-site comparison was carried out at the Physikalisch-Technische Bundesanstalt (PTB). It was the third time the PTB participated in this comparison program (previous comparisons in 1996 and 2013). Measurements of the 100 Ω transfer standard in terms of the ohm realized from the QHE-based standards of the PTB and the BIPM agreed to 4 parts in 1010 with a relative combined standard uncertainty uc = 21 × 10−10. Measurements of 10 kΩ/100 Ω and 100 Ω/1 Ω ratios agreed to 7 parts in 1010 with uc = 21 × 10−10 and to 7 parts in 1010 with uc = 24 × 10−10, respectively. The above result for the comparison of the 100 transfer standard was obtained when both the BIPM and the PTB used a GaAs-based QHRS. For the first time in this key comparison program, the 100 Ω comparison was repeated using the same GaAs-based QHRS for the BIPM and a graphene-based QHRS for the PTB. These additional comparative measurements also gave excellent results and confirmed the equivalence between the two types of QHRS. This confirms the possibility of using a graphene-based QHRS as a primary standard for realizing the ohm under relaxed experimental conditions (typically 5 T and 4.2 K). |
| --- | --- |
| KEY WORDS FOR SEARCH separated by ; | Quantum Hall effect; Quantum Hall resistance standard; Graphene; Key comparison; International comparison; Resistance standard; Resistance ratio |

| AUTHOR (e.g. J. P. Johnson – one author for each line) Please add lines if necessary. | AFFILIATION AND POSTAL ADDRESS | ORCID ID (optional) |
| --- | --- | --- |
| Pierre GOURNAY | Bureau International de Poids et Mesures (BIPM). 12bis Grande Rue, F-92310 Sèvres, France. pierre.gournay@bipm.org | 0000-0002-3040-3324 |
| Benjamin ROLLAND | Bureau International de Poids et Mesures (BIPM). 12bis Grande Rue, F-92310 Sèvres, France. Benjamin.rolland@bipm.org | 0000-0002-8701-2199 |
| Mattias KRUSKOPF | Physikalisch-Technische Bundesanstalt (PTB) Bundesallee 100, 38116 Braunschweig mattias.kruskopf@ptb.de | 0000-0003-2846-3157 |
| Eckart PESEL | Physikalisch-Technische Bundesanstalt (PTB) Bundesallee 100, 38116 Braunschweig Eckart.Pesel@ptb.de | 0000-0003-4112-7377 |
| Johanna LENGSFELD | Physikalisch-Technische Bundesanstalt (PTB) Bundesallee 100, 38116 Braunschweig johanna.lengsfeld@ptb.de | 0009-0005-1381-0795 |
| Dinesh Kumar PATEL | Physikalisch-Technische Bundesanstalt (PTB) Bundesallee 100, 38116 Braunschweig dinesh.patel@ptb.de | 0000-0002-4934-1039 |