## Introduction
EURAMET.AUV.A-K5 Draft A Tables of Data
This Microsoft Excel spreadsheet contains supplementary data for Report on Key Comparison EURAMET.AUV.A-K5
The size and number of these tables made it impractical to include them within the body of the report.
The 'Parameters' tab lists all declared microphone paramaters used by each participant.
For each of the two reference microphones (4160 811014 and 4160 2036126) there are the following tabs:
Level | Tables of declared sensitivity level measurements and associated expanded uncertainties for all participating laboratories
Phase | Tables of declared sensitivity phase measurements and associated expanded uncertainties for all participating laboratories
Additionally for 4160 2036126 there are the following tabs:
DoE Level | Degrees of equivalence for sensitivity level relative to the KCRV value derived through CCAUV.A-K5
DoE Phase | Degrees of equivalence for sensitivity phase, relative to the KCRV value derived through CCAUV.A-K5
Spreadsheet protection
This spreadsheet is password protected. The password can be requested from janine.avison@npl.co.uk.

## Parameters
Declared Parameters
Derived by minimising difference in results for multiple couplers
Declared as nominal values
Derived from other data
811014 | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG
total volume (mm3) | UK: 682.94; FR: 687.7; DE: 688
front volume (mm3) | 532.7 | 533.7 | 527.2 | 535 | 550 | 543 | 543 | 531.1 | 548 |  | 532.06 | 529.185
cavity depth (mm) | 1.9465 | 1.96 | 1.942 | 1.9449 | 1.95 | 1.95 | 1.96 | 1.935 | 1.948 | 1.944 | 1.95 | 1.947
acoustic mass (kg m-4) | UK: 385.1; EG: 392.5
acoustic compliance (kg-1 m4 s2) | UK: 1.059e-12; EG: 9.51e-13
acoustic resistance (kg m-4 s-1) | EG: 21400000
loss factor |  | 0.98 |  | 1.0132 | 1.05 | 1.05 | 1.05 | 1.05 | 1.05 | 1.05 | 1.05 | 1
res. frequency (kHz) |  | 7.95 | 7.834 | 7.7433 | 8.2 | 8.2 | 8.3 | 8.1 | 7.875 | 7.84 | 8.62 | 7.9
equivalent volume (mm3) |  | 154 | 160 | 150 | 145 | 137 | 132 | 153.9 | 136 | 150 | 126.85 | 142.6829
temp. coeff. at 250 Hz (dB/K)(Level) | UK: -0.0028; FR: -0.002; BG: -0.002; AT: -0.0019; ES: -0.00213; EG: -0.002
press. coeff. at 250 Hz (dB/kPa)(Level) | UK: -0.0153; FR: -0.0152; BG: -0.0015; AT: -0.0153; ES: -0.0158; EG: -0.016
temp. coeff. at 250 Hz (deg/K)(Phase) | UK: -0.0032
press. coeff. at 250 Hz (deg/kPa)(Phase) | UK: 0.0012
diaphragm diameter (mm) | IT: 17.9; CH: 17.9; AT: 17.9
damping | IT: 0.98
front cavity excessive surface/added thread area (mm2)
2036126 | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG
total volume (mm3) | UK: 663.79; FR: 668.7; DE: 666
front volume (mm3) | 531 | 533.7 | 531.21 | 535 | 535 | 535 | 543 | 541 | 529 |  | 540.74 | 532.717
cavity depth (mm) | 1.958 | 1.96 | 1.955 | 1.9449 | 1.95 | 1.95 | 1.96 | 1.973 | 1.966 | 1.956 | 1.97 | 1.96
acoustic mass (kg m-4) | UK: 406; EG: 405
acoustic compliance (kg-1 m4 s2) | UK: 9.36e-13; EG: -9.4e-13
acoustic resistance (kg m-4 s-1) | UK: 22000000; EG: 23000000
loss factor |  | 0.94 |  | 1.0561 | 1.05 | 1.05 | 1.05 | 1.05 | 1.08 | 1.06 | 1.05 | 1
res. frequency (kHz) |  | 8.85 | 8.69 | 8.6761 | 8.2 | 8.2 | 8.3 | 8.75 | 8.725 | 8.68 | 8.62 | 8.25
equivalent volume (mm3) |  | 135 | 135.68 | 136 | 136 | 136 | 132 | 123.7 | 136 | 128 | 126.85 | 142.44
temp. coeff. at 250 Hz (dB/K)(Level) | UK: -0.0028; FR: -0.002; BG: -0.002; AT: -0.0019; ES: -0.00213; EG: -0.0025
press. coeff. at 250 Hz (dB/kPa)(Level) | UK: -0.0153; FR: -0.0152; BG: -0.0015; AT: -0.0153; ES: -0.0158; EG: -0.015
temp. coeff. at 250 Hz (deg/K)(Phase) | UK: -0.0032
press. coeff. at 250 Hz (deg/kPa)(Phase) | UK: 0.0012
diaphragm diameter (mm) | IT: 17.9; CH: 17.9
damping | IT: 1.01
front cavity excessive surface/added thread area (mm2)

## 2036126 Level
[25 empty] | Estimated reference sensitivity level | Std Unc (ERSL)
| Sensitivity Level (4160 2036126) dB re 1 V/Pa | [11 empty] | Expanded Uncertainties (dB) | [13 empty] | Level difference relative to estimated reference sensitivity level (dB)
Frequency (Hz) | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG | dB re 1 V/Pa | dB | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG
1.9953 | Sensitivity Level (4160 2036126) dB re 1 V/Pa UK: -26.69; Sensitivity Level (4160 2036126) dB re 1 V/Pa FR: -26.808; Sensitivity Level (4160 2036126) dB re 1 V/Pa IT: -26.76; Sensitivity Level (4160 2036126) dB re 1 V/Pa CH: -26.754; Sensitivity Level (4160 2036126) dB re 1 V/Pa AT: -26.763; Sensitivity Level (4160 2036126) dB re 1 V/Pa DE: -26.8; Expanded Uncertainties (dB) UK: 0.25; Expanded Uncertainties (dB) FR: 0.14; Expanded Uncertainties (dB) IT: 0.14; Expanded Uncertainties (dB) CH: 0.127; Expanded Uncertainties (dB) AT: 0.06; Expanded Uncertainties (dB) DE: 0.23 | -26.7346 | 0.036895 | 0.044555 | -0.073445 | -0.025445 | -0.019445 |  |  |  | -0.028445 |  | -0.065445
2.5119 | Sensitivity Level (4160 2036126) dB re 1 V/Pa UK: -26.82; Sensitivity Level (4160 2036126) dB re 1 V/Pa FR: -26.876; Sensitivity Level (4160 2036126) dB re 1 V/Pa IT: -26.84; Sensitivity Level (4160 2036126) dB re 1 V/Pa CH: -26.854; Sensitivity Level (4160 2036126) dB re 1 V/Pa AT: -26.864; Sensitivity Level (4160 2036126) dB re 1 V/Pa DE: -26.89; Expanded Uncertainties (dB) UK: 0.25; Expanded Uncertainties (dB) FR: 0.08; Expanded Uncertainties (dB) IT: 0.11; Expanded Uncertainties (dB) CH: 0.104; Expanded Uncertainties (dB) AT: 0.048; Expanded Uncertainties (dB) DE: 0.23 | -26.8627 | 0.025279 | 0.042711 | -0.013289 | 0.022711 | 0.0087109 |  |  |  | -0.0012891 |  | -0.027289
3.1623 | Sensitivity Level (4160 2036126) dB re 1 V/Pa UK: -26.88; Sensitivity Level (4160 2036126) dB re 1 V/Pa FR: -26.963; Sensitivity Level (4160 2036126) dB re 1 V/Pa IT: -26.92; Sensitivity Level (4160 2036126) dB re 1 V/Pa CH: -26.938; Sensitivity Level (4160 2036126) dB re 1 V/Pa AT: -26.954; Sensitivity Level (4160 2036126) dB re 1 V/Pa DE: -26.97; Expanded Uncertainties (dB) UK: 0.25; Expanded Uncertainties (dB) FR: 0.06; Expanded Uncertainties (dB) IT: 0.1; Expanded Uncertainties (dB) CH: 0.086; Expanded Uncertainties (dB) AT: 0.044; Expanded Uncertainties (dB) DE: 0.23 | -26.9333 | 0.030719 | 0.053323 | -0.029677 | 0.013323 | -0.0046765 |  |  |  | -0.020677 |  | -0.036677
3.9811 | Sensitivity Level (4160 2036126) dB re 1 V/Pa UK: -26.97; Sensitivity Level (4160 2036126) dB re 1 V/Pa FR: -27.036; Sensitivity Level (4160 2036126) dB re 1 V/Pa IT: -27; Sensitivity Level (4160 2036126) dB re 1 V/Pa CH: -27.016; Sensitivity Level (4160 2036126) dB re 1 V/Pa AT: -27.033; Sensitivity Level (4160 2036126) dB re 1 V/Pa DE: -27.04; Expanded Uncertainties (dB) UK: 0.07; Expanded Uncertainties (dB) FR: 0.05; Expanded Uncertainties (dB) IT: 0.09; Expanded Uncertainties (dB) CH: 0.072; Expanded Uncertainties (dB) AT: 0.043; Expanded Uncertainties (dB) DE: 0.12 | -27.0104 | 0.033883 | 0.040448 | -0.025552 | 0.010448 | -0.0055522 |  |  |  | -0.022552 |  | -0.029552
5.0119 | Sensitivity Level (4160 2036126) dB re 1 V/Pa UK: -27.05; Sensitivity Level (4160 2036126) dB re 1 V/Pa FR: -27.103; Sensitivity Level (4160 2036126) dB re 1 V/Pa IT: -27.06; Sensitivity Level (4160 2036126) dB re 1 V/Pa CH: -27.081; Sensitivity Level (4160 2036126) dB re 1 V/Pa AT: -27.101; Sensitivity Level (4160 2036126) dB re 1 V/Pa DE: -27.11; Expanded Uncertainties (dB) UK: 0.07; Expanded Uncertainties (dB) FR: 0.04; Expanded Uncertainties (dB) IT: 0.08; Expanded Uncertainties (dB) CH: 0.063; Expanded Uncertainties (dB) AT: 0.041; Expanded Uncertainties (dB) DE: 0.12 | -27.0742 | 0.031555 | 0.024204 | -0.028796 | 0.014204 | -0.0067959 |  |  |  | -0.026796 |  | -0.035796
6.3096 | Sensitivity Level (4160 2036126) dB re 1 V/Pa UK: -27.11; Sensitivity Level (4160 2036126) dB re 1 V/Pa FR: -27.16; Sensitivity Level (4160 2036126) dB re 1 V/Pa IT: -27.12; Sensitivity Level (4160 2036126) dB re 1 V/Pa CH: -27.14; Sensitivity Level (4160 2036126) dB re 1 V/Pa AT: -27.16; Sensitivity Level (4160 2036126) dB re 1 V/Pa DE: -27.17; Expanded Uncertainties (dB) UK: 0.07; Expanded Uncertainties (dB) FR: 0.04; Expanded Uncertainties (dB) IT: 0.07; Expanded Uncertainties (dB) CH: 0.056; Expanded Uncertainties (dB) AT: 0.041; Expanded Uncertainties (dB) DE: 0.12 | -27.1281 | 0.028919 | 0.018104 | -0.031896 | 0.0081044 | -0.011896 |  |  |  | -0.031896 |  | -0.041896
7.9433 | Sensitivity Level (4160 2036126) dB re 1 V/Pa UK: -27.16; Sensitivity Level (4160 2036126) dB re 1 V/Pa FR: -27.206; Sensitivity Level (4160 2036126) dB re 1 V/Pa IT: -27.17; Sensitivity Level (4160 2036126) dB re 1 V/Pa CH: -27.19; Sensitivity Level (4160 2036126) dB re 1 V/Pa AT: -27.211; Sensitivity Level (4160 2036126) dB re 1 V/Pa DE: -27.21; Expanded Uncertainties (dB) UK: 0.06; Expanded Uncertainties (dB) FR: 0.03; Expanded Uncertainties (dB) IT: 0.06; Expanded Uncertainties (dB) CH: 0.051; Expanded Uncertainties (dB) AT: 0.04; Expanded Uncertainties (dB) DE: 0.08 | -27.1787 | 0.02749 | 0.018694 | -0.027306 | 0.0086943 | -0.011306 |  |  |  | -0.032306 |  | -0.031306
10 | Sensitivity Level (4160 2036126) dB re 1 V/Pa UK: -27.21; Sensitivity Level (4160 2036126) dB re 1 V/Pa FR: -27.247; Sensitivity Level (4160 2036126) dB re 1 V/Pa IT: -27.22; Sensitivity Level (4160 2036126) dB re 1 V/Pa CH: -27.233; Sensitivity Level (4160 2036126) dB re 1 V/Pa AT: -27.254; Sensitivity Level (4160 2036126) dB re 1 V/Pa DE: -27.25; Sensitivity Level (4160 2036126) dB re 1 V/Pa ES: -27.273; Expanded Uncertainties (dB) UK: 0.06; Expanded Uncertainties (dB) FR: 0.03; Expanded Uncertainties (dB) IT: 0.06; Expanded Uncertainties (dB) CH: 0.048; Expanded Uncertainties (dB) AT: 0.039; Expanded Uncertainties (dB) DE: 0.08; Expanded Uncertainties (dB) ES: 0.08 | -27.231 | 0.027516 | 0.020953 | -0.016047 | 0.010953 | -0.002047 |  |  |  | -0.023047 |  | -0.019047 | -0.042047
12.5893 | Sensitivity Level (4160 2036126) dB re 1 V/Pa UK: -27.25; Sensitivity Level (4160 2036126) dB re 1 V/Pa FR: -27.281; Sensitivity Level (4160 2036126) dB re 1 V/Pa IT: -27.26; Sensitivity Level (4160 2036126) dB re 1 V/Pa CH: -27.269; Sensitivity Level (4160 2036126) dB re 1 V/Pa AT: -27.291; Sensitivity Level (4160 2036126) dB re 1 V/Pa DE: -27.29; Sensitivity Level (4160 2036126) dB re 1 V/Pa ES: -27.306; Expanded Uncertainties (dB) UK: 0.06; Expanded Uncertainties (dB) FR: 0.03; Expanded Uncertainties (dB) IT: 0.05; Expanded Uncertainties (dB) CH: 0.046; Expanded Uncertainties (dB) AT: 0.039; Expanded Uncertainties (dB) DE: 0.08; Expanded Uncertainties (dB) ES: 0.08 | -27.2717 | 0.02354 | 0.021721 | -0.0092795 | 0.011721 | 0.0027205 |  |  |  | -0.019279 |  | -0.018279 | -0.034279
15.8489 | Sensitivity Level (4160 2036126) dB re 1 V/Pa UK: -27.29; Sensitivity Level (4160 2036126) dB re 1 V/Pa FR: -27.313; Sensitivity Level (4160 2036126) dB re 1 V/Pa IT: -27.29; Sensitivity Level (4160 2036126) dB re 1 V/Pa CH: -27.299; Sensitivity Level (4160 2036126) dB re 1 V/Pa AT: -27.321; Sensitivity Level (4160 2036126) dB re 1 V/Pa DE: -27.32; Sensitivity Level (4160 2036126) dB re 1 V/Pa ES: -27.334; Expanded Uncertainties (dB) UK: 0.05; Expanded Uncertainties (dB) FR: 0.03; Expanded Uncertainties (dB) IT: 0.05; Expanded Uncertainties (dB) CH: 0.044; Expanded Uncertainties (dB) AT: 0.039; Expanded Uncertainties (dB) DE: 0.05; Expanded Uncertainties (dB) ES: 0.08 | -27.2981 | 0.024048 | 0.0080841 | -0.014916 | 0.0080841 | -0.00091592 |  |  |  | -0.022916 |  | -0.021916 | -0.035916
19.9526 | -27.33 | -27.341 | -27.32 | -27.325 | -27.362 | -27.367 | -26.5742 | -27.348 | -27.326 | -27.35 | -27.358 | -27.334 | 0.03 | 0.03 | 0.03 | 0.044 | 0.04 | 0.07 | 0.09 | 0.039 | 0.05 | 0.05 | 0.08 | 0.0633 | -27.3281 | 0.0075676 | -0.0019418 | -0.012942 | 0.0080582 | 0.0030582 | -0.033942 | -0.038942 | 0.75382 | -0.019942 | 0.0020582 | -0.021942 | -0.029942 | -0.0059418
21.1349 | -27.34 | -27.347 | -27.32 | -27.331 |  |  | -26.5773 | -27.356 | -27.336 | -27.35 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.09 | 0.039 | 0.05 | 0.05 |  |  | -27.3322 | 0.0076203 | -0.0077511 | -0.014751 | 0.012249 | 0.0012489 |  |  | 0.7549 | -0.023751 | -0.0037511 | -0.017751
22.3872 | -27.34 | -27.35 | -27.33 | -27.335 |  |  | -26.58 | -27.362 | -27.34 | -27.36 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.09 | 0.039 | 0.04 | 0.05 |  |  | -27.3377 | 0.0076203 | -0.0022713 | -0.012271 | 0.0077287 | 0.0027287 |  |  | 0.75771 | -0.024271 | -0.0022713 | -0.022271
23.7137 | -27.35 | -27.354 | -27.34 | -27.342 |  |  | -26.5825 | -27.367 | -27.346 | -27.36 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.09 | 0.038 | 0.04 | 0.05 |  |  | -27.3467 | 0.0076203 | -0.0032775 | -0.0072775 | 0.0067225 | 0.0047225 |  |  | 0.76421 | -0.020278 | 0.00072248 | -0.013278
25.1189 | -27.35 | -27.362 | -27.34 | -27.346 | -27.383 | -27.39 | -26.5851 | -27.372 | -27.351 | -27.37 | -27.379 | -27.361 | 0.03 | 0.03 | 0.03 | 0.044 | 0.04 | 0.06 | 0.09 | 0.038 | 0.04 | 0.05 | 0.05 | 0.0585 | -27.3457 | 0.0074344 | -0.0043484 | -0.016348 | 0.0056516 | -0.00034841 | -0.037348 | -0.044348 | 0.76051 | -0.026348 | -0.0053484 | -0.024348 | -0.033348 | -0.015348
26.6073 | -27.36 | -27.366 | -27.35 | -27.353 |  |  | -26.5874 | -27.377 | -27.359 | -27.37 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.05 |  |  | -27.3562 | 0.0074344 | -0.0038456 | -0.0098456 | 0.0061544 | 0.0031544 |  |  | 0.76872 | -0.020846 | -0.0028456 | -0.013846
28.1838 | -27.37 | -27.372 | -27.35 | -27.367 |  |  | -26.5898 | -27.383 | -27.362 | -27.38 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.09 | 0.038 | 0.04 | 0.05 |  |  | -27.3613 | 0.0074344 | -0.008727 | -0.010727 | 0.011273 | -0.005727 |  |  | 0.77145 | -0.021727 | -0.00072699 | -0.018727
29.8538 | -27.37 | -27.377 | -27.36 | -27.361 |  |  | -26.592 | -27.386 | -27.368 | -27.38 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.09 | 0.038 | 0.04 | 0.05 |  |  | -27.3657 | 0.0074344 | -0.0042771 | -0.011277 | 0.0057229 | 0.0047229 |  |  | 0.77369 | -0.020277 | -0.0022771 | -0.014277
31.6228 | -27.37 | -27.379 | -27.36 | -27.366 | -27.4 | -27.4 | -26.5941 | -27.39 | -27.375 | -27.39 | -27.398 | -27.381 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.09 | 0.038 | 0.04 | 0.04 | 0.05 | 0.0555 | -27.3663 | 0.0073166 | -0.0036827 | -0.012683 | 0.0063173 | 0.00031731 | -0.033683 | -0.033683 | 0.77218 | -0.023683 | -0.0086827 | -0.023683 | -0.031683 | -0.014683
33.4965 | -27.38 | -27.386 | -27.37 | -27.37 |  |  | -26.5959 | -27.395 | -27.378 | -27.39 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | -27.3784 | 0.0072639 | -0.0015691 | -0.0075691 | 0.0084309 | 0.0084309 |  |  | 0.78258 | -0.016569 | 0.00043093 | -0.011569
35.4813 | -27.39 | -27.387 | -27.37 | -27.375 |  |  | -26.5976 | -27.401 | -27.383 | -27.4 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | -27.384 | 0.0072639 | -0.0059939 | -0.0029939 | 0.014006 | 0.0090061 |  |  | 0.78639 | -0.016994 | 0.0010061 | -0.015994
37.5837 | -27.39 | -27.395 | -27.38 | -27.38 |  |  | -26.5992 | -27.404 | -27.389 | -27.4 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | -27.3851 | 0.0072639 | -0.0048847 | -0.0098847 | 0.0051153 | 0.0051153 |  |  | 0.78596 | -0.018885 | -0.0038847 | -0.014885
39.8107 | -27.39 | -27.401 | -27.38 | -27.382 | -27.414 | -27.418 | -26.6005 | -27.408 | -27.393 | -27.4 | -27.414 | -27.399 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.09 | 0.038 | 0.04 | 0.04 | 0.05 | 0.0554 | -27.3862 | 0.0072608 | -0.0038332 | -0.014833 | 0.0061668 | 0.0041668 | -0.027833 | -0.031833 | 0.78562 | -0.021833 | -0.0068332 | -0.013833 | -0.027833 | -0.012833
42.1697 | -27.4 | -27.401 | -27.39 | -27.387 |  |  | -26.6018 | -27.412 | -27.398 | -27.41 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | -27.3959 | 0.0072608 | -0.0041356 | -0.0051356 | 0.0058644 | 0.0088644 |  |  | 0.7941 | -0.016136 | -0.0021356 | -0.014136
44.6684 | -27.4 | -27.401 | -27.39 | -27.39 |  |  | -26.6028 | -27.416 | -27.402 | -27.41 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | -27.3976 | 0.0072608 | -0.0023636 | -0.0033636 | 0.0076364 | 0.0076364 |  |  | 0.79483 | -0.018364 | -0.0043636 | -0.012364
47.3151 | -27.41 | -27.406 | -27.4 | -27.4 |  |  | -26.6037 | -27.419 | -27.407 | -27.42 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | -27.4065 | 0.007288 | -0.0034696 | 0.00053036 | 0.0065304 | 0.0065304 |  |  | 0.80283 | -0.01247 | -0.00046964 | -0.01347
50.1187 | -27.41 | -27.409 | -27.4 | -27.38 | -27.428 | -27.431 | -26.6046 | -27.423 | -27.41 | -27.42 | -27.428 | -27.414 | 0.03 | 0.03 | 0.03 | 0.05 | 0.04 | 0.05 | 0.09 | 0.038 | 0.04 | 0.04 | 0.05 | 0.0555 | -27.4104 | 0.0091955 | 0.00044003 | 0.00144 | 0.01044 | 0.03044 | -0.01756 | -0.02056 | 0.80588 | -0.01256 | 0.00044003 | -0.00956 | -0.01756 | -0.00356
53.0884 | -27.42 | -27.411 | -27.41 | -27.402 |  |  | -26.6052 | -27.426 | -27.413 | -27.42 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | -27.4165 | 0.007288 | -0.0034666 | 0.0055334 | 0.0065334 | 0.014533 |  |  | 0.81137 | -0.0094666 | 0.0035334 | -0.0034666
56.2341 | -27.42 | -27.415 | -27.41 | -27.405 |  |  | -26.6056 | -27.43 | -27.418 | -27.43 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | -27.4154 | 0.007288 | -0.0045907 | 0.00040935 | 0.0054093 | 0.010409 |  |  | 0.80976 | -0.014591 | -0.0025907 | -0.014591
59.5662 | -27.42 | -27.417 | -27.41 | -27.408 |  |  | -26.6061 | -27.433 | -27.421 | -27.43 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | -27.4153 | 0.007288 | -0.0047089 | -0.0017089 | 0.0052911 | 0.0072911 |  |  | 0.80921 | -0.017709 | -0.0057089 | -0.014709
63.0957 | -27.43 | -27.421 | -27.42 | -27.411 | -27.439 | -27.439 | -26.6063 | -27.435 | -27.425 | -27.43 | -27.441 | -27.428 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.038 | 0.04 | 0.04 | 0.03 | 0.0512 | -27.4232 | 0.007288 | -0.0068036 | 0.0021964 | 0.0031964 | 0.012196 | -0.015804 | -0.015804 | 0.8169 | -0.011804 | -0.0018036 | -0.0068036 | -0.017804 | -0.0048036
66.8344 | -27.43 | -27.424 | -27.42 | -27.415 |  |  | -26.6064 | -27.438 | -27.429 | -27.44 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | -27.4238 | 0.007288 | -0.0061789 | -0.00017887 | 0.0038211 | 0.0088211 |  |  | 0.81746 | -0.014179 | -0.0051789 | -0.016179
70.7946 | -27.44 | -27.428 | -27.42 | -27.417 |  |  | -26.6065 | -27.443 | -27.432 | -27.44 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | -27.4282 | 0.0072032 | -0.011813 | 0.00018702 | 0.008187 | 0.011187 |  |  | 0.82171 | -0.014813 | -0.003813 | -0.011813
74.9894 | -27.44 | -27.431 | -27.43 | -27.42 |  |  | -26.6063 | -27.445 | -27.435 | -27.44 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | -27.4346 | 0.0072032 | -0.0054389 | 0.0035611 | 0.0045611 | 0.014561 |  |  | 0.82822 | -0.010439 | -0.00043893 | -0.0054389
79.4328 | -27.44 | -27.434 | -27.43 | -27.423 | -27.449 | -27.453 | -26.6063 | -27.447 | -27.439 | -27.45 | -27.453 | -27.441 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.038 | 0.04 | 0.04 | 0.03 | 0.0511 | -27.4355 | 0.0072032 | -0.0045401 | 0.0014599 | 0.0054599 | 0.01246 | -0.01354 | -0.01754 | 0.82911 | -0.01154 | -0.0035401 | -0.01454 | -0.01754 | -0.0055401
84.1395 | -27.44 | -27.435 | -27.43 | -27.426 |  |  | -26.6058 | -27.45 | -27.442 | -27.45 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | -27.4332 | 0.0072032 | -0.0068041 | -0.0018041 | 0.0031959 | 0.0071959 |  |  | 0.82741 | -0.016804 | -0.0088041 | -0.016804
89.1251 | -27.45 | -27.438 | -27.44 | -27.428 |  |  | -26.6055 | -27.454 | -27.444 | -27.45 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | -27.4436 | 0.0072032 | -0.0063774 | 0.0056226 | 0.0036226 | 0.015623 |  |  | 0.83815 | -0.010377 | -0.00037737 | -0.0063774
94.4061 | -27.45 | -27.442 | -27.44 | -27.43 |  |  | -26.6048 | -27.456 | -27.447 | -27.45 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | -27.4448 | 0.0072151 | -0.0051873 | 0.0028127 | 0.0048127 | 0.014813 |  |  | 0.83998 | -0.011187 | -0.0021873 | -0.0051873
100 | -27.45 | -27.445 | -27.44 | -27.433 | -27.459 | -27.462 | -26.6046 | -27.458 | -27.449 | -27.46 | -27.462 | -27.451 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.038 | 0.04 | 0.04 | 0.03 | 0.0511 | -27.4436 | 0.0072151 | -0.0063653 | -0.0013653 | 0.0036347 | 0.010635 | -0.015365 | -0.018365 | 0.83908 | -0.014365 | -0.0053653 | -0.016365 | -0.018365 | -0.0073653
105.9254 | -27.46 | -27.444 | -27.44 | -27.436 |  |  | -26.6043 | -27.46 | -27.452 | -27.46 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | -27.4472 | 0.0072109 | -0.01282 | 0.0031801 | 0.0071801 | 0.01118 |  |  | 0.8429 | -0.01282 | -0.0048199 | -0.01282
112.2018 | -27.46 | -27.447 | -27.45 | -27.438 |  |  | -26.6037 | -27.463 | -27.454 | -27.46 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | -27.4547 | 0.0072109 | -0.0052913 | 0.0077087 | 0.0047087 | 0.016709 |  |  | 0.85098 | -0.0082913 | 0.00070875 | -0.0052913
118.8502 | -27.46 | -27.451 | -27.45 | -27.441 |  |  | -26.6033 | -27.465 | -27.46 | -27.46 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | -27.4534 | 0.0071705 | -0.0066422 | 0.0023578 | 0.0033578 | 0.012358 |  |  | 0.85004 | -0.011642 | -0.0066422 | -0.0066422
125.8925 | -27.46 | -27.454 | -27.45 | -27.443 | -27.469 | -27.466 | -26.6029 | -27.466 | -27.461 | -27.46 | -27.472 | -27.46 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.038 | 0.04 | 0.03 | 0.03 | 0.0491 | -27.4519 | 0.0071705 | -0.0080799 | -0.0020799 | 0.0019201 | 0.0089201 | -0.01708 | -0.01408 | 0.84904 | -0.01408 | -0.0090799 | -0.0080799 | -0.02008 | -0.0080799
133.3521 | -27.47 | -27.455 | -27.46 | -27.444 |  |  | -26.6023 | -27.468 | -27.464 | -27.47 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.03 |  |  | -27.4641 | 0.0071705 | -0.005887 | 0.009113 | 0.004113 | 0.020113 |  |  | 0.86185 | -0.003887 | 0.00011303 | -0.005887
141.2538 | -27.47 | -27.461 | -27.46 | -27.446 |  |  | -26.6018 | -27.471 | -27.466 | -27.47 |  |  | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4612 | 0.0070702 | -0.008752 | 0.00024799 | 0.001248 | 0.015248 |  |  | 0.85943 | -0.009752 | -0.004752 | -0.008752
149.6236 | -27.47 | -27.465 | -27.46 | -27.445 |  |  | -26.6013 | -27.472 | -27.468 | -27.47 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4628 | 0.0070702 | -0.0071726 | -0.0021726 | 0.0028274 | 0.017827 |  |  | 0.86155 | -0.0091726 | -0.0051726 | -0.0071726
158.4893 | -27.47 | -27.464 | -27.46 | -27.452 | -27.475 | -27.476 | -26.6009 | -27.474 | -27.471 | -27.47 | -27.48 | -27.469 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0491 | -27.4613 | 0.0070702 | -0.0087322 | -0.0027322 | 0.0012678 | 0.0092678 | -0.013732 | -0.014732 | 0.86041 | -0.012732 | -0.0097322 | -0.0087322 | -0.018732 | -0.0077322
167.8804 | -27.48 | -27.464 | -27.46 | -27.453 |  |  | -26.6003 | -27.475 | -27.472 | -27.47 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4634 | 0.0070702 | -0.016623 | -0.00062278 | 0.0033772 | 0.010377 |  |  | 0.86303 | -0.011623 | -0.0086228 | -0.0066228
177.8279 | -27.48 | -27.465 | -27.47 | -27.455 |  |  | -26.5997 | -27.478 | -27.474 | -27.48 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4715 | 0.0070702 | -0.0085373 | 0.0064627 | 0.0014627 | 0.016463 |  |  | 0.87175 | -0.0065373 | -0.0025373 | -0.0085373
188.3649 | -27.48 | -27.47 | -27.47 | -27.456 |  |  | -26.5992 | -27.478 | -27.477 | -27.48 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4728 | 0.0070702 | -0.0072482 | 0.0027518 | 0.0027518 | 0.016752 |  |  | 0.87356 | -0.0052482 | -0.0042482 | -0.0072482
199.5262 | -27.48 | -27.474 | -27.47 | -27.458 | -27.479 | -27.483 | -26.5977 | -27.479 | -27.479 | -27.48 | -27.487 | -27.476 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.049 | -27.4707 | 0.0070528 | -0.0093053 | -0.0033053 | 0.00069471 | 0.012695 | -0.0083053 | -0.012305 | 0.87302 | -0.0083053 | -0.0083053 | -0.0093053 | -0.016305 | -0.0053053
211.3489 | -27.48 | -27.476 | -27.47 | -27.459 |  |  | -26.5971 | -27.48 | -27.481 | -27.48 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4715 | 0.0070528 | -0.0085103 | -0.0045103 | 0.0014897 | 0.01249 |  |  | 0.87442 | -0.0085103 | -0.0095103 | -0.0085103
223.8721 | -27.48 | -27.477 | -27.47 | -27.461 |  |  | -26.5964 | -27.482 | -27.482 | -27.48 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4708 | 0.0070528 | -0.0091987 | -0.0061987 | 0.00080132 | 0.0098013 |  |  | 0.87437 | -0.011199 | -0.011199 | -0.0091987
237.1374 | -27.49 | -27.477 | -27.47 | -27.462 |  |  | -26.5914 | -27.483 | -27.484 | -27.48 |  |  | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4771 | 0.0070528 | -0.012931 | 6.9409e-05 | 0.0070694 | 0.015069 |  |  | 0.88566 | -0.0059306 | -0.0069306 | -0.0029306
251.1886 | -27.49 | -27.479 | -27.48 | -27.469 | -27.485 | -27.484 | -26.5907 | -27.476 | -27.486 | -27.49 | -27.492 | -27.482 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0479 | -27.4832 | 0.0070528 | -0.0067524 | 0.0042476 | 0.0032476 | 0.014248 | -0.0017524 | -0.00075242 | 0.89254 | 0.0072476 | -0.0027524 | -0.0067524 | -0.0087524 | 0.0012476
266.0725 | -27.49 | -27.481 | -27.48 | -27.465 |  |  | -26.5899 | -27.478 | -27.487 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.48 | 0.0070528 | -0.0099728 | -0.00097278 | 2.7222e-05 | 0.015027 |  |  | 0.89015 | 0.0020272 | -0.0069728 | -0.0099728
281.8383 | -27.49 | -27.48 | -27.48 | -27.466 |  |  | -26.5889 | -27.479 | -27.489 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4818 | 0.0069484 | -0.0082304 | 0.0017696 | 0.0017696 | 0.01577 |  |  | 0.89285 | 0.0027696 | -0.0072304 | -0.0082304
298.5383 | -27.49 | -27.483 | -27.48 | -27.467 |  |  | -26.5879 | -27.48 | -27.491 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4825 | 0.0069484 | -0.0074743 | -0.00047425 | 0.0025257 | 0.015526 |  |  | 0.89463 | 0.0025257 | -0.0084743 | -0.0074743
316.2278 | -27.49 | -27.484 | -27.48 | -27.468 | -27.488 | -27.489 | -26.5867 | -27.482 | -27.491 | -27.49 | -27.495 | -27.486 | 0.03 | 0.03 | 0.03 | 0.044 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0479 | -27.4785 | 0.0069484 | -0.011548 | -0.0055479 | -0.0015479 | 0.010452 | -0.0095479 | -0.010548 | 0.8918 | -0.0035479 | -0.012548 | -0.011548 | -0.016548 | -0.0075479
334.9654 | -27.49 | -27.487 | -27.48 | -27.469 |  |  | -26.5855 | -27.483 | -27.492 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4807 | 0.0069484 | -0.0092714 | -0.0062714 | 0.00072861 | 0.011729 |  |  | 0.89528 | -0.0022714 | -0.011271 | -0.0092714
354.8134 | -27.49 | -27.487 | -27.48 | -27.47 |  |  | -26.584 | -27.484 | -27.493 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4811 | 0.0069484 | -0.0089363 | -0.0059363 | 0.0010637 | 0.011064 |  |  | 0.89705 | -0.0029363 | -0.011936 | -0.0089363
375.8374 | -27.49 | -27.487 | -27.48 | -27.471 |  |  | -26.5823 | -27.484 | -27.493 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4814 | 0.0069568 | -0.0086013 | -0.0056013 | 0.0013987 | 0.010399 |  |  | 0.89911 | -0.0026013 | -0.011601 | -0.0086013
398.1072 | -27.49 | -27.489 | -27.48 | -27.47 | -27.49 | -27.491 | -26.5807 | -27.485 | -27.494 | -27.49 | -27.497 | -27.488 | 0.03 | 0.03 | 0.03 | 0.044 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.048 | -27.4821 | 0.0069568 | -0.0078831 | -0.0068831 | 0.0021169 | 0.012117 | -0.0078831 | -0.0088831 | 0.90141 | -0.0028831 | -0.011883 | -0.0078831 | -0.014883 | -0.0058831
421.6965 | -27.49 | -27.488 | -27.48 | -27.471 |  |  | -26.5787 | -27.486 | -27.495 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4821 | 0.0069568 | -0.0078675 | -0.0058675 | 0.0021325 | 0.011132 |  |  | 0.90342 | -0.0038675 | -0.012868 | -0.0078675
446.6836 | -27.49 | -27.488 | -27.48 | -27.471 |  |  | -26.5765 | -27.486 | -27.495 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4816 | 0.0069568 | -0.0083705 | -0.0063705 | 0.0016295 | 0.010629 |  |  | 0.90512 | -0.0043705 | -0.013371 | -0.0083705
473.1513 | -27.5 | -27.489 | -27.48 | -27.471 |  |  | -26.574 | -27.486 | -27.495 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4866 | 0.0069568 | -0.01344 | -0.0024402 | 0.0065598 | 0.01556 |  |  | 0.91255 | 0.0005598 | -0.0084402 | -0.0034402
501.1872 | -27.49 | -27.487 | -27.48 | -27.471 | -27.489 | -27.487 | -26.5715 | -27.486 | -27.494 | -27.49 | -27.497 | -27.488 | 0.03 | 0.03 | 0.03 | 0.044 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0475 | -27.4807 | 0.0069568 | -0.0092782 | -0.0062782 | 0.00072175 | 0.0097218 | -0.0082782 | -0.0062782 | 0.90925 | -0.0052782 | -0.013278 | -0.0092782 | -0.016278 | -0.0072782
530.8844 | -27.49 | -27.483 | -27.48 | -27.471 |  |  | -26.5685 | -27.486 | -27.494 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4803 | 0.0069568 | -0.0096674 | -0.0026674 | 0.00033256 | 0.0093326 |  |  | 0.91188 | -0.0056674 | -0.013667 | -0.0096674
562.3413 | -27.5 | -27.484 | -27.48 | -27.47 |  |  | -26.5649 | -27.485 | -27.495 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4852 | 0.0069568 | -0.014831 | 0.0011686 | 0.0051686 | 0.015169 |  |  | 0.92025 | 0.00016861 | -0.0098314 | -0.0048314
595.6621 | -27.49 | -27.483 | -27.48 | -27.47 |  |  | -26.5612 | -27.484 | -27.494 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4778 | 0.0069568 | -0.012186 | -0.0051859 | -0.0021859 | 0.0078141 |  |  | 0.91658 | -0.0061859 | -0.016186 | -0.012186
630.9573 | -27.49 | -27.482 | -27.48 | -27.469 | -27.484 | -27.487 | -26.5574 | -27.484 | -27.492 | -27.49 | -27.495 | -27.485 | 0.03 | 0.03 | 0.03 | 0.044 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0474 | -27.4763 | 0.0069568 | -0.013735 | -0.0057351 | -0.0037351 | 0.0072649 | -0.0077351 | -0.010735 | 0.91891 | -0.0077351 | -0.015735 | -0.013735 | -0.018735 | -0.0087351
668.3439 | -27.49 | -27.48 | -27.48 | -27.468 |  |  | -26.5529 | -27.483 | -27.491 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4753 | 0.0069568 | -0.014678 | -0.0046778 | -0.0046778 | 0.0073222 |  |  | 0.92238 | -0.0076778 | -0.015678 | -0.014678
707.9458 | -27.49 | -27.478 | -27.48 | -27.466 |  |  | -26.5479 | -27.481 | -27.488 | -27.49 |  |  | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4777 | 0.0069568 | -0.012334 | -0.00033439 | -0.0023344 | 0.011666 |  |  | 0.92973 | -0.0033344 | -0.010334 | -0.012334
749.8942 | -27.49 | -27.475 | -27.48 | -27.464 |  |  | -26.5423 | -27.478 | -27.486 | -27.48 |  |  | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4784 | 0.0069568 | -0.011648 | 0.0033521 | -0.0016479 | 0.014352 |  |  | 0.93601 | 0.00035207 | -0.0076479 | -0.0016479
794.3282 | -27.48 | -27.471 | -27.47 | -27.461 | -27.477 | -27.476 | -26.5361 | -27.476 | -27.484 | -27.48 | -27.487 | -27.476 | 0.03 | 0.03 | 0.03 | 0.045 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0475 | -27.4661 | 0.0070358 | -0.013906 | -0.0049062 | -0.0039062 | 0.0050938 | -0.010906 | -0.0099062 | 0.92998 | -0.0099062 | -0.017906 | -0.013906 | -0.020906 | -0.0099062
841.3951 | -27.48 | -27.471 | -27.47 | -27.459 |  |  | -26.5284 | -27.474 | -27.481 | -27.48 |  |  | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4672 | 0.0070358 | -0.012762 | -0.0037621 | -0.0027621 | 0.0082379 |  |  | 0.93882 | -0.0067621 | -0.013762 | -0.012762
891.2509 | -27.48 | -27.467 | -27.47 | -27.455 |  |  | -26.5197 | -27.471 | -27.477 | -27.48 |  |  | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4668 | 0.0070358 | -0.013244 | -0.00024399 | -0.003244 | 0.011756 |  |  | 0.9471 | -0.004244 | -0.010244 | -0.013244
944.0609 | -27.47 | -27.464 | -27.46 | -27.451 |  |  | -26.5111 | -27.467 | -27.473 | -27.47 |  |  | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4575 | 0.0070358 | -0.0125 | -0.0065002 | -0.0025002 | 0.0064998 |  |  | 0.94639 | -0.0095002 | -0.0155 | -0.0125
1000 | -27.47 | -27.456 | -27.46 | -27.447 | -27.461 | -27.459 | -26.5023 | -27.463 | -27.469 | -27.47 | -27.471 | -27.461 | 0.03 | 0.03 | 0.03 | 0.045 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0474 | -27.4603 | 0.0070358 | -0.0096964 | 0.0043036 | 0.00030365 | 0.013304 | -0.00069635 | 0.0013036 | 0.95802 | -0.0026964 | -0.0086964 | -0.0096964 | -0.010696 | -0.00069635
1059.2537 | -27.46 | -27.451 | -27.45 | -27.442 |  |  | -26.4918 | -27.458 | -27.463 | -27.46 |  |  | 0.03 | 0.03 | 0.03 | 0.046 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4483 | 0.0070358 | -0.011741 | -0.0027405 | -0.0017405 | 0.0062595 |  |  | 0.9565 | -0.0097405 | -0.014741 | -0.011741
1122.0185 | -27.46 | -27.447 | -27.45 | -27.435 |  |  | -26.4807 | -27.452 | -27.457 | -27.46 |  |  | 0.03 | 0.03 | 0.03 | 0.046 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4505 | 0.0070372 | -0.0094899 | 0.0035101 | 0.00051014 | 0.01551 |  |  | 0.96982 | -0.0014899 | -0.0064899 | -0.0094899
1188.5022 | -27.45 | -27.434 | -27.44 | -27.429 |  |  | -26.4685 | -27.445 | -27.452 | -27.45 |  |  | 0.03 | 0.03 | 0.03 | 0.046 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4376 | 0.0070385 | -0.012437 | 0.0035625 | -0.0024375 | 0.0085625 |  |  | 0.96907 | -0.0074375 | -0.014437 | -0.012437
1258.9254 | -27.44 | -27.428 | -27.43 | -27.421 | -27.435 | -27.433 | -26.4547 | -27.438 | -27.444 | -27.44 | -27.446 | -27.434 | 0.03 | 0.03 | 0.03 | 0.046 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0475 | -27.4273 | 0.0070397 | -0.01273 | -0.00073046 | -0.0027305 | 0.0062695 | -0.0077305 | -0.0057305 | 0.97255 | -0.01073 | -0.01673 | -0.01273 | -0.01873 | -0.0067305
1333.5214 | -27.43 | -27.419 | -27.43 | -27.412 |  |  | -26.4395 | -27.431 | -27.435 | -27.43 |  |  | 0.03 | 0.03 | 0.03 | 0.046 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -27.4242 | 0.0070397 | -0.0057758 | 0.0052242 | -0.0057758 | 0.012224 |  |  | 0.98476 | -0.0067758 | -0.010776 | -0.0057758
1412.5375 | -27.43 | -27.412 | -27.42 | -27.403 |  |  | -26.4224 | -27.423 | -27.425 | -27.42 |  |  | 0.03 | 0.03 | 0.03 | 0.046 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -27.42 | 0.0070397 | -0.0099595 | 0.0080405 | 4.0473e-05 | 0.01704 |  |  | 0.99759 | -0.0029595 | -0.0049595 | 4.0473e-05
1496.2357 | -27.41 | -27.404 | -27.4 | -27.392 |  |  | -26.403 | -27.413 | -27.414 | -27.41 |  |  | 0.03 | 0.03 | 0.03 | 0.046 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -27.3955 | 0.0070397 | -0.014507 | -0.0085072 | -0.0045072 | 0.0034928 |  |  | 0.99249 | -0.017507 | -0.018507 | -0.014507
1584.8932 | -27.4 | -27.393 | -27.39 | -27.379 | -27.392 | -27.388 | -26.3823 | -27.401 | -27.402 | -27.4 | -27.405 | -27.391 | 0.03 | 0.03 | 0.03 | 0.046 | 0.04 | 0.05 | 0.08 | 0.041 | 0.04 | 0.03 | 0.03 | 0.0475 | -27.3885 | 0.0080944 | -0.011547 | -0.0045472 | -0.0015472 | 0.0094528 | -0.0035472 | 0.0004528 | 1.0061 | -0.012547 | -0.013547 | -0.011547 | -0.016547 | -0.0025472
1678.804 | -27.39 | -27.38 | -27.38 | -27.366 |  |  | -26.3592 | -27.388 | -27.388 | -27.38 |  |  | 0.03 | 0.03 | 0.03 | 0.046 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -27.3798 | 0.0080944 | -0.01015 | -0.00015017 | -0.00015017 | 0.01385 |  |  | 1.0207 | -0.0081502 | -0.0081502 | -0.00015017
1778.2794 | -27.37 | -27.362 | -27.36 | -27.35 |  |  | -26.3334 | -27.373 | -27.372 | -27.37 |  |  | 0.03 | 0.03 | 0.03 | 0.047 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -27.3589 | 0.0080944 | -0.011077 | -0.0030768 | -0.0010768 | 0.0089232 |  |  | 1.0255 | -0.014077 | -0.013077 | -0.011077
1883.6491 | -27.35 | -27.343 | -27.35 | -27.332 |  |  | -26.305 | -27.355 | -27.355 | -27.35 |  |  | 0.03 | 0.03 | 0.03 | 0.047 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -27.34 | 0.0080944 | -0.009997 | -0.002997 | -0.009997 | 0.008003 |  |  | 1.035 | -0.014997 | -0.014997 | -0.009997
1995.2623 | -27.33 | -27.322 | -27.33 | -27.312 | -27.325 | -27.321 | -26.2752 | -27.338 | -27.335 | -27.33 | -27.338 | -27.322 | 0.03 | 0.03 | 0.03 | 0.047 | 0.04 | 0.05 | 0.08 | 0.041 | 0.04 | 0.03 | 0.03 | 0.0486 | -27.3244 | 0.0080944 | -0.0056192 | 0.0023808 | -0.0056192 | 0.012381 | -0.00061916 | 0.0033808 | 1.0492 | -0.013619 | -0.010619 | -0.0056192 | -0.013619 | 0.0023808
2113.489 | -27.31 | -27.301 | -27.3 | -27.29 |  |  | -26.2404 | -27.318 | -27.313 | -27.31 |  |  | 0.03 | 0.035 | 0.03 | 0.047 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -27.2996 | 0.0092484 | -0.010409 | -0.0014086 | -0.00040862 | 0.0095914 |  |  | 1.0592 | -0.018409 | -0.013409 | -0.010409
2238.7211 | -27.29 | -27.276 | -27.28 | -27.266 |  |  | -26.2016 | -27.297 | -27.289 | -27.29 |  |  | 0.03 | 0.035 | 0.03 | 0.048 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -27.2779 | 0.0092484 | -0.012148 | 0.0018522 | -0.0021478 | 0.011852 |  |  | 1.0762 | -0.019148 | -0.011148 | -0.012148
2371.3737 | -27.26 | -27.251 | -27.25 | -27.237 |  |  | -26.1591 | -27.271 | -27.262 | -27.26 |  |  | 0.03 | 0.035 | 0.03 | 0.048 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -27.2495 | 0.0092484 | -0.010536 | -0.0015362 | -0.00053617 | 0.012464 |  |  | 1.0903 | -0.021536 | -0.012536 | -0.010536
2511.8864 | -27.23 | -27.219 | -27.22 | -27.209 | -27.221 | -27.219 | -26.1117 | -27.243 | -27.232 | -27.23 | -27.236 | -27.213 | 0.03 | 0.035 | 0.03 | 0.049 | 0.04 | 0.06 | 0.08 | 0.041 | 0.04 | 0.03 | 0.03 | 0.0487 | -27.2205 | 0.0092484 | -0.0095139 | 0.0014861 | 0.0004861 | 0.011486 | -0.0005139 | 0.0014861 | 1.1087 | -0.022514 | -0.011514 | -0.0095139 | -0.015514 | 0.0074861
2660.7251 | -27.2 | -27.187 | -27.19 | -27.176 |  |  | -26.0596 | -27.211 | -27.199 | -27.2 |  |  | 0.04 | 0.04 | 0.03 | 0.049 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -27.184 | 0.012211 | -0.01601 | -0.0030104 | -0.0060104 | 0.0079896 |  |  | 1.1244 | -0.02701 | -0.01501 | -0.01601
2818.3829 | -27.16 | -27.152 | -27.15 | -27.138 |  |  | -26.0032 | -27.172 | -27.163 | -27.16 |  |  | 0.04 | 0.04 | 0.03 | 0.05 |  |  | 0.08 | 0.042 | 0.04 | 0.03 |  |  | -27.146 | 0.012208 | -0.013951 | -0.0059512 | -0.0039512 | 0.0080488 |  |  | 1.1428 | -0.025951 | -0.016951 | -0.013951
2985.3826 | -27.12 | -27.111 | -27.11 | -27.098 |  |  | -25.9414 | -27.136 | -27.123 | -27.12 |  |  | 0.04 | 0.04 | 0.03 | 0.051 |  |  | 0.08 | 0.042 | 0.04 | 0.03 |  |  | -27.1031 | 0.012208 | -0.016874 | -0.0078744 | -0.0068744 | 0.0051256 |  |  | 1.1617 | -0.032874 | -0.019874 | -0.016874
3162.2777 | -27.08 | -27.069 | -27.07 | -27.054 | -27.068 | -27.065 | -25.8739 | -27.095 | -27.079 | -27.08 | -27.086 | -27.056 | 0.04 | 0.04 | 0.03 | 0.053 | 0.04 | 0.06 | 0.08 | 0.042 | 0.04 | 0.03 | 0.03 | 0.0492 | -27.0642 | 0.012207 | -0.015832 | -0.0048321 | -0.0058321 | 0.010168 | -0.0038321 | -0.00083213 | 1.1902 | -0.030832 | -0.014832 | -0.015832 | -0.021832 | 0.0081679
3349.6544 | -27.03 | -27.021 | -27.03 | -27.003 |  |  | -25.8012 | -27.046 | -27.032 | -27.03 |  |  | 0.04 | 0.04 | 0.03 | 0.055 |  |  | 0.08 | 0.043 | 0.05 | 0.03 |  |  | -27.0185 | 0.013869 | -0.011524 | -0.0025236 | -0.011524 | 0.015476 |  |  | 1.2172 | -0.027524 | -0.013524 | -0.011524
3548.1339 | -26.98 | -26.972 | -26.98 | -26.957 |  |  | -25.7234 | -27.005 | -26.981 | -26.98 |  |  | 0.04 | 0.04 | 0.03 | 0.057 |  |  | 0.08 | 0.043 | 0.05 | 0.03 |  |  | -26.9696 | 0.013867 | -0.010436 | -0.0024356 | -0.010436 | 0.012564 |  |  | 1.2462 | -0.035436 | -0.011436 | -0.010436
3758.374 | -26.93 | -26.919 | -26.92 | -26.901 |  |  | -25.64 | -26.952 | -26.927 | -26.93 |  |  | 0.04 | 0.04 | 0.03 | 0.06 |  |  | 0.08 | 0.044 | 0.05 | 0.03 |  |  | -26.9127 | 0.013848 | -0.017286 | -0.0062856 | -0.0072856 | 0.011714 |  |  | 1.2727 | -0.039286 | -0.014286 | -0.017286
3981.0717 | -26.87 | -26.865 | -26.87 | -26.844 | -26.862 | -26.855 | -25.552 | -26.897 | -26.87 | -26.87 | -26.883 | -26.842 | 0.04 | 0.04 | 0.03 | 0.064 | 0.04 | 0.08 | 0.09 | 0.045 | 0.05 | 0.03 | 0.05 | 0.0625 | -26.8561 | 0.013846 | -0.013885 | -0.0088846 | -0.013885 | 0.012115 | -0.0058846 | 0.0011154 | 1.3041 | -0.040885 | -0.013885 | -0.013885 | -0.026885 | 0.014115
4216.965 | -26.81 | -26.809 | -26.81 | -26.786 |  |  | -25.4607 | -26.839 | -26.813 | -26.81 |  |  | 0.05 | 0.05 | 0.03 | 0.068 |  |  | 0.09 | 0.046 | 0.05 | 0.03 |  |  | -26.7955 | 0.0157 | -0.014548 | -0.013548 | -0.014548 | 0.0094525 |  |  | 1.3348 | -0.043548 | -0.017548 | -0.014548
4466.8359 | -26.75 | -26.744 | -26.75 | -26.727 |  |  | -25.368 | -26.777 | -26.755 | -26.75 |  |  | 0.05 | 0.05 | 0.03 | 0.073 |  |  | 0.09 | 0.048 | 0.06 | 0.03 |  |  | -26.7334 | 0.015696 | -0.016642 | -0.010642 | -0.016642 | 0.006358 |  |  | 1.3654 | -0.043642 | -0.021642 | -0.016642
4731.5126 | -26.7 | -26.688 | -26.7 | -26.671 |  |  | -25.2771 | -26.72 | -26.696 | -26.7 |  |  | 0.05 | 0.05 | 0.03 | 0.078 |  |  | 0.09 | 0.049 | 0.06 | 0.03 |  |  | -26.6834 | 0.015678 | -0.01656 | -0.0045605 | -0.01656 | 0.01244 |  |  | 1.4063 | -0.03656 | -0.01256 | -0.01656
5011.8723 | -26.65 | -26.638 | -26.63 | -26.618 | -26.641 | -26.635 | -25.1933 | -26.668 | -26.647 | -26.65 | -26.667 | -26.614 | 0.05 | 0.05 | 0.02 | 0.085 | 0.04 | 0.1 | 0.09 | 0.05 | 0.07 | 0.05 | 0.05 | 0.0732 | -26.6288 | 0.014079 | -0.021203 | -0.0092034 | -0.0012034 | 0.010797 | -0.012203 | -0.0062034 | 1.4355 | -0.039203 | -0.018203 | -0.021203 | -0.038203 | 0.014797
5308.8444 | -26.61 | -26.599 | -26.6 | -26.578 |  |  | -25.1229 | -26.621 | -26.61 | -26.61 |  |  | 0.05 | 0.06 | 0.03 | 0.091 |  |  | 0.1 | 0.052 | 0.07 | 0.05 |  |  | -26.5915 | 0.014773 | -0.018478 | -0.0074783 | -0.0084783 | 0.013522 |  |  | 1.4687 | -0.029478 | -0.018478 | -0.018478
5623.4133 | -26.58 | -26.574 | -26.57 | -26.547 |  |  | -25.0761 | -26.599 | -26.583 | -26.59 |  |  | 0.05 | 0.06 | 0.03 | 0.096 |  |  | 0.1 | 0.053 | 0.08 | 0.05 |  |  | -26.5667 | 0.014763 | -0.013334 | -0.0073335 | -0.0033335 | 0.019666 |  |  | 1.4906 | -0.032334 | -0.016334 | -0.023334
5956.6214 | -26.58 | -26.58 | -26.57 | -26.55 |  |  | -25.067 | -26.592 | -26.581 | -26.59 |  |  | 0.05 | 0.06 | 0.03 | 0.1 |  |  | 0.1 | 0.055 | 0.08 | 0.05 |  |  | -26.5577 | 0.017142 | -0.022277 | -0.022277 | -0.012277 | 0.0077228 |  |  | 1.4907 | -0.034277 | -0.023277 | -0.032277
6309.5734 | -26.62 | -26.61 | -26.61 | -26.589 | -26.606 | -26.629 | -25.1169 | -26.619 | -26.618 | -26.62 | -26.653 | -26.582 | 0.05 | 0.06 | 0.03 | 0.102 | 0.05 | 0.1 | 0.11 | 0.056 | 0.08 | 0.05 | 0.05 | 0.0786 | -26.5979 | 0.014756 | -0.022118 | -0.012118 | -0.012118 | 0.0088817 | -0.0081183 | -0.031118 | 1.4809 | -0.021118 | -0.020118 | -0.022118 | -0.055118 | 0.015882
6683.4392 | -26.7 | -26.696 | -26.69 | -26.658 |  |  | -25.2291 | -26.694 | -26.701 | -26.71 |  |  | 0.06 | 0.07 | 0.03 | 0.098 |  |  | 0.11 | 0.057 | 0.08 | 0.05 |  |  | -26.677 | 0.01471 | -0.023029 | -0.019029 | -0.013029 | 0.018971 |  |  | 1.4479 | -0.017029 | -0.024029 | -0.033029
7079.4578 | -26.85 | -26.841 | -26.84 | -26.815 |  |  | -25.4462 | -26.83 | -26.849 | -26.85 |  |  | 0.06 | 0.07 | 0.03 | 0.094 |  |  | 0.11 | 0.058 | 0.09 | 0.05 |  |  | -26.8252 | 0.017617 | -0.024813 | -0.015813 | -0.014813 | 0.010187 |  |  | 1.379 | -0.0048127 | -0.023813 | -0.024813
7498.9421 | -27.08 | -27.072 | -27.08 | -27.049 |  |  | -25.7856 | -27.054 | -27.079 | -27.08 |  |  | 0.06 | 0.07 | 0.03 | 0.087 |  |  | 0.11 | 0.058 | 0.09 | 0.05 |  |  | -27.0681 | 0.017206 | -0.011892 | -0.0038924 | -0.011892 | 0.019108 |  |  | 1.2825 | 0.014108 | -0.010892 | -0.011892
7943.2823 | -27.4 | -27.41 | -27.42 | -27.388 | -27.429 | -27.443 | -26.2673 | -27.389 | -27.416 | -27.42 | -27.477 | -27.374 | 0.06 | 0.07 | 0.04 | 0.082 | 0.05 | 0.12 | 0.14 | 0.057 | 0.1 | 0.05 | 0.07 | 0.0945 | -27.4108 | 0.020279 | 0.010812 | 0.00081233 | -0.0091877 | 0.022812 | -0.018188 | -0.032188 | 1.1435 | 0.021812 | -0.0051877 | -0.0091877 | -0.066188 | 0.036812
8413.9514 | -27.86 | -27.877 | -27.88 | -27.849 |  |  | -26.92 | -27.829 | -27.872 | -27.88 |  |  | 0.11 | 0.08 | 0.06 | 0.08 |  |  | 0.14 | 0.056 | 0.12 | 0.08 |  |  | -27.8731 | 0.032081 | 0.013146 | -0.0038537 | -0.0068537 | 0.024146 |  |  | 0.95314 | 0.044146 | 0.0011463 | -0.0068537
8912.5094 | -28.45 | -28.483 | -28.49 | -28.448 |  |  | -27.7351 | -28.437 | -28.467 | -28.47 |  |  | 0.11 | 0.09 | 0.06 | 0.084 |  |  | 0.16 | 0.054 | 0.15 | 0.08 |  |  | -28.4854 | 0.031814 | 0.035411 | 0.002411 | -0.004589 | 0.037411 |  |  | 0.7503 | 0.048411 | 0.018411 | 0.015411
9440.6088 | Sensitivity Level (4160 2036126) dB re 1 V/Pa UK: -29.14; Sensitivity Level (4160 2036126) dB re 1 V/Pa FR: -29.181; Sensitivity Level (4160 2036126) dB re 1 V/Pa IT: -29.18; Sensitivity Level (4160 2036126) dB re 1 V/Pa CH: -29.141; Sensitivity Level (4160 2036126) dB re 1 V/Pa RS: -28.7186; Sensitivity Level (4160 2036126) dB re 1 V/Pa SE: -29.188; Sensitivity Level (4160 2036126) dB re 1 V/Pa DE: -29.18; Expanded Uncertainties (dB) UK: 0.11; Expanded Uncertainties (dB) FR: 0.09; Expanded Uncertainties (dB) IT: 0.06; Expanded Uncertainties (dB) CH: 0.094; Expanded Uncertainties (dB) RS: 0.18; Expanded Uncertainties (dB) SE: 0.17; Expanded Uncertainties (dB) DE: 0.08 | -29.1806 | 0.032096 | 0.040564 | -0.00043646 | 0.00056354 | 0.039564 |  |  | 0.46192 |  | -0.0074365 | 0.00056354
10000 | -29.97 | -29.998 | -30.02 | -29.965 | -30.057 | -30.079 | -29.8481 | -30.071 | -30.03 | -30.01 | -30.12 | -29.978 | 0.11 | 0.1 | 0.06 | 0.11 | 0.06 | 0.14 | 0.2 | 0.102 | 0.2 | 0.08 | 0.08 | 0.1488 | -30.0319 | 0.031618 | 0.061874 | 0.033874 | 0.011874 | 0.066874 | -0.025126 | -0.047126 | 0.18376 | -0.039126 | 0.0018737 | 0.021874 | -0.088126 | 0.053874

## 2036126 DoE Level
| Linking DoEs from CCAUV.A-K5 |  | 4160 2036126 | [11 empty] | Linking U(DoE)s from CCAUV.A-K5
|  |  | DoE (dB) | [13 empty] | U(DoE) (dB)
Frequency (Hz) | UK | IT | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG | UK | IT | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG
1.9953 | UK: -0.042047; IT: 0.017953; DoE (dB) UK: 0.034463; DoE (dB) FR: -0.073445; DoE (dB) IT: -0.022829; DoE (dB) CH: -0.019445; DoE (dB) AT: -0.028445; DoE (dB) DE: -0.065445; DoE (dB) UK: 0.26713; DoE (dB) IT: 0.14476; U(DoE) (dB) UK: 0.25331; U(DoE) (dB) FR: 0.15826; U(DoE) (dB) IT: 0.13753; U(DoE) (dB) CH: 0.14688; U(DoE) (dB) AT: 0.095106; U(DoE) (dB) DE: 0.24155
2.5119 | UK: 0.033818; IT: 0.023818; DoE (dB) UK: 0.041645; DoE (dB) FR: -0.013289; DoE (dB) IT: 0.023527; DoE (dB) CH: 0.0087109; DoE (dB) AT: -0.0012891; DoE (dB) DE: -0.027289; DoE (dB) UK: 0.26755; DoE (dB) IT: 0.093173; U(DoE) (dB) UK: 0.24965; U(DoE) (dB) FR: 0.094637; U(DoE) (dB) IT: 0.093104; U(DoE) (dB) CH: 0.11564; U(DoE) (dB) AT: 0.069715; U(DoE) (dB) DE: 0.23549
3.1623 | UK: -0.00055047; IT: 0.02945; DoE (dB) UK: 0.046897; DoE (dB) FR: -0.029677; DoE (dB) IT: 0.028338; DoE (dB) CH: -0.0046765; DoE (dB) AT: -0.020677; DoE (dB) DE: -0.036677; DoE (dB) UK: 0.26848; DoE (dB) IT: 0.074711; U(DoE) (dB) UK: 0.25302; U(DoE) (dB) FR: 0.085875; U(DoE) (dB) IT: 0.074681; U(DoE) (dB) CH: 0.10569; U(DoE) (dB) AT: 0.075568; U(DoE) (dB) DE: 0.23806
3.9811 | UK: 0.021366; IT: 0.021366; DoE (dB) UK: 0.037799; DoE (dB) FR: -0.025552; DoE (dB) IT: 0.019083; DoE (dB) CH: -0.0055522; DoE (dB) AT: -0.022552; DoE (dB) DE: -0.029552; DoE (dB) UK: 0.11773; DoE (dB) IT: 0.055315; U(DoE) (dB) UK: 0.088341; U(DoE) (dB) FR: 0.084216; U(DoE) (dB) IT: 0.054248; U(DoE) (dB) CH: 0.098875; U(DoE) (dB) AT: 0.080258; U(DoE) (dB) DE: 0.13781
5.0119 | UK: 0.010488; IT: 0.020488; DoE (dB) UK: 0.022307; DoE (dB) FR: -0.028796; DoE (dB) IT: 0.019325; DoE (dB) CH: -0.0067959; DoE (dB) AT: -0.026796; DoE (dB) DE: -0.035796; DoE (dB) UK: 0.11817; DoE (dB) IT: 0.045439; U(DoE) (dB) UK: 0.086057; U(DoE) (dB) FR: 0.074719; U(DoE) (dB) IT: 0.044736; U(DoE) (dB) CH: 0.089173; U(DoE) (dB) AT: 0.075259; U(DoE) (dB) DE: 0.13558
6.3096 | UK: -0.011393; IT: 0.018607; DoE (dB) UK: 0.014038; DoE (dB) FR: -0.031896; DoE (dB) IT: 0.016963; DoE (dB) CH: -0.011896; DoE (dB) AT: -0.031896; DoE (dB) DE: -0.041896; DoE (dB) UK: 0.11862; DoE (dB) IT: 0.035641; U(DoE) (dB) UK: 0.08358; U(DoE) (dB) FR: 0.070322; U(DoE) (dB) IT: 0.035224; U(DoE) (dB) CH: 0.080506; U(DoE) (dB) AT: 0.070896; U(DoE) (dB) DE: 0.13321
7.9433 | UK: -0.0050233; IT: 0.014977; DoE (dB) UK: 0.015315; DoE (dB) FR: -0.027306; DoE (dB) IT: 0.013358; DoE (dB) CH: -0.011306; DoE (dB) AT: -0.032306; DoE (dB) DE: -0.031306; DoE (dB) UK: 0.11899; DoE (dB) IT: 0.036858; U(DoE) (dB) UK: 0.07374; U(DoE) (dB) FR: 0.062632; U(DoE) (dB) IT: 0.036102; U(DoE) (dB) CH: 0.074991; U(DoE) (dB) AT: 0.067991; U(DoE) (dB) DE: 0.097071
10 | -0.0027733 | 0.017227 | 0.017573 | -0.016047 | 0.015617 | -0.002047 |  |  |  | -0.023047 |  | -0.019047 | -0.042047 |  | 0.11907 | 0.037104 | 0.073812 | 0.062677 | 0.03636 | 0.073024 |  |  |  | 0.06745 |  | 0.0971 | 0.0971
12.5893 | -0.012138 | 0.017862 | 0.016902 | -0.0092795 | 0.016532 | 0.0027205 |  |  |  | -0.019279 |  | -0.018279 | -0.034279 |  | 0.11924 | 0.026796 | 0.069754 | 0.055825 | 0.026428 | 0.065821 |  |  |  | 0.061135 |  | 0.092825 | 0.092825
15.8489 | -0.0082591 | 0.011741 | 0.0057977 | -0.014916 | 0.010876 | -0.00091592 |  |  |  | -0.022916 |  | -0.021916 | -0.035916 |  | 0.10924 | 0.027073 | 0.062479 | 0.056685 | 0.026506 | 0.065186 |  |  |  | 0.061921 |  | 0.069377 | 0.093345
19.9526 | -0.0019418 | 0.0080582 | -0.0019418 | -0.012942 | 0.0080582 | 0.0030582 | -0.033942 | -0.038942 | 0.75382 | -0.019942 | 0.0020582 | -0.021942 | -0.029942 | -0.0059418 | 0.027853 | 0.027853 | 0.027195 | 0.033602 | 0.027195 | 0.04653 | 0.042768 | 0.071618 | 0.091264 | 0.041834 | 0.052241 | 0.052241 | 0.081419 | 0.065084
21.1349 | -0.0027511 | 0.0072489 | -0.0052511 | -0.014751 | 0.0097489 | 0.0012489 |  |  | 0.7549 | -0.023751 | -0.0037511 | -0.017751 |  |  | 0.027769 | 0.027769 | 0.027109 | 0.033649 | 0.027109 | 0.046565 |  |  | 0.091281 | 0.041872 | 0.052271 | 0.052271
22.3872 | -0.0072713 | 0.012729 | -0.0047713 | -0.012271 | 0.010229 | 0.0027287 |  |  | 0.75771 | -0.024271 | -0.0022713 | -0.022271 |  |  | 0.027769 | 0.027769 | 0.027109 | 0.033649 | 0.027109 | 0.046565 |  |  | 0.091281 | 0.041872 | 0.042805 | 0.052271
23.7137 | -0.0082775 | 0.011722 | -0.0057775 | -0.0072775 | 0.0092225 | 0.0047225 |  |  | 0.76421 | -0.020278 | 0.00072248 | -0.013278 |  |  | 0.027769 | 0.027769 | 0.027109 | 0.033649 | 0.027109 | 0.046565 |  |  | 0.091281 | 0.040942 | 0.042805 | 0.052271
25.1189 | -0.0093484 | 0.010652 | -0.0068484 | -0.016348 | 0.0081516 | -0.00034841 | -0.037348 | -0.044348 | 0.76051 | -0.026348 | -0.0053484 | -0.024348 | -0.033348 | -0.015348 | 0.028092 | 0.028092 | 0.027433 | 0.033483 | 0.027433 | 0.046444 | 0.042674 | 0.061815 | 0.09122 | 0.040805 | 0.042674 | 0.052164 | 0.052164 | 0.06036
26.6073 | -0.0038456 | 0.0061544 | -0.0038456 | -0.0098456 | 0.0061544 | 0.0031544 |  |  | 0.76872 | -0.020846 | -0.0028456 | -0.013846 |  |  | 0.028092 | 0.028092 | 0.027433 | 0.033483 | 0.027433 | 0.045498 |  |  | 0.09122 | 0.040805 | 0.042674 | 0.052164
28.1838 | -0.008727 | 0.011273 | -0.008727 | -0.010727 | 0.011273 | -0.005727 |  |  | 0.77145 | -0.021727 | -0.00072699 | -0.018727 |  |  | 0.028092 | 0.028092 | 0.027433 | 0.033483 | 0.027433 | 0.046444 |  |  | 0.09122 | 0.040805 | 0.042674 | 0.052164
29.8538 | -0.0092771 | 0.010723 | -0.0067771 | -0.011277 | 0.0082229 | 0.0047229 |  |  | 0.77369 | -0.020277 | -0.0022771 | -0.014277 |  |  | 0.028092 | 0.028092 | 0.027433 | 0.033483 | 0.027433 | 0.046444 |  |  | 0.09122 | 0.040805 | 0.042674 | 0.052164
31.6228 | -0.0036827 | 0.0063173 | -0.0036827 | -0.012683 | 0.0063173 | 0.00031731 | -0.033683 | -0.033683 | 0.77218 | -0.023683 | -0.0086827 | -0.023683 | -0.031683 | -0.014683 | 0.0283 | 0.0283 | 0.027641 | 0.033379 | 0.027641 | 0.045422 | 0.042593 | 0.052097 | 0.091182 | 0.04072 | 0.042593 | 0.042593 | 0.052097 | 0.057397
33.4965 | -0.0065691 | 0.013431 | -0.0040691 | -0.0075691 | 0.010931 | 0.0084309 |  |  | 0.78258 | -0.016569 | 0.00043093 | -0.011569 |  |  | 0.028381 | 0.028381 | 0.027724 | 0.033332 | 0.027724 | 0.045388 |  |  | 0.091165 | 0.040682 | 0.042556 | 0.042556
35.4813 | -0.00099394 | 0.0090061 | -0.0034939 | -0.0029939 | 0.011506 | 0.0090061 |  |  | 0.78639 | -0.016994 | 0.0010061 | -0.015994 |  |  | 0.028381 | 0.028381 | 0.027724 | 0.033332 | 0.027724 | 0.045388 |  |  | 0.091165 | 0.040682 | 0.042556 | 0.042556
37.5837 | -0.0048847 | 0.0051153 | -0.0048847 | -0.0098847 | 0.0051153 | 0.0051153 |  |  | 0.78596 | -0.018885 | -0.0038847 | -0.014885 |  |  | 0.028381 | 0.028381 | 0.027724 | 0.033332 | 0.027724 | 0.045388 |  |  | 0.091165 | 0.040682 | 0.042556 | 0.042556
39.8107 | -0.0088332 | 0.011167 | -0.0063332 | -0.014833 | 0.0086668 | 0.0041668 | -0.027833 | -0.031833 | 0.78562 | -0.021833 | -0.0068332 | -0.013833 | -0.027833 | -0.012833 | 0.028386 | 0.028386 | 0.027729 | 0.03333 | 0.027729 | 0.045386 | 0.042554 | 0.052066 | 0.091164 | 0.04068 | 0.042554 | 0.042554 | 0.052066 | 0.057272
42.1697 | -0.0041356 | 0.0058644 | -0.0041356 | -0.0051356 | 0.0058644 | 0.0088644 |  |  | 0.7941 | -0.016136 | -0.0021356 | -0.014136 |  |  | 0.028386 | 0.028386 | 0.027729 | 0.03333 | 0.027729 | 0.045386 |  |  | 0.091164 | 0.04068 | 0.042554 | 0.042554
44.6684 | -0.0073636 | 0.012636 | -0.0048636 | -0.0033636 | 0.010136 | 0.0076364 |  |  | 0.79483 | -0.018364 | -0.0043636 | -0.012364 |  |  | 0.028386 | 0.028386 | 0.027729 | 0.03333 | 0.027729 | 0.045386 |  |  | 0.091164 | 0.04068 | 0.042554 | 0.042554
47.3151 | -0.0034696 | 0.0065304 | -0.0034696 | 0.00053036 | 0.0065304 | 0.0065304 |  |  | 0.80283 | -0.01247 | -0.00046964 | -0.01347 |  |  | 0.028386 | 0.028386 | 0.02772 | 0.033354 | 0.02772 | 0.045403 |  |  | 0.091173 | 0.0407 | 0.042573 | 0.042573
50.1187 | -0.008757 | 0.011243 | 0.00011482 | 0.00144 | 0.010765 | 0.03044 | -0.01756 | -0.02056 | 0.80588 | -0.01256 | 0.00044003 | -0.00956 | -0.01756 | -0.00356 | 0.06926 | 0.02823 | 0.034319 | 0.035189 | 0.028044 | 0.053275 | 0.044025 | 0.053275 | 0.09186 | 0.042216 | 0.044025 | 0.044025 | 0.053275 | 0.058468
53.0884 | -0.0034666 | 0.0065334 | -0.0034666 | 0.0055334 | 0.0065334 | 0.014533 |  |  | 0.81137 | -0.0094666 | 0.0035334 | -0.0034666 |  |  | 0.028386 | 0.028386 | 0.02772 | 0.033354 | 0.02772 | 0.045403 |  |  | 0.091173 | 0.0407 | 0.042573 | 0.042573
56.2341 | -0.0095907 | 0.010409 | -0.0070907 | 0.00040935 | 0.0079093 | 0.010409 |  |  | 0.80976 | -0.014591 | -0.0025907 | -0.014591 |  |  | 0.028386 | 0.028386 | 0.02772 | 0.033354 | 0.02772 | 0.045403 |  |  | 0.091173 | 0.0407 | 0.042573 | 0.042573
59.5662 | -0.0047089 | 0.0052911 | -0.0047089 | -0.0017089 | 0.0052911 | 0.0072911 |  |  | 0.80921 | -0.017709 | -0.0057089 | -0.014709 |  |  | 0.028386 | 0.028386 | 0.02772 | 0.033354 | 0.02772 | 0.045403 |  |  | 0.091173 | 0.0407 | 0.042573 | 0.042573
63.0957 | -0.011804 | 0.0081964 | -0.0093036 | 0.0021964 | 0.0056964 | 0.012196 | -0.015804 | -0.015804 | 0.8169 | -0.011804 | -0.0018036 | -0.0068036 | -0.017804 | -0.0048036 | 0.028386 | 0.028386 | 0.02772 | 0.033354 | 0.02772 | 0.045403 | 0.042573 | 0.052081 | 0.081317 | 0.0407 | 0.042573 | 0.042573 | 0.033354 | 0.053234
66.8344 | -0.0061789 | 0.0038211 | -0.0061789 | -0.00017887 | 0.0038211 | 0.0088211 |  |  | 0.81746 | -0.014179 | -0.0051789 | -0.016179 |  |  | 0.028386 | 0.028386 | 0.02772 | 0.033354 | 0.02772 | 0.045403 |  |  | 0.081317 | 0.0407 | 0.042573 | 0.042573
70.7946 | -0.011813 | 0.008187 | -0.011813 | 0.00018702 | 0.008187 | 0.011187 |  |  | 0.82171 | -0.014813 | -0.003813 | -0.011813 |  |  | 0.028515 | 0.028515 | 0.027852 | 0.03328 | 0.027852 | 0.045349 |  |  | 0.081287 | 0.040639 | 0.042515 | 0.042515
74.9894 | -0.010439 | 0.0095611 | -0.0079389 | 0.0035611 | 0.0070611 | 0.014561 |  |  | 0.82822 | -0.010439 | -0.00043893 | -0.0054389 |  |  | 0.028515 | 0.028515 | 0.027852 | 0.03328 | 0.027852 | 0.045349 |  |  | 0.081287 | 0.040639 | 0.042515 | 0.042515
79.4328 | -0.0045401 | 0.0054599 | -0.0045401 | 0.0014599 | 0.0054599 | 0.01246 | -0.01354 | -0.01754 | 0.82911 | -0.01154 | -0.0035401 | -0.01454 | -0.01754 | -0.0055401 | 0.028515 | 0.028515 | 0.027852 | 0.03328 | 0.027852 | 0.045349 | 0.042515 | 0.052034 | 0.081287 | 0.040639 | 0.042515 | 0.042515 | 0.03328 | 0.053092
84.1395 | -0.011804 | 0.0081959 | -0.0093041 | -0.0018041 | 0.0056959 | 0.0071959 |  |  | 0.82741 | -0.016804 | -0.0088041 | -0.016804 |  |  | 0.028515 | 0.028515 | 0.027852 | 0.03328 | 0.027852 | 0.045349 |  |  | 0.081287 | 0.040639 | 0.042515 | 0.042515
89.1251 | -0.0063774 | 0.0036226 | -0.0063774 | 0.0056226 | 0.0036226 | 0.015623 |  |  | 0.83815 | -0.010377 | -0.00037737 | -0.0063774 |  |  | 0.028515 | 0.028515 | 0.027852 | 0.03328 | 0.027852 | 0.045349 |  |  | 0.081287 | 0.040639 | 0.042515 | 0.042515
94.4061 | -0.0051873 | 0.0048127 | -0.0051873 | 0.0028127 | 0.0048127 | 0.014813 |  |  | 0.83998 | -0.011187 | -0.0021873 | -0.0051873 |  |  | 0.028515 | 0.028515 | 0.027848 | 0.03329 | 0.027848 | 0.045357 |  |  | 0.081291 | 0.040648 | 0.042523 | 0.042523
100 | -0.011365 | 0.0086347 | -0.0088653 | -0.0013653 | 0.0061347 | 0.010635 | -0.015365 | -0.018365 | 0.83908 | -0.014365 | -0.0053653 | -0.016365 | -0.018365 | -0.0073653 | 0.028515 | 0.028515 | 0.027848 | 0.03329 | 0.027848 | 0.045357 | 0.042523 | 0.052041 | 0.081291 | 0.040648 | 0.042523 | 0.042523 | 0.03329 | 0.053098
105.9254 | -0.0078199 | 0.0021801 | -0.01032 | 0.0031801 | 0.0046801 | 0.01118 |  |  | 0.8429 | -0.01282 | -0.0048199 | -0.01282 |  |  | 0.028522 | 0.028522 | 0.027855 | 0.033286 | 0.027855 | 0.045354 |  |  | 0.08129 | 0.040645 | 0.04252 | 0.04252
112.2018 | -0.0052913 | 0.0047087 | -0.0052913 | 0.0077087 | 0.0047087 | 0.016709 |  |  | 0.85098 | -0.0082913 | 0.00070875 | -0.0052913 |  |  | 0.028522 | 0.028522 | 0.027855 | 0.033286 | 0.027855 | 0.045354 |  |  | 0.08129 | 0.040645 | 0.04252 | 0.04252
118.8502 | -0.011642 | 0.0083578 | -0.0091422 | 0.0023578 | 0.0058578 | 0.012358 |  |  | 0.85004 | -0.011642 | -0.0066422 | -0.0066422 |  |  | 0.028583 | 0.028583 | 0.027917 | 0.033252 | 0.027917 | 0.045328 |  |  | 0.081275 | 0.040616 | 0.042493 | 0.042493
125.8925 | -0.0080799 | 0.0019201 | -0.0080799 | -0.0020799 | 0.0019201 | 0.0089201 | -0.01708 | -0.01408 | 0.84904 | -0.01408 | -0.0090799 | -0.0080799 | -0.02008 | -0.0080799 | 0.028583 | 0.028583 | 0.027917 | 0.033252 | 0.027917 | 0.045328 | 0.042493 | 0.052016 | 0.081275 | 0.040616 | 0.042493 | 0.033252 | 0.033252 | 0.051151
133.3521 | -0.005887 | 0.004113 | -0.005887 | 0.009113 | 0.004113 | 0.020113 |  |  | 0.86185 | -0.003887 | 0.00011303 | -0.005887 |  |  | 0.028583 | 0.028583 | 0.027917 | 0.033252 | 0.027917 | 0.045328 |  |  | 0.081275 | 0.040616 | 0.042493 | 0.033252
141.2538 | -0.013752 | 0.006248 | -0.011252 | 0.00024799 | 0.003748 | 0.015248 |  |  | 0.85943 | -0.009752 | -0.004752 | -0.008752 |  |  | 0.028733 | 0.028733 | 0.028071 | 0.033166 | 0.028071 | 0.047169 |  |  | 0.08124 | 0.03961 | 0.042426 | 0.033166
149.6236 | -0.012173 | 0.0078274 | -0.0096726 | -0.0021726 | 0.0053274 | 0.017827 |  |  | 0.86155 | -0.0091726 | -0.0051726 | -0.0071726 |  |  | 0.028733 | 0.028733 | 0.028071 | 0.033166 | 0.028071 | 0.046216 |  |  | 0.08124 | 0.03961 | 0.042426 | 0.033166
158.4893 | -0.0087322 | 0.0012678 | -0.0087322 | -0.0027322 | 0.0012678 | 0.0092678 | -0.013732 | -0.014732 | 0.86041 | -0.012732 | -0.0097322 | -0.0087322 | -0.018732 | -0.0077322 | 0.028733 | 0.028733 | 0.028071 | 0.033166 | 0.028071 | 0.045265 | 0.042426 | 0.051961 | 0.08124 | 0.03961 | 0.042426 | 0.033166 | 0.033166 | 0.051096
167.8804 | -0.016623 | 0.0033772 | -0.016623 | -0.00062278 | 0.0033772 | 0.010377 |  |  | 0.86303 | -0.011623 | -0.0086228 | -0.0066228 |  |  | 0.028733 | 0.028733 | 0.028071 | 0.033166 | 0.028071 | 0.045265 |  |  | 0.08124 | 0.03961 | 0.042426 | 0.033166
177.8279 | -0.013537 | 0.0064627 | -0.011037 | 0.0064627 | 0.0039627 | 0.016463 |  |  | 0.87175 | -0.0065373 | -0.0025373 | -0.0085373 |  |  | 0.028733 | 0.028733 | 0.028071 | 0.033166 | 0.028071 | 0.045265 |  |  | 0.08124 | 0.03961 | 0.042426 | 0.033166
188.3649 | -0.012248 | 0.0077518 | -0.0097482 | 0.0027518 | 0.0052518 | 0.016752 |  |  | 0.87356 | -0.0052482 | -0.0042482 | -0.0072482 |  |  | 0.028733 | 0.028733 | 0.028071 | 0.033166 | 0.028071 | 0.045265 |  |  | 0.08124 | 0.03961 | 0.042426 | 0.033166
199.5262 | -0.0093053 | 0.00069471 | -0.0093053 | -0.0033053 | 0.00069471 | 0.012695 | -0.0083053 | -0.012305 | 0.87302 | -0.0083053 | -0.0083053 | -0.0093053 | -0.016305 | -0.0053053 | 0.028778 | 0.028778 | 0.028113 | 0.033151 | 0.028113 | 0.045254 | 0.042414 | 0.051952 | 0.081234 | 0.039598 | 0.042414 | 0.033151 | 0.033151 | 0.05099
211.3489 | -0.0085103 | 0.0014897 | -0.0085103 | -0.0045103 | 0.0014897 | 0.01249 |  |  | 0.87442 | -0.0085103 | -0.0095103 | -0.0085103 |  |  | 0.028778 | 0.028778 | 0.028113 | 0.033151 | 0.028113 | 0.045254 |  |  | 0.081234 | 0.039598 | 0.042414 | 0.033151
223.8721 | -0.014199 | 0.0058013 | -0.011699 | -0.0061987 | 0.0033013 | 0.0098013 |  |  | 0.87437 | -0.011199 | -0.011199 | -0.0091987 |  |  | 0.028778 | 0.028778 | 0.028113 | 0.033151 | 0.028113 | 0.045254 |  |  | 0.081234 | 0.039598 | 0.042414 | 0.033151
237.1374 | -0.012931 | 0.0070694 | -0.012931 | 6.9409e-05 | 0.0070694 | 0.015069 |  |  | 0.88566 | -0.0059306 | -0.0069306 | -0.0029306 |  |  | 0.028778 | 0.028778 | 0.028113 | 0.033151 | 0.028113 | 0.045254 |  |  | 0.081234 | 0.039598 | 0.042414 | 0.033151
251.1886 | -0.011752 | 0.0082476 | -0.0092524 | 0.0042476 | 0.0057476 | 0.014248 | -0.0017524 | -0.00075242 | 0.89254 | 0.0072476 | -0.0027524 | -0.0067524 | -0.0087524 | 0.0012476 | 0.028778 | 0.028778 | 0.028113 | 0.033151 | 0.028113 | 0.045254 | 0.042414 | 0.051952 | 0.081234 | 0.039598 | 0.042414 | 0.033151 | 0.033151 | 0.049934
266.0725 | -0.0099728 | 2.7222e-05 | -0.0099728 | -0.00097278 | 2.7222e-05 | 0.015027 |  |  | 0.89015 | 0.0020272 | -0.0069728 | -0.0099728 |  |  | 0.028778 | 0.028778 | 0.028113 | 0.033151 | 0.028113 | 0.046206 |  |  | 0.081234 | 0.039598 | 0.042414 | 0.033151
281.8383 | -0.0082304 | 0.0017696 | -0.0082304 | 0.0017696 | 0.0017696 | 0.01577 |  |  | 0.89285 | 0.0027696 | -0.0072304 | -0.0082304 |  |  | 0.02893 | 0.02893 | 0.028269 | 0.033062 | 0.028269 | 0.046142 |  |  | 0.081198 | 0.039524 | 0.042345 | 0.033062
298.5383 | -0.0074743 | 0.0025257 | -0.0074743 | -0.00047425 | 0.0025257 | 0.015526 |  |  | 0.89463 | 0.0025257 | -0.0084743 | -0.0074743 |  |  | 0.02893 | 0.02893 | 0.028269 | 0.033062 | 0.028269 | 0.046142 |  |  | 0.081198 | 0.039524 | 0.042345 | 0.033062
316.2278 | -0.016548 | 0.0034521 | -0.014048 | -0.0055479 | 0.00095209 | 0.010452 | -0.0095479 | -0.010548 | 0.8918 | -0.0035479 | -0.012548 | -0.011548 | -0.016548 | -0.0075479 | 0.02893 | 0.02893 | 0.028269 | 0.033062 | 0.028269 | 0.046142 | 0.042345 | 0.051895 | 0.081198 | 0.039524 | 0.042345 | 0.033062 | 0.033062 | 0.049875
334.9654 | -0.014271 | 0.0057286 | -0.011771 | -0.0062714 | 0.0032286 | 0.011729 |  |  | 0.89528 | -0.0022714 | -0.011271 | -0.0092714 |  |  | 0.02893 | 0.02893 | 0.028269 | 0.033062 | 0.028269 | 0.046142 |  |  | 0.081198 | 0.039524 | 0.042345 | 0.033062
354.8134 | -0.013936 | 0.0060637 | -0.011436 | -0.0059363 | 0.0035637 | 0.011064 |  |  | 0.89705 | -0.0029363 | -0.011936 | -0.0089363 |  |  | 0.02893 | 0.02893 | 0.028269 | 0.033062 | 0.028269 | 0.046142 |  |  | 0.081198 | 0.039524 | 0.042345 | 0.033062
375.8374 | -0.013601 | 0.0063987 | -0.011101 | -0.0056013 | 0.0038987 | 0.010399 |  |  | 0.89911 | -0.0026013 | -0.011601 | -0.0086013 |  |  | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.046147 |  |  | 0.081201 | 0.03953 | 0.042351 | 0.033069
398.1072 | -0.012883 | 0.0071169 | -0.010383 | -0.0068831 | 0.0046169 | 0.012117 | -0.0078831 | -0.0088831 | 0.90141 | -0.0028831 | -0.011883 | -0.0078831 | -0.014883 | -0.0058831 | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.046147 | 0.042351 | 0.0519 | 0.081201 | 0.03953 | 0.042351 | 0.033069 | 0.033069 | 0.049976
421.6965 | -0.012868 | 0.0071325 | -0.010368 | -0.0058675 | 0.0046325 | 0.011132 |  |  | 0.90342 | -0.0038675 | -0.012868 | -0.0078675 |  |  | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.046147 |  |  | 0.081201 | 0.03953 | 0.042351 | 0.033069
446.6836 | -0.013371 | 0.0066295 | -0.010871 | -0.0063705 | 0.0041295 | 0.010629 |  |  | 0.90512 | -0.0043705 | -0.013371 | -0.0083705 |  |  | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.046147 |  |  | 0.081201 | 0.03953 | 0.042351 | 0.033069
473.1513 | -0.01344 | 0.0065598 | -0.01344 | -0.0024402 | 0.0065598 | 0.01556 |  |  | 0.91255 | 0.0005598 | -0.0084402 | -0.0034402 |  |  | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.046147 |  |  | 0.081201 | 0.03953 | 0.042351 | 0.033069
501.1872 | -0.014278 | 0.0057218 | -0.011778 | -0.0062782 | 0.0032218 | 0.0097218 | -0.0082782 | -0.0062782 | 0.90925 | -0.0052782 | -0.013278 | -0.0092782 | -0.016278 | -0.0072782 | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.046147 | 0.042351 | 0.0519 | 0.081201 | 0.03953 | 0.042351 | 0.033069 | 0.033069 | 0.049496
530.8844 | -0.014667 | 0.0053326 | -0.012167 | -0.0026674 | 0.0028326 | 0.0093326 |  |  | 0.91188 | -0.0056674 | -0.013667 | -0.0096674 |  |  | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.046147 |  |  | 0.081201 | 0.03953 | 0.042351 | 0.033069
562.3413 | -0.014831 | 0.0051686 | -0.014831 | 0.0011686 | 0.0051686 | 0.015169 |  |  | 0.92025 | 0.00016861 | -0.0098314 | -0.0048314 |  |  | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.046147 |  |  | 0.081201 | 0.03953 | 0.042351 | 0.033069
595.6621 | -0.017186 | 0.0028141 | -0.014686 | -0.0051859 | 0.00031405 | 0.0078141 |  |  | 0.91658 | -0.0061859 | -0.016186 | -0.012186 |  |  | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.047102 |  |  | 0.081201 | 0.03953 | 0.042351 | 0.033069
630.9573 | -0.018735 | 0.0012649 | -0.016235 | -0.0057351 | -0.0012351 | 0.0072649 | -0.0077351 | -0.010735 | 0.91891 | -0.0077351 | -0.015735 | -0.013735 | -0.018735 | -0.0087351 | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.046147 | 0.042351 | 0.0519 | 0.081201 | 0.03953 | 0.042351 | 0.033069 | 0.033069 | 0.0494
668.3439 | -0.019678 | 0.00032219 | -0.017178 | -0.0046778 | -0.0021778 | 0.0073222 |  |  | 0.92238 | -0.0076778 | -0.015678 | -0.014678 |  |  | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.047102 |  |  | 0.081201 | 0.03953 | 0.042351 | 0.033069
707.9458 | -0.012334 | -0.0023344 | -0.012334 | -0.00033439 | -0.0023344 | 0.011666 |  |  | 0.92973 | -0.0033344 | -0.010334 | -0.012334 |  |  | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.047102 |  |  | 0.081201 | 0.03953 | 0.042351 | 0.033069
749.8942 | -0.016648 | 0.0033521 | -0.014148 | 0.0033521 | 0.00085207 | 0.014352 |  |  | 0.93601 | 0.00035207 | -0.0076479 | -0.0016479 |  |  | 0.02893 | 0.02893 | 0.028267 | 0.033069 | 0.028267 | 0.047102 |  |  | 0.081201 | 0.03953 | 0.042351 | 0.033069
794.3282 | -0.018906 | 0.0010938 | -0.016406 | -0.0049062 | -0.0014062 | 0.0050938 | -0.010906 | -0.0099062 | 0.92998 | -0.0099062 | -0.017906 | -0.013906 | -0.020906 | -0.0099062 | 0.02893 | 0.02893 | 0.028244 | 0.033136 | 0.028244 | 0.047149 | 0.042403 | 0.051942 | 0.081228 | 0.039585 | 0.042403 | 0.033136 | 0.033136 | 0.04954
841.3951 | -0.012762 | -0.0027621 | -0.012762 | -0.0037621 | -0.0027621 | 0.0082379 |  |  | 0.93882 | -0.0067621 | -0.013762 | -0.012762 |  |  | 0.02893 | 0.02893 | 0.028244 | 0.033136 | 0.028244 | 0.047149 |  |  | 0.081228 | 0.039585 | 0.042403 | 0.033136
891.2509 | -0.018244 | 0.001756 | -0.015744 | -0.00024399 | -0.00074399 | 0.011756 |  |  | 0.9471 | -0.004244 | -0.010244 | -0.013244 |  |  | 0.02893 | 0.02893 | 0.028244 | 0.033136 | 0.028244 | 0.047149 |  |  | 0.081228 | 0.039585 | 0.042403 | 0.033136
944.0609 | -0.0125 | -0.0025002 | -0.0125 | -0.0065002 | -0.0025002 | 0.0064998 |  |  | 0.94639 | -0.0095002 | -0.0155 | -0.0125 |  |  | 0.02893 | 0.02893 | 0.028244 | 0.033136 | 0.028244 | 0.047149 |  |  | 0.081228 | 0.039585 | 0.042403 | 0.033136
1000 | -0.0096964 | 0.00030365 | -0.0096964 | 0.0043036 | 0.00030365 | 0.013304 | -0.00069635 | 0.0013036 | 0.95802 | -0.0026964 | -0.0086964 | -0.0096964 | -0.010696 | -0.00069635 | 0.02893 | 0.02893 | 0.028244 | 0.033136 | 0.028244 | 0.047149 | 0.042403 | 0.051942 | 0.081228 | 0.039585 | 0.042403 | 0.033136 | 0.033136 | 0.049445
1059.2537 | -0.016741 | 0.0032595 | -0.014241 | -0.0027405 | 0.0007595 | 0.0062595 |  |  | 0.9565 | -0.0097405 | -0.014741 | -0.011741 |  |  | 0.02893 | 0.02893 | 0.028244 | 0.033136 | 0.028244 | 0.048104 |  |  | 0.081228 | 0.039585 | 0.042403 | 0.033136
1122.0185 | -0.01449 | 0.0055101 | -0.01199 | 0.0035101 | 0.0030101 | 0.01551 |  |  | 0.96982 | -0.0014899 | -0.0064899 | -0.0094899 |  |  | 0.028928 | 0.028928 | 0.028242 | 0.033137 | 0.028242 | 0.048105 |  |  | 0.081229 | 0.039586 | 0.042404 | 0.033137
1188.5022 | -0.012437 | -0.0024375 | -0.012437 | 0.0035625 | -0.0024375 | 0.0085625 |  |  | 0.96907 | -0.0074375 | -0.014437 | -0.012437 |  |  | 0.028926 | 0.028926 | 0.02824 | 0.033139 | 0.02824 | 0.048106 |  |  | 0.081229 | 0.039587 | 0.042405 | 0.033139
1258.9254 | -0.01273 | -0.0027305 | -0.01273 | -0.00073046 | -0.0027305 | 0.0062695 | -0.0077305 | -0.0057305 | 0.97255 | -0.01073 | -0.01673 | -0.01273 | -0.01873 | -0.0067305 | 0.028925 | 0.028925 | 0.028238 | 0.03314 | 0.028238 | 0.048106 | 0.042406 | 0.051944 | 0.081229 | 0.039588 | 0.042406 | 0.03314 | 0.03314 | 0.049543
1333.5214 | -0.015776 | 0.0042242 | -0.010776 | 0.0052242 | -0.00077578 | 0.012224 |  |  | 0.98476 | -0.0067758 | -0.010776 | -0.0057758 |  |  | 0.028925 | 0.028925 | 0.028238 | 0.03314 | 0.028238 | 0.048106 |  |  | 0.081229 | 0.039588 | 0.042406 | 0.03314
1412.5375 | -0.0099595 | 4.0473e-05 | -0.0099595 | 0.0080405 | 4.0473e-05 | 0.01704 |  |  | 0.99759 | -0.0029595 | -0.0049595 | 4.0473e-05 |  |  | 0.028925 | 0.028925 | 0.028238 | 0.03314 | 0.028238 | 0.048106 |  |  | 0.081229 | 0.04335 | 0.042406 | 0.03314
1496.2357 | -0.014507 | -0.0045072 | -0.014507 | -0.0085072 | -0.0045072 | 0.0034928 |  |  | 0.99249 | -0.017507 | -0.018507 | -0.014507 |  |  | 0.028925 | 0.028925 | 0.028238 | 0.03314 | 0.028238 | 0.048106 |  |  | 0.081229 | 0.04335 | 0.042406 | 0.03314
1584.8932 | -0.011547 | -0.0015472 | -0.011547 | -0.0045472 | -0.0015472 | 0.0094528 | -0.0035472 | 0.0004528 | 1.0061 | -0.012547 | -0.013547 | -0.011547 | -0.016547 | -0.0025472 | 0.028926 | 0.028926 | 0.027909 | 0.034089 | 0.027909 | 0.048766 | 0.043152 | 0.052555 | 0.081622 | 0.04408 | 0.043152 | 0.034089 | 0.034089 | 0.050183
1678.804 | -0.01015 | -0.00015017 | -0.01015 | -0.00015017 | -0.00015017 | 0.01385 |  |  | 1.0207 | -0.0081502 | -0.0081502 | -0.00015017 |  |  | 0.028926 | 0.028926 | 0.027909 | 0.034089 | 0.027909 | 0.048766 |  |  | 0.081622 | 0.04408 | 0.043152 | 0.034089
1778.2794 | -0.011077 | -0.0010768 | -0.011077 | -0.0030768 | -0.0010768 | 0.0089232 |  |  | 1.0255 | -0.014077 | -0.013077 | -0.011077 |  |  | 0.028926 | 0.028926 | 0.027909 | 0.034089 | 0.027909 | 0.04971 |  |  | 0.081622 | 0.04408 | 0.043152 | 0.034089
1883.6491 | -0.014997 | -0.004997 | -0.012497 | -0.002997 | -0.007497 | 0.008003 |  |  | 1.035 | -0.014997 | -0.014997 | -0.009997 |  |  | 0.028926 | 0.028926 | 0.027909 | 0.034089 | 0.027909 | 0.04971 |  |  | 0.081622 | 0.04408 | 0.043152 | 0.034089
1995.2623 | -0.010619 | -0.00061916 | -0.0081192 | 0.0023808 | -0.0031192 | 0.012381 | -0.00061916 | 0.0033808 | 1.0492 | -0.013619 | -0.010619 | -0.0056192 | -0.013619 | 0.0023808 | 0.028926 | 0.028926 | 0.027909 | 0.034089 | 0.027909 | 0.04971 | 0.043152 | 0.052555 | 0.081622 | 0.04408 | 0.043152 | 0.034089 | 0.034089 | 0.051225
2113.489 | -0.010409 | -0.00040862 | -0.010409 | -0.0014086 | -0.00040862 | 0.0095914 |  |  | 1.0592 | -0.018409 | -0.013409 | -0.010409 |  |  | 0.028892 | 0.028892 | 0.027455 | 0.039587 | 0.027455 | 0.050509 |  |  | 0.08211 | 0.044979 | 0.04407 | 0.035244
2238.7211 | -0.012148 | -0.0021478 | -0.012148 | 0.0018522 | -0.0021478 | 0.011852 |  |  | 1.0762 | -0.019148 | -0.011148 | -0.012148 |  |  | 0.028892 | 0.028892 | 0.027455 | 0.039587 | 0.027455 | 0.051441 |  |  | 0.08211 | 0.044979 | 0.04407 | 0.035244
2371.3737 | -0.010536 | -0.00053617 | -0.010536 | -0.0015362 | -0.00053617 | 0.012464 |  |  | 1.0903 | -0.021536 | -0.012536 | -0.010536 |  |  | 0.028892 | 0.028892 | 0.027455 | 0.039587 | 0.027455 | 0.051441 |  |  | 0.08211 | 0.044979 | 0.04407 | 0.035244
2511.8864 | -0.0095139 | 0.0004861 | -0.0095139 | 0.0014861 | 0.0004861 | 0.011486 | -0.0005139 | 0.0014861 | 1.1087 | -0.022514 | -0.011514 | -0.0095139 | -0.015514 | 0.0074861 | 0.028892 | 0.028892 | 0.027455 | 0.039587 | 0.027455 | 0.052375 | 0.04407 | 0.062786 | 0.08211 | 0.044979 | 0.04407 | 0.035244 | 0.035244 | 0.052094
2660.7251 | -0.01601 | -0.0060104 | -0.01601 | -0.0030104 | -0.0060104 | 0.0079896 |  |  | 1.1244 | -0.02701 | -0.01501 | -0.01601 |  |  | 0.039119 | 0.028815 | 0.035553 | 0.046866 | 0.027314 | 0.054748 |  |  | 0.083644 | 0.047722 | 0.046866 | 0.038683
2818.3829 | -0.013951 | -0.0039512 | -0.013951 | -0.0059512 | -0.0039512 | 0.0080488 |  |  | 1.1428 | -0.025951 | -0.016951 | -0.013951 |  |  | 0.039024 | 0.028685 | 0.035447 | 0.046863 | 0.027177 | 0.055643 |  |  | 0.083643 | 0.048581 | 0.046863 | 0.038679
2985.3826 | -0.016874 | -0.0068744 | -0.016874 | -0.0078744 | -0.0068744 | 0.0051256 |  |  | 1.1617 | -0.032874 | -0.019874 | -0.016874 |  |  | 0.039024 | 0.028685 | 0.035447 | 0.046863 | 0.027177 | 0.056543 |  |  | 0.083643 | 0.048581 | 0.046863 | 0.038679
3162.2777 | -0.015832 | -0.0058321 | -0.015832 | -0.0048321 | -0.0058321 | 0.010168 | -0.0038321 | -0.00083213 | 1.1902 | -0.030832 | -0.014832 | -0.015832 | -0.021832 | 0.0081679 | 0.039018 | 0.028677 | 0.035441 | 0.046862 | 0.027168 | 0.058353 | 0.046862 | 0.064777 | 0.083643 | 0.048581 | 0.046862 | 0.038679 | 0.038679 | 0.054925
3349.6544 | -0.017924 | -0.0079236 | -0.014724 | -0.0025236 | -0.0097236 | 0.015476 |  |  | 1.2172 | -0.027524 | -0.013524 | -0.011524 |  |  | 0.038998 | 0.028651 | 0.034135 | 0.048676 | 0.026614 | 0.061599 |  |  | 0.084672 | 0.05117 | 0.057179 | 0.040858
3548.1339 | -0.016836 | -0.0068356 | -0.013636 | -0.0024356 | -0.0086356 | 0.012564 |  |  | 1.2462 | -0.035436 | -0.011436 | -0.010436 |  |  | 0.038991 | 0.028641 | 0.034127 | 0.048674 | 0.026604 | 0.063389 |  |  | 0.084671 | 0.051168 | 0.057177 | 0.040856
3758.374 | -0.017286 | -0.0072856 | -0.017286 | -0.0062856 | -0.0072856 | 0.011714 |  |  | 1.2727 | -0.039286 | -0.014286 | -0.017286 |  |  | 0.038922 | 0.028547 | 0.034048 | 0.048653 | 0.026503 | 0.066084 |  |  | 0.084659 | 0.051991 | 0.057158 | 0.04083
3981.0717 | -0.020285 | -0.010285 | -0.017085 | -0.0088846 | -0.012085 | 0.012115 | -0.0058846 | 0.0011154 | 1.3041 | -0.040885 | -0.013885 | -0.013885 | -0.026885 | 0.014115 | 0.038914 | 0.028536 | 0.034039 | 0.04865 | 0.026491 | 0.069734 | 0.04865 | 0.084657 | 0.094164 | 0.052838 | 0.057156 | 0.040827 | 0.057156 | 0.06836
4216.965 | -0.014548 | -0.014548 | -0.014548 | -0.013548 | -0.014548 | 0.0094525 |  |  | 1.3348 | -0.043548 | -0.017548 | -0.014548 |  |  | 0.048911 | 0.028148 | 0.040791 | 0.059042 | 0.026418 | 0.0749 |  |  | 0.09532 | 0.055695 | 0.059042 | 0.043428
4466.8359 | -0.016642 | -0.016642 | -0.016642 | -0.010642 | -0.016642 | 0.006358 |  |  | 1.3654 | -0.043642 | -0.021642 | -0.016642 |  |  | 0.048901 | 0.028131 | 0.040778 | 0.059037 | 0.026399 | 0.079463 |  |  | 0.095317 | 0.057353 | 0.067716 | 0.043421
4731.5126 | -0.023913 | -0.013913 | -0.020237 | -0.0045605 | -0.015237 | 0.01244 |  |  | 1.4063 | -0.03656 | -0.01256 | -0.01656 |  |  | 0.048863 | 0.028064 | 0.040732 | 0.059019 | 0.026328 | 0.084067 |  |  | 0.095306 | 0.058174 | 0.067699 | 0.043396
5011.8723 | -0.021203 | -0.0012034 | -0.021203 | -0.0092034 | -0.0012034 | 0.010797 | -0.012203 | -0.0062034 | 1.4355 | -0.039203 | -0.018203 | -0.021203 | -0.038203 | 0.014797 | 0.048824 | 0.027996 | 0.040038 | 0.057383 | 0.025633 | 0.089542 | 0.048917 | 0.10389 | 0.094302 | 0.057383 | 0.075451 | 0.057383 | 0.057383 | 0.078429
5308.8444 | -0.025831 | -0.0058313 | -0.022155 | -0.0074783 | -0.0071548 | 0.013522 |  |  | 1.4687 | -0.029478 | -0.018478 | -0.018478 |  |  | 0.048782 | 0.027923 | 0.041724 | 0.06688 | 0.026399 | 0.095676 |  |  | 0.10427 | 0.059808 | 0.07598 | 0.058077
5623.4133 | -0.020686 | -0.00068646 | -0.01701 | -0.0073335 | -0.00201 | 0.019666 |  |  | 1.4906 | -0.032334 | -0.016334 | -0.023334 |  |  | 0.048753 | 0.027873 | 0.041691 | 0.066872 | 0.026346 | 0.10044 |  |  | 0.10427 | 0.06067 | 0.085275 | 0.058068
5956.6214 | -0.022277 | -0.012277 | -0.022277 | -0.022277 | -0.012277 | 0.0077228 |  |  | 1.4907 | -0.034277 | -0.023277 | -0.032277 |  |  | 0.048663 | 0.038315 | 0.042431 | 0.069104 | 0.033953 | 0.10571 |  |  | 0.10571 | 0.064811 | 0.087037 | 0.060625
6309.5734 | -0.022118 | -0.012118 | -0.022118 | -0.012118 | -0.012118 | 0.0088817 | -0.0081183 | -0.031118 | 1.4809 | -0.021118 | -0.020118 | -0.022118 | -0.055118 | 0.015882 | 0.048732 | 0.027835 | 0.041665 | 0.066866 | 0.026306 | 0.10618 | 0.05806 | 0.10426 | 0.11389 | 0.063301 | 0.08527 | 0.05806 | 0.05806 | 0.083958
6683.4392 | -0.015137 | -0.015137 | -0.018065 | -0.019029 | -0.014083 | 0.018971 |  |  | 1.4479 | -0.017029 | -0.024029 | -0.033029 |  |  | 0.048725 | 0.027822 | 0.043713 | 0.075931 | 0.026722 | 0.10232 |  |  | 0.11387 | 0.064144 | 0.085238 | 0.058013
7079.4578 | -0.017653 | -0.017653 | -0.021251 | -0.015813 | -0.015703 | 0.010187 |  |  | 1.379 | -0.0048127 | -0.023813 | -0.024813 |  |  | 0.058831 | 0.038224 | 0.051318 | 0.078368 | 0.034898 | 0.10039 |  |  | 0.11551 | 0.067864 | 0.096651 | 0.061168
7498.9421 | -0.019158 | -0.0091585 | -0.0155 | -0.0038924 | -0.011291 | 0.019108 |  |  | 1.2825 | 0.014108 | -0.010892 | -0.011892 |  |  | 0.058803 | 0.038181 | 0.051111 | 0.078001 | 0.034212 | 0.093559 |  |  | 0.11526 | 0.06744 | 0.096355 | 0.060698
7943.2823 | -0.0016936 | -0.0016936 | 0.0046044 | 0.00081233 | -0.007349 | 0.022812 | -0.018188 | -0.032188 | 1.1435 | 0.021812 | -0.0051877 | -0.0091877 | -0.066188 | 0.036812 | 0.058677 | 0.048405 | 0.052107 | 0.080901 | 0.041906 | 0.091482 | 0.064381 | 0.12667 | 0.14576 | 0.069957 | 0.10791 | 0.064381 | 0.080901 | 0.10284
8413.9514 | -0.0027903 | -0.0027903 | 0.0051812 | -0.0038537 | -0.0052745 | 0.024146 |  |  | 0.95314 | 0.044146 | 0.0011463 | -0.0068537 |  |  | 0.10912 | 0.058376 | 0.087869 | 0.10255 | 0.054783 | 0.10255 |  |  | 0.154 | 0.085163 | 0.13608 | 0.10255
8912.5094 | 0.003532 | 0.003532 | 0.019481 | 0.002411 | -0.0014306 | 0.037411 |  |  | 0.7503 | 0.048411 | 0.018411 | 0.015411 |  |  | 0.10873 | 0.057646 | 0.087369 | 0.11022 | 0.054013 | 0.10538 |  |  | 0.17219 | 0.083454 | 0.16294 | 0.10222
9440.6088 | 0.00092297 | 0.010923 | 0.020754 | -0.00043646 | 0.0047964 | 0.039564 |  |  | 0.46192 |  | -0.0074365 | 0.00056354 |  |  | 0.10863 | 0.057453 | 0.087371 | 0.11055 | 0.053881 | 0.11383 |  |  | 0.1911 |  | 0.18172 | 0.10257
10000 | -0.0015686 | 0.028431 | 0.030178 | 0.033874 | 0.018646 | 0.066874 | -0.025126 | -0.047126 | 0.18376 | -0.039126 | 0.0018737 | 0.021874 | -0.088126 | 0.053874 | 0.10792 | 0.056103 | 0.086462 | 0.11832 | 0.052451 | 0.12688 | 0.087171 | 0.15362 | 0.20976 | 0.12001 | 0.20976 | 0.10197 | 0.10197 | 0.16168

## 2036126 Phase
[25 empty] | Estimated reference sensitivity phase | Std Unc (ERSP)
| Sensitivity Phase (4160 2036126) degrees | [11 empty] | Expanded Uncertainties (degrees) | [13 empty] | Phase difference relative to estimated reference sensitivity phase (degrees)
Frequency (Hz) | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG | (degrees) | (degrees) | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG
1.9953 | Sensitivity Phase (4160 2036126) degrees UK: 177.09; Sensitivity Phase (4160 2036126) degrees FR: 175.75; Sensitivity Phase (4160 2036126) degrees CH: 176.328; Sensitivity Phase (4160 2036126) degrees DE: 175.8; Expanded Uncertainties (degrees) UK: 3.3; Expanded Uncertainties (degrees) FR: 3.1; Expanded Uncertainties (degrees) CH: 0.495; Expanded Uncertainties (degrees) DE: 4.5 | 174.6993 | 0.33599 | 2.3907 | 1.0507 |  | 1.6287 | [5 empty] | 1.1007
2.5119 | Sensitivity Phase (4160 2036126) degrees UK: 176.87; Sensitivity Phase (4160 2036126) degrees FR: 176.38; Sensitivity Phase (4160 2036126) degrees CH: 176.538; Sensitivity Phase (4160 2036126) degrees DE: 176.2; Expanded Uncertainties (degrees) UK: 3.3; Expanded Uncertainties (degrees) FR: 2.7; Expanded Uncertainties (degrees) CH: 0.402; Expanded Uncertainties (degrees) DE: 4.5 | 175.6458 | 0.29947 | 1.2242 | 0.73419 |  | 0.89219 | [5 empty] | 0.55419
3.1623 | Sensitivity Phase (4160 2036126) degrees UK: 177.18; Sensitivity Phase (4160 2036126) degrees FR: 176.67; Sensitivity Phase (4160 2036126) degrees CH: 176.767; Sensitivity Phase (4160 2036126) degrees DE: 176.5; Expanded Uncertainties (degrees) UK: 3.3; Expanded Uncertainties (degrees) FR: 2.4; Expanded Uncertainties (degrees) CH: 0.32; Expanded Uncertainties (degrees) DE: 4.5 | 175.9973 | 0.26411 | 1.1827 | 0.67273 |  | 0.76973 | [5 empty] | 0.50273
3.9811 | Sensitivity Phase (4160 2036126) degrees UK: 177.28; Sensitivity Phase (4160 2036126) degrees FR: 176.94; Sensitivity Phase (4160 2036126) degrees CH: 177.01; Sensitivity Phase (4160 2036126) degrees DE: 176.8; Expanded Uncertainties (degrees) UK: 1.5; Expanded Uncertainties (degrees) FR: 2.1; Expanded Uncertainties (degrees) CH: 0.257; Expanded Uncertainties (degrees) DE: 2.5 | 176.4623 | 0.21179 | 0.81765 | 0.47765 |  | 0.54765 | [5 empty] | 0.33765
5.0119 | Sensitivity Phase (4160 2036126) degrees UK: 177.43; Sensitivity Phase (4160 2036126) degrees FR: 177.21; Sensitivity Phase (4160 2036126) degrees CH: 177.252; Sensitivity Phase (4160 2036126) degrees DE: 177.2; Expanded Uncertainties (degrees) UK: 1.5; Expanded Uncertainties (degrees) FR: 1.9; Expanded Uncertainties (degrees) CH: 0.207; Expanded Uncertainties (degrees) DE: 2.5 | 176.8749 | 0.18496 | 0.55513 | 0.33513 |  | 0.37713 | [5 empty] | 0.32513
6.3096 | Sensitivity Phase (4160 2036126) degrees UK: 177.57; Sensitivity Phase (4160 2036126) degrees FR: 177.47; Sensitivity Phase (4160 2036126) degrees CH: 177.501; Sensitivity Phase (4160 2036126) degrees DE: 177.4; Expanded Uncertainties (degrees) UK: 1.5; Expanded Uncertainties (degrees) FR: 1.7; Expanded Uncertainties (degrees) CH: 0.167; Expanded Uncertainties (degrees) DE: 2.5 | 177.2602 | 0.13934 | 0.30979 | 0.20979 |  | 0.24079 | [5 empty] | 0.13979
7.9433 | Sensitivity Phase (4160 2036126) degrees UK: 177.78; Sensitivity Phase (4160 2036126) degrees FR: 177.73; Sensitivity Phase (4160 2036126) degrees CH: 177.746; Sensitivity Phase (4160 2036126) degrees DE: 177.7; Expanded Uncertainties (degrees) UK: 0.8; Expanded Uncertainties (degrees) FR: 1.6; Expanded Uncertainties (degrees) CH: 0.132; Expanded Uncertainties (degrees) DE: 1.5 | 177.538 | 0.11436 | 0.24199 | 0.19199 |  | 0.20799 | [5 empty] | 0.16199
10 | Sensitivity Phase (4160 2036126) degrees UK: 177.97; Sensitivity Phase (4160 2036126) degrees FR: 177.96; Sensitivity Phase (4160 2036126) degrees CH: 177.969; Sensitivity Phase (4160 2036126) degrees DE: 177.9; Expanded Uncertainties (degrees) UK: 0.8; Expanded Uncertainties (degrees) FR: 1.4; Expanded Uncertainties (degrees) CH: 0.108; Expanded Uncertainties (degrees) DE: 1.5 | 177.8348 | 0.10056 | 0.13519 | 0.12519 |  | 0.13419 | [5 empty] | 0.065195
12.5893 | Sensitivity Phase (4160 2036126) degrees UK: 178.14; Sensitivity Phase (4160 2036126) degrees FR: 178.17; Sensitivity Phase (4160 2036126) degrees CH: 178.18; Sensitivity Phase (4160 2036126) degrees DE: 178.1; Expanded Uncertainties (degrees) UK: 0.8; Expanded Uncertainties (degrees) FR: 1.2; Expanded Uncertainties (degrees) CH: 0.084; Expanded Uncertainties (degrees) DE: 1.5 | 178.0301 | 0.089149 | 0.1099 | 0.1399 |  | 0.1499 | [5 empty] | 0.069898
15.8489 | Sensitivity Phase (4160 2036126) degrees UK: 178.3; Sensitivity Phase (4160 2036126) degrees FR: 178.34; Sensitivity Phase (4160 2036126) degrees CH: 178.355; Sensitivity Phase (4160 2036126) degrees DE: 178.3; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 1.1; Expanded Uncertainties (degrees) CH: 0.069; Expanded Uncertainties (degrees) DE: 1 | 178.2133 | 0.085014 | 0.086693 | 0.12669 |  | 0.14169 | [5 empty] | 0.086693
19.9526 | Sensitivity Phase (4160 2036126) degrees UK: 178.51; Sensitivity Phase (4160 2036126) degrees FR: 178.5; Sensitivity Phase (4160 2036126) degrees IT: 178.7; Sensitivity Phase (4160 2036126) degrees CH: 178.507; Sensitivity Phase (4160 2036126) degrees SE: 178.21; Sensitivity Phase (4160 2036126) degrees DE: 178.5; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 1; Expanded Uncertainties (degrees) IT: 1.6; Expanded Uncertainties (degrees) CH: 0.058; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 1 | 178.4315 | 0.16822 | 0.078523 | 0.068523 | 0.26852 | 0.075523 |  |  |  |  | -0.22148 | 0.068523
21.1349 | Sensitivity Phase (4160 2036126) degrees UK: 178.54; Sensitivity Phase (4160 2036126) degrees FR: 178.53; Sensitivity Phase (4160 2036126) degrees IT: 178.7; Sensitivity Phase (4160 2036126) degrees CH: 178.538; Sensitivity Phase (4160 2036126) degrees SE: 178.28; Sensitivity Phase (4160 2036126) degrees DE: 178.5; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 1; Expanded Uncertainties (degrees) IT: 1.6; Expanded Uncertainties (degrees) CH: 0.058; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 1 | 178.5344 | 0.16753 | 0.0056415 | -0.0043585 | 0.16564 | 0.0036415 |  |  |  |  | -0.25436 | -0.034358
22.3872 | Sensitivity Phase (4160 2036126) degrees UK: 178.57; Sensitivity Phase (4160 2036126) degrees FR: 178.56; Sensitivity Phase (4160 2036126) degrees IT: 178.7; Sensitivity Phase (4160 2036126) degrees CH: 178.571; Sensitivity Phase (4160 2036126) degrees SE: 178.3; Sensitivity Phase (4160 2036126) degrees DE: 178.6; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 1; Expanded Uncertainties (degrees) IT: 1.5; Expanded Uncertainties (degrees) CH: 0.048; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 1 | 178.562 | 0.16638 | 0.007989 | -0.002011 | 0.13799 | 0.008989 |  |  |  |  | -0.26201 | 0.037989
23.7137 | Sensitivity Phase (4160 2036126) degrees UK: 178.58; Sensitivity Phase (4160 2036126) degrees FR: 178.59; Sensitivity Phase (4160 2036126) degrees IT: 178.7; Sensitivity Phase (4160 2036126) degrees CH: 178.606; Sensitivity Phase (4160 2036126) degrees SE: 178.36; Sensitivity Phase (4160 2036126) degrees DE: 178.6; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 1; Expanded Uncertainties (degrees) IT: 1.4; Expanded Uncertainties (degrees) CH: 0.045; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 1 | 178.5859 | 0.16517 | -0.0059482 | 0.0040518 | 0.11405 | 0.020052 |  |  |  |  | -0.22595 | 0.014052
25.1189 | Sensitivity Phase (4160 2036126) degrees UK: 178.61; Sensitivity Phase (4160 2036126) degrees FR: 178.62; Sensitivity Phase (4160 2036126) degrees IT: 178.7; Sensitivity Phase (4160 2036126) degrees CH: 178.614; Sensitivity Phase (4160 2036126) degrees SE: 178.38; Sensitivity Phase (4160 2036126) degrees DE: 178.6; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.9; Expanded Uncertainties (degrees) IT: 1.3; Expanded Uncertainties (degrees) CH: 0.063; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 1 | 178.6153 | 0.15444 | -0.0052887 | 0.0047113 | 0.084711 | -0.0012887 |  |  |  |  | -0.23529 | -0.015289
26.6073 | Sensitivity Phase (4160 2036126) degrees UK: 178.63; Sensitivity Phase (4160 2036126) degrees FR: 178.64; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.664; Sensitivity Phase (4160 2036126) degrees SE: 178.4; Sensitivity Phase (4160 2036126) degrees DE: 178.6; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.9; Expanded Uncertainties (degrees) IT: 1.3; Expanded Uncertainties (degrees) CH: 0.036; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 1 | 178.6549 | 0.15441 | -0.02487 | -0.01487 | 0.14513 | 0.0091305 |  |  |  |  | -0.25487 | -0.05487
28.1838 | Sensitivity Phase (4160 2036126) degrees UK: 178.65; Sensitivity Phase (4160 2036126) degrees FR: 178.67; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.654; Sensitivity Phase (4160 2036126) degrees SE: 178.44; Sensitivity Phase (4160 2036126) degrees DE: 178.7; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.9; Expanded Uncertainties (degrees) IT: 1.2; Expanded Uncertainties (degrees) CH: 0.052; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 1 | 178.6742 | 0.15329 | -0.024244 | -0.0042439 | 0.12576 | -0.020244 |  |  |  |  | -0.23424 | 0.025756
29.8538 | Sensitivity Phase (4160 2036126) degrees UK: 178.68; Sensitivity Phase (4160 2036126) degrees FR: 178.69; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.708; Sensitivity Phase (4160 2036126) degrees SE: 178.48; Sensitivity Phase (4160 2036126) degrees DE: 178.7; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.9; Expanded Uncertainties (degrees) IT: 1.1; Expanded Uncertainties (degrees) CH: 0.039; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 1 | 178.6871 | 0.15184 | -0.0071139 | 0.0028861 | 0.11289 | 0.020886 |  |  |  |  | -0.20711 | 0.012886
31.6228 | Sensitivity Phase (4160 2036126) degrees UK: 178.69; Sensitivity Phase (4160 2036126) degrees FR: 178.71; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.735; Sensitivity Phase (4160 2036126) degrees SE: 178.5; Sensitivity Phase (4160 2036126) degrees DE: 178.7; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.8; Expanded Uncertainties (degrees) IT: 1.1; Expanded Uncertainties (degrees) CH: 0.031; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.6988 | 0.15197 | -0.0088162 | 0.011184 | 0.10118 | 0.036184 |  |  |  |  | -0.19882 | 0.0011838
33.4965 | Sensitivity Phase (4160 2036126) degrees UK: 178.7; Sensitivity Phase (4160 2036126) degrees FR: 178.73; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.743; Sensitivity Phase (4160 2036126) degrees SE: 178.53; Sensitivity Phase (4160 2036126) degrees DE: 178.7; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.8; Expanded Uncertainties (degrees) IT: 1; Expanded Uncertainties (degrees) CH: 0.029; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7098 | 0.1497 | -0.0097713 | 0.020229 | 0.090229 | 0.033229 |  |  |  |  | -0.17977 | -0.0097713
35.4813 | Sensitivity Phase (4160 2036126) degrees UK: 178.73; Sensitivity Phase (4160 2036126) degrees FR: 178.75; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.766; Sensitivity Phase (4160 2036126) degrees SE: 178.55; Sensitivity Phase (4160 2036126) degrees DE: 178.7; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.8; Expanded Uncertainties (degrees) IT: 1; Expanded Uncertainties (degrees) CH: 0.033; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7214 | 0.14967 | 0.0085893 | 0.028589 | 0.078589 | 0.044589 |  |  |  |  | -0.17141 | -0.021411
37.5837 | Sensitivity Phase (4160 2036126) degrees UK: 178.74; Sensitivity Phase (4160 2036126) degrees FR: 178.77; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.783; Sensitivity Phase (4160 2036126) degrees SE: 178.57; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.8; Expanded Uncertainties (degrees) IT: 0.9; Expanded Uncertainties (degrees) CH: 0.026; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7397 | 0.14735 | 0.00026502 | 0.030265 | 0.060265 | 0.043265 |  |  |  |  | -0.16973 | 0.060265
39.8107 | Sensitivity Phase (4160 2036126) degrees UK: 178.75; Sensitivity Phase (4160 2036126) degrees FR: 178.78; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.788; Sensitivity Phase (4160 2036126) degrees SE: 178.59; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.8; Expanded Uncertainties (degrees) IT: 0.9; Expanded Uncertainties (degrees) CH: 0.029; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7707 | 0.1472 | -0.020705 | 0.0092953 | 0.029295 | 0.017295 |  |  |  |  | -0.1807 | 0.029295
42.1697 | Sensitivity Phase (4160 2036126) degrees UK: 178.77; Sensitivity Phase (4160 2036126) degrees FR: 178.79; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.801; Sensitivity Phase (4160 2036126) degrees SE: 178.62; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.8; Expanded Uncertainties (degrees) IT: 0.8; Expanded Uncertainties (degrees) CH: 0.027; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.762 | 0.1441 | 0.0079602 | 0.02796 | 0.03796 | 0.03896 |  |  |  |  | -0.14204 | 0.03796
44.6684 | Sensitivity Phase (4160 2036126) degrees UK: 178.78; Sensitivity Phase (4160 2036126) degrees FR: 178.8; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.81; Sensitivity Phase (4160 2036126) degrees SE: 178.63; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.8; Expanded Uncertainties (degrees) IT: 0.8; Expanded Uncertainties (degrees) CH: 0.026; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7604 | 0.14406 | 0.019578 | 0.039578 | 0.039578 | 0.049578 |  |  |  |  | -0.13042 | 0.039578
47.3151 | Sensitivity Phase (4160 2036126) degrees UK: 178.78; Sensitivity Phase (4160 2036126) degrees FR: 178.81; Sensitivity Phase (4160 2036126) degrees IT: 178.9; Sensitivity Phase (4160 2036126) degrees CH: 178.819; Sensitivity Phase (4160 2036126) degrees SE: 178.64; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.8; Expanded Uncertainties (degrees) IT: 0.7; Expanded Uncertainties (degrees) CH: 0.037; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7834 | 0.14059 | -0.0034184 | 0.026582 | 0.11658 | 0.035582 |  |  |  |  | -0.14342 | 0.016582
50.1187 | Sensitivity Phase (4160 2036126) degrees UK: 178.79; Sensitivity Phase (4160 2036126) degrees FR: 178.82; Sensitivity Phase (4160 2036126) degrees IT: 178.9; Sensitivity Phase (4160 2036126) degrees CH: 178.787; Sensitivity Phase (4160 2036126) degrees SE: 178.64; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.7; Expanded Uncertainties (degrees) IT: 0.7; Expanded Uncertainties (degrees) CH: 0.036; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 179.1479 | 0.14055 | -0.35793 | -0.32793 | -0.24793 | -0.36093 |  |  |  |  | -0.50793 | -0.34793
53.0884 | Sensitivity Phase (4160 2036126) degrees UK: 178.78; Sensitivity Phase (4160 2036126) degrees FR: 178.82; Sensitivity Phase (4160 2036126) degrees IT: 178.9; Sensitivity Phase (4160 2036126) degrees CH: 178.84; Sensitivity Phase (4160 2036126) degrees SE: 178.65; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.7; Expanded Uncertainties (degrees) IT: 0.7; Expanded Uncertainties (degrees) CH: 0.028; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.8055 | 0.14034 | -0.025546 | 0.014454 | 0.094454 | 0.034454 |  |  |  |  | -0.15555 | -0.0055462
56.2341 | Sensitivity Phase (4160 2036126) degrees UK: 178.78; Sensitivity Phase (4160 2036126) degrees FR: 178.83; Sensitivity Phase (4160 2036126) degrees IT: 178.9; Sensitivity Phase (4160 2036126) degrees CH: 178.834; Sensitivity Phase (4160 2036126) degrees SE: 178.67; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.7; Expanded Uncertainties (degrees) IT: 0.7; Expanded Uncertainties (degrees) CH: 0.022; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.8205 | 0.14012 | -0.040549 | 0.0094513 | 0.079451 | 0.013451 |  |  |  |  | -0.15055 | -0.020549
59.5662 | Sensitivity Phase (4160 2036126) degrees UK: 178.78; Sensitivity Phase (4160 2036126) degrees FR: 178.83; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.839; Sensitivity Phase (4160 2036126) degrees SE: 178.67; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.7; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.02; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7899 | 0.13441 | -0.0098836 | 0.040116 | 0.010116 | 0.049116 |  |  |  |  | -0.11988 | 0.010116
63.0957 | Sensitivity Phase (4160 2036126) degrees UK: 178.79; Sensitivity Phase (4160 2036126) degrees FR: 178.83; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.832; Sensitivity Phase (4160 2036126) degrees SE: 178.68; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.7; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.018; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.8114 | 0.13422 | -0.021397 | 0.018603 | -0.011397 | 0.020603 |  |  |  |  | -0.1314 | -0.011397
66.8344 | Sensitivity Phase (4160 2036126) degrees UK: 178.78; Sensitivity Phase (4160 2036126) degrees FR: 178.82; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.83; Sensitivity Phase (4160 2036126) degrees SE: 178.68; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.7; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.017; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.8037 | 0.13419 | -0.023714 | 0.016286 | -0.0037139 | 0.026286 |  |  |  |  | -0.12371 | -0.0037139
70.7946 | Sensitivity Phase (4160 2036126) degrees UK: 178.78; Sensitivity Phase (4160 2036126) degrees FR: 178.82; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.825; Sensitivity Phase (4160 2036126) degrees SE: 178.67; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.7; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.018; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.8017 | 0.13395 | -0.021652 | 0.018348 | -0.0016519 | 0.023348 |  |  |  |  | -0.13165 | -0.0016519
74.9894 | Sensitivity Phase (4160 2036126) degrees UK: 178.79; Sensitivity Phase (4160 2036126) degrees FR: 178.81; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.816; Sensitivity Phase (4160 2036126) degrees SE: 178.67; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.7; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.017; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.8055 | 0.1339 | -0.015481 | 0.0045189 | -0.0054811 | 0.010519 |  |  |  |  | -0.13548 | -0.0054811
79.4328 | Sensitivity Phase (4160 2036126) degrees UK: 178.78; Sensitivity Phase (4160 2036126) degrees FR: 178.8; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.806; Sensitivity Phase (4160 2036126) degrees SE: 178.66; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.6; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.019; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7966 | 0.13385 | -0.016624 | 0.0033764 | 0.0033764 | 0.0093764 |  |  |  |  | -0.13662 | 0.0033764
84.1395 | Sensitivity Phase (4160 2036126) degrees UK: 178.75; Sensitivity Phase (4160 2036126) degrees FR: 178.79; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.794; Sensitivity Phase (4160 2036126) degrees SE: 178.65; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.6; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.017; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7651 | 0.13377 | -0.015145 | 0.024855 | 0.034855 | 0.028855 |  |  |  |  | -0.11515 | 0.034855
89.1251 | Sensitivity Phase (4160 2036126) degrees UK: 178.73; Sensitivity Phase (4160 2036126) degrees FR: 178.77; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.781; Sensitivity Phase (4160 2036126) degrees SE: 178.65; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.6; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.015; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7516 | 0.13376 | -0.021551 | 0.018449 | 0.048449 | 0.029449 |  |  |  |  | -0.10155 | 0.048449
94.4061 | Sensitivity Phase (4160 2036126) degrees UK: 178.72; Sensitivity Phase (4160 2036126) degrees FR: 178.75; Sensitivity Phase (4160 2036126) degrees IT: 178.8; Sensitivity Phase (4160 2036126) degrees CH: 178.767; Sensitivity Phase (4160 2036126) degrees SE: 178.63; Sensitivity Phase (4160 2036126) degrees DE: 178.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.6; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.014; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7608 | 0.13421 | -0.040791 | -0.010791 | 0.039209 | 0.006209 |  |  |  |  | -0.13079 | 0.039209
100 | Sensitivity Phase (4160 2036126) degrees UK: 178.69; Sensitivity Phase (4160 2036126) degrees FR: 178.74; Sensitivity Phase (4160 2036126) degrees IT: 178.7; Sensitivity Phase (4160 2036126) degrees CH: 178.744; Sensitivity Phase (4160 2036126) degrees SE: 178.57; Sensitivity Phase (4160 2036126) degrees DE: 178.7; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.016; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7095 | 0.13416 | -0.019495 | 0.030505 | -0.0094952 | 0.034505 |  |  |  |  | -0.1395 | -0.0094952
105.9254 | Sensitivity Phase (4160 2036126) degrees UK: 178.69; Sensitivity Phase (4160 2036126) degrees FR: 178.71; Sensitivity Phase (4160 2036126) degrees IT: 178.7; Sensitivity Phase (4160 2036126) degrees CH: 178.722; Sensitivity Phase (4160 2036126) degrees SE: 178.59; Sensitivity Phase (4160 2036126) degrees DE: 178.7; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.017; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.7113 | 0.13403 | -0.021307 | -0.0013073 | -0.011307 | 0.010693 |  |  |  |  | -0.12131 | -0.011307
112.2018 | Sensitivity Phase (4160 2036126) degrees UK: 178.66; Sensitivity Phase (4160 2036126) degrees FR: 178.69; Sensitivity Phase (4160 2036126) degrees IT: 178.7; Sensitivity Phase (4160 2036126) degrees CH: 178.704; Sensitivity Phase (4160 2036126) degrees SE: 178.57; Sensitivity Phase (4160 2036126) degrees DE: 178.7; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.013; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.673 | 0.134 | -0.012991 | 0.017009 | 0.027009 | 0.031009 |  |  |  |  | -0.10299 | 0.027009
118.8502 | Sensitivity Phase (4160 2036126) degrees UK: 178.62; Sensitivity Phase (4160 2036126) degrees FR: 178.66; Sensitivity Phase (4160 2036126) degrees IT: 178.7; Sensitivity Phase (4160 2036126) degrees CH: 178.67; Sensitivity Phase (4160 2036126) degrees SE: 178.55; Sensitivity Phase (4160 2036126) degrees DE: 178.7; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.025; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.6709 | 0.134 | -0.05091 | -0.01091 | 0.02909 | -0.0009096 |  |  |  |  | -0.12091 | 0.02909
125.8925 | Sensitivity Phase (4160 2036126) degrees UK: 178.59; Sensitivity Phase (4160 2036126) degrees FR: 178.63; Sensitivity Phase (4160 2036126) degrees IT: 178.6; Sensitivity Phase (4160 2036126) degrees CH: 178.635; Sensitivity Phase (4160 2036126) degrees SE: 178.52; Sensitivity Phase (4160 2036126) degrees DE: 178.6; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.016; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.6026 | 0.134 | -0.01255 | 0.02745 | -0.0025505 | 0.03245 |  |  |  |  | -0.08255 | -0.0025505
133.3521 | Sensitivity Phase (4160 2036126) degrees UK: 178.55; Sensitivity Phase (4160 2036126) degrees FR: 178.59; Sensitivity Phase (4160 2036126) degrees IT: 178.6; Sensitivity Phase (4160 2036126) degrees CH: 178.592; Sensitivity Phase (4160 2036126) degrees SE: 178.49; Sensitivity Phase (4160 2036126) degrees DE: 178.6; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.02; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.5791 | 0.13396 | -0.029087 | 0.010913 | 0.020913 | 0.012913 |  |  |  |  | -0.089087 | 0.020913
141.2538 | Sensitivity Phase (4160 2036126) degrees UK: 178.5; Sensitivity Phase (4160 2036126) degrees FR: 178.56; Sensitivity Phase (4160 2036126) degrees IT: 178.5; Sensitivity Phase (4160 2036126) degrees CH: 178.546; Sensitivity Phase (4160 2036126) degrees SE: 178.45; Sensitivity Phase (4160 2036126) degrees DE: 178.6; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.061; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.5246 | 0.13396 | -0.024604 | 0.035396 | -0.024604 | 0.021396 |  |  |  |  | -0.074604 | 0.075396
149.6236 | Sensitivity Phase (4160 2036126) degrees UK: 178.49; Sensitivity Phase (4160 2036126) degrees FR: 178.51; Sensitivity Phase (4160 2036126) degrees IT: 178.5; Sensitivity Phase (4160 2036126) degrees CH: 178.489; Sensitivity Phase (4160 2036126) degrees SE: 178.41; Sensitivity Phase (4160 2036126) degrees DE: 178.5; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.042; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.5091 | 0.1338 | -0.019072 | 0.00092767 | -0.0090723 | -0.020072 |  |  |  |  | -0.099072 | -0.0090723
158.4893 | Sensitivity Phase (4160 2036126) degrees UK: 178.44; Sensitivity Phase (4160 2036126) degrees FR: 178.47; Sensitivity Phase (4160 2036126) degrees IT: 178.5; Sensitivity Phase (4160 2036126) degrees CH: 178.474; Sensitivity Phase (4160 2036126) degrees SE: 178.37; Sensitivity Phase (4160 2036126) degrees DE: 178.5; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.017; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.4566 | 0.1338 | -0.016561 | 0.013439 | 0.043439 | 0.017439 |  |  |  |  | -0.086561 | 0.043439
167.8804 | Sensitivity Phase (4160 2036126) degrees UK: 178.37; Sensitivity Phase (4160 2036126) degrees FR: 178.42; Sensitivity Phase (4160 2036126) degrees IT: 178.4; Sensitivity Phase (4160 2036126) degrees CH: 178.427; Sensitivity Phase (4160 2036126) degrees SE: 178.32; Sensitivity Phase (4160 2036126) degrees DE: 178.4; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.018; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.3945 | 0.1338 | -0.024505 | 0.025495 | 0.0054949 | 0.032495 |  |  |  |  | -0.074505 | 0.0054949
177.8279 | Sensitivity Phase (4160 2036126) degrees UK: 178.35; Sensitivity Phase (4160 2036126) degrees FR: 178.36; Sensitivity Phase (4160 2036126) degrees IT: 178.4; Sensitivity Phase (4160 2036126) degrees CH: 178.372; Sensitivity Phase (4160 2036126) degrees SE: 178.27; Sensitivity Phase (4160 2036126) degrees DE: 178.4; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.018; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.3666 | 0.13375 | -0.016555 | -0.006555 | 0.033445 | 0.005445 |  |  |  |  | -0.096555 | 0.033445
188.3649 | Sensitivity Phase (4160 2036126) degrees UK: 178.26; Sensitivity Phase (4160 2036126) degrees FR: 178.31; Sensitivity Phase (4160 2036126) degrees IT: 178.3; Sensitivity Phase (4160 2036126) degrees CH: 178.314; Sensitivity Phase (4160 2036126) degrees SE: 178.21; Sensitivity Phase (4160 2036126) degrees DE: 178.3; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.02; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.2878 | 0.13375 | -0.027821 | 0.022179 | 0.012179 | 0.026179 |  |  |  |  | -0.077821 | 0.012179
199.5262 | Sensitivity Phase (4160 2036126) degrees UK: 178.2; Sensitivity Phase (4160 2036126) degrees FR: 178.25; Sensitivity Phase (4160 2036126) degrees IT: 178.2; Sensitivity Phase (4160 2036126) degrees CH: 178.256; Sensitivity Phase (4160 2036126) degrees SE: 178.16; Sensitivity Phase (4160 2036126) degrees DE: 178.2; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.02; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.2241 | 0.1355 | -0.0241 | 0.0259 | -0.0241 | 0.0319 |  |  |  |  | -0.0641 | -0.0241
211.3489 | Sensitivity Phase (4160 2036126) degrees UK: 178.12; Sensitivity Phase (4160 2036126) degrees FR: 178.18; Sensitivity Phase (4160 2036126) degrees IT: 178.2; Sensitivity Phase (4160 2036126) degrees CH: 178.185; Sensitivity Phase (4160 2036126) degrees SE: 178.09; Sensitivity Phase (4160 2036126) degrees DE: 178.2; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.021; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.1496 | 0.1355 | -0.029644 | 0.030356 | 0.050356 | 0.035356 |  |  |  |  | -0.059644 | 0.050356
223.8721 | Sensitivity Phase (4160 2036126) degrees UK: 178.06; Sensitivity Phase (4160 2036126) degrees FR: 178.11; Sensitivity Phase (4160 2036126) degrees IT: 178.1; Sensitivity Phase (4160 2036126) degrees CH: 178.118; Sensitivity Phase (4160 2036126) degrees SE: 178.02; Sensitivity Phase (4160 2036126) degrees DE: 178.1; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.022; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 178.0836 | 0.1355 | -0.023567 | 0.026433 | 0.016433 | 0.034433 |  |  |  |  | -0.063567 | 0.016433
237.1374 | Sensitivity Phase (4160 2036126) degrees UK: 177.97; Sensitivity Phase (4160 2036126) degrees FR: 178.03; Sensitivity Phase (4160 2036126) degrees IT: 178; Sensitivity Phase (4160 2036126) degrees CH: 178.037; Sensitivity Phase (4160 2036126) degrees SE: 177.94; Sensitivity Phase (4160 2036126) degrees DE: 178; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.024; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 177.9968 | 0.1355 | -0.026754 | 0.033246 | 0.0032458 | 0.040246 |  |  |  |  | -0.056754 | 0.0032458
251.1886 | Sensitivity Phase (4160 2036126) degrees UK: 177.9; Sensitivity Phase (4160 2036126) degrees FR: 177.95; Sensitivity Phase (4160 2036126) degrees IT: 177.9; Sensitivity Phase (4160 2036126) degrees CH: 177.956; Sensitivity Phase (4160 2036126) degrees SE: 177.86; Sensitivity Phase (4160 2036126) degrees DE: 177.9; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.04; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 177.9455 | 0.13549 | -0.045539 | 0.0044609 | -0.045539 | 0.010461 |  |  |  |  | -0.085539 | -0.045539
266.0725 | Sensitivity Phase (4160 2036126) degrees UK: 177.8; Sensitivity Phase (4160 2036126) degrees FR: 177.86; Sensitivity Phase (4160 2036126) degrees IT: 177.8; Sensitivity Phase (4160 2036126) degrees CH: 177.858; Sensitivity Phase (4160 2036126) degrees SE: 177.78; Sensitivity Phase (4160 2036126) degrees DE: 177.9; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.032; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 177.8332 | 0.13546 | -0.033209 | 0.026791 | -0.033209 | 0.024791 |  |  |  |  | -0.053209 | 0.066791
281.8383 | Sensitivity Phase (4160 2036126) degrees UK: 177.7; Sensitivity Phase (4160 2036126) degrees FR: 177.76; Sensitivity Phase (4160 2036126) degrees IT: 177.7; Sensitivity Phase (4160 2036126) degrees CH: 177.769; Sensitivity Phase (4160 2036126) degrees SE: 177.68; Sensitivity Phase (4160 2036126) degrees DE: 177.8; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.027; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 177.7266 | 0.13546 | -0.026606 | 0.033394 | -0.026606 | 0.042394 |  |  |  |  | -0.046606 | 0.073394
298.5383 | Sensitivity Phase (4160 2036126) degrees UK: 177.58; Sensitivity Phase (4160 2036126) degrees FR: 177.66; Sensitivity Phase (4160 2036126) degrees IT: 177.6; Sensitivity Phase (4160 2036126) degrees CH: 177.666; Sensitivity Phase (4160 2036126) degrees SE: 177.58; Sensitivity Phase (4160 2036126) degrees DE: 177.7; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.03; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 177.6203 | 0.13541 | -0.040293 | 0.039707 | -0.020293 | 0.045707 |  |  |  |  | -0.040293 | 0.079707
316.2278 | Sensitivity Phase (4160 2036126) degrees UK: 177.5; Sensitivity Phase (4160 2036126) degrees FR: 177.55; Sensitivity Phase (4160 2036126) degrees IT: 177.5; Sensitivity Phase (4160 2036126) degrees CH: 177.556; Sensitivity Phase (4160 2036126) degrees SE: 177.48; Sensitivity Phase (4160 2036126) degrees DE: 177.6; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.032; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 177.5292 | 0.13541 | -0.029228 | 0.020772 | -0.029228 | 0.026772 |  |  |  |  | -0.049228 | 0.070772
334.9654 | Sensitivity Phase (4160 2036126) degrees UK: 177.38; Sensitivity Phase (4160 2036126) degrees FR: 177.44; Sensitivity Phase (4160 2036126) degrees IT: 177.4; Sensitivity Phase (4160 2036126) degrees CH: 177.439; Sensitivity Phase (4160 2036126) degrees SE: 177.36; Sensitivity Phase (4160 2036126) degrees DE: 177.4; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.033; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 177.4086 | 0.13541 | -0.028572 | 0.031428 | -0.0085716 | 0.030428 |  |  |  |  | -0.048572 | -0.0085716
354.8134 | Sensitivity Phase (4160 2036126) degrees UK: 177.27; Sensitivity Phase (4160 2036126) degrees FR: 177.31; Sensitivity Phase (4160 2036126) degrees IT: 177.3; Sensitivity Phase (4160 2036126) degrees CH: 177.318; Sensitivity Phase (4160 2036126) degrees SE: 177.24; Sensitivity Phase (4160 2036126) degrees DE: 177.3; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.035; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 177.2778 | 0.13541 | -0.0077767 | 0.032223 | 0.022223 | 0.040223 |  |  |  |  | -0.037777 | 0.022223
375.8374 | Sensitivity Phase (4160 2036126) degrees UK: 177.15; Sensitivity Phase (4160 2036126) degrees FR: 177.18; Sensitivity Phase (4160 2036126) degrees IT: 177.2; Sensitivity Phase (4160 2036126) degrees CH: 177.182; Sensitivity Phase (4160 2036126) degrees SE: 177.11; Sensitivity Phase (4160 2036126) degrees DE: 177.2; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.037; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 177.1843 | 0.14161 | -0.034285 | -0.004285 | 0.015715 | -0.002285 |  |  |  |  | -0.074285 | 0.015715
398.1072 | Sensitivity Phase (4160 2036126) degrees UK: 176.99; Sensitivity Phase (4160 2036126) degrees FR: 177.04; Sensitivity Phase (4160 2036126) degrees IT: 177; Sensitivity Phase (4160 2036126) degrees CH: 177.044; Sensitivity Phase (4160 2036126) degrees SE: 176.97; Sensitivity Phase (4160 2036126) degrees DE: 177; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.04; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 176.9889 | 0.14161 | 0.0011365 | 0.051136 | 0.011136 | 0.055136 |  |  |  |  | -0.018864 | 0.011136
421.6965 | Sensitivity Phase (4160 2036126) degrees UK: 176.87; Sensitivity Phase (4160 2036126) degrees FR: 176.89; Sensitivity Phase (4160 2036126) degrees IT: 176.9; Sensitivity Phase (4160 2036126) degrees CH: 176.89; Sensitivity Phase (4160 2036126) degrees SE: 176.82; Sensitivity Phase (4160 2036126) degrees DE: 176.9; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.042; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 176.902 | 0.14161 | -0.032045 | -0.012045 | -0.0020446 | -0.012045 |  |  |  |  | -0.082045 | -0.0020446
446.6836 | Sensitivity Phase (4160 2036126) degrees UK: 176.7; Sensitivity Phase (4160 2036126) degrees FR: 176.73; Sensitivity Phase (4160 2036126) degrees IT: 176.7; Sensitivity Phase (4160 2036126) degrees CH: 176.737; Sensitivity Phase (4160 2036126) degrees SE: 176.66; Sensitivity Phase (4160 2036126) degrees DE: 176.7; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.045; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 176.7352 | 0.14161 | -0.035223 | -0.0052232 | -0.035223 | 0.0017768 |  |  |  |  | -0.075223 | -0.035223
473.1513 | Sensitivity Phase (4160 2036126) degrees UK: 176.51; Sensitivity Phase (4160 2036126) degrees FR: 176.56; Sensitivity Phase (4160 2036126) degrees IT: 176.5; Sensitivity Phase (4160 2036126) degrees CH: 176.564; Sensitivity Phase (4160 2036126) degrees SE: 176.49; Sensitivity Phase (4160 2036126) degrees DE: 176.6; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.047; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 176.5127 | 0.14161 | -0.0026682 | 0.047332 | -0.012668 | 0.051332 |  |  |  |  | -0.022668 | 0.087332
501.1872 | Sensitivity Phase (4160 2036126) degrees UK: 176.34; Sensitivity Phase (4160 2036126) degrees FR: 176.38; Sensitivity Phase (4160 2036126) degrees IT: 176.4; Sensitivity Phase (4160 2036126) degrees CH: 176.383; Sensitivity Phase (4160 2036126) degrees SE: 176.32; Sensitivity Phase (4160 2036126) degrees DE: 176.4; Expanded Uncertainties (degrees) UK: 0.4; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.05; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 176.3673 | 0.14161 | -0.027281 | 0.012719 | 0.032719 | 0.015719 |  |  |  |  | -0.047281 | 0.032719
530.8844 | Sensitivity Phase (4160 2036126) degrees UK: 176.14; Sensitivity Phase (4160 2036126) degrees FR: 176.19; Sensitivity Phase (4160 2036126) degrees IT: 176.2; Sensitivity Phase (4160 2036126) degrees CH: 176.191; Sensitivity Phase (4160 2036126) degrees SE: 176.13; Sensitivity Phase (4160 2036126) degrees DE: 176.2; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.053; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 176.1731 | 0.16338 | -0.033063 | 0.016937 | 0.026937 | 0.017937 |  |  |  |  | -0.043063 | 0.026937
562.3413 | Sensitivity Phase (4160 2036126) degrees UK: 175.93; Sensitivity Phase (4160 2036126) degrees FR: 175.98; Sensitivity Phase (4160 2036126) degrees IT: 176; Sensitivity Phase (4160 2036126) degrees CH: 175.982; Sensitivity Phase (4160 2036126) degrees SE: 175.92; Sensitivity Phase (4160 2036126) degrees DE: 176; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.057; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 175.962 | 0.16338 | -0.031981 | 0.018019 | 0.038019 | 0.020019 |  |  |  |  | -0.041981 | 0.038019
595.6621 | Sensitivity Phase (4160 2036126) degrees UK: 175.7; Sensitivity Phase (4160 2036126) degrees FR: 175.77; Sensitivity Phase (4160 2036126) degrees IT: 175.7; Sensitivity Phase (4160 2036126) degrees CH: 175.768; Sensitivity Phase (4160 2036126) degrees SE: 175.71; Sensitivity Phase (4160 2036126) degrees DE: 175.8; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.059; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 175.7304 | 0.16338 | -0.030399 | 0.039601 | -0.030399 | 0.037601 |  |  |  |  | -0.020399 | 0.069601
630.9573 | Sensitivity Phase (4160 2036126) degrees UK: 175.47; Sensitivity Phase (4160 2036126) degrees FR: 175.53; Sensitivity Phase (4160 2036126) degrees IT: 175.5; Sensitivity Phase (4160 2036126) degrees CH: 175.535; Sensitivity Phase (4160 2036126) degrees SE: 175.48; Sensitivity Phase (4160 2036126) degrees DE: 175.5; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.063; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 175.4941 | 0.16338 | -0.02406 | 0.03594 | 0.0059401 | 0.04094 |  |  |  |  | -0.01406 | 0.0059401
668.3439 | Sensitivity Phase (4160 2036126) degrees UK: 175.22; Sensitivity Phase (4160 2036126) degrees FR: 175.29; Sensitivity Phase (4160 2036126) degrees IT: 175.3; Sensitivity Phase (4160 2036126) degrees CH: 175.289; Sensitivity Phase (4160 2036126) degrees SE: 175.23; Sensitivity Phase (4160 2036126) degrees DE: 175.3; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.067; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 175.287 | 0.16338 | -0.067003 | 0.0029972 | 0.012997 | 0.0019972 |  |  |  |  | -0.057003 | 0.012997
707.9458 | Sensitivity Phase (4160 2036126) degrees UK: 174.96; Sensitivity Phase (4160 2036126) degrees FR: 175.03; Sensitivity Phase (4160 2036126) degrees IT: 175; Sensitivity Phase (4160 2036126) degrees CH: 175.025; Sensitivity Phase (4160 2036126) degrees SE: 174.97; Sensitivity Phase (4160 2036126) degrees DE: 175; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.072; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 175.0213 | 0.16338 | -0.061349 | 0.0086514 | -0.021349 | 0.0036514 |  |  |  |  | -0.051349 | -0.021349
749.8942 | Sensitivity Phase (4160 2036126) degrees UK: 174.73; Sensitivity Phase (4160 2036126) degrees FR: 174.75; Sensitivity Phase (4160 2036126) degrees IT: 174.7; Sensitivity Phase (4160 2036126) degrees CH: 174.754; Sensitivity Phase (4160 2036126) degrees SE: 174.69; Sensitivity Phase (4160 2036126) degrees DE: 174.8; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.077; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 174.7697 | 0.16338 | -0.039677 | -0.019677 | -0.069677 | -0.015677 |  |  |  |  | -0.079677 | 0.030323
794.3282 | Sensitivity Phase (4160 2036126) degrees UK: 174.38; Sensitivity Phase (4160 2036126) degrees FR: 174.45; Sensitivity Phase (4160 2036126) degrees IT: 174.4; Sensitivity Phase (4160 2036126) degrees CH: 174.456; Sensitivity Phase (4160 2036126) degrees SE: 174.4; Sensitivity Phase (4160 2036126) degrees DE: 174.5; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.081; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 174.434 | 0.19795 | -0.054008 | 0.015992 | -0.034008 | 0.021992 |  |  |  |  | -0.034008 | 0.065992
841.3951 | Sensitivity Phase (4160 2036126) degrees UK: 174.07; Sensitivity Phase (4160 2036126) degrees FR: 174.14; Sensitivity Phase (4160 2036126) degrees IT: 174.1; Sensitivity Phase (4160 2036126) degrees CH: 174.141; Sensitivity Phase (4160 2036126) degrees SE: 174.09; Sensitivity Phase (4160 2036126) degrees DE: 174.1; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.086; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 174.1095 | 0.19795 | -0.039487 | 0.030513 | -0.0094867 | 0.031513 |  |  |  |  | -0.019487 | -0.0094867
891.2509 | Sensitivity Phase (4160 2036126) degrees UK: 173.75; Sensitivity Phase (4160 2036126) degrees FR: 173.8; Sensitivity Phase (4160 2036126) degrees IT: 173.8; Sensitivity Phase (4160 2036126) degrees CH: 173.808; Sensitivity Phase (4160 2036126) degrees SE: 173.76; Sensitivity Phase (4160 2036126) degrees DE: 173.8; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.091; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 173.8507 | 0.19795 | -0.10072 | -0.050718 | -0.050718 | -0.042718 |  |  |  |  | -0.090718 | -0.050718
944.0609 | Sensitivity Phase (4160 2036126) degrees UK: 173.44; Sensitivity Phase (4160 2036126) degrees FR: 173.45; Sensitivity Phase (4160 2036126) degrees IT: 173.4; Sensitivity Phase (4160 2036126) degrees CH: 173.449; Sensitivity Phase (4160 2036126) degrees SE: 173.4; Sensitivity Phase (4160 2036126) degrees DE: 173.5; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.096; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 173.4415 | 0.19795 | -0.0015401 | 0.0084599 | -0.04154 | 0.0074599 |  |  |  |  | -0.04154 | 0.05846
1000 | Sensitivity Phase (4160 2036126) degrees UK: 173.03; Sensitivity Phase (4160 2036126) degrees FR: 173.07; Sensitivity Phase (4160 2036126) degrees IT: 173.1; Sensitivity Phase (4160 2036126) degrees CH: 173.076; Sensitivity Phase (4160 2036126) degrees SE: 173.02; Sensitivity Phase (4160 2036126) degrees DE: 173.1; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 0.4; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.103; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 173.1097 | 0.19795 | -0.07965 | -0.03965 | -0.0096503 | -0.03365 |  |  |  |  | -0.08965 | -0.0096503
1059.2537 | Sensitivity Phase (4160 2036126) degrees UK: 172.59; Sensitivity Phase (4160 2036126) degrees FR: 172.66; Sensitivity Phase (4160 2036126) degrees IT: 172.7; Sensitivity Phase (4160 2036126) degrees CH: 172.67; Sensitivity Phase (4160 2036126) degrees SE: 172.62; Sensitivity Phase (4160 2036126) degrees DE: 172.7; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.11; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 172.6861 | 0.21865 | -0.096118 | -0.026118 | 0.013882 | -0.016118 |  |  |  |  | -0.066118 | 0.013882
1122.0185 | Sensitivity Phase (4160 2036126) degrees UK: 172.17; Sensitivity Phase (4160 2036126) degrees FR: 172.24; Sensitivity Phase (4160 2036126) degrees IT: 172.2; Sensitivity Phase (4160 2036126) degrees CH: 172.243; Sensitivity Phase (4160 2036126) degrees SE: 172.19; Sensitivity Phase (4160 2036126) degrees DE: 172.3; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.116; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 172.2249 | 0.21865 | -0.054884 | 0.015116 | -0.024884 | 0.018116 |  |  |  |  | -0.034884 | 0.075116
1188.5022 | Sensitivity Phase (4160 2036126) degrees UK: 171.74; Sensitivity Phase (4160 2036126) degrees FR: 171.77; Sensitivity Phase (4160 2036126) degrees IT: 171.8; Sensitivity Phase (4160 2036126) degrees CH: 171.788; Sensitivity Phase (4160 2036126) degrees SE: 171.74; Sensitivity Phase (4160 2036126) degrees DE: 171.8; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.124; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.4; Expanded Uncertainties (degrees) DE: 0.5 | 171.7953 | 0.21865 | -0.055266 | -0.025266 | 0.0047336 | -0.0072664 |  |  |  |  | -0.055266 | 0.0047336
1258.9254 | Sensitivity Phase (4160 2036126) degrees UK: 171.21; Sensitivity Phase (4160 2036126) degrees FR: 171.29; Sensitivity Phase (4160 2036126) degrees IT: 171.3; Sensitivity Phase (4160 2036126) degrees CH: 171.304; Sensitivity Phase (4160 2036126) degrees SE: 171.26; Sensitivity Phase (4160 2036126) degrees DE: 171.3; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.13; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 171.2822 | 0.21865 | -0.072234 | 0.0077664 | 0.017766 | 0.021766 |  |  |  |  | -0.022234 | 0.017766
1333.5214 | Sensitivity Phase (4160 2036126) degrees UK: 170.69; Sensitivity Phase (4160 2036126) degrees FR: 170.77; Sensitivity Phase (4160 2036126) degrees IT: 170.8; Sensitivity Phase (4160 2036126) degrees CH: 170.781; Sensitivity Phase (4160 2036126) degrees SE: 170.74; Sensitivity Phase (4160 2036126) degrees DE: 170.8; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.14; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 170.804 | 0.21865 | -0.11402 | -0.034024 | -0.0040241 | -0.023024 |  |  |  |  | -0.064024 | -0.0040241
1412.5375 | Sensitivity Phase (4160 2036126) degrees UK: 170.15; Sensitivity Phase (4160 2036126) degrees FR: 170.23; Sensitivity Phase (4160 2036126) degrees IT: 170.2; Sensitivity Phase (4160 2036126) degrees CH: 170.233; Sensitivity Phase (4160 2036126) degrees SE: 170.19; Sensitivity Phase (4160 2036126) degrees DE: 170.2; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.149; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 170.2394 | 0.21865 | -0.089429 | -0.0094291 | -0.039429 | -0.0064291 |  |  |  |  | -0.049429 | -0.039429
1496.2357 | Sensitivity Phase (4160 2036126) degrees UK: 169.54; Sensitivity Phase (4160 2036126) degrees FR: 169.65; Sensitivity Phase (4160 2036126) degrees IT: 169.6; Sensitivity Phase (4160 2036126) degrees CH: 169.645; Sensitivity Phase (4160 2036126) degrees SE: 169.61; Sensitivity Phase (4160 2036126) degrees DE: 169.7; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.158; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 169.6247 | 0.21864 | -0.08466 | 0.02534 | -0.02466 | 0.02034 |  |  |  |  | -0.01466 | 0.07534
1584.8932 | Sensitivity Phase (4160 2036126) degrees UK: 168.94; Sensitivity Phase (4160 2036126) degrees FR: 169.03; Sensitivity Phase (4160 2036126) degrees IT: 169; Sensitivity Phase (4160 2036126) degrees CH: 169.019; Sensitivity Phase (4160 2036126) degrees SE: 168.98; Sensitivity Phase (4160 2036126) degrees DE: 169; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.5; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.168; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 169.0564 | 0.25335 | -0.11644 | -0.026441 | -0.056441 | -0.037441 |  |  |  |  | -0.076441 | -0.056441
1678.804 | Sensitivity Phase (4160 2036126) degrees UK: 168.25; Sensitivity Phase (4160 2036126) degrees FR: 168.36; Sensitivity Phase (4160 2036126) degrees IT: 168.3; Sensitivity Phase (4160 2036126) degrees CH: 168.349; Sensitivity Phase (4160 2036126) degrees SE: 168.32; Sensitivity Phase (4160 2036126) degrees DE: 168.4; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.6; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.179; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 168.3421 | 0.25335 | -0.092055 | 0.017945 | -0.042055 | 0.0069447 |  |  |  |  | -0.022055 | 0.057945
1778.2794 | Sensitivity Phase (4160 2036126) degrees UK: 167.54; Sensitivity Phase (4160 2036126) degrees FR: 167.64; Sensitivity Phase (4160 2036126) degrees IT: 167.6; Sensitivity Phase (4160 2036126) degrees CH: 167.642; Sensitivity Phase (4160 2036126) degrees SE: 167.6; Sensitivity Phase (4160 2036126) degrees DE: 167.7; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.6; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.19; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 167.6509 | 0.25334 | -0.11093 | -0.010934 | -0.050934 | -0.0089338 |  |  |  |  | -0.050934 | 0.049066
1883.6491 | Sensitivity Phase (4160 2036126) degrees UK: 166.78; Sensitivity Phase (4160 2036126) degrees FR: 166.87; Sensitivity Phase (4160 2036126) degrees IT: 166.9; Sensitivity Phase (4160 2036126) degrees CH: 166.873; Sensitivity Phase (4160 2036126) degrees SE: 166.84; Sensitivity Phase (4160 2036126) degrees DE: 166.9; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.6; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.203; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 166.909 | 0.25334 | -0.12895 | -0.038954 | -0.0089537 | -0.035954 |  |  |  |  | -0.068954 | -0.0089537
1995.2623 | Sensitivity Phase (4160 2036126) degrees UK: 165.95; Sensitivity Phase (4160 2036126) degrees FR: 166.05; Sensitivity Phase (4160 2036126) degrees IT: 166; Sensitivity Phase (4160 2036126) degrees CH: 166.053; Sensitivity Phase (4160 2036126) degrees SE: 166.03; Sensitivity Phase (4160 2036126) degrees DE: 166.1; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.6; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.216; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 166.0534 | 0.25332 | -0.10338 | -0.0033822 | -0.053382 | -0.00038224 |  |  |  |  | -0.023382 | 0.046618
2113.489 | Sensitivity Phase (4160 2036126) degrees UK: 165.07; Sensitivity Phase (4160 2036126) degrees FR: 165.18; Sensitivity Phase (4160 2036126) degrees IT: 165.2; Sensitivity Phase (4160 2036126) degrees CH: 165.182; Sensitivity Phase (4160 2036126) degrees SE: 165.15; Sensitivity Phase (4160 2036126) degrees DE: 165.2; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.7; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.229; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 165.2347 | 0.27654 | -0.16466 | -0.054661 | -0.034661 | -0.052661 |  |  |  |  | -0.084661 | -0.034661
2238.7211 | Sensitivity Phase (4160 2036126) degrees UK: 164.14; Sensitivity Phase (4160 2036126) degrees FR: 164.23; Sensitivity Phase (4160 2036126) degrees IT: 164.2; Sensitivity Phase (4160 2036126) degrees CH: 164.236; Sensitivity Phase (4160 2036126) degrees SE: 164.21; Sensitivity Phase (4160 2036126) degrees DE: 164.3; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.7; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.244; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 164.263 | 0.27646 | -0.123 | -0.032996 | -0.062996 | -0.026996 |  |  |  |  | -0.052996 | 0.037004
2371.3737 | Sensitivity Phase (4160 2036126) degrees UK: 163.13; Sensitivity Phase (4160 2036126) degrees FR: 163.23; Sensitivity Phase (4160 2036126) degrees IT: 163.2; Sensitivity Phase (4160 2036126) degrees CH: 163.226; Sensitivity Phase (4160 2036126) degrees SE: 163.2; Sensitivity Phase (4160 2036126) degrees DE: 163.3; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.7; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.261; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 163.2427 | 0.27646 | -0.11266 | -0.012656 | -0.042656 | -0.016656 |  |  |  |  | -0.042656 | 0.057344
2511.8864 | Sensitivity Phase (4160 2036126) degrees UK: 162.06; Sensitivity Phase (4160 2036126) degrees FR: 162.13; Sensitivity Phase (4160 2036126) degrees IT: 162.1; Sensitivity Phase (4160 2036126) degrees CH: 162.145; Sensitivity Phase (4160 2036126) degrees SE: 162.12; Sensitivity Phase (4160 2036126) degrees DE: 162.2; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 0.7; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.277; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 162.1613 | 0.27646 | -0.10135 | -0.031349 | -0.061349 | -0.016349 |  |  |  |  | -0.041349 | 0.038651
2660.7251 | Sensitivity Phase (4160 2036126) degrees UK: 160.87; Sensitivity Phase (4160 2036126) degrees FR: 160.96; Sensitivity Phase (4160 2036126) degrees IT: 161; Sensitivity Phase (4160 2036126) degrees CH: 160.973; Sensitivity Phase (4160 2036126) degrees SE: 160.95; Sensitivity Phase (4160 2036126) degrees DE: 161; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 0.8; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.294; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 161.0307 | 0.29038 | -0.16066 | -0.070665 | -0.030665 | -0.057665 |  |  |  |  | -0.080665 | -0.030665
2818.3829 | Sensitivity Phase (4160 2036126) degrees UK: 159.61; Sensitivity Phase (4160 2036126) degrees FR: 159.71; Sensitivity Phase (4160 2036126) degrees IT: 159.7; Sensitivity Phase (4160 2036126) degrees CH: 159.708; Sensitivity Phase (4160 2036126) degrees SE: 159.69; Sensitivity Phase (4160 2036126) degrees DE: 159.8; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 0.8; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.313; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 159.7567 | 0.29036 | -0.14673 | -0.046733 | -0.056733 | -0.048733 |  |  |  |  | -0.066733 | 0.043267
2985.3826 | Sensitivity Phase (4160 2036126) degrees UK: 158.17; Sensitivity Phase (4160 2036126) degrees FR: 158.32; Sensitivity Phase (4160 2036126) degrees IT: 158.3; Sensitivity Phase (4160 2036126) degrees CH: 158.344; Sensitivity Phase (4160 2036126) degrees SE: 158.33; Sensitivity Phase (4160 2036126) degrees DE: 158.4; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 0.8; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.332; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.5; Expanded Uncertainties (degrees) DE: 0.5 | 158.3955 | 0.29028 | -0.22552 | -0.075521 | -0.095521 | -0.051521 |  |  |  |  | -0.065521 | 0.0044792
3162.2777 | Sensitivity Phase (4160 2036126) degrees UK: 156.73; Sensitivity Phase (4160 2036126) degrees FR: 156.85; Sensitivity Phase (4160 2036126) degrees IT: 156.8; Sensitivity Phase (4160 2036126) degrees CH: 156.861; Sensitivity Phase (4160 2036126) degrees SE: 156.85; Sensitivity Phase (4160 2036126) degrees DE: 156.9; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 0.8; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.354; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 0.5 | 156.9122 | 0.29027 | -0.18217 | -0.062166 | -0.11217 | -0.051166 |  |  |  |  | -0.062166 | -0.012166
3349.6544 | Sensitivity Phase (4160 2036126) degrees UK: 155.1; Sensitivity Phase (4160 2036126) degrees FR: 155.22; Sensitivity Phase (4160 2036126) degrees IT: 155.2; Sensitivity Phase (4160 2036126) degrees CH: 155.272; Sensitivity Phase (4160 2036126) degrees SE: 155.24; Sensitivity Phase (4160 2036126) degrees DE: 155.3; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 0.9; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.37; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 0.5 | 155.2962 | 0.29393 | -0.1962 | -0.076202 | -0.096202 | -0.024202 |  |  |  |  | -0.056202 | 0.003798
3548.1339 | Sensitivity Phase (4160 2036126) degrees UK: 153.33; Sensitivity Phase (4160 2036126) degrees FR: 153.47; Sensitivity Phase (4160 2036126) degrees IT: 153.5; Sensitivity Phase (4160 2036126) degrees CH: 153.486; Sensitivity Phase (4160 2036126) degrees SE: 153.49; Sensitivity Phase (4160 2036126) degrees DE: 153.5; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 0.9; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.398; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 0.5 | 153.5736 | 0.2939 | -0.24356 | -0.10356 | -0.073558 | -0.087558 |  |  |  |  | -0.083558 | -0.073558
3758.374 | Sensitivity Phase (4160 2036126) degrees UK: 151.4; Sensitivity Phase (4160 2036126) degrees FR: 151.55; Sensitivity Phase (4160 2036126) degrees IT: 151.5; Sensitivity Phase (4160 2036126) degrees CH: 151.574; Sensitivity Phase (4160 2036126) degrees SE: 151.58; Sensitivity Phase (4160 2036126) degrees DE: 151.6; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 0.9; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.416; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 0.5 | 151.6063 | 0.29389 | -0.20625 | -0.056251 | -0.10625 | -0.032251 |  |  |  |  | -0.026251 | -0.0062507
3981.0717 | Sensitivity Phase (4160 2036126) degrees UK: 149.31; Sensitivity Phase (4160 2036126) degrees FR: 149.49; Sensitivity Phase (4160 2036126) degrees IT: 149.4; Sensitivity Phase (4160 2036126) degrees CH: 149.478; Sensitivity Phase (4160 2036126) degrees SE: 149.48; Sensitivity Phase (4160 2036126) degrees DE: 149.5; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 0.9; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.435; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 0.5 | 149.4783 | 0.29388 | -0.16826 | 0.011742 | -0.078258 | -0.00025816 |  |  |  |  | 0.0017418 | 0.021742
4216.965 | Sensitivity Phase (4160 2036126) degrees UK: 146.94; Sensitivity Phase (4160 2036126) degrees FR: 147.21; Sensitivity Phase (4160 2036126) degrees IT: 147.1; Sensitivity Phase (4160 2036126) degrees CH: 147.171; Sensitivity Phase (4160 2036126) degrees SE: 147.18; Sensitivity Phase (4160 2036126) degrees DE: 147.2; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 1; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.457; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 0.5 | 147.1841 | 0.28912 | -0.24414 | 0.025861 | -0.084139 | -0.013139 |  |  |  |  | -0.0041393 | 0.015861
4466.8359 | Sensitivity Phase (4160 2036126) degrees UK: 144.47; Sensitivity Phase (4160 2036126) degrees FR: 144.7; Sensitivity Phase (4160 2036126) degrees IT: 144.6; Sensitivity Phase (4160 2036126) degrees CH: 144.639; Sensitivity Phase (4160 2036126) degrees SE: 144.67; Sensitivity Phase (4160 2036126) degrees DE: 144.7; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 1; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.472; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 0.5 | 144.7142 | 0.28912 | -0.24423 | -0.014226 | -0.11423 | -0.075226 |  |  |  |  | -0.044226 | -0.014226
4731.5126 | Sensitivity Phase (4160 2036126) degrees UK: 141.72; Sensitivity Phase (4160 2036126) degrees FR: 141.92; Sensitivity Phase (4160 2036126) degrees IT: 141.8; Sensitivity Phase (4160 2036126) degrees CH: 141.856; Sensitivity Phase (4160 2036126) degrees SE: 141.86; Sensitivity Phase (4160 2036126) degrees DE: 141.9; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 1; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.482; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 0.5 | 141.8928 | 0.28902 | -0.17283 | 0.027172 | -0.092828 | -0.036828 |  |  |  |  | -0.032828 | 0.007172
5011.8723 | Sensitivity Phase (4160 2036126) degrees UK: 138.64; Sensitivity Phase (4160 2036126) degrees FR: 138.82; Sensitivity Phase (4160 2036126) degrees IT: 138.7; Sensitivity Phase (4160 2036126) degrees CH: 138.741; Sensitivity Phase (4160 2036126) degrees SE: 138.77; Sensitivity Phase (4160 2036126) degrees DE: 138.8; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 1; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.488; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 1 | 138.855 | 0.28877 | -0.21495 | -0.034953 | -0.15495 | -0.11395 |  |  |  |  | -0.084953 | -0.054953
5308.8444 | Sensitivity Phase (4160 2036126) degrees UK: 135.19; Sensitivity Phase (4160 2036126) degrees FR: 135.42; Sensitivity Phase (4160 2036126) degrees IT: 135.2; Sensitivity Phase (4160 2036126) degrees CH: 135.343; Sensitivity Phase (4160 2036126) degrees SE: 135.39; Sensitivity Phase (4160 2036126) degrees DE: 135.4; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 1.1; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.499; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 1 | 135.3829 | 0.26014 | -0.19293 | 0.037073 | -0.18293 | -0.039927 |  |  |  |  | 0.0070732 | 0.017073
5623.4133 | Sensitivity Phase (4160 2036126) degrees UK: 131.46; Sensitivity Phase (4160 2036126) degrees FR: 131.61; Sensitivity Phase (4160 2036126) degrees IT: 131.5; Sensitivity Phase (4160 2036126) degrees CH: 131.513; Sensitivity Phase (4160 2036126) degrees SE: 131.64; Sensitivity Phase (4160 2036126) degrees DE: 131.7; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 1.1; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.5; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 1 | 131.6179 | 0.26014 | -0.15792 | -0.0079215 | -0.11792 | -0.10492 |  |  |  |  | 0.022079 | 0.082079
5956.6214 | Sensitivity Phase (4160 2036126) degrees UK: 127.16; Sensitivity Phase (4160 2036126) degrees FR: 127.44; Sensitivity Phase (4160 2036126) degrees IT: 127.3; Sensitivity Phase (4160 2036126) degrees CH: 127.32; Sensitivity Phase (4160 2036126) degrees SE: 127.47; Sensitivity Phase (4160 2036126) degrees DE: 127.5; Expanded Uncertainties (degrees) UK: 0.7; Expanded Uncertainties (degrees) FR: 1.1; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.491; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 1 | 127.435 | 0.26013 | -0.27496 | 0.0050401 | -0.13496 | -0.11496 |  |  |  |  | 0.03504 | 0.06504
6309.5734 | Sensitivity Phase (4160 2036126) degrees UK: 122.6; Sensitivity Phase (4160 2036126) degrees FR: 122.89; Sensitivity Phase (4160 2036126) degrees IT: 122.7; Sensitivity Phase (4160 2036126) degrees CH: 122.726; Sensitivity Phase (4160 2036126) degrees SE: 122.88; Sensitivity Phase (4160 2036126) degrees DE: 122.9; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 1.1; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.49; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 1 | 122.7355 | 0.24214 | -0.13554 | 0.15446 | -0.035543 | -0.0095428 |  |  |  |  | 0.14446 | 0.16446
6683.4392 | Sensitivity Phase (4160 2036126) degrees UK: 117.59; Sensitivity Phase (4160 2036126) degrees FR: 117.85; Sensitivity Phase (4160 2036126) degrees IT: 117.6; Sensitivity Phase (4160 2036126) degrees CH: 117.643; Sensitivity Phase (4160 2036126) degrees SE: 117.85; Sensitivity Phase (4160 2036126) degrees DE: 117.9; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 1.2; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.504; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 1 | 117.6689 | 0.19616 | -0.078919 | 0.18108 | -0.068919 | -0.025919 |  |  |  |  | 0.18108 | 0.23108
7079.4578 | Sensitivity Phase (4160 2036126) degrees UK: 112.04; Sensitivity Phase (4160 2036126) degrees FR: 112.38; Sensitivity Phase (4160 2036126) degrees IT: 112.1; Sensitivity Phase (4160 2036126) degrees CH: 112.155; Sensitivity Phase (4160 2036126) degrees SE: 112.39; Sensitivity Phase (4160 2036126) degrees DE: 112.4; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 1.2; Expanded Uncertainties (degrees) IT: 0.6; Expanded Uncertainties (degrees) CH: 0.511; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 1 | 112.1399 | 0.21223 | -0.09994 | 0.24006 | -0.03994 | 0.01506 |  |  |  |  | 0.25006 | 0.26006
7498.9421 | Sensitivity Phase (4160 2036126) degrees UK: 106.12; Sensitivity Phase (4160 2036126) degrees FR: 106.49; Sensitivity Phase (4160 2036126) degrees IT: 106.2; Sensitivity Phase (4160 2036126) degrees CH: 106.247; Sensitivity Phase (4160 2036126) degrees SE: 106.51; Sensitivity Phase (4160 2036126) degrees DE: 106.5; Expanded Uncertainties (degrees) UK: 0.6; Expanded Uncertainties (degrees) FR: 1.2; Expanded Uncertainties (degrees) IT: 0.7; Expanded Uncertainties (degrees) CH: 0.532; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 1 | 106.1948 | 0.21139 | -0.07482 | 0.29518 | 0.00518 | 0.05218 |  |  |  |  | 0.31518 | 0.30518
7943.2823 | Sensitivity Phase (4160 2036126) degrees UK: 99.96; Sensitivity Phase (4160 2036126) degrees FR: 100.27; Sensitivity Phase (4160 2036126) degrees IT: 100; Sensitivity Phase (4160 2036126) degrees CH: 99.982; Sensitivity Phase (4160 2036126) degrees SE: 100.31; Sensitivity Phase (4160 2036126) degrees DE: 100.3; Expanded Uncertainties (degrees) UK: 0.5; Expanded Uncertainties (degrees) FR: 1.3; Expanded Uncertainties (degrees) IT: 0.8; Expanded Uncertainties (degrees) CH: 0.563; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 1 | 99.9147 | 0.19677 | 0.045349 | 0.35535 | 0.085349 | 0.067349 |  |  |  |  | 0.39535 | 0.38535
8413.9514 | Sensitivity Phase (4160 2036126) degrees UK: 93.5; Sensitivity Phase (4160 2036126) degrees FR: 93.81; Sensitivity Phase (4160 2036126) degrees IT: 93.6; Sensitivity Phase (4160 2036126) degrees CH: 93.523; Sensitivity Phase (4160 2036126) degrees SE: 93.9; Sensitivity Phase (4160 2036126) degrees DE: 93.8; Expanded Uncertainties (degrees) UK: 0.8; Expanded Uncertainties (degrees) FR: 1.3; Expanded Uncertainties (degrees) IT: 0.8; Expanded Uncertainties (degrees) CH: 0.614; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.6; Expanded Uncertainties (degrees) DE: 1 | 93.451 | 0.3209 | 0.048984 | 0.35898 | 0.14898 | 0.071984 |  |  |  |  | 0.44898 | 0.34898
8912.5094 | Sensitivity Phase (4160 2036126) degrees UK: 87.03; Sensitivity Phase (4160 2036126) degrees FR: 87.39; Sensitivity Phase (4160 2036126) degrees IT: 87.2; Sensitivity Phase (4160 2036126) degrees CH: 87.094; Sensitivity Phase (4160 2036126) degrees SE: 87.44; Sensitivity Phase (4160 2036126) degrees DE: 87.4; Expanded Uncertainties (degrees) UK: 0.8; Expanded Uncertainties (degrees) FR: 1.3; Expanded Uncertainties (degrees) IT: 0.8; Expanded Uncertainties (degrees) CH: 0.699; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.7; Expanded Uncertainties (degrees) DE: 1 | 87.0535 | 0.32035 | -0.023464 | 0.33654 | 0.14654 | 0.040536 |  |  |  |  | 0.38654 | 0.34654
9440.6088 | Sensitivity Phase (4160 2036126) degrees UK: 80.73; Sensitivity Phase (4160 2036126) degrees FR: 81.41; Sensitivity Phase (4160 2036126) degrees IT: 81; Sensitivity Phase (4160 2036126) degrees CH: 80.922; Sensitivity Phase (4160 2036126) degrees SE: 81.27; Sensitivity Phase (4160 2036126) degrees DE: 81.2; Expanded Uncertainties (degrees) UK: 0.8; Expanded Uncertainties (degrees) FR: 1.4; Expanded Uncertainties (degrees) IT: 0.9; Expanded Uncertainties (degrees) CH: 0.838; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.7; Expanded Uncertainties (degrees) DE: 1 | 80.6212 | 0.33488 | 0.1088 | 0.7888 | 0.3788 | 0.3008 |  |  |  |  | 0.6488 | 0.5788
10000 | Sensitivity Phase (4160 2036126) degrees UK: 74.26; Sensitivity Phase (4160 2036126) degrees FR: 74.94; Sensitivity Phase (4160 2036126) degrees IT: 74.4; Sensitivity Phase (4160 2036126) degrees CH: 74.285; Sensitivity Phase (4160 2036126) degrees SE: 75.19; Sensitivity Phase (4160 2036126) degrees DE: 74.7; Expanded Uncertainties (degrees) UK: 0.8; Expanded Uncertainties (degrees) FR: 1.6; Expanded Uncertainties (degrees) IT: 1.1; Expanded Uncertainties (degrees) CH: 1.035; Expanded Uncertainties (degrees) AT: 0; Expanded Uncertainties (degrees) SE: 0.7; Expanded Uncertainties (degrees) DE: 1 | 74.2856 | 0.36338 | -0.02558 | 0.65442 | 0.11442 | -0.00058025 |  |  |  |  | 0.90442 | 0.41442

## 2036126 DoE Phase
| Linking DoEs from CCAUV.A-K5 |  | 4160 2036126 | [11 empty] | Linking U(DoE)s from CCAUV.A-K5
|  |  | DoE (degrees) | [13 empty] | U(DoE) (degrees)
Frequency (Hz) | UK | IT | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG | UK | IT | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG
1.9953 | UK: 2.3907; DoE (degrees) UK: 2.3907; DoE (degrees) FR: 1.0507; DoE (degrees) CH: 1.6287; DoE (degrees) DE: 1.1007; DoE (degrees) UK: 3.2349; U(DoE) (degrees) UK: 3.2349; U(DoE) (degrees) FR: 3.172; U(DoE) (degrees) CH: 0.83461; U(DoE) (degrees) DE: 4.5499
2.5119 | UK: 1.2242; DoE (degrees) UK: 1.2242; DoE (degrees) FR: 0.73419; DoE (degrees) CH: 0.89219; DoE (degrees) DE: 0.55419; DoE (degrees) UK: 3.2492; U(DoE) (degrees) UK: 3.2492; U(DoE) (degrees) FR: 2.7656; U(DoE) (degrees) CH: 0.72135; U(DoE) (degrees) DE: 4.5397
3.1623 | UK: 1.1827; DoE (degrees) UK: 1.1827; DoE (degrees) FR: 0.67273; DoE (degrees) CH: 0.76973; DoE (degrees) DE: 0.50273; DoE (degrees) UK: 3.2615; U(DoE) (degrees) UK: 3.2615; U(DoE) (degrees) FR: 2.4574; U(DoE) (degrees) CH: 0.61759; U(DoE) (degrees) DE: 4.5309
3.9811 | UK: 0.81765; DoE (degrees) UK: 0.81765; DoE (degrees) FR: 0.47765; DoE (degrees) CH: 0.54765; DoE (degrees) DE: 0.33765; DoE (degrees) UK: 1.4448; U(DoE) (degrees) UK: 1.4448; U(DoE) (degrees) FR: 2.1423; U(DoE) (degrees) CH: 0.49545; U(DoE) (degrees) DE: 2.5356
5.0119 | UK: 0.55513; DoE (degrees) UK: 0.55513; DoE (degrees) FR: 0.33513; DoE (degrees) CH: 0.37713; DoE (degrees) DE: 0.32513; DoE (degrees) UK: 1.4596; U(DoE) (degrees) UK: 1.4596; U(DoE) (degrees) FR: 1.9357; U(DoE) (degrees) CH: 0.4239; U(DoE) (degrees) DE: 2.5272
6.3096 | UK: 0.30979; DoE (degrees) UK: 0.30979; DoE (degrees) FR: 0.20979; DoE (degrees) CH: 0.24079; DoE (degrees) DE: 0.13979; DoE (degrees) UK: 1.4799; U(DoE) (degrees) UK: 1.4799; U(DoE) (degrees) FR: 1.7227; U(DoE) (degrees) CH: 0.3249; U(DoE) (degrees) DE: 2.5155
7.9433 | UK: 0.24199; DoE (degrees) UK: 0.24199; DoE (degrees) FR: 0.19199; DoE (degrees) CH: 0.20799; DoE (degrees) DE: 0.16199; DoE (degrees) UK: 0.77595; U(DoE) (degrees) UK: 0.77595; U(DoE) (degrees) FR: 1.6163; U(DoE) (degrees) CH: 0.26408; U(DoE) (degrees) DE: 1.5173
10 | UK: 0.13519; DoE (degrees) UK: 0.13519; DoE (degrees) FR: 0.12519; DoE (degrees) CH: 0.13419; DoE (degrees) DE: 0.065195; DoE (degrees) UK: 0.78374; U(DoE) (degrees) UK: 0.78374; U(DoE) (degrees) FR: 1.4144; U(DoE) (degrees) CH: 0.22828; U(DoE) (degrees) DE: 1.5134
12.5893 | UK: 0.1099; DoE (degrees) UK: 0.1099; DoE (degrees) FR: 0.1399; DoE (degrees) CH: 0.1499; DoE (degrees) DE: 0.069898; DoE (degrees) UK: 0.78938; U(DoE) (degrees) UK: 0.78938; U(DoE) (degrees) FR: 1.2132; U(DoE) (degrees) CH: 0.19709; U(DoE) (degrees) DE: 1.5106
15.8489 | UK: 0.086693; DoE (degrees) UK: 0.086693; DoE (degrees) FR: 0.12669; DoE (degrees) CH: 0.14169; DoE (degrees) DE: 0.086693; DoE (degrees) UK: 0.48476; U(DoE) (degrees) UK: 0.48476; U(DoE) (degrees) FR: 1.1131; U(DoE) (degrees) CH: 0.1835; U(DoE) (degrees) DE: 1.0144
19.9526 | UK: 0.078523; IT: 0.26852; DoE (degrees) UK: 0.078523; DoE (degrees) FR: 0.068523; DoE (degrees) IT: 0.26852; DoE (degrees) CH: 0.075523; DoE (degrees) SE: -0.22148; DoE (degrees) DE: 0.068523; DoE (degrees) UK: 0.48723; DoE (degrees) IT: 1.5961; U(DoE) (degrees) UK: 0.48455; U(DoE) (degrees) FR: 1.0551; U(DoE) (degrees) IT: 1.5082; U(DoE) (degrees) CH: 0.34141; U(DoE) (degrees) SE: 0.60266; U(DoE) (degrees) DE: 1.0551
21.1349 | UK: 0.0056415; IT: 0.16564; DoE (degrees) UK: 0.0056415; DoE (degrees) FR: -0.0043585; DoE (degrees) IT: 0.16564; DoE (degrees) CH: 0.0036415; DoE (degrees) SE: -0.25436; DoE (degrees) DE: -0.034358; DoE (degrees) UK: 0.48902; DoE (degrees) IT: 1.5966; U(DoE) (degrees) UK: 0.48635; U(DoE) (degrees) FR: 1.0546; U(DoE) (degrees) IT: 1.5088; U(DoE) (degrees) CH: 0.34005; U(DoE) (degrees) SE: 0.60189; U(DoE) (degrees) DE: 1.0546
22.3872 | UK: 0.008989; IT: 0.12899; DoE (degrees) UK: 0.008489; DoE (degrees) FR: -0.002011; DoE (degrees) IT: 0.13349; DoE (degrees) CH: 0.008989; DoE (degrees) SE: -0.26201; DoE (degrees) DE: 0.037989; DoE (degrees) UK: 0.48948; DoE (degrees) IT: 1.4965; U(DoE) (degrees) UK: 0.48649; U(DoE) (degrees) FR: 1.0539; U(DoE) (degrees) IT: 1.4152; U(DoE) (degrees) CH: 0.33621; U(DoE) (degrees) SE: 0.60061; U(DoE) (degrees) DE: 1.0539
23.7137 | UK: -0.0025545; IT: 0.087445; DoE (degrees) UK: -0.0042514; DoE (degrees) FR: 0.0040518; DoE (degrees) IT: 0.10075; DoE (degrees) CH: 0.020052; DoE (degrees) SE: -0.22595; DoE (degrees) DE: 0.014052; DoE (degrees) UK: 0.48964; DoE (degrees) IT: 1.3963; U(DoE) (degrees) UK: 0.48625; U(DoE) (degrees) FR: 1.0531; U(DoE) (degrees) IT: 1.3216; U(DoE) (degrees) CH: 0.33338; U(DoE) (degrees) SE: 0.59927; U(DoE) (degrees) DE: 1.0531
25.1189 | UK: -0.009613; IT: 0.13039; DoE (degrees) UK: -0.0074508; DoE (degrees) FR: 0.0047113; DoE (degrees) IT: 0.10755; DoE (degrees) CH: -0.0012887; DoE (degrees) SE: -0.23529; DoE (degrees) DE: -0.015289; DoE (degrees) UK: 0.38728; DoE (degrees) IT: 1.2961; U(DoE) (degrees) UK: 0.38447; U(DoE) (degrees) FR: 0.95153; U(DoE) (degrees) IT: 1.1991; U(DoE) (degrees) CH: 0.31524; U(DoE) (degrees) SE: 0.58772; U(DoE) (degrees) DE: 1.0466
26.6073 | UK: -0.020545; IT: 0.099455; DoE (degrees) UK: -0.022707; DoE (degrees) FR: -0.01487; DoE (degrees) IT: 0.12229; DoE (degrees) CH: 0.0091305; DoE (degrees) SE: -0.25487; DoE (degrees) DE: -0.05487; DoE (degrees) UK: 0.38742; DoE (degrees) IT: 1.2962; U(DoE) (degrees) UK: 0.38461; U(DoE) (degrees) FR: 0.95151; U(DoE) (degrees) IT: 1.1992; U(DoE) (degrees) CH: 0.31091; U(DoE) (degrees) SE: 0.58768; U(DoE) (degrees) DE: 1.0466
28.1838 | UK: -0.019244; IT: 0.080756; DoE (degrees) UK: -0.021744; DoE (degrees) FR: -0.0042439; DoE (degrees) IT: 0.10326; DoE (degrees) CH: -0.020244; DoE (degrees) SE: -0.23424; DoE (degrees) DE: 0.025756; DoE (degrees) UK: 0.38751; DoE (degrees) IT: 1.1959; U(DoE) (degrees) UK: 0.38426; U(DoE) (degrees) FR: 0.95078; U(DoE) (degrees) IT: 1.1077; U(DoE) (degrees) CH: 0.31095; U(DoE) (degrees) SE: 0.5865; U(DoE) (degrees) DE: 1.0459
29.8538 | UK: -0.011785; IT: 0.14821; DoE (degrees) UK: -0.0094497; DoE (degrees) FR: 0.0028861; DoE (degrees) IT: 0.13055; DoE (degrees) CH: 0.020886; DoE (degrees) SE: -0.20711; DoE (degrees) DE: 0.012886; DoE (degrees) UK: 0.3878; DoE (degrees) IT: 1.0956; U(DoE) (degrees) UK: 0.38401; U(DoE) (degrees) FR: 0.94985; U(DoE) (degrees) IT: 1.0163; U(DoE) (degrees) CH: 0.30617; U(DoE) (degrees) SE: 0.58499; U(DoE) (degrees) DE: 1.0451
31.6228 | UK: -0.011152; IT: 0.11885; DoE (degrees) UK: -0.0099841; DoE (degrees) FR: 0.011184; DoE (degrees) IT: 0.11002; DoE (degrees) CH: 0.036184; DoE (degrees) SE: -0.19882; DoE (degrees) DE: 0.0011838; DoE (degrees) UK: 0.38843; DoE (degrees) IT: 1.0958; U(DoE) (degrees) UK: 0.38462; U(DoE) (degrees) FR: 0.85579; U(DoE) (degrees) IT: 1.0162; U(DoE) (degrees) CH: 0.30553; U(DoE) (degrees) SE: 0.50238; U(DoE) (degrees) DE: 0.58514
33.4965 | UK: -0.011151; IT: 0.098849; DoE (degrees) UK: -0.010461; DoE (degrees) FR: 0.020229; DoE (degrees) IT: 0.094539; DoE (degrees) CH: 0.033229; DoE (degrees) SE: -0.17977; DoE (degrees) DE: -0.0097713; DoE (degrees) UK: 0.39057; DoE (degrees) IT: 0.99626; U(DoE) (degrees) UK: 0.38609; U(DoE) (degrees) FR: 0.85419; U(DoE) (degrees) IT: 0.92563; U(DoE) (degrees) CH: 0.3008; U(DoE) (degrees) SE: 0.49964; U(DoE) (degrees) DE: 0.58278
35.4813 | UK: 0.0085893; IT: 0.078589; DoE (degrees) UK: 0.0085893; DoE (degrees) FR: 0.028589; DoE (degrees) IT: 0.078589; DoE (degrees) CH: 0.044589; DoE (degrees) SE: -0.17141; DoE (degrees) DE: -0.021411; DoE (degrees) UK: 0.39066; DoE (degrees) IT: 0.9963; U(DoE) (degrees) UK: 0.38619; U(DoE) (degrees) FR: 0.85417; U(DoE) (degrees) IT: 0.92567; U(DoE) (degrees) CH: 0.30116; U(DoE) (degrees) SE: 0.49961; U(DoE) (degrees) DE: 0.58276
37.5837 | UK: 0.0019145; IT: 0.051915; DoE (degrees) UK: 0.0010898; DoE (degrees) FR: 0.030265; DoE (degrees) IT: 0.05609; DoE (degrees) CH: 0.043265; DoE (degrees) SE: -0.16973; DoE (degrees) DE: 0.060265; DoE (degrees) UK: 0.39079; DoE (degrees) IT: 0.89594; U(DoE) (degrees) UK: 0.38543; U(DoE) (degrees) FR: 0.85256; U(DoE) (degrees) IT: 0.83439; U(DoE) (degrees) CH: 0.29585; U(DoE) (degrees) SE: 0.49684; U(DoE) (degrees) DE: 0.58039
39.8107 | UK: -0.019055; IT: 0.020945; DoE (degrees) UK: -0.01988; DoE (degrees) FR: 0.0092953; DoE (degrees) IT: 0.02512; DoE (degrees) CH: 0.017295; DoE (degrees) SE: -0.1807; DoE (degrees) DE: 0.029295; DoE (degrees) UK: 0.39141; DoE (degrees) IT: 0.89622; U(DoE) (degrees) UK: 0.38607; U(DoE) (degrees) FR: 0.85245; U(DoE) (degrees) IT: 0.83468; U(DoE) (degrees) CH: 0.29582; U(DoE) (degrees) SE: 0.49666; U(DoE) (degrees) DE: 0.58023
42.1697 | UK: -0.01004; IT: 0.10996; DoE (degrees) UK: -0.0010398; DoE (degrees) FR: 0.02796; DoE (degrees) IT: 0.07396; DoE (degrees) CH: 0.03896; DoE (degrees) SE: -0.14204; DoE (degrees) DE: 0.03796; DoE (degrees) UK: 0.39168; DoE (degrees) IT: 0.79587; U(DoE) (degrees) UK: 0.38519; U(DoE) (degrees) FR: 0.85033; U(DoE) (degrees) IT: 0.7435; U(DoE) (degrees) CH: 0.28946; U(DoE) (degrees) SE: 0.49301; U(DoE) (degrees) DE: 0.57711
44.6684 | UK: 0.0055775; IT: 0.095578; DoE (degrees) UK: 0.012578; DoE (degrees) FR: 0.039578; DoE (degrees) IT: 0.067578; DoE (degrees) CH: 0.049578; DoE (degrees) SE: -0.13042; DoE (degrees) DE: 0.039578; DoE (degrees) UK: 0.39183; DoE (degrees) IT: 0.79595; U(DoE) (degrees) UK: 0.38534; U(DoE) (degrees) FR: 0.8503; U(DoE) (degrees) IT: 0.74358; U(DoE) (degrees) CH: 0.2893; U(DoE) (degrees) SE: 0.49297; U(DoE) (degrees) DE: 0.57708
47.3151 | UK: 0.0064277; IT: 0.086428; DoE (degrees) UK: 0.0015047; DoE (degrees) FR: 0.026582; DoE (degrees) IT: 0.1015; DoE (degrees) CH: 0.035582; DoE (degrees) SE: -0.14342; DoE (degrees) DE: 0.016582; DoE (degrees) UK: 0.39185; DoE (degrees) IT: 0.69538; U(DoE) (degrees) UK: 0.38378; U(DoE) (degrees) FR: 0.84798; U(DoE) (degrees) IT: 0.65178; U(DoE) (degrees) CH: 0.28361; U(DoE) (degrees) SE: 0.48894; U(DoE) (degrees) DE: 0.57364
50.1187 | UK: -0.46624; IT: 0.083763; DoE (degrees) UK: -0.41208; DoE (degrees) FR: -0.32793; DoE (degrees) IT: -0.082083; DoE (degrees) CH: -0.36093; DoE (degrees) SE: -0.50793; DoE (degrees) DE: -0.34793; DoE (degrees) UK: 0.39202; DoE (degrees) IT: 0.69547; U(DoE) (degrees) UK: 0.38395; U(DoE) (degrees) FR: 0.75433; U(DoE) (degrees) IT: 0.65188; U(DoE) (degrees) CH: 0.28339; U(DoE) (degrees) SE: 0.48889; U(DoE) (degrees) DE: 0.5736
53.0884 | UK: -0.010777; IT: 0.049223; DoE (degrees) UK: -0.018162; DoE (degrees) FR: 0.014454; DoE (degrees) IT: 0.071838; DoE (degrees) CH: 0.034454; DoE (degrees) SE: -0.15555; DoE (degrees) DE: -0.0055462; DoE (degrees) UK: 0.39283; DoE (degrees) IT: 0.69593; U(DoE) (degrees) UK: 0.38477; U(DoE) (degrees) FR: 0.75418; U(DoE) (degrees) IT: 0.65236; U(DoE) (degrees) CH: 0.28208; U(DoE) (degrees) SE: 0.48866; U(DoE) (degrees) DE: 0.5734
56.2341 | UK: -0.025779; IT: 0.034221; DoE (degrees) UK: -0.033164; DoE (degrees) FR: 0.0094513; DoE (degrees) IT: 0.056836; DoE (degrees) CH: 0.013451; DoE (degrees) SE: -0.15055; DoE (degrees) DE: -0.020549; DoE (degrees) UK: 0.39369; DoE (degrees) IT: 0.69641; U(DoE) (degrees) UK: 0.38566; U(DoE) (degrees) FR: 0.75401; U(DoE) (degrees) IT: 0.65288; U(DoE) (degrees) CH: 0.28111; U(DoE) (degrees) SE: 0.4884; U(DoE) (degrees) DE: 0.57318
59.5662 | UK: -0.019114; IT: 0.030886; DoE (degrees) UK: -0.014499; DoE (degrees) FR: 0.040116; DoE (degrees) IT: 0.020501; DoE (degrees) CH: 0.049116; DoE (degrees) SE: -0.11988; DoE (degrees) DE: 0.010116; DoE (degrees) UK: 0.39371; DoE (degrees) IT: 0.59583; U(DoE) (degrees) UK: 0.38364; U(DoE) (degrees) FR: 0.74985; U(DoE) (degrees) IT: 0.56159; U(DoE) (degrees) CH: 0.26957; U(DoE) (degrees) SE: 0.48194; U(DoE) (degrees) DE: 0.56769
63.0957 | UK: -0.033705; IT: 0.016295; DoE (degrees) UK: -0.027551; DoE (degrees) FR: 0.018603; DoE (degrees) IT: 0.0024488; DoE (degrees) CH: 0.020603; DoE (degrees) SE: -0.1314; DoE (degrees) DE: -0.011397; DoE (degrees) UK: 0.39443; DoE (degrees) IT: 0.5963; U(DoE) (degrees) UK: 0.38437; U(DoE) (degrees) FR: 0.74971; U(DoE) (degrees) IT: 0.56209; U(DoE) (degrees) CH: 0.26905; U(DoE) (degrees) SE: 0.48173; U(DoE) (degrees) DE: 0.56751
66.8344 | UK: -0.032945; IT: 0.017055; DoE (degrees) UK: -0.028329; DoE (degrees) FR: 0.016286; DoE (degrees) IT: 0.0066707; DoE (degrees) CH: 0.026286; DoE (degrees) SE: -0.12371; DoE (degrees) DE: -0.0037139; DoE (degrees) UK: 0.39456; DoE (degrees) IT: 0.59638; U(DoE) (degrees) UK: 0.38451; U(DoE) (degrees) FR: 0.74968; U(DoE) (degrees) IT: 0.56218; U(DoE) (degrees) CH: 0.26892; U(DoE) (degrees) SE: 0.48169; U(DoE) (degrees) DE: 0.56747
70.7946 | UK: -0.030883; IT: 0.019117; DoE (degrees) UK: -0.026267; DoE (degrees) FR: 0.018348; DoE (degrees) IT: 0.0087327; DoE (degrees) CH: 0.023348; DoE (degrees) SE: -0.13165; DoE (degrees) DE: -0.0016519; DoE (degrees) UK: 0.39545; DoE (degrees) IT: 0.59698; U(DoE) (degrees) UK: 0.38543; U(DoE) (degrees) FR: 0.74951; U(DoE) (degrees) IT: 0.56281; U(DoE) (degrees) CH: 0.2685; U(DoE) (degrees) SE: 0.48142; U(DoE) (degrees) DE: 0.56725
74.9894 | UK: -0.027789; IT: 0.022211; DoE (degrees) UK: -0.021635; DoE (degrees) FR: 0.0045189; DoE (degrees) IT: 0.0083651; DoE (degrees) CH: 0.010519; DoE (degrees) SE: -0.13548; DoE (degrees) DE: -0.0054811; DoE (degrees) UK: 0.39562; DoE (degrees) IT: 0.59709; U(DoE) (degrees) UK: 0.3856; U(DoE) (degrees) FR: 0.74948; U(DoE) (degrees) IT: 0.56292; U(DoE) (degrees) CH: 0.26835; U(DoE) (degrees) SE: 0.48137; U(DoE) (degrees) DE: 0.56721
79.4328 | UK: -0.028931; IT: 0.031069; DoE (degrees) UK: -0.022777; DoE (degrees) FR: 0.0033764; DoE (degrees) IT: 0.017223; DoE (degrees) CH: 0.0093764; DoE (degrees) SE: -0.13662; DoE (degrees) DE: 0.0033764; DoE (degrees) UK: 0.39581; DoE (degrees) IT: 0.59722; U(DoE) (degrees) UK: 0.3858; U(DoE) (degrees) FR: 0.65701; U(DoE) (degrees) IT: 0.56306; U(DoE) (degrees) CH: 0.26838; U(DoE) (degrees) SE: 0.48132; U(DoE) (degrees) DE: 0.56716
84.1395 | UK: -0.015145; IT: 0.034855; DoE (degrees) UK: -0.015145; DoE (degrees) FR: 0.024855; DoE (degrees) IT: 0.034855; DoE (degrees) CH: 0.028855; DoE (degrees) SE: -0.11515; DoE (degrees) DE: 0.034855; DoE (degrees) UK: 0.39612; DoE (degrees) IT: 0.59742; U(DoE) (degrees) UK: 0.38611; U(DoE) (degrees) FR: 0.65695; U(DoE) (degrees) IT: 0.56327; U(DoE) (degrees) CH: 0.26808; U(DoE) (degrees) SE: 0.48123; U(DoE) (degrees) DE: 0.56708
89.1251 | UK: -0.021551; IT: 0.048449; DoE (degrees) UK: -0.021551; DoE (degrees) FR: 0.018449; DoE (degrees) IT: 0.048449; DoE (degrees) CH: 0.029449; DoE (degrees) SE: -0.10155; DoE (degrees) DE: 0.048449; DoE (degrees) UK: 0.39617; DoE (degrees) IT: 0.59745; U(DoE) (degrees) UK: 0.38616; U(DoE) (degrees) FR: 0.65693; U(DoE) (degrees) IT: 0.56331; U(DoE) (degrees) CH: 0.26793; U(DoE) (degrees) SE: 0.48121; U(DoE) (degrees) DE: 0.56707
94.4061 | UK: -0.0069448; IT: -0.036945; DoE (degrees) UK: -0.023868; DoE (degrees) FR: -0.010791; DoE (degrees) IT: 0.0011321; DoE (degrees) CH: 0.006209; DoE (degrees) SE: -0.13079; DoE (degrees) DE: 0.039209; DoE (degrees) UK: 0.39705; DoE (degrees) IT: 0.59804; U(DoE) (degrees) UK: 0.38696; U(DoE) (degrees) FR: 0.65731; U(DoE) (degrees) IT: 0.56355; U(DoE) (degrees) CH: 0.26879; U(DoE) (degrees) SE: 0.48172; U(DoE) (degrees) DE: 0.5675
100 | UK: -0.016418; IT: -0.016418; DoE (degrees) UK: -0.017957; DoE (degrees) FR: 0.030505; DoE (degrees) IT: -0.012957; DoE (degrees) CH: 0.034505; DoE (degrees) SE: -0.1395; DoE (degrees) DE: -0.0094952; DoE (degrees) UK: 0.39724; DoE (degrees) IT: 0.59816; U(DoE) (degrees) UK: 0.38715; U(DoE) (degrees) FR: 0.56745; U(DoE) (degrees) IT: 0.56368; U(DoE) (degrees) CH: 0.2688; U(DoE) (degrees) SE: 0.48166; U(DoE) (degrees) DE: 0.56745
105.9254 | UK: -0.027461; IT: 0.0025389; DoE (degrees) UK: -0.024384; DoE (degrees) FR: -0.0013073; DoE (degrees) IT: -0.0043842; DoE (degrees) CH: 0.010693; DoE (degrees) SE: -0.12131; DoE (degrees) DE: -0.011307; DoE (degrees) UK: 0.39774; DoE (degrees) IT: 0.59849; U(DoE) (degrees) UK: 0.38766; U(DoE) (degrees) FR: 0.56733; U(DoE) (degrees) IT: 0.56404; U(DoE) (degrees) CH: 0.2686; U(DoE) (degrees) SE: 0.48152; U(DoE) (degrees) DE: 0.56733
112.2018 | UK: -0.016067; IT: 0.033933; DoE (degrees) UK: -0.014529; DoE (degrees) FR: 0.017009; DoE (degrees) IT: 0.030471; DoE (degrees) CH: 0.031009; DoE (degrees) SE: -0.10299; DoE (degrees) DE: 0.027009; DoE (degrees) UK: 0.39786; DoE (degrees) IT: 0.59858; U(DoE) (degrees) UK: 0.38779; U(DoE) (degrees) FR: 0.56729; U(DoE) (degrees) IT: 0.56413; U(DoE) (degrees) CH: 0.26831; U(DoE) (degrees) SE: 0.48148; U(DoE) (degrees) DE: 0.56729
118.8502 | UK: -0.02014; IT: -0.04014; DoE (degrees) UK: -0.035525; DoE (degrees) FR: -0.01091; DoE (degrees) IT: -0.005525; DoE (degrees) CH: -0.0009096; DoE (degrees) SE: -0.12091; DoE (degrees) DE: 0.02909; DoE (degrees) UK: 0.39786; DoE (degrees) IT: 0.59858; U(DoE) (degrees) UK: 0.38779; U(DoE) (degrees) FR: 0.56729; U(DoE) (degrees) IT: 0.56413; U(DoE) (degrees) CH: 0.26916; U(DoE) (degrees) SE: 0.48148; U(DoE) (degrees) DE: 0.56729
125.8925 | UK: -0.0094735; IT: -0.0094735; DoE (degrees) UK: -0.011012; DoE (degrees) FR: 0.02745; DoE (degrees) IT: -0.006012; DoE (degrees) CH: 0.03245; DoE (degrees) SE: -0.08255; DoE (degrees) DE: -0.0025505; DoE (degrees) UK: 0.39787; DoE (degrees) IT: 0.59858; U(DoE) (degrees) UK: 0.38779; U(DoE) (degrees) FR: 0.56729; U(DoE) (degrees) IT: 0.56413; U(DoE) (degrees) CH: 0.26847; U(DoE) (degrees) SE: 0.48148; U(DoE) (degrees) DE: 0.56729
133.3521 | UK: -0.029087; IT: 0.020913; DoE (degrees) UK: -0.029087; DoE (degrees) FR: 0.010913; DoE (degrees) IT: 0.020913; DoE (degrees) CH: 0.012913; DoE (degrees) SE: -0.089087; DoE (degrees) DE: 0.020913; DoE (degrees) UK: 0.39803; DoE (degrees) IT: 0.59869; U(DoE) (degrees) UK: 0.38796; U(DoE) (degrees) FR: 0.56725; U(DoE) (degrees) IT: 0.56424; U(DoE) (degrees) CH: 0.26866; U(DoE) (degrees) SE: 0.48143; U(DoE) (degrees) DE: 0.56725
141.2538 | UK: -0.01845; IT: -0.03845; DoE (degrees) UK: -0.021527; DoE (degrees) FR: 0.035396; DoE (degrees) IT: -0.031527; DoE (degrees) CH: 0.021396; DoE (degrees) SE: -0.074604; DoE (degrees) DE: 0.075396; DoE (degrees) UK: 0.39803; DoE (degrees) IT: 0.59869; U(DoE) (degrees) UK: 0.38796; U(DoE) (degrees) FR: 0.56725; U(DoE) (degrees) IT: 0.56424; U(DoE) (degrees) CH: 0.27477; U(DoE) (degrees) SE: 0.48143; U(DoE) (degrees) DE: 0.56725
149.6236 | UK: -0.022149; IT: -0.0021493; DoE (degrees) UK: -0.020611; DoE (degrees) FR: 0.00092767; DoE (degrees) IT: -0.0056108; DoE (degrees) CH: -0.020072; DoE (degrees) SE: -0.099072; DoE (degrees) DE: -0.0090723; DoE (degrees) UK: 0.39862; DoE (degrees) IT: 0.59908; U(DoE) (degrees) UK: 0.38856; U(DoE) (degrees) FR: 0.56711; U(DoE) (degrees) IT: 0.56466; U(DoE) (degrees) CH: 0.27087; U(DoE) (degrees) SE: 0.48126; U(DoE) (degrees) DE: 0.56711
158.4893 | UK: -0.016561; IT: 0.043439; DoE (degrees) UK: -0.016561; DoE (degrees) FR: 0.013439; DoE (degrees) IT: 0.043439; DoE (degrees) CH: 0.017439; DoE (degrees) SE: -0.086561; DoE (degrees) DE: 0.043439; DoE (degrees) UK: 0.39862; DoE (degrees) IT: 0.59908; U(DoE) (degrees) UK: 0.38856; U(DoE) (degrees) FR: 0.56711; U(DoE) (degrees) IT: 0.56466; U(DoE) (degrees) CH: 0.26814; U(DoE) (degrees) SE: 0.48126; U(DoE) (degrees) DE: 0.56711
167.8804 | UK: -0.018351; IT: -0.0083512; DoE (degrees) UK: -0.021428; DoE (degrees) FR: 0.025495; DoE (degrees) IT: -0.0014281; DoE (degrees) CH: 0.032495; DoE (degrees) SE: -0.074505; DoE (degrees) DE: 0.0054949; DoE (degrees) UK: 0.39862; DoE (degrees) IT: 0.59908; U(DoE) (degrees) UK: 0.38856; U(DoE) (degrees) FR: 0.56711; U(DoE) (degrees) IT: 0.56466; U(DoE) (degrees) CH: 0.2682; U(DoE) (degrees) SE: 0.48126; U(DoE) (degrees) DE: 0.56711
177.8279 | UK: -0.022709; IT: 0.047291; DoE (degrees) UK: -0.019632; DoE (degrees) FR: -0.006555; DoE (degrees) IT: 0.040368; DoE (degrees) CH: 0.005445; DoE (degrees) SE: -0.096555; DoE (degrees) DE: 0.033445; DoE (degrees) UK: 0.3988; DoE (degrees) IT: 0.5992; U(DoE) (degrees) UK: 0.38875; U(DoE) (degrees) FR: 0.56706; U(DoE) (degrees) IT: 0.56478; U(DoE) (degrees) CH: 0.26811; U(DoE) (degrees) SE: 0.4812; U(DoE) (degrees) DE: 0.56706
188.3649 | UK: -0.024744; IT: 0.005256; DoE (degrees) UK: -0.026282; DoE (degrees) FR: 0.022179; DoE (degrees) IT: 0.0087175; DoE (degrees) CH: 0.026179; DoE (degrees) SE: -0.077821; DoE (degrees) DE: 0.012179; DoE (degrees) UK: 0.3988; DoE (degrees) IT: 0.5992; U(DoE) (degrees) UK: 0.38875; U(DoE) (degrees) FR: 0.56706; U(DoE) (degrees) IT: 0.56478; U(DoE) (degrees) CH: 0.26825; U(DoE) (degrees) SE: 0.4812; U(DoE) (degrees) DE: 0.56706
199.5262 | UK: -0.021023; IT: -0.031023; DoE (degrees) UK: -0.022561; DoE (degrees) FR: 0.0259; DoE (degrees) IT: -0.027561; DoE (degrees) CH: 0.0319; DoE (degrees) SE: -0.0641; DoE (degrees) DE: -0.0241; DoE (degrees) UK: 0.39901; DoE (degrees) IT: 0.59934; U(DoE) (degrees) UK: 0.38869; U(DoE) (degrees) FR: 0.56872; U(DoE) (degrees) IT: 0.56396; U(DoE) (degrees) CH: 0.27173; U(DoE) (degrees) SE: 0.48315; U(DoE) (degrees) DE: 0.56872
211.3489 | UK: -0.02349; IT: 0.03651; DoE (degrees) UK: -0.026567; DoE (degrees) FR: 0.030356; DoE (degrees) IT: 0.043433; DoE (degrees) CH: 0.035356; DoE (degrees) SE: -0.059644; DoE (degrees) DE: 0.050356; DoE (degrees) UK: 0.39901; DoE (degrees) IT: 0.59934; U(DoE) (degrees) UK: 0.38869; U(DoE) (degrees) FR: 0.56872; U(DoE) (degrees) IT: 0.56396; U(DoE) (degrees) CH: 0.27181; U(DoE) (degrees) SE: 0.48315; U(DoE) (degrees) DE: 0.56872
223.8721 | UK: -0.02049; IT: 0.0095104; DoE (degrees) UK: -0.022028; DoE (degrees) FR: 0.026433; DoE (degrees) IT: 0.012972; DoE (degrees) CH: 0.034433; DoE (degrees) SE: -0.063567; DoE (degrees) DE: 0.016433; DoE (degrees) UK: 0.39901; DoE (degrees) IT: 0.59934; U(DoE) (degrees) UK: 0.38869; U(DoE) (degrees) FR: 0.56872; U(DoE) (degrees) IT: 0.56396; U(DoE) (degrees) CH: 0.27188; U(DoE) (degrees) SE: 0.48315; U(DoE) (degrees) DE: 0.56872
237.1374 | UK: -0.0206; IT: -0.0106; DoE (degrees) UK: -0.023677; DoE (degrees) FR: 0.033246; DoE (degrees) IT: -0.0036772; DoE (degrees) CH: 0.040246; DoE (degrees) SE: -0.056754; DoE (degrees) DE: 0.0032458; DoE (degrees) UK: 0.39901; DoE (degrees) IT: 0.59934; U(DoE) (degrees) UK: 0.38869; U(DoE) (degrees) FR: 0.56872; U(DoE) (degrees) IT: 0.56396; U(DoE) (degrees) CH: 0.27205; U(DoE) (degrees) SE: 0.48315; U(DoE) (degrees) DE: 0.56872
251.1886 | UK: -0.05477; IT: -0.02477; DoE (degrees) UK: -0.050154; DoE (degrees) FR: 0.0044609; DoE (degrees) IT: -0.035154; DoE (degrees) CH: 0.010461; DoE (degrees) SE: -0.085539; DoE (degrees) DE: -0.045539; DoE (degrees) UK: 0.39903; DoE (degrees) IT: 0.59936; U(DoE) (degrees) UK: 0.38871; U(DoE) (degrees) FR: 0.48315; U(DoE) (degrees) IT: 0.56398; U(DoE) (degrees) CH: 0.27392; U(DoE) (degrees) SE: 0.48315; U(DoE) (degrees) DE: 0.56871
266.0725 | UK: -0.033209; IT: -0.033209; DoE (degrees) UK: -0.033209; DoE (degrees) FR: 0.026791; DoE (degrees) IT: -0.033209; DoE (degrees) CH: 0.024791; DoE (degrees) SE: -0.053209; DoE (degrees) DE: 0.066791; DoE (degrees) UK: 0.39918; DoE (degrees) IT: 0.59945; U(DoE) (degrees) UK: 0.38886; U(DoE) (degrees) FR: 0.48311; U(DoE) (degrees) IT: 0.56408; U(DoE) (degrees) CH: 0.27279; U(DoE) (degrees) SE: 0.48311; U(DoE) (degrees) DE: 0.56868
281.8383 | UK: -0.02353; IT: -0.03353; DoE (degrees) UK: -0.025068; DoE (degrees) FR: 0.033394; DoE (degrees) IT: -0.030068; DoE (degrees) CH: 0.042394; DoE (degrees) SE: -0.046606; DoE (degrees) DE: 0.073394; DoE (degrees) UK: 0.39918; DoE (degrees) IT: 0.59945; U(DoE) (degrees) UK: 0.38886; U(DoE) (degrees) FR: 0.48311; U(DoE) (degrees) IT: 0.56408; U(DoE) (degrees) CH: 0.27225; U(DoE) (degrees) SE: 0.48311; U(DoE) (degrees) DE: 0.56868
298.5383 | UK: -0.037216; IT: -0.027216; DoE (degrees) UK: -0.038755; DoE (degrees) FR: 0.039707; DoE (degrees) IT: -0.023755; DoE (degrees) CH: 0.045707; DoE (degrees) SE: -0.040293; DoE (degrees) DE: 0.079707; DoE (degrees) UK: 0.39938; DoE (degrees) IT: 0.59958; U(DoE) (degrees) UK: 0.38907; U(DoE) (degrees) FR: 0.48305; U(DoE) (degrees) IT: 0.56422; U(DoE) (degrees) CH: 0.27247; U(DoE) (degrees) SE: 0.48305; U(DoE) (degrees) DE: 0.56863
316.2278 | UK: -0.035382; IT: -0.015382; DoE (degrees) UK: -0.032305; DoE (degrees) FR: 0.020772; DoE (degrees) IT: -0.022305; DoE (degrees) CH: 0.026772; DoE (degrees) SE: -0.049228; DoE (degrees) DE: 0.070772; DoE (degrees) UK: 0.39938; DoE (degrees) IT: 0.59958; U(DoE) (degrees) UK: 0.38907; U(DoE) (degrees) FR: 0.48305; U(DoE) (degrees) IT: 0.56422; U(DoE) (degrees) CH: 0.27269; U(DoE) (degrees) SE: 0.48305; U(DoE) (degrees) DE: 0.56863
334.9654 | UK: -0.034725; IT: 0.0052746; DoE (degrees) UK: -0.031649; DoE (degrees) FR: 0.031428; DoE (degrees) IT: -0.0016485; DoE (degrees) CH: 0.030428; DoE (degrees) SE: -0.048572; DoE (degrees) DE: -0.0085716; DoE (degrees) UK: 0.39938; DoE (degrees) IT: 0.59958; U(DoE) (degrees) UK: 0.38907; U(DoE) (degrees) FR: 0.48305; U(DoE) (degrees) IT: 0.56422; U(DoE) (degrees) CH: 0.27281; U(DoE) (degrees) SE: 0.48305; U(DoE) (degrees) DE: 0.56863
354.8134 | UK: -0.013931; IT: 0.036069; DoE (degrees) UK: -0.010854; DoE (degrees) FR: 0.032223; DoE (degrees) IT: 0.029146; DoE (degrees) CH: 0.040223; DoE (degrees) SE: -0.037777; DoE (degrees) DE: 0.022223; DoE (degrees) UK: 0.39938; DoE (degrees) IT: 0.59958; U(DoE) (degrees) UK: 0.38907; U(DoE) (degrees) FR: 0.48305; U(DoE) (degrees) IT: 0.56422; U(DoE) (degrees) CH: 0.27306; U(DoE) (degrees) SE: 0.48305; U(DoE) (degrees) DE: 0.56863
375.8374 | UK: -0.015824; IT: -0.025824; DoE (degrees) UK: -0.025054; DoE (degrees) FR: -0.004285; DoE (degrees) IT: -0.0050543; DoE (degrees) CH: -0.002285; DoE (degrees) SE: -0.074285; DoE (degrees) DE: 0.015715; DoE (degrees) UK: 0.39938; DoE (degrees) IT: 0.59958; U(DoE) (degrees) UK: 0.38808; U(DoE) (degrees) FR: 0.49012; U(DoE) (degrees) IT: 0.56077; U(DoE) (degrees) CH: 0.28563; U(DoE) (degrees) SE: 0.49012; U(DoE) (degrees) DE: 0.57464
398.1072 | UK: -0.0019404; IT: 0.01806; DoE (degrees) UK: -0.00040196; DoE (degrees) FR: 0.051136; DoE (degrees) IT: 0.014598; DoE (degrees) CH: 0.055136; DoE (degrees) SE: -0.018864; DoE (degrees) DE: 0.011136; DoE (degrees) UK: 0.39938; DoE (degrees) IT: 0.59958; U(DoE) (degrees) UK: 0.38808; U(DoE) (degrees) FR: 0.49012; U(DoE) (degrees) IT: 0.56077; U(DoE) (degrees) CH: 0.28604; U(DoE) (degrees) SE: 0.49012; U(DoE) (degrees) DE: 0.57464
421.6965 | UK: -0.022814; IT: -0.022814; DoE (degrees) UK: -0.027429; DoE (degrees) FR: -0.012045; DoE (degrees) IT: -0.012429; DoE (degrees) CH: -0.012045; DoE (degrees) SE: -0.082045; DoE (degrees) DE: -0.0020446; DoE (degrees) UK: 0.39938; DoE (degrees) IT: 0.59958; U(DoE) (degrees) UK: 0.38808; U(DoE) (degrees) FR: 0.49012; U(DoE) (degrees) IT: 0.56077; U(DoE) (degrees) CH: 0.28632; U(DoE) (degrees) SE: 0.49012; U(DoE) (degrees) DE: 0.57464
446.6836 | UK: -0.025992; IT: -0.055992; DoE (degrees) UK: -0.030608; DoE (degrees) FR: -0.0052232; DoE (degrees) IT: -0.045608; DoE (degrees) CH: 0.0017768; DoE (degrees) SE: -0.075223; DoE (degrees) DE: -0.035223; DoE (degrees) UK: 0.39938; DoE (degrees) IT: 0.59958; U(DoE) (degrees) UK: 0.38808; U(DoE) (degrees) FR: 0.49012; U(DoE) (degrees) IT: 0.56077; U(DoE) (degrees) CH: 0.28678; U(DoE) (degrees) SE: 0.49012; U(DoE) (degrees) DE: 0.57464
473.1513 | UK: -0.018053; IT: 0.021947; DoE (degrees) UK: -0.010361; DoE (degrees) FR: 0.047332; DoE (degrees) IT: 0.0046395; DoE (degrees) CH: 0.051332; DoE (degrees) SE: -0.022668; DoE (degrees) DE: 0.087332; DoE (degrees) UK: 0.39938; DoE (degrees) IT: 0.59958; U(DoE) (degrees) UK: 0.38808; U(DoE) (degrees) FR: 0.49012; U(DoE) (degrees) IT: 0.56077; U(DoE) (degrees) CH: 0.2871; U(DoE) (degrees) SE: 0.49012; U(DoE) (degrees) DE: 0.57464
501.1872 | UK: -0.01805; IT: 0.01195; DoE (degrees) UK: -0.022666; DoE (degrees) FR: 0.012719; DoE (degrees) IT: 0.022334; DoE (degrees) CH: 0.015719; DoE (degrees) SE: -0.047281; DoE (degrees) DE: 0.032719; DoE (degrees) UK: 0.39938; DoE (degrees) IT: 0.59958; U(DoE) (degrees) UK: 0.38808; U(DoE) (degrees) FR: 0.49012; U(DoE) (degrees) IT: 0.56077; U(DoE) (degrees) CH: 0.28761; U(DoE) (degrees) SE: 0.49012; U(DoE) (degrees) DE: 0.57464
530.8844 | UK: -0.024867; IT: 0.015133; DoE (degrees) UK: -0.028965; DoE (degrees) FR: 0.016937; DoE (degrees) IT: 0.021035; DoE (degrees) CH: 0.017937; DoE (degrees) SE: -0.043063; DoE (degrees) DE: 0.026937; DoE (degrees) UK: 0.49959; DoE (degrees) IT: 0.59966; U(DoE) (degrees) UK: 0.4807; U(DoE) (degrees) FR: 0.5165; U(DoE) (degrees) IT: 0.56674; U(DoE) (degrees) CH: 0.33103; U(DoE) (degrees) SE: 0.5165; U(DoE) (degrees) DE: 0.59731
562.3413 | UK: -0.027883; IT: 0.032117; DoE (degrees) UK: -0.029932; DoE (degrees) FR: 0.018019; DoE (degrees) IT: 0.035068; DoE (degrees) CH: 0.020019; DoE (degrees) SE: -0.041981; DoE (degrees) DE: 0.038019; DoE (degrees) UK: 0.49959; DoE (degrees) IT: 0.59966; U(DoE) (degrees) UK: 0.4807; U(DoE) (degrees) FR: 0.5165; U(DoE) (degrees) IT: 0.56674; U(DoE) (degrees) CH: 0.3317; U(DoE) (degrees) SE: 0.5165; U(DoE) (degrees) DE: 0.59731
595.6621 | UK: -0.026301; IT: -0.036301; DoE (degrees) UK: -0.02835; DoE (degrees) FR: 0.039601; DoE (degrees) IT: -0.03335; DoE (degrees) CH: 0.037601; DoE (degrees) SE: -0.020399; DoE (degrees) DE: 0.069601; DoE (degrees) UK: 0.49959; DoE (degrees) IT: 0.59966; U(DoE) (degrees) UK: 0.4807; U(DoE) (degrees) FR: 0.5165; U(DoE) (degrees) IT: 0.56674; U(DoE) (degrees) CH: 0.33205; U(DoE) (degrees) SE: 0.5165; U(DoE) (degrees) DE: 0.59731
630.9573 | UK: -0.02406; IT: 0.0059401; DoE (degrees) UK: -0.02406; DoE (degrees) FR: 0.03594; DoE (degrees) IT: 0.0059401; DoE (degrees) CH: 0.04094; DoE (degrees) SE: -0.01406; DoE (degrees) DE: 0.0059401; DoE (degrees) UK: 0.49959; DoE (degrees) IT: 0.59966; U(DoE) (degrees) UK: 0.4807; U(DoE) (degrees) FR: 0.5165; U(DoE) (degrees) IT: 0.56674; U(DoE) (degrees) CH: 0.33278; U(DoE) (degrees) SE: 0.5165; U(DoE) (degrees) DE: 0.59731
668.3439 | UK: -0.034216; IT: -0.034216; DoE (degrees) UK: -0.050609; DoE (degrees) FR: 0.0029972; DoE (degrees) IT: -0.010609; DoE (degrees) CH: 0.0019972; DoE (degrees) SE: -0.057003; DoE (degrees) DE: 0.012997; DoE (degrees) UK: 0.49959; DoE (degrees) IT: 0.59966; U(DoE) (degrees) UK: 0.4807; U(DoE) (degrees) FR: 0.5165; U(DoE) (degrees) IT: 0.56674; U(DoE) (degrees) CH: 0.33356; U(DoE) (degrees) SE: 0.5165; U(DoE) (degrees) DE: 0.59731
707.9458 | UK: -0.036758; IT: -0.056758; DoE (degrees) UK: -0.049054; DoE (degrees) FR: 0.0086514; DoE (degrees) IT: -0.039054; DoE (degrees) CH: 0.0036514; DoE (degrees) SE: -0.051349; DoE (degrees) DE: -0.021349; DoE (degrees) UK: 0.49959; DoE (degrees) IT: 0.59966; U(DoE) (degrees) UK: 0.4807; U(DoE) (degrees) FR: 0.5165; U(DoE) (degrees) IT: 0.56674; U(DoE) (degrees) CH: 0.3346; U(DoE) (degrees) SE: 0.5165; U(DoE) (degrees) DE: 0.59731
749.8942 | UK: -0.043776; IT: -0.063776; DoE (degrees) UK: -0.041727; DoE (degrees) FR: -0.019677; DoE (degrees) IT: -0.066727; DoE (degrees) CH: -0.015677; DoE (degrees) SE: -0.079677; DoE (degrees) DE: 0.030323; DoE (degrees) UK: 0.49959; DoE (degrees) IT: 0.59966; U(DoE) (degrees) UK: 0.4807; U(DoE) (degrees) FR: 0.5165; U(DoE) (degrees) IT: 0.56674; U(DoE) (degrees) CH: 0.33571; U(DoE) (degrees) SE: 0.5165; U(DoE) (degrees) DE: 0.59731
794.3282 | UK: -0.041713; IT: -0.051713; DoE (degrees) UK: -0.04786; DoE (degrees) FR: 0.015992; DoE (degrees) IT: -0.04286; DoE (degrees) CH: 0.021992; DoE (degrees) SE: -0.034008; DoE (degrees) DE: 0.065992; DoE (degrees) UK: 0.49959; DoE (degrees) IT: 0.59966; U(DoE) (degrees) UK: 0.47157; U(DoE) (degrees) FR: 0.5628; U(DoE) (degrees) IT: 0.5506; U(DoE) (degrees) CH: 0.40411; U(DoE) (degrees) SE: 0.5628; U(DoE) (degrees) DE: 0.63776
841.3951 | UK: -0.03129; IT: -0.02129; DoE (degrees) UK: -0.035388; DoE (degrees) FR: 0.030513; DoE (degrees) IT: -0.015388; DoE (degrees) CH: 0.031513; DoE (degrees) SE: -0.019487; DoE (degrees) DE: -0.0094867; DoE (degrees) UK: 0.49959; DoE (degrees) IT: 0.59966; U(DoE) (degrees) UK: 0.47157; U(DoE) (degrees) FR: 0.5628; U(DoE) (degrees) IT: 0.5506; U(DoE) (degrees) CH: 0.40514; U(DoE) (degrees) SE: 0.5628; U(DoE) (degrees) DE: 0.63776
891.2509 | UK: -0.092521; IT: -0.062521; DoE (degrees) UK: -0.096619; DoE (degrees) FR: -0.050718; DoE (degrees) IT: -0.056619; DoE (degrees) CH: -0.042718; DoE (degrees) SE: -0.090718; DoE (degrees) DE: -0.050718; DoE (degrees) UK: 0.49959; DoE (degrees) IT: 0.59966; U(DoE) (degrees) UK: 0.47157; U(DoE) (degrees) FR: 0.5628; U(DoE) (degrees) IT: 0.5506; U(DoE) (degrees) CH: 0.40623; U(DoE) (degrees) SE: 0.5628; U(DoE) (degrees) DE: 0.63776
944.0609 | UK: -0.038425; IT: 0.011575; DoE (degrees) UK: -0.019983; DoE (degrees) FR: 0.0084599; DoE (degrees) IT: -0.014983; DoE (degrees) CH: 0.0074599; DoE (degrees) SE: -0.04154; DoE (degrees) DE: 0.05846; DoE (degrees) UK: 0.49959; DoE (degrees) IT: 0.59966; U(DoE) (degrees) UK: 0.47157; U(DoE) (degrees) FR: 0.5628; U(DoE) (degrees) IT: 0.5506; U(DoE) (degrees) CH: 0.40738; U(DoE) (degrees) SE: 0.5628; U(DoE) (degrees) DE: 0.63776
1000 | UK: -0.096044; IT: 0.013956; DoE (degrees) UK: -0.087847; DoE (degrees) FR: -0.03965; DoE (degrees) IT: 0.002153; DoE (degrees) CH: -0.03365; DoE (degrees) SE: -0.08965; DoE (degrees) DE: -0.0096503; DoE (degrees) UK: 0.49943; DoE (degrees) IT: 0.59952; U(DoE) (degrees) UK: 0.47139; U(DoE) (degrees) FR: 0.56279; U(DoE) (degrees) IT: 0.55045; U(DoE) (degrees) CH: 0.40908; U(DoE) (degrees) SE: 0.56279; U(DoE) (degrees) DE: 0.63776
1059.2537 | UK: -0.026118; IT: -0.056118; DoE (degrees) UK: -0.061118; DoE (degrees) FR: -0.026118; DoE (degrees) IT: -0.021118; DoE (degrees) CH: -0.016118; DoE (degrees) SE: -0.066118; DoE (degrees) DE: 0.013882; DoE (degrees) UK: 0.59952; DoE (degrees) IT: 0.59952; U(DoE) (degrees) UK: 0.55822; U(DoE) (degrees) FR: 0.66425; U(DoE) (degrees) IT: 0.55822; U(DoE) (degrees) CH: 0.45091; U(DoE) (degrees) SE: 0.59264; U(DoE) (degrees) DE: 0.66425
1122.0185 | UK: -0.079884; IT: 0.0001162; DoE (degrees) UK: -0.067384; DoE (degrees) FR: 0.015116; DoE (degrees) IT: -0.012384; DoE (degrees) CH: 0.018116; DoE (degrees) SE: -0.034884; DoE (degrees) DE: 0.075116; DoE (degrees) UK: 0.59952; DoE (degrees) IT: 0.59952; U(DoE) (degrees) UK: 0.55822; U(DoE) (degrees) FR: 0.66425; U(DoE) (degrees) IT: 0.55822; U(DoE) (degrees) CH: 0.45241; U(DoE) (degrees) SE: 0.59264; U(DoE) (degrees) DE: 0.66425
1188.5022 | UK: -0.035266; IT: -0.015266; DoE (degrees) UK: -0.045266; DoE (degrees) FR: -0.025266; DoE (degrees) IT: -0.0052664; DoE (degrees) CH: -0.0072664; DoE (degrees) SE: -0.055266; DoE (degrees) DE: 0.0047336; DoE (degrees) UK: 0.59952; DoE (degrees) IT: 0.59952; U(DoE) (degrees) UK: 0.55822; U(DoE) (degrees) FR: 0.66425; U(DoE) (degrees) IT: 0.55822; U(DoE) (degrees) CH: 0.45453; U(DoE) (degrees) SE: 0.59264; U(DoE) (degrees) DE: 0.66425
1258.9254 | UK: -0.057234; IT: 0.0027664; DoE (degrees) UK: -0.064734; DoE (degrees) FR: 0.0077664; DoE (degrees) IT: 0.010266; DoE (degrees) CH: 0.021766; DoE (degrees) SE: -0.022234; DoE (degrees) DE: 0.017766; DoE (degrees) UK: 0.59952; DoE (degrees) IT: 0.59952; U(DoE) (degrees) UK: 0.55822; U(DoE) (degrees) FR: 0.66425; U(DoE) (degrees) IT: 0.55822; U(DoE) (degrees) CH: 0.45621; U(DoE) (degrees) SE: 0.66425; U(DoE) (degrees) DE: 0.66425
1333.5214 | UK: -0.074024; IT: -0.044024; DoE (degrees) UK: -0.094024; DoE (degrees) FR: -0.034024; DoE (degrees) IT: -0.024024; DoE (degrees) CH: -0.023024; DoE (degrees) SE: -0.064024; DoE (degrees) DE: -0.0040241; DoE (degrees) UK: 0.59952; DoE (degrees) IT: 0.59952; U(DoE) (degrees) UK: 0.55822; U(DoE) (degrees) FR: 0.66425; U(DoE) (degrees) IT: 0.55822; U(DoE) (degrees) CH: 0.45915; U(DoE) (degrees) SE: 0.66425; U(DoE) (degrees) DE: 0.66425
1412.5375 | UK: -0.074429; IT: -0.054429; DoE (degrees) UK: -0.081929; DoE (degrees) FR: -0.0094291; DoE (degrees) IT: -0.046929; DoE (degrees) CH: -0.0064291; DoE (degrees) SE: -0.049429; DoE (degrees) DE: -0.039429; DoE (degrees) UK: 0.59952; DoE (degrees) IT: 0.59952; U(DoE) (degrees) UK: 0.55822; U(DoE) (degrees) FR: 0.66425; U(DoE) (degrees) IT: 0.55822; U(DoE) (degrees) CH: 0.46198; U(DoE) (degrees) SE: 0.66425; U(DoE) (degrees) DE: 0.66425
1496.2357 | UK: -0.08466; IT: -0.02466; DoE (degrees) UK: -0.08466; DoE (degrees) FR: 0.02534; DoE (degrees) IT: -0.02466; DoE (degrees) CH: 0.02034; DoE (degrees) SE: -0.01466; DoE (degrees) DE: 0.07534; DoE (degrees) UK: 0.59934; DoE (degrees) IT: 0.59934; U(DoE) (degrees) UK: 0.55803; U(DoE) (degrees) FR: 0.66424; U(DoE) (degrees) IT: 0.55803; U(DoE) (degrees) CH: 0.46494; U(DoE) (degrees) SE: 0.66424; U(DoE) (degrees) DE: 0.66424
1584.8932 | UK: -0.12144; IT: -0.051441; DoE (degrees) UK: -0.11894; DoE (degrees) FR: -0.026441; DoE (degrees) IT: -0.053941; DoE (degrees) CH: -0.037441; DoE (degrees) SE: -0.076441; DoE (degrees) DE: -0.056441; DoE (degrees) UK: 0.59927; DoE (degrees) IT: 0.59927; U(DoE) (degrees) UK: 0.543; U(DoE) (degrees) FR: 0.71185; U(DoE) (degrees) IT: 0.543; U(DoE) (degrees) CH: 0.53382; U(DoE) (degrees) SE: 0.71185; U(DoE) (degrees) DE: 0.71185
1678.804 | UK: -0.10206; IT: -0.032055; DoE (degrees) UK: -0.097055; DoE (degrees) FR: 0.017945; DoE (degrees) IT: -0.037055; DoE (degrees) CH: 0.0069447; DoE (degrees) SE: -0.022055; DoE (degrees) DE: 0.057945; DoE (degrees) UK: 0.59927; DoE (degrees) IT: 0.59927; U(DoE) (degrees) UK: 0.543; U(DoE) (degrees) FR: 0.78533; U(DoE) (degrees) IT: 0.543; U(DoE) (degrees) CH: 0.53738; U(DoE) (degrees) SE: 0.71185; U(DoE) (degrees) DE: 0.71185
1778.2794 | UK: -0.10093; IT: -0.060934; DoE (degrees) UK: -0.10593; DoE (degrees) FR: -0.010934; DoE (degrees) IT: -0.055934; DoE (degrees) CH: -0.0089338; DoE (degrees) SE: -0.050934; DoE (degrees) DE: 0.049066; DoE (degrees) UK: 0.59926; DoE (degrees) IT: 0.59926; U(DoE) (degrees) UK: 0.54298; U(DoE) (degrees) FR: 0.78532; U(DoE) (degrees) IT: 0.54298; U(DoE) (degrees) CH: 0.54114; U(DoE) (degrees) SE: 0.71185; U(DoE) (degrees) DE: 0.71185
1883.6491 | UK: -0.10395; IT: -0.033954; DoE (degrees) UK: -0.11645; DoE (degrees) FR: -0.038954; DoE (degrees) IT: -0.021454; DoE (degrees) CH: -0.035954; DoE (degrees) SE: -0.068954; DoE (degrees) DE: -0.0089537; DoE (degrees) UK: 0.59926; DoE (degrees) IT: 0.59926; U(DoE) (degrees) UK: 0.54298; U(DoE) (degrees) FR: 0.78532; U(DoE) (degrees) IT: 0.54298; U(DoE) (degrees) CH: 0.54584; U(DoE) (degrees) SE: 0.71185; U(DoE) (degrees) DE: 0.71185
1995.2623 | UK: -0.11338; IT: -0.043382; DoE (degrees) UK: -0.10838; DoE (degrees) FR: -0.0033822; DoE (degrees) IT: -0.048382; DoE (degrees) CH: -0.00038224; DoE (degrees) SE: -0.023382; DoE (degrees) DE: 0.046618; DoE (degrees) UK: 0.59917; DoE (degrees) IT: 0.59917; U(DoE) (degrees) UK: 0.54288; U(DoE) (degrees) FR: 0.78529; U(DoE) (degrees) IT: 0.54288; U(DoE) (degrees) CH: 0.55076; U(DoE) (degrees) SE: 0.71182; U(DoE) (degrees) DE: 0.71182
2113.489 | UK: -0.092851; IT: -0.092851; DoE (degrees) UK: -0.13557; DoE (degrees) FR: -0.054661; DoE (degrees) IT: -0.063752; DoE (degrees) CH: -0.052661; DoE (degrees) SE: -0.084661; DoE (degrees) DE: -0.034661; DoE (degrees) UK: 0.69928; DoE (degrees) IT: 0.59916; U(DoE) (degrees) UK: 0.59594; U(DoE) (degrees) FR: 0.89213; U(DoE) (degrees) IT: 0.54488; U(DoE) (degrees) CH: 0.59861; U(DoE) (degrees) SE: 0.74558; U(DoE) (degrees) DE: 0.74558
2238.7211 | UK: -0.11747; IT: -0.067472; DoE (degrees) UK: -0.12076; DoE (degrees) FR: -0.032996; DoE (degrees) IT: -0.065234; DoE (degrees) CH: -0.026996; DoE (degrees) SE: -0.052996; DoE (degrees) DE: 0.037004; DoE (degrees) UK: 0.69907; DoE (degrees) IT: 0.59892; U(DoE) (degrees) UK: 0.5957; U(DoE) (degrees) FR: 0.89203; U(DoE) (degrees) IT: 0.5446; U(DoE) (degrees) CH: 0.60436; U(DoE) (degrees) SE: 0.74546; U(DoE) (degrees) DE: 0.74546
2371.3737 | UK: -0.096085; IT: -0.056085; DoE (degrees) UK: -0.10594; DoE (degrees) FR: -0.012656; DoE (degrees) IT: -0.04937; DoE (degrees) CH: -0.016656; DoE (degrees) SE: -0.042656; DoE (degrees) DE: 0.057344; DoE (degrees) UK: 0.69907; DoE (degrees) IT: 0.59892; U(DoE) (degrees) UK: 0.5957; U(DoE) (degrees) FR: 0.89203; U(DoE) (degrees) IT: 0.5446; U(DoE) (degrees) CH: 0.61142; U(DoE) (degrees) SE: 0.74546; U(DoE) (degrees) DE: 0.74546
2511.8864 | UK: -0.10135; IT: -0.061349; DoE (degrees) UK: -0.10135; DoE (degrees) FR: -0.031349; DoE (degrees) IT: -0.061349; DoE (degrees) CH: -0.016349; DoE (degrees) SE: -0.041349; DoE (degrees) DE: 0.038651; DoE (degrees) UK: 0.69907; DoE (degrees) IT: 0.59892; U(DoE) (degrees) UK: 0.5957; U(DoE) (degrees) FR: 0.89203; U(DoE) (degrees) IT: 0.5446; U(DoE) (degrees) CH: 0.61842; U(DoE) (degrees) SE: 0.74546; U(DoE) (degrees) DE: 0.74546
2660.7251 | UK: -0.11455; IT: -0.064547; DoE (degrees) UK: -0.13761; DoE (degrees) FR: -0.070665; DoE (degrees) IT: -0.047606; DoE (degrees) CH: -0.057665; DoE (degrees) SE: -0.080665; DoE (degrees) DE: -0.030665; DoE (degrees) UK: 0.69901; DoE (degrees) IT: 0.59885; U(DoE) (degrees) UK: 0.61119; U(DoE) (degrees) FR: 0.98857; U(DoE) (degrees) IT: 0.54453; U(DoE) (degrees) CH: 0.65093; U(DoE) (degrees) SE: 0.76634; U(DoE) (degrees) DE: 0.76634
2818.3829 | UK: -0.14673; IT: -0.056733; DoE (degrees) UK: -0.14673; DoE (degrees) FR: -0.046733; DoE (degrees) IT: -0.056733; DoE (degrees) CH: -0.048733; DoE (degrees) SE: -0.066733; DoE (degrees) DE: 0.043267; DoE (degrees) UK: 0.69898; DoE (degrees) IT: 0.59881; U(DoE) (degrees) UK: 0.61115; U(DoE) (degrees) FR: 0.98856; U(DoE) (degrees) IT: 0.54448; U(DoE) (degrees) CH: 0.65971; U(DoE) (degrees) SE: 0.76632; U(DoE) (degrees) DE: 0.76632
2985.3826 | UK: -0.18517; IT: -0.12517; DoE (degrees) UK: -0.20534; DoE (degrees) FR: -0.075521; DoE (degrees) IT: -0.11034; DoE (degrees) CH: -0.051521; DoE (degrees) SE: -0.065521; DoE (degrees) DE: 0.0044792; DoE (degrees) UK: 0.69876; DoE (degrees) IT: 0.59855; U(DoE) (degrees) UK: 0.6109; U(DoE) (degrees) FR: 0.98846; U(DoE) (degrees) IT: 0.5442; U(DoE) (degrees) CH: 0.66879; U(DoE) (degrees) SE: 0.76619; U(DoE) (degrees) DE: 0.76619
3162.2777 | UK: -0.11875; IT: -0.15875; DoE (degrees) UK: -0.15046; DoE (degrees) FR: -0.062166; DoE (degrees) IT: -0.13546; DoE (degrees) CH: -0.051166; DoE (degrees) SE: -0.062166; DoE (degrees) DE: -0.012166; DoE (degrees) UK: 0.69874; DoE (degrees) IT: 0.59853; U(DoE) (degrees) UK: 0.61088; U(DoE) (degrees) FR: 0.98845; U(DoE) (degrees) IT: 0.54418; U(DoE) (degrees) CH: 0.67997; U(DoE) (degrees) SE: 0.83489; U(DoE) (degrees) DE: 0.76618
3349.6544 | UK: -0.15008; IT: -0.13008; DoE (degrees) UK: -0.17314; DoE (degrees) FR: -0.076202; DoE (degrees) IT: -0.11314; DoE (degrees) CH: -0.024202; DoE (degrees) SE: -0.056202; DoE (degrees) DE: 0.003798; DoE (degrees) UK: 0.6987; DoE (degrees) IT: 0.59848; U(DoE) (degrees) UK: 0.60842; U(DoE) (degrees) FR: 1.075; U(DoE) (degrees) IT: 0.54266; U(DoE) (degrees) CH: 0.6946; U(DoE) (degrees) SE: 0.83998; U(DoE) (degrees) DE: 0.77173
3548.1339 | UK: -0.16862; IT: -0.12862; DoE (degrees) UK: -0.20609; DoE (degrees) FR: -0.10356; DoE (degrees) IT: -0.10109; DoE (degrees) CH: -0.087558; DoE (degrees) SE: -0.083558; DoE (degrees) DE: -0.073558; DoE (degrees) UK: 0.69864; DoE (degrees) IT: 0.59842; U(DoE) (degrees) UK: 0.60836; U(DoE) (degrees) FR: 1.075; U(DoE) (degrees) IT: 0.54259; U(DoE) (degrees) CH: 0.70988; U(DoE) (degrees) SE: 0.83995; U(DoE) (degrees) DE: 0.7717
3758.374 | UK: -0.17743; IT: -0.12743; DoE (degrees) UK: -0.19184; DoE (degrees) FR: -0.056251; DoE (degrees) IT: -0.11684; DoE (degrees) CH: -0.032251; DoE (degrees) SE: -0.026251; DoE (degrees) DE: -0.0062507; DoE (degrees) UK: 0.69861; DoE (degrees) IT: 0.59838; U(DoE) (degrees) UK: 0.60832; U(DoE) (degrees) FR: 1.0749; U(DoE) (degrees) IT: 0.54255; U(DoE) (degrees) CH: 0.7201; U(DoE) (degrees) SE: 0.83993; U(DoE) (degrees) DE: 0.77168
3981.0717 | UK: -0.1452; IT: -0.095199; DoE (degrees) UK: -0.15673; DoE (degrees) FR: 0.011742; DoE (degrees) IT: -0.086729; DoE (degrees) CH: -0.00025816; DoE (degrees) SE: 0.0017418; DoE (degrees) DE: 0.021742; DoE (degrees) UK: 0.69858; DoE (degrees) IT: 0.59834; U(DoE) (degrees) UK: 0.60828; U(DoE) (degrees) FR: 1.0749; U(DoE) (degrees) IT: 0.54251; U(DoE) (degrees) CH: 0.73122; U(DoE) (degrees) SE: 0.83992; U(DoE) (degrees) DE: 0.77166
4216.965 | UK: -0.20379; IT: -0.11379; DoE (degrees) UK: -0.22396; DoE (degrees) FR: 0.025861; DoE (degrees) IT: -0.098963; DoE (degrees) CH: -0.013139; DoE (degrees) SE: -0.0041393; DoE (degrees) DE: 0.015861; DoE (degrees) UK: 0.69857; DoE (degrees) IT: 0.59834; U(DoE) (degrees) UK: 0.6114; U(DoE) (degrees) FR: 1.1551; U(DoE) (degrees) IT: 0.5444; U(DoE) (degrees) CH: 0.73703; U(DoE) (degrees) SE: 0.83328; U(DoE) (degrees) DE: 0.76444
4466.8359 | UK: -0.2154; IT: -0.1354; DoE (degrees) UK: -0.22981; DoE (degrees) FR: -0.014226; DoE (degrees) IT: -0.12481; DoE (degrees) CH: -0.075226; DoE (degrees) SE: -0.044226; DoE (degrees) DE: -0.014226; DoE (degrees) UK: 0.69857; DoE (degrees) IT: 0.59833; U(DoE) (degrees) UK: 0.6114; U(DoE) (degrees) FR: 1.1551; U(DoE) (degrees) IT: 0.54439; U(DoE) (degrees) CH: 0.74642; U(DoE) (degrees) SE: 0.83328; U(DoE) (degrees) DE: 0.76443
4731.5126 | UK: -0.13248; IT: -0.12248; DoE (degrees) UK: -0.15265; DoE (degrees) FR: 0.027172; DoE (degrees) IT: -0.10765; DoE (degrees) CH: -0.036828; DoE (degrees) SE: -0.032828; DoE (degrees) DE: 0.007172; DoE (degrees) UK: 0.69832; DoE (degrees) IT: 0.59804; U(DoE) (degrees) UK: 0.61111; U(DoE) (degrees) FR: 1.1551; U(DoE) (degrees) IT: 0.54407; U(DoE) (degrees) CH: 0.75264; U(DoE) (degrees) SE: 0.83315; U(DoE) (degrees) DE: 0.76429
5011.8723 | UK: -0.11695; IT: -0.22695; DoE (degrees) UK: -0.16595; DoE (degrees) FR: -0.034953; DoE (degrees) IT: -0.19095; DoE (degrees) CH: -0.11395; DoE (degrees) SE: -0.084953; DoE (degrees) DE: -0.054953; DoE (degrees) UK: 0.69764; DoE (degrees) IT: 0.59724; U(DoE) (degrees) UK: 0.61033; U(DoE) (degrees) FR: 1.1548; U(DoE) (degrees) IT: 0.54319; U(DoE) (degrees) CH: 0.75611; U(DoE) (degrees) SE: 0.8328; U(DoE) (degrees) DE: 1.1548
5308.8444 | UK: -0.14681; IT: -0.21681; DoE (degrees) UK: -0.16987; DoE (degrees) FR: 0.037073; DoE (degrees) IT: -0.19987; DoE (degrees) CH: -0.039927; DoE (degrees) SE: 0.0070732; DoE (degrees) DE: 0.017073; DoE (degrees) UK: 0.69753; DoE (degrees) IT: 0.59711; U(DoE) (degrees) UK: 0.62775; U(DoE) (degrees) FR: 1.2168; U(DoE) (degrees) IT: 0.55374; U(DoE) (degrees) CH: 0.7209; U(DoE) (degrees) SE: 0.79416; U(DoE) (degrees) DE: 1.1272
5623.4133 | UK: -0.071451; IT: -0.18145; DoE (degrees) UK: -0.11469; DoE (degrees) FR: -0.0079215; DoE (degrees) IT: -0.14969; DoE (degrees) CH: -0.10492; DoE (degrees) SE: 0.022079; DoE (degrees) DE: 0.082079; DoE (degrees) UK: 0.69752; DoE (degrees) IT: 0.59711; U(DoE) (degrees) UK: 0.62774; U(DoE) (degrees) FR: 1.2168; U(DoE) (degrees) IT: 0.55373; U(DoE) (degrees) CH: 0.72159; U(DoE) (degrees) SE: 0.79416; U(DoE) (degrees) DE: 1.1272
5956.6214 | UK: -0.13084; IT: -0.24084; DoE (degrees) UK: -0.2029; DoE (degrees) FR: 0.0050401; DoE (degrees) IT: -0.1879; DoE (degrees) CH: -0.11496; DoE (degrees) SE: 0.03504; DoE (degrees) DE: 0.06504; DoE (degrees) UK: 0.69746; DoE (degrees) IT: 0.59704; U(DoE) (degrees) UK: 0.62768; U(DoE) (degrees) FR: 1.2168; U(DoE) (degrees) IT: 0.55366; U(DoE) (degrees) CH: 0.71536; U(DoE) (degrees) SE: 0.79414; U(DoE) (degrees) DE: 1.1272
6309.5734 | UK: -0.055543; IT: -0.11554; DoE (degrees) UK: -0.095543; DoE (degrees) FR: 0.15446; DoE (degrees) IT: -0.075543; DoE (degrees) CH: -0.0095428; DoE (degrees) SE: 0.14446; DoE (degrees) DE: 0.16446; DoE (degrees) UK: 0.5968; DoE (degrees) IT: 0.5968; U(DoE) (degrees) UK: 0.5452; U(DoE) (degrees) FR: 1.2019; U(DoE) (degrees) IT: 0.5452; U(DoE) (degrees) CH: 0.68893; U(DoE) (degrees) SE: 0.77105; U(DoE) (degrees) DE: 1.1111
6683.4392 | UK: 0.0010809; IT: -0.14892; DoE (degrees) UK: -0.038919; DoE (degrees) FR: 0.18108; DoE (degrees) IT: -0.10892; DoE (degrees) CH: -0.025919; DoE (degrees) SE: 0.18108; DoE (degrees) DE: 0.23108; DoE (degrees) UK: 0.59623; DoE (degrees) IT: 0.59623; U(DoE) (degrees) UK: 0.56318; U(DoE) (degrees) FR: 1.2625; U(DoE) (degrees) IT: 0.56318; U(DoE) (degrees) CH: 0.63869; U(DoE) (degrees) SE: 0.71688; U(DoE) (degrees) DE: 1.0742
7079.4578 | UK: -0.021379; IT: -0.15138; DoE (degrees) UK: -0.060751; DoE (degrees) FR: 0.24006; DoE (degrees) IT: -0.079129; DoE (degrees) CH: 0.01506; DoE (degrees) SE: 0.25006; DoE (degrees) DE: 0.26006; DoE (degrees) UK: 0.59607; DoE (degrees) IT: 0.69663; U(DoE) (degrees) UK: 0.56868; U(DoE) (degrees) FR: 1.2729; U(DoE) (degrees) IT: 0.6147; U(DoE) (degrees) CH: 0.66429; U(DoE) (degrees) SE: 0.73496; U(DoE) (degrees) DE: 1.0864
7498.9421 | UK: -0.032951; IT: -0.052951; DoE (degrees) UK: -0.05391; DoE (degrees) FR: 0.29518; DoE (degrees) IT: -0.019679; DoE (degrees) CH: 0.05218; DoE (degrees) SE: 0.31518; DoE (degrees) DE: 0.30518; DoE (degrees) UK: 0.5959; DoE (degrees) IT: 0.69649; U(DoE) (degrees) UK: 0.56825; U(DoE) (degrees) FR: 1.2723; U(DoE) (degrees) IT: 0.63558; U(DoE) (degrees) CH: 0.67953; U(DoE) (degrees) SE: 0.73399; U(DoE) (degrees) DE: 1.0857
7943.2823 | UK: 0.084168; IT: -0.015832; DoE (degrees) UK: 0.064704; DoE (degrees) FR: 0.35535; DoE (degrees) IT: 0.04207; DoE (degrees) CH: 0.067349; DoE (degrees) SE: 0.39535; DoE (degrees) DE: 0.38535; DoE (degrees) UK: 0.49174; DoE (degrees) IT: 0.79487; U(DoE) (degrees) UK: 0.47638; U(DoE) (degrees) FR: 1.3583; U(DoE) (degrees) IT: 0.70721; U(DoE) (degrees) CH: 0.68691; U(DoE) (degrees) SE: 0.71755; U(DoE) (degrees) DE: 1.0747
8413.9514 | UK: 0.16074; IT: 0.040738; DoE (degrees) UK: 0.10484; DoE (degrees) FR: 0.35898; DoE (degrees) IT: 0.1002; DoE (degrees) CH: 0.071984; DoE (degrees) SE: 0.44898; DoE (degrees) DE: 0.34898; DoE (degrees) UK: 0.79444; DoE (degrees) IT: 0.79444; U(DoE) (degrees) UK: 0.72381; U(DoE) (degrees) FR: 1.4498; U(DoE) (degrees) IT: 0.71401; U(DoE) (degrees) CH: 0.8882; U(DoE) (degrees) SE: 0.87858; U(DoE) (degrees) DE: 1.1882
8912.5094 | UK: 0.10356; IT: 0.023556; DoE (degrees) UK: 0.040005; DoE (degrees) FR: 0.33654; DoE (degrees) IT: 0.091101; DoE (degrees) CH: 0.040536; DoE (degrees) SE: 0.38654; DoE (degrees) DE: 0.34654; DoE (degrees) UK: 0.79197; DoE (degrees) IT: 0.79197; U(DoE) (degrees) UK: 0.72103; U(DoE) (degrees) FR: 1.4493; U(DoE) (degrees) IT: 0.71132; U(DoE) (degrees) CH: 0.9482; U(DoE) (degrees) SE: 0.94894; U(DoE) (degrees) DE: 1.1876
9440.6088 | UK: 0.28353; IT: 0.17353; DoE (degrees) UK: 0.19597; DoE (degrees) FR: 0.7888; DoE (degrees) IT: 0.29957; DoE (degrees) CH: 0.3008; DoE (degrees) SE: 0.6488; DoE (degrees) DE: 0.5788; DoE (degrees) UK: 0.78789; DoE (degrees) IT: 0.88925; U(DoE) (degrees) UK: 0.72332; U(DoE) (degrees) FR: 1.552; U(DoE) (degrees) IT: 0.76716; U(DoE) (degrees) CH: 1.0728; U(DoE) (degrees) SE: 0.96881; U(DoE) (degrees) DE: 1.2036
10000 | UK: 0.18501; IT: -0.25499; DoE (degrees) UK: 0.079409; DoE (degrees) FR: 0.65442; DoE (degrees) IT: -0.028137; DoE (degrees) CH: -0.00058025; DoE (degrees) SE: 0.90442; DoE (degrees) DE: 0.41442; DoE (degrees) UK: 0.784; DoE (degrees) IT: 1.0884; U(DoE) (degrees) UK: 0.73311; U(DoE) (degrees) FR: 1.7573; U(DoE) (degrees) IT: 0.91011; U(DoE) (degrees) CH: 1.2647; U(DoE) (degrees) SE: 1.009; U(DoE) (degrees) DE: 1.2362

## 811014 Level
| Sensitivity Level (4160 811014) dB re 1 V/Pa | [12 empty] | Expanded Uncertainties (dB) | [11 empty] | Level difference relative to estimated sensitivity level (dB)
Frequency (Hz) | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG | unweighted mean | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG
1.9953 | -25.25 | -25.347 | -25.31 | -25.328 |  |  |  | -25.312 |  | -25.34 |  |  | -25.3145 | Expanded Uncertainties (dB) UK: 0.25; Expanded Uncertainties (dB) FR: 0.14; Expanded Uncertainties (dB) IT: 0.14; Expanded Uncertainties (dB) CH: 0.127; Expanded Uncertainties (dB) AT: 0.058; Expanded Uncertainties (dB) DE: 0.23; Level difference relative to estimated sensitivity level (dB) UK: 0.0645; Level difference relative to estimated sensitivity level (dB) FR: -0.0325; Level difference relative to estimated sensitivity level (dB) IT: 0.0045; Level difference relative to estimated sensitivity level (dB) CH: -0.0135; Level difference relative to estimated sensitivity level (dB) AT: 0.0025; Level difference relative to estimated sensitivity level (dB) DE: -0.0255
2.5119 | -25.38 | -25.491 | -25.43 | -25.465 |  |  |  | -25.45 |  | -25.47 |  |  | -25.4477 | Expanded Uncertainties (dB) UK: 0.25; Expanded Uncertainties (dB) FR: 0.08; Expanded Uncertainties (dB) IT: 0.11; Expanded Uncertainties (dB) CH: 0.104; Expanded Uncertainties (dB) AT: 0.051; Expanded Uncertainties (dB) DE: 0.23; Level difference relative to estimated sensitivity level (dB) UK: 0.067667; Level difference relative to estimated sensitivity level (dB) FR: -0.043333; Level difference relative to estimated sensitivity level (dB) IT: 0.017667; Level difference relative to estimated sensitivity level (dB) CH: -0.017333; Level difference relative to estimated sensitivity level (dB) AT: -0.0023333; Level difference relative to estimated sensitivity level (dB) DE: -0.022333
3.1623 | -25.53 | -25.608 | -25.55 | -25.589 |  |  |  | -25.578 |  | -25.59 |  |  | -25.5742 | Expanded Uncertainties (dB) UK: 0.25; Expanded Uncertainties (dB) FR: 0.06; Expanded Uncertainties (dB) IT: 0.1; Expanded Uncertainties (dB) CH: 0.086; Expanded Uncertainties (dB) AT: 0.046; Expanded Uncertainties (dB) DE: 0.23; Level difference relative to estimated sensitivity level (dB) UK: 0.044167; Level difference relative to estimated sensitivity level (dB) FR: -0.033833; Level difference relative to estimated sensitivity level (dB) IT: 0.024167; Level difference relative to estimated sensitivity level (dB) CH: -0.014833; Level difference relative to estimated sensitivity level (dB) AT: -0.0038333; Level difference relative to estimated sensitivity level (dB) DE: -0.015833
3.9811 | -25.65 | -25.715 | -25.66 | -25.7 |  |  |  | -25.69 |  | -25.69 |  |  | -25.6842 | Expanded Uncertainties (dB) UK: 0.07; Expanded Uncertainties (dB) FR: 0.05; Expanded Uncertainties (dB) IT: 0.09; Expanded Uncertainties (dB) CH: 0.072; Expanded Uncertainties (dB) AT: 0.043; Expanded Uncertainties (dB) DE: 0.12; Level difference relative to estimated sensitivity level (dB) UK: 0.034167; Level difference relative to estimated sensitivity level (dB) FR: -0.030833; Level difference relative to estimated sensitivity level (dB) IT: 0.024167; Level difference relative to estimated sensitivity level (dB) CH: -0.015833; Level difference relative to estimated sensitivity level (dB) AT: -0.0058333; Level difference relative to estimated sensitivity level (dB) DE: -0.0058333
5.0119 | -25.74 | -25.808 | -25.75 | -25.797 |  |  |  | -25.788 |  | -25.79 |  |  | -25.7788 | Expanded Uncertainties (dB) UK: 0.07; Expanded Uncertainties (dB) FR: 0.04; Expanded Uncertainties (dB) IT: 0.08; Expanded Uncertainties (dB) CH: 0.063; Expanded Uncertainties (dB) AT: 0.042; Expanded Uncertainties (dB) DE: 0.12; Level difference relative to estimated sensitivity level (dB) UK: 0.038833; Level difference relative to estimated sensitivity level (dB) FR: -0.029167; Level difference relative to estimated sensitivity level (dB) IT: 0.028833; Level difference relative to estimated sensitivity level (dB) CH: -0.018167; Level difference relative to estimated sensitivity level (dB) AT: -0.0091667; Level difference relative to estimated sensitivity level (dB) DE: -0.011167
6.3096 | -25.84 | -25.89 | -25.84 | -25.88 |  |  |  | -25.873 |  | -25.87 |  |  | -25.8655 | Expanded Uncertainties (dB) UK: 0.07; Expanded Uncertainties (dB) FR: 0.04; Expanded Uncertainties (dB) IT: 0.07; Expanded Uncertainties (dB) CH: 0.056; Expanded Uncertainties (dB) AT: 0.041; Expanded Uncertainties (dB) DE: 0.12; Level difference relative to estimated sensitivity level (dB) UK: 0.0255; Level difference relative to estimated sensitivity level (dB) FR: -0.0245; Level difference relative to estimated sensitivity level (dB) IT: 0.0255; Level difference relative to estimated sensitivity level (dB) CH: -0.0145; Level difference relative to estimated sensitivity level (dB) AT: -0.0075; Level difference relative to estimated sensitivity level (dB) DE: -0.0045
7.9433 | -25.91 | -25.952 | -25.91 | -25.951 |  |  |  | -25.945 |  | -25.94 |  |  | -25.9347 | Expanded Uncertainties (dB) UK: 0.06; Expanded Uncertainties (dB) FR: 0.03; Expanded Uncertainties (dB) IT: 0.06; Expanded Uncertainties (dB) CH: 0.051; Expanded Uncertainties (dB) AT: 0.04; Expanded Uncertainties (dB) DE: 0.08; Level difference relative to estimated sensitivity level (dB) UK: 0.024667; Level difference relative to estimated sensitivity level (dB) FR: -0.017333; Level difference relative to estimated sensitivity level (dB) IT: 0.024667; Level difference relative to estimated sensitivity level (dB) CH: -0.016333; Level difference relative to estimated sensitivity level (dB) AT: -0.010333; Level difference relative to estimated sensitivity level (dB) DE: -0.0053333
10 | -25.97 | -26.01 | -25.97 | -26.011 |  |  |  | -26.005 |  | -26 | -26.077 |  | -26.0061 | Expanded Uncertainties (dB) UK: 0.06; Expanded Uncertainties (dB) FR: 0.03; Expanded Uncertainties (dB) IT: 0.06; Expanded Uncertainties (dB) CH: 0.048; Expanded Uncertainties (dB) AT: 0.04; Expanded Uncertainties (dB) DE: 0.08; Expanded Uncertainties (dB) ES: 0.08; Level difference relative to estimated sensitivity level (dB) UK: 0.036143; Level difference relative to estimated sensitivity level (dB) FR: -0.0038571; Level difference relative to estimated sensitivity level (dB) IT: 0.036143; Level difference relative to estimated sensitivity level (dB) CH: -0.0048571; Level difference relative to estimated sensitivity level (dB) AT: 0.0011429; Level difference relative to estimated sensitivity level (dB) DE: 0.0061429; Level difference relative to estimated sensitivity level (dB) ES: -0.070857
12.5893 | -26.03 | -26.057 | -26.02 | -26.06 |  |  |  | -26.055 |  | -26.04 | -26.123 |  | -26.055 | Expanded Uncertainties (dB) UK: 0.06; Expanded Uncertainties (dB) FR: 0.03; Expanded Uncertainties (dB) IT: 0.05; Expanded Uncertainties (dB) CH: 0.046; Expanded Uncertainties (dB) AT: 0.039; Expanded Uncertainties (dB) DE: 0.08; Expanded Uncertainties (dB) ES: 0.08; Level difference relative to estimated sensitivity level (dB) UK: 0.025; Level difference relative to estimated sensitivity level (dB) FR: -0.002; Level difference relative to estimated sensitivity level (dB) IT: 0.035; Level difference relative to estimated sensitivity level (dB) CH: -0.005; Level difference relative to estimated sensitivity level (dB) AT: 0; Level difference relative to estimated sensitivity level (dB) DE: 0.015; Level difference relative to estimated sensitivity level (dB) ES: -0.068
15.8489 | -26.07 | -26.093 | -26.07 | -26.101 |  |  |  | -26.096 |  | -26.09 | -26.162 |  | -26.0974 | Expanded Uncertainties (dB) UK: 0.05; Expanded Uncertainties (dB) FR: 0.03; Expanded Uncertainties (dB) IT: 0.05; Expanded Uncertainties (dB) CH: 0.044; Expanded Uncertainties (dB) AT: 0.039; Expanded Uncertainties (dB) DE: 0.05; Expanded Uncertainties (dB) ES: 0.08; Level difference relative to estimated sensitivity level (dB) UK: 0.027429; Level difference relative to estimated sensitivity level (dB) FR: 0.0044286; Level difference relative to estimated sensitivity level (dB) IT: 0.027429; Level difference relative to estimated sensitivity level (dB) CH: -0.0035714; Level difference relative to estimated sensitivity level (dB) AT: 0.0014286; Level difference relative to estimated sensitivity level (dB) DE: 0.0074286; Level difference relative to estimated sensitivity level (dB) ES: -0.064571
19.9526 | -26.12 | -26.128 | -26.1 | -26.134 | -26.136 | -26.156 | -26.2148 | -26.131 | -26.109 | -26.12 | -26.196 | -26.137 | -26.1402 | 0.03 | 0.03 | 0.04 | 0.044 | 0.04 | 0.07 | 0.09 | 0.039 | 0.05 | 0.05 | 0.08 | 0.064 | 0.020154 | 0.012154 | 0.040154 | 0.0061538 | 0.0041538 | -0.015846 | -0.074691 | 0.0091538 | 0.031154 | 0.020154 | -0.055846 | 0.0031538
21.1349 | -26.13 | -26.14 | -26.11 | -26.142 |  |  | -26.2209 | -26.141 | -26.11 | -26.13 |  |  | -26.1405 | 0.03 | 0.03 | 0.04 | 0.044 |  |  | 0.09 | 0.039 | 0.05 | 0.05 |  |  | 0.010493 | 0.00049333 | 0.030493 | -0.0015067 |  |  | -0.080453 | -0.00050667 | 0.030493 | 0.010493
22.3872 | -26.13 | -26.146 | -26.12 | -26.15 |  |  | -26.2267 | -26.15 | -26.123 | -26.14 |  |  | -26.1482 | 0.03 | 0.03 | 0.04 | 0.044 |  |  | 0.09 | 0.039 | 0.04 | 0.05 |  |  | 0.018206 | 0.0022064 | 0.028206 | -0.0017936 |  |  | -0.078445 | -0.0017936 | 0.025206 | 0.0082064
23.7137 | -26.14 | -26.148 | -26.12 | -26.156 |  |  | -26.2325 | -26.156 | -26.129 | -26.14 |  |  | -26.1527 | 0.03 | 0.03 | 0.04 | 0.044 |  |  | 0.09 | 0.039 | 0.04 | 0.05 |  |  | 0.01269 | 0.0046896 | 0.03269 | -0.0033104 |  |  | -0.079827 | -0.0033104 | 0.02369 | 0.01269
25.1189 | -26.15 | -26.156 | -26.13 | -26.166 | -26.163 | -26.182 | -26.2383 | -26.163 | -26.135 | -26.15 | -26.225 | -26.161 | -26.1683 | 0.03 | 0.03 | 0.03 | 0.044 | 0.04 | 0.06 | 0.09 | 0.038 | 0.04 | 0.05 | 0.05 | 0.0596 | 0.018278 | 0.012278 | 0.038278 | 0.0022778 | 0.0052778 | -0.013722 | -0.070056 | 0.0052778 | 0.033278 | 0.018278 | -0.056722 | 0.0072778
26.6073 | -26.16 | -26.16 | -26.14 | -26.169 |  |  | -26.2434 | -26.17 | -26.143 | -26.15 |  |  | -26.1669 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.05 |  |  | 0.0069282 | 0.0069282 | 0.026928 | -0.0020718 |  |  | -0.076497 | -0.0030718 | 0.023928 | 0.016928
28.1838 | -26.16 | -26.169 | -26.15 | -26.17 |  |  | -26.2488 | -26.177 | -26.148 | -26.16 |  |  | -26.1728 | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.09 | 0.038 | 0.04 | 0.05 |  |  | 0.01285 | 0.0038496 | 0.02285 | 0.0028496 |  |  | -0.075947 | -0.0041504 | 0.02485 | 0.01285
29.8538 | -26.17 | -26.177 | -26.15 | -26.183 |  |  | -26.2539 | -26.182 | -26.159 | -26.17 |  |  | -26.1806 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.09 | 0.038 | 0.04 | 0.05 |  |  | 0.010613 | 0.0036129 | 0.030613 | -0.0023871 |  |  | -0.07329 | -0.0013871 | 0.021613 | 0.010613
31.6228 | -26.18 | -26.182 | -26.16 | -26.188 | -26.186 | -26.207 | -26.2587 | -26.186 | -26.165 | -26.17 | -26.249 | -26.181 | -26.1927 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.09 | 0.038 | 0.04 | 0.04 | 0.05 | 0.0581 | 0.012721 | 0.010721 | 0.032721 | 0.0047212 | 0.0067212 | -0.014279 | -0.065933 | 0.0067212 | 0.027721 | 0.022721 | -0.056279 | 0.011721
33.4965 | -26.18 | -26.19 | -26.17 | -26.194 |  |  | -26.2635 | -26.193 | -26.17 | -26.18 |  |  | -26.1926 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | 0.012559 | 0.0025592 | 0.022559 | -0.0014408 |  |  | -0.070914 | -0.00044081 | 0.022559 | 0.012559
35.4813 | -26.19 | -26.192 | -26.17 | -26.199 |  |  | -26.2679 | -26.199 | -26.177 | -26.18 |  |  | -26.1969 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | 0.0068673 | 0.0048673 | 0.026867 | -0.0021327 |  |  | -0.071071 | -0.0021327 | 0.019867 | 0.016867
37.5837 | -26.2 | -26.2 | -26.18 | -26.206 |  |  | -26.2725 | -26.204 | -26.182 | -26.19 |  |  | -26.2043 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | 0.0043127 | 0.0043127 | 0.024313 | -0.0016873 |  |  | -0.068189 | 0.0003127 | 0.022313 | 0.014313
39.8107 | -26.2 | -26.202 | -26.18 | -26.211 | -26.206 | -26.228 | -26.2768 | -26.209 | -26.188 | -26.2 | -26.27 | -26.204 | -26.2146 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.09 | 0.038 | 0.04 | 0.04 | 0.05 | 0.0577 | 0.014567 | 0.012567 | 0.034567 | 0.0035674 | 0.0085674 | -0.013433 | -0.062241 | 0.0055674 | 0.026567 | 0.014567 | -0.055433 | 0.010567
42.1697 | -26.21 | -26.206 | -26.19 | -26.214 |  |  | -26.281 | -26.214 | -26.191 | -26.2 |  |  | -26.2132 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | 0.0032441 | 0.0072441 | 0.023244 | -0.00075586 |  |  | -0.067709 | -0.00075586 | 0.022244 | 0.013244
44.6684 | -26.21 | -26.21 | -26.19 | -26.22 |  |  | -26.2851 | -26.219 | -26.198 | -26.21 |  |  | -26.2178 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | 0.0077586 | 0.0077586 | 0.027759 | -0.0022414 |  |  | -0.06731 | -0.0012414 | 0.019759 | 0.0077586
47.3151 | -26.22 | -26.22 | -26.2 | -26.224 |  |  | -26.289 | -26.223 | -26.203 | -26.21 |  |  | -26.2236 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | 0.0036218 | 0.0036218 | 0.023622 | -0.00037823 |  |  | -0.065352 | 0.00062177 | 0.020622 | 0.013622
50.1187 | -26.22 | -26.225 | -26.2 | -26.226 | -26.223 | -26.245 | -26.2929 | -26.227 | -26.21 | -26.22 | -26.289 | -26.226 | -26.2337 | 0.03 | 0.03 | 0.03 | 0.047 | 0.04 | 0.05 | 0.09 | 0.038 | 0.04 | 0.04 | 0.05 | 0.0567 | 0.013655 | 0.0086549 | 0.033655 | 0.0076549 | 0.010655 | -0.011345 | -0.059204 | 0.0066549 | 0.023655 | 0.013655 | -0.055345 | 0.0076549
53.0884 | -26.23 | -26.228 | -26.21 | -26.227 |  |  | -26.2966 | -26.231 | -26.213 | -26.22 |  |  | -26.232 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | 0.0019528 | 0.0039528 | 0.021953 | 0.0049528 |  |  | -0.06467 | 0.00095282 | 0.018953 | 0.011953
56.2341 | -26.23 | -26.229 | -26.21 | -26.237 |  |  | -26.3002 | -26.237 | -26.218 | -26.23 |  |  | -26.2364 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | 0.0064047 | 0.0074047 | 0.026405 | -0.00059529 |  |  | -0.063833 | -0.00059529 | 0.018405 | 0.0064047
59.5662 | -26.24 | -26.231 | -26.22 | -26.241 |  |  | -26.3038 | -26.24 | -26.223 | -26.23 |  |  | -26.2411 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.09 | 0.038 | 0.04 | 0.04 |  |  | 0.0010959 | 0.010096 | 0.021096 | 9.5861e-05 |  |  | -0.062671 | 0.0010959 | 0.018096 | 0.011096
63.0957 | -26.24 | -26.238 | -26.22 | -26.245 | -26.239 | -26.257 | -26.3072 | -26.243 | -26.228 | -26.23 | -26.305 | -26.241 | -26.2494 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.038 | 0.04 | 0.04 | 0.03 | 0.0538 | 0.0094345 | 0.011435 | 0.029435 | 0.0044345 | 0.010435 | -0.0075655 | -0.05778 | 0.0064345 | 0.021435 | 0.019435 | -0.055565 | 0.0084345
66.8344 | -26.24 | -26.241 | -26.23 | -26.249 |  |  | -26.3106 | -26.247 | -26.231 | -26.24 |  |  | -26.2486 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | 0.0085693 | 0.0075693 | 0.018569 | -0.00043069 |  |  | -0.061985 | 0.0015693 | 0.017569 | 0.0085693
70.7946 | -26.25 | -26.244 | -26.23 | -26.253 |  |  | -26.3139 | -26.252 | -26.236 | -26.24 |  |  | -26.2524 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | 0.00236 | 0.00836 | 0.02236 | -0.00063996 |  |  | -0.06152 | 0.00036004 | 0.01636 | 0.01236
74.9894 | -26.25 | -26.25 | -26.24 | -26.257 |  |  | -26.3171 | -26.255 | -26.24 | -26.24 |  |  | -26.2561 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | 0.0061403 | 0.0061403 | 0.01614 | -0.00085973 |  |  | -0.060982 | 0.0011403 | 0.01614 | 0.01614
79.4328 | -26.26 | -26.252 | -26.24 | -26.26 | -26.252 | -26.274 | -26.3201 | -26.258 | -26.244 | -26.25 | -26.32 | -26.256 | -26.2655 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.038 | 0.04 | 0.04 | 0.03 | 0.0531 | 0.005509 | 0.013509 | 0.025509 | 0.005509 | 0.013509 | -0.008491 | -0.054599 | 0.007509 | 0.021509 | 0.015509 | -0.054491 | 0.009509
84.1395 | -26.26 | -26.254 | -26.24 | -26.263 |  |  | -26.3233 | -26.261 | -26.247 | -26.25 |  |  | -26.2623 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | 0.0022862 | 0.0082862 | 0.022286 | -0.00071377 |  |  | -0.061004 | 0.0012862 | 0.015286 | 0.012286
89.1251 | -26.26 | -26.261 | -26.25 | -26.267 |  |  | -26.326 | -26.266 | -26.25 | -26.26 |  |  | -26.2675 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | 0.0075038 | 0.0065038 | 0.017504 | 0.00050376 |  |  | -0.058526 | 0.0015038 | 0.017504 | 0.0075038
94.4061 | -26.27 | -26.264 | -26.25 | -26.27 |  |  | -26.3288 | -26.269 | -26.256 | -26.26 |  |  | -26.271 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | 0.00097375 | 0.0069738 | 0.020974 | 0.00097375 |  |  | -0.057816 | 0.0019738 | 0.014974 | 0.010974
100 | -26.27 | -26.266 | -26.26 | -26.274 | -26.264 | -26.287 | -26.3317 | -26.271 | -26.259 | -26.26 | -26.333 | -26.27 | -26.2788 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.038 | 0.04 | 0.04 | 0.03 | 0.0531 | 0.0088097 | 0.01281 | 0.01881 | 0.0048097 | 0.01481 | -0.0081903 | -0.052907 | 0.0078097 | 0.01981 | 0.01881 | -0.05419 | 0.0088097
105.9254 | -26.28 | -26.27 | -26.26 | -26.277 |  |  | -26.3342 | -26.274 | -26.262 | -26.27 |  |  | -26.2784 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | -0.0015939 | 0.0084061 | 0.018406 | 0.0014061 |  |  | -0.055843 | 0.0044061 | 0.016406 | 0.0084061
112.2018 | -26.28 | -26.273 | -26.26 | -26.279 |  |  | -26.337 | -26.278 | -26.265 | -26.27 |  |  | -26.2803 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | 0.00025436 | 0.0072544 | 0.020254 | 0.0012544 |  |  | -0.056781 | 0.0022544 | 0.015254 | 0.010254
118.8502 | -26.28 | -26.277 | -26.27 | -26.283 |  |  | -26.3395 | -26.28 | -26.268 | -26.27 |  |  | -26.2834 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.04 |  |  | 0.0034341 | 0.0064341 | 0.013434 | 0.00043414 |  |  | -0.056039 | 0.0034341 | 0.015434 | 0.013434
125.8925 | -26.28 | -26.276 | -26.27 | -26.285 | -26.274 | -26.292 | -26.3419 | -26.281 | -26.272 | -26.27 | -26.345 | -26.283 | -26.2892 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.038 | 0.04 | 0.03 | 0.03 | 0.0515 | 0.0091621 | 0.013162 | 0.019162 | 0.0041621 | 0.015162 | -0.0028379 | -0.052783 | 0.0081621 | 0.017162 | 0.019162 | -0.055838 | 0.0061621
133.3521 | -26.29 | -26.278 | -26.27 | -26.286 |  |  | -26.3443 | -26.284 | -26.274 | -26.28 |  |  | -26.2883 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.03 |  |  | -0.0017104 | 0.01029 | 0.01829 | 0.0022896 |  |  | -0.056027 | 0.0042896 | 0.01429 | 0.0082896
141.2538 | -26.29 | -26.284 | -26.27 | -26.29 |  |  | -26.3465 | -26.288 | -26.277 | -26.28 |  |  | -26.2907 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.038 | 0.04 | 0.03 |  |  | 0.00068222 | 0.0066822 | 0.020682 | 0.00068222 |  |  | -0.055776 | 0.0026822 | 0.013682 | 0.010682
149.6236 | -26.3 | -26.286 | -26.28 | -26.302 |  |  | -26.3487 | -26.29 | -26.281 | -26.28 |  |  | -26.296 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.038 | 0.04 | 0.03 |  |  | -0.0040405 | 0.0099595 | 0.01596 | -0.0060405 |  |  | -0.052717 | 0.0059595 | 0.01496 | 0.01596
158.4893 | -26.3 | -26.288 | -26.28 | -26.295 | -26.283 | -26.304 | -26.3506 | -26.292 | -26.283 | -26.29 | -26.355 | -26.293 | -26.3011 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.038 | 0.04 | 0.03 | 0.03 | 0.0513 | 0.0011371 | 0.013137 | 0.021137 | 0.0061371 | 0.018137 | -0.0028629 | -0.049508 | 0.0091371 | 0.018137 | 0.011137 | -0.053863 | 0.0081371
167.8804 | -26.3 | -26.292 | -26.28 | -26.297 |  |  | -26.3527 | -26.294 | -26.287 | -26.29 |  |  | -26.2991 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.03 |  |  | -0.00091482 | 0.0070852 | 0.019085 | 0.0020852 |  |  | -0.053596 | 0.0050852 | 0.012085 | 0.0090852
177.8279 | -26.3 | -26.293 | -26.28 | -26.299 |  |  | -26.3547 | -26.297 | -26.289 | -26.29 |  |  | -26.3003 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | 0.0003356 | 0.0073356 | 0.020336 | 0.0013356 |  |  | -0.054349 | 0.0033356 | 0.011336 | 0.010336
188.3649 | -26.3 | -26.293 | -26.29 | -26.301 |  |  | -26.3565 | -26.298 | -26.292 | -26.29 |  |  | -26.3026 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.03 |  |  | 0.00256 | 0.00956 | 0.01256 | 0.00156 |  |  | -0.05392 | 0.00456 | 0.01056 | 0.01256
199.5262 | -26.31 | -26.296 | -26.29 | -26.304 | -26.292 | -26.312 | -26.3501 | -26.3 | -26.294 | -26.29 | -26.364 | -26.302 | -26.3087 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0514 | -0.001328 | 0.012672 | 0.018672 | 0.004672 | 0.016672 | -0.003328 | -0.041392 | 0.008672 | 0.014672 | 0.018672 | -0.055328 | 0.006672
211.3489 | -26.31 | -26.297 | -26.29 | -26.306 |  |  | -26.3516 | -26.301 | -26.296 | -26.3 |  |  | -26.3065 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.03 |  |  | -0.003547 | 0.009453 | 0.016453 | 0.00045302 |  |  | -0.045171 | 0.005453 | 0.010453 | 0.006453
223.8721 | -26.31 | -26.297 | -26.29 | -26.308 |  |  | -26.3532 | -26.303 | -26.297 | -26.3 |  |  | -26.3073 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.03 |  |  | -0.0027288 | 0.010271 | 0.017271 | -0.00072878 |  |  | -0.045899 | 0.0042712 | 0.010271 | 0.0072712
237.1374 | -26.31 | -26.298 | -26.3 | -26.309 |  |  | -26.3589 | -26.305 | -26.3 | -26.3 |  |  | -26.3101 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.038 | 0.04 | 0.03 |  |  | 0.00011521 | 0.012115 | 0.010115 | 0.0011152 |  |  | -0.048806 | 0.0051152 | 0.010115 | 0.010115
251.1886 | -26.31 | -26.302 | -26.3 | -26.308 | -26.298 | -26.316 | -26.3602 | -26.298 | -26.301 | -26.3 | -26.371 | -26.309 | -26.3144 | 0.03 | 0.03 | 0.03 | 0.044 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0507 | 0.0044366 | 0.012437 | 0.014437 | 0.0064366 | 0.016437 | -0.0015634 | -0.045802 | 0.016437 | 0.013437 | 0.014437 | -0.056563 | 0.0054366
266.0725 | -26.31 | -26.308 | -26.3 | -26.312 |  |  | -26.3615 | -26.3 | -26.303 | -26.3 |  |  | -26.3118 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | 0.0018095 | 0.0038095 | 0.011809 | -0.00019055 |  |  | -0.049666 | 0.011809 | 0.0088095 | 0.011809
281.8383 | -26.32 | -26.308 | -26.3 | -26.313 |  |  | -26.3626 | -26.302 | -26.305 | -26.3 |  |  | -26.3138 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.006175 | 0.005825 | 0.013825 | 0.00082504 |  |  | -0.048775 | 0.011825 | 0.008825 | 0.013825
298.5383 | -26.32 | -26.307 | -26.3 | -26.315 |  |  | -26.3636 | -26.303 | -26.305 | -26.3 |  |  | -26.3142 | 0.03 | 0.03 | 0.03 | 0.043 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.0058002 | 0.0071998 | 0.0142 | -0.00080019 |  |  | -0.049399 | 0.0112 | 0.0091998 | 0.0142
316.2278 | -26.32 | -26.309 | -26.3 | -26.316 | -26.303 | -26.324 | -26.3644 | -26.305 | -26.308 | -26.31 | -26.376 | -26.315 | -26.3209 | 0.03 | 0.03 | 0.03 | 0.043 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0498 | 0.00086865 | 0.011869 | 0.020869 | 0.0048686 | 0.017869 | -0.0031314 | -0.043555 | 0.015869 | 0.012869 | 0.010869 | -0.055131 | 0.0058686
334.9654 | -26.32 | -26.312 | -26.3 | -26.317 |  |  | -26.3651 | -26.306 | -26.309 | -26.31 |  |  | -26.3174 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.0026063 | 0.0053937 | 0.017394 | 0.00039373 |  |  | -0.047756 | 0.011394 | 0.0083937 | 0.0073937
354.8134 | -26.32 | -26.308 | -26.31 | -26.317 |  |  | -26.3657 | -26.307 | -26.31 | -26.31 |  |  | -26.3185 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.001538 | 0.010462 | 0.008462 | 0.001462 |  |  | -0.047234 | 0.011462 | 0.008462 | 0.008462
375.8374 | -26.32 | -26.309 | -26.31 | -26.316 |  |  | -26.3661 | -26.308 | -26.311 | -26.31 |  |  | -26.3188 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.0012316 | 0.0097684 | 0.0087684 | 0.0027684 |  |  | -0.047379 | 0.010768 | 0.0077684 | 0.0087684
398.1072 | -26.32 | -26.312 | -26.31 | -26.32 | -26.304 | -26.326 | -26.3664 | -26.309 | -26.312 | -26.31 | -26.38 | -26.318 | -26.3239 | 0.03 | 0.03 | 0.03 | 0.044 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0498 | 0.0039473 | 0.011947 | 0.013947 | 0.0039473 | 0.019947 | -0.0020527 | -0.04242 | 0.014947 | 0.011947 | 0.013947 | -0.056053 | 0.0059473
421.6965 | -26.32 | -26.31 | -26.31 | -26.319 |  |  | -26.3665 | -26.309 | -26.312 | -26.31 |  |  | -26.3196 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.00043851 | 0.0095615 | 0.0095615 | 0.00056149 |  |  | -0.04693 | 0.010561 | 0.0075615 | 0.0095615
446.6836 | -26.32 | -26.312 | -26.31 | -26.32 |  |  | -26.3663 | -26.31 | -26.312 | -26.31 |  |  | -26.32 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | 3.5666e-05 | 0.0080357 | 0.010036 | 3.5666e-05 |  |  | -0.04625 | 0.010036 | 0.0080357 | 0.010036
473.1513 | -26.32 | -26.312 | -26.31 | -26.32 |  |  | -26.366 | -26.31 | -26.312 | -26.31 |  |  | -26.32 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -4.3888e-06 | 0.0079956 | 0.0099956 | -4.3888e-06 |  |  | -0.045969 | 0.0099956 | 0.0079956 | 0.0099956
501.1872 | -26.32 | -26.312 | -26.31 | -26.32 | -26.305 | -26.322 | -26.3653 | -26.31 | -26.313 | -26.31 | -26.379 | -26.318 | -26.3237 | 0.03 | 0.03 | 0.03 | 0.044 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.05 | 0.0036911 | 0.011691 | 0.013691 | 0.0036911 | 0.018691 | 0.0016911 | -0.041602 | 0.013691 | 0.010691 | 0.013691 | -0.055309 | 0.0056911
530.8844 | -26.32 | -26.309 | -26.31 | -26.319 |  |  | -26.3645 | -26.31 | -26.312 | -26.31 |  |  | -26.3193 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.00068821 | 0.010312 | 0.0093118 | 0.00031179 |  |  | -0.045183 | 0.0093118 | 0.0073118 | 0.0093118
562.3413 | -26.32 | -26.309 | -26.31 | -26.319 |  |  | -26.3634 | -26.309 | -26.31 | -26.31 |  |  | -26.3188 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.0012006 | 0.0097994 | 0.0087994 | -0.00020062 |  |  | -0.044596 | 0.0097994 | 0.0087994 | 0.0087994
595.6621 | -26.32 | -26.309 | -26.31 | -26.317 |  |  | -26.362 | -26.308 | -26.31 | -26.31 |  |  | -26.3182 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.0017549 | 0.0092451 | 0.0082451 | 0.0012451 |  |  | -0.043716 | 0.010245 | 0.0082451 | 0.0082451
630.9573 | -26.32 | -26.311 | -26.3 | -26.316 | -26.299 | -26.32 | -26.3601 | -26.307 | -26.308 | -26.3 | -26.375 | -26.313 | -26.3191 | 0.03 | 0.03 | 0.03 | 0.044 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0493 | -0.00091126 | 0.0080887 | 0.019089 | 0.0030887 | 0.020089 | -0.00091126 | -0.040976 | 0.012089 | 0.011089 | 0.019089 | -0.055911 | 0.0060887
668.3439 | -26.32 | -26.308 | -26.3 | -26.314 |  |  | -26.3579 | -26.305 | -26.306 | -26.3 |  |  | -26.3139 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.0061324 | 0.0058676 | 0.013868 | -0.00013242 |  |  | -0.044073 | 0.0088676 | 0.0078676 | 0.013868
707.9458 | -26.31 | -26.302 | -26.3 | -26.311 |  |  | -26.3553 | -26.303 | -26.304 | -26.3 |  |  | -26.3107 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | 0.00066309 | 0.0086631 | 0.010663 | -0.00033691 |  |  | -0.044642 | 0.0076631 | 0.0066631 | 0.010663
749.8942 | -26.31 | -26.298 | -26.3 | -26.309 |  |  | -26.3523 | -26.3 | -26.301 | -26.3 |  |  | -26.3088 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.0012172 | 0.010783 | 0.0087828 | -0.00021719 |  |  | -0.04348 | 0.0087828 | 0.0077828 | 0.0087828
794.3282 | -26.31 | -26.293 | -26.29 | -26.306 | -26.288 | -26.308 | -26.349 | -26.297 | -26.298 | -26.29 | -26.365 | -26.303 | -26.3081 | 0.03 | 0.03 | 0.03 | 0.044 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0493 | -0.0019168 | 0.015083 | 0.018083 | 0.0020832 | 0.020083 | 8.3199e-05 | -0.040915 | 0.011083 | 0.010083 | 0.018083 | -0.056917 | 0.0050832
841.3951 | -26.3 | -26.291 | -26.29 | -26.302 |  |  | -26.3439 | -26.294 | -26.294 | -26.29 |  |  | -26.3006 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | 0.00060718 | 0.0096072 | 0.010607 | -0.0013928 |  |  | -0.04325 | 0.0066072 | 0.0066072 | 0.010607
891.2509 | -26.3 | -26.29 | -26.28 | -26.297 |  |  | -26.3392 | -26.29 | -26.289 | -26.29 |  |  | -26.2969 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.0030987 | 0.0069013 | 0.016901 | -9.8686e-05 |  |  | -0.042309 | 0.0069013 | 0.0079013 | 0.0069013
944.0609 | -26.29 | -26.288 | -26.28 | -26.291 |  |  | -26.3335 | -26.285 | -26.283 | -26.28 |  |  | -26.2913 | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | 0.0013183 | 0.0033183 | 0.011318 | 0.00031829 |  |  | -0.042228 | 0.0063183 | 0.0083183 | 0.011318
1000 | -26.29 | -26.277 | -26.27 | -26.286 | -26.269 | -26.286 | -26.3274 | -26.28 | -26.276 | -26.27 | -26.345 | -26.282 | -26.2882 | 0.03 | 0.03 | 0.03 | 0.045 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0499 | -0.0018014 | 0.011199 | 0.018199 | 0.0021986 | 0.019199 | 0.0021986 | -0.039185 | 0.0081986 | 0.012199 | 0.018199 | -0.056801 | 0.0061986
1059.2537 | -26.28 | -26.269 | -26.27 | -26.278 |  |  | -26.32 | -26.274 | -26.27 | -26.27 |  |  | -26.2789 | 0.03 | 0.03 | 0.03 | 0.044 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | -0.0011248 | 0.0098752 | 0.0088752 | 0.0008752 |  |  | -0.041126 | 0.0048752 | 0.0088752 | 0.0088752
1122.0185 | -26.27 | -26.265 | -26.26 | -26.271 |  |  | -26.3118 | -26.267 | -26.263 | -26.26 |  |  | -26.271 | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | 0.00097937 | 0.0059794 | 0.010979 | -2.0629e-05 |  |  | -0.040856 | 0.0039794 | 0.0079794 | 0.010979
1188.5022 | -26.26 | -26.252 | -26.25 | -26.262 |  |  | -26.3024 | -26.258 | -26.254 | -26.25 |  |  | -26.2611 | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | 0.0010506 | 0.0090506 | 0.011051 | -0.00094937 |  |  | -0.041354 | 0.0030506 | 0.0070506 | 0.011051
1258.9254 | -26.25 | -26.243 | -26.24 | -26.252 | -26.233 | -26.25 | -26.2919 | -26.249 | -26.244 | -26.24 | -26.311 | -26.248 | -26.2543 | 0.03 | 0.03 | 0.03 | 0.045 | 0.04 | 0.05 | 0.08 | 0.037 | 0.04 | 0.03 | 0.03 | 0.0488 | 0.0043265 | 0.011327 | 0.014327 | 0.0023265 | 0.021327 | 0.0043265 | -0.037592 | 0.0053265 | 0.010327 | 0.014327 | -0.056673 | 0.0063265
1333.5214 | -26.24 | -26.23 | -26.23 | -26.24 |  |  | -26.2797 | -26.239 | -26.233 | -26.23 |  |  | -26.2402 | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.037 | 0.04 | 0.03 |  |  | 0.00021456 | 0.010215 | 0.010215 | 0.00021456 |  |  | -0.039502 | 0.0012146 | 0.0072146 | 0.010215
1412.5375 | -26.23 | -26.215 | -26.21 | -26.227 |  |  | -26.2668 | -26.229 | -26.22 | -26.22 |  |  | -26.2272 | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -0.0027777 | 0.012222 | 0.017222 | 0.00022229 |  |  | -0.039556 | -0.0017777 | 0.0072223 | 0.0072223
1496.2357 | -26.21 | -26.203 | -26.2 | -26.212 |  |  | -26.2514 | -26.216 | -26.205 | -26.2 |  |  | -26.2122 | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | 0.0021744 | 0.0091744 | 0.012174 | 0.00017436 |  |  | -0.039221 | -0.0038256 | 0.0071744 | 0.012174
1584.8932 | -26.2 | -26.188 | -26.18 | -26.195 | -26.178 | -26.193 | -26.2345 | -26.2 | -26.189 | -26.18 | -26.255 | -26.191 | -26.1986 | 0.03 | 0.03 | 0.03 | 0.045 | 0.04 | 0.05 | 0.08 | 0.041 | 0.04 | 0.03 | 0.03 | 0.0484 | -0.0013761 | 0.010624 | 0.018624 | 0.0036239 | 0.020624 | 0.0056239 | -0.035863 | -0.0013761 | 0.0096239 | 0.018624 | -0.056376 | 0.0076239
1678.804 | -26.18 | -26.17 | -26.16 | -26.176 |  |  | -26.2154 | -26.184 | -26.17 | -26.17 |  |  | -26.1782 | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -0.0018203 | 0.0081797 | 0.01818 | 0.0021797 |  |  | -0.037258 | -0.0058203 | 0.0081797 | 0.0081797
1778.2794 | -26.16 | -26.149 | -26.14 | -26.155 |  |  | -26.1941 | -26.165 | -26.15 | -26.14 |  |  | -26.1566 | 0.03 | 0.03 | 0.03 | 0.045 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -0.003367 | 0.007633 | 0.016633 | 0.001633 |  |  | -0.037431 | -0.008367 | 0.006633 | 0.016633
1883.6491 | -26.13 | -26.125 | -26.12 | -26.132 |  |  | -26.1702 | -26.142 | -26.126 | -26.12 |  |  | -26.1331 | 0.03 | 0.03 | 0.03 | 0.046 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | 0.0031459 | 0.0081459 | 0.013146 | 0.0011459 |  |  | -0.037021 | -0.0088541 | 0.0071459 | 0.013146
1995.2623 | -26.11 | -26.1 | -26.09 | -26.105 | -26.09 | -26.102 | -26.1455 | -26.119 | -26.102 | -26.1 | -26.166 | -26.101 | -26.1109 | 0.03 | 0.03 | 0.03 | 0.046 | 0.04 | 0.05 | 0.08 | 0.041 | 0.04 | 0.03 | 0.03 | 0.0503 | 0.00087136 | 0.010871 | 0.020871 | 0.0058714 | 0.020871 | 0.0088714 | -0.034585 | -0.0081286 | 0.0088714 | 0.010871 | -0.055129 | 0.0098714
2113.489 | -26.08 | -26.074 | -26.07 | -26.076 |  |  | -26.1159 | -26.093 | -26.072 | -26.07 |  |  | -26.0814 | 0.03 | 0.035 | 0.03 | 0.046 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | 0.0013666 | 0.0073666 | 0.011367 | 0.0053666 |  |  | -0.034566 | -0.011633 | 0.0093666 | 0.011367
2238.7211 | -26.05 | -26.04 | -26.03 | -26.043 |  |  | -26.0831 | -26.064 | -26.04 | -26.04 |  |  | -26.0488 | 0.03 | 0.035 | 0.03 | 0.046 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | -0.0012327 | 0.0087673 | 0.018767 | 0.0057673 |  |  | -0.034371 | -0.015233 | 0.0087673 | 0.0087673
2371.3737 | -26.01 | -26.006 | -26 | -26.007 |  |  | -26.0467 | -26.031 | -26.004 | -26 |  |  | -26.0131 | 0.03 | 0.035 | 0.03 | 0.047 |  |  | 0.08 | 0.041 | 0.04 | 0.03 |  |  | 0.0030824 | 0.0070824 | 0.013082 | 0.0060824 |  |  | -0.033577 | -0.017918 | 0.0090824 | 0.013082
2511.8864 | -25.97 | -25.962 | -25.96 | -25.967 | -25.953 | -25.966 | -26.0063 | -25.994 | -25.965 | -25.96 | -26.029 | -25.966 | -25.9749 | 0.03 | 0.035 | 0.03 | 0.047 | 0.04 | 0.06 | 0.08 | 0.041 | 0.04 | 0.03 | 0.03 | 0.0489 | 0.0048625 | 0.012862 | 0.014862 | 0.0078625 | 0.021862 | 0.0088625 | -0.031487 | -0.019138 | 0.0098625 | 0.014862 | -0.054138 | 0.0088625
2660.7251 | -25.93 | -25.918 | -25.91 | -25.923 |  |  | -25.962 | -25.953 | -25.922 | -25.92 |  |  | -25.9298 | 0.04 | 0.04 | 0.03 | 0.048 |  |  | 0.08 | 0.042 | 0.04 | 0.03 |  |  | -0.00024739 | 0.011753 | 0.019753 | 0.0067526 |  |  | -0.032268 | -0.023247 | 0.0077526 | 0.0097526
2818.3829 | -25.88 | -25.874 | -25.87 | -25.874 |  |  | -25.9134 | -25.908 | -25.874 | -25.87 |  |  | -25.8829 | 0.04 | 0.04 | 0.03 | 0.049 |  |  | 0.08 | 0.042 | 0.04 | 0.03 |  |  | 0.0029292 | 0.0089292 | 0.012929 | 0.0089292 |  |  | -0.030504 | -0.025071 | 0.0089292 | 0.012929
2985.3826 | -25.83 | -25.82 | -25.81 | -25.822 |  |  | -25.8601 | -25.858 | -25.821 | -25.82 |  |  | -25.8301 | 0.04 | 0.04 | 0.03 | 0.05 |  |  | 0.08 | 0.042 | 0.04 | 0.03 |  |  | 0.00014336 | 0.010143 | 0.020143 | 0.0081434 |  |  | -0.030003 | -0.027857 | 0.0091434 | 0.010143
3162.2777 | -25.77 | -25.766 | -25.76 | -25.763 | -25.754 | -25.762 | -25.8028 | -25.805 | -25.766 | -25.76 | -25.828 | -25.765 | -25.7751 | 0.04 | 0.04 | 0.03 | 0.052 | 0.04 | 0.06 | 0.08 | 0.043 | 0.04 | 0.03 | 0.03 | 0.0505 | 0.0051482 | 0.0091482 | 0.015148 | 0.012148 | 0.021148 | 0.013148 | -0.027631 | -0.029852 | 0.0091482 | 0.015148 | -0.052852 | 0.010148
3349.6544 | -25.71 | -25.704 | -25.7 | -25.703 |  |  | -25.7406 | -25.746 | -25.704 | -25.7 |  |  | -25.7134 | 0.04 | 0.04 | 0.03 | 0.054 |  |  | 0.08 | 0.044 | 0.05 | 0.03 |  |  | 0.0034454 | 0.0094454 | 0.013445 | 0.010445 |  |  | -0.027118 | -0.032555 | 0.0094454 | 0.013445
3548.1339 | -25.64 | -25.636 | -25.63 | -25.637 |  |  | -25.6742 | -25.689 | -25.641 | -25.64 |  |  | -25.6484 | 0.04 | 0.04 | 0.03 | 0.056 |  |  | 0.08 | 0.044 | 0.05 | 0.03 |  |  | 0.0083986 | 0.012399 | 0.018399 | 0.011399 |  |  | -0.02579 | -0.040601 | 0.0073986 | 0.0083986
3758.374 | -25.57 | -25.571 | -25.57 | -25.568 |  |  | -25.6046 | -25.624 | -25.573 | -25.58 |  |  | -25.5826 | 0.04 | 0.04 | 0.03 | 0.059 |  |  | 0.08 | 0.045 | 0.05 | 0.03 |  |  | 0.01257 | 0.01157 | 0.01257 | 0.01457 |  |  | -0.021991 | -0.04143 | 0.0095702 | 0.0025702
3981.0717 | -25.5 | -25.502 | -25.5 | -25.497 | -25.494 | -25.495 | -25.5325 | -25.556 | -25.503 | -25.51 | -25.562 | -25.497 | -25.5124 | 0.04 | 0.04 | 0.03 | 0.063 | 0.04 | 0.08 | 0.09 | 0.046 | 0.05 | 0.03 | 0.05 | 0.0705 | 0.012371 | 0.010371 | 0.012371 | 0.015371 | 0.018371 | 0.017371 | -0.020079 | -0.043629 | 0.0093709 | 0.0023709 | -0.049629 | 0.015371
4216.965 | -25.43 | -25.433 | -25.43 | -25.426 |  |  | -25.4601 | -25.487 | -25.432 | -25.44 |  |  | -25.4423 | 0.05 | 0.05 | 0.03 | 0.067 |  |  | 0.09 | 0.048 | 0.05 | 0.03 |  |  | 0.01226 | 0.00926 | 0.01226 | 0.01626 |  |  | -0.01782 | -0.04474 | 0.01026 | 0.00226
4466.8359 | -25.37 | -25.366 | -25.36 | -25.357 |  |  | -25.3903 | -25.42 | -25.367 | -25.38 |  |  | -25.3763 | 0.05 | 0.05 | 0.03 | 0.073 |  |  | 0.09 | 0.049 | 0.06 | 0.03 |  |  | 0.0062885 | 0.010289 | 0.016289 | 0.019289 |  |  | -0.01402 | -0.043711 | 0.0092885 | -0.0037115
4731.5126 | -25.31 | -25.307 | -25.3 | -25.295 |  |  | -25.3269 | -25.357 | -25.307 | -25.32 |  |  | -25.3154 | 0.05 | 0.05 | 0.03 | 0.079 |  |  | 0.09 | 0.051 | 0.06 | 0.03 |  |  | 0.0053672 | 0.0083672 | 0.015367 | 0.020367 |  |  | -0.011571 | -0.041633 | 0.0083672 | -0.0046328
5011.8723 | -25.26 | -25.261 | -25.24 | -25.248 | -25.255 | -25.259 | -25.2761 | -25.307 | -25.262 | -25.27 | -25.311 | -25.239 | -25.2657 | 0.05 | 0.05 | 0.03 | 0.085 | 0.04 | 0.1 | 0.09 | 0.053 | 0.07 | 0.05 | 0.05 | 0.0828 | 0.0056768 | 0.0046768 | 0.025677 | 0.017677 | 0.010677 | 0.0066768 | -0.010445 | -0.041323 | 0.0036768 | -0.0043232 | -0.045323 | 0.026677
5308.8444 | -25.23 | -25.238 | -25.22 | -25.22 |  |  | -25.2468 | -25.275 | -25.241 | -25.25 |  |  | -25.2401 | 0.05 | 0.06 | 0.03 | 0.091 |  |  | 0.1 | 0.055 | 0.07 | 0.05 |  |  | 0.010095 | 0.0020945 | 0.020095 | 0.020095 |  |  | -0.0066617 | -0.034905 | -0.00090547 | -0.0099055
5623.4133 | -25.24 | -25.248 | -25.23 | -25.226 |  |  | -25.2489 | -25.275 | -25.248 | -25.26 |  |  | -25.247 | 0.05 | 0.06 | 0.03 | 0.096 |  |  | 0.1 | 0.055 | 0.08 | 0.05 |  |  | 0.0069881 | -0.0010119 | 0.016988 | 0.020988 |  |  | -0.001917 | -0.028012 | -0.0010119 | -0.013012
5956.6214 | -25.29 | -25.303 | -25.29 | -25.272 |  |  | -25.297 | -25.317 | -25.304 | -25.31 |  |  | -25.2979 | 0.05 | 0.06 | 0.03 | 0.099 |  |  | 0.1 | 0.056 | 0.08 | 0.05 |  |  | 0.0078711 | -0.0051289 | 0.0078711 | 0.025871 |  |  | 0.00090204 | -0.019129 | -0.0061289 | -0.012129
6309.5734 | -25.41 | -25.42 | -25.4 | -25.387 | -25.411 | -25.397 | -25.4094 | -25.418 | -25.42 | -25.43 | -25.443 | -25.394 | -25.4116 | 0.05 | 0.06 | 0.03 | 0.1 | 0.05 | 0.1 | 0.11 | 0.057 | 0.08 | 0.05 | 0.05 | 0.0933 | 0.0016193 | -0.0083807 | 0.011619 | 0.024619 | 0.00061929 | 0.014619 | 0.0021878 | -0.0063807 | -0.0083807 | -0.018381 | -0.031381 | 0.017619
6683.4392 | -25.6 | -25.614 | -25.6 | -25.585 |  |  | -25.6027 | -25.6 | -25.617 | -25.63 |  |  | -25.6061 | 0.06 | 0.07 | 0.03 | 0.099 |  |  | 0.11 | 0.057 | 0.08 | 0.05 |  |  | 0.0060852 | -0.0079148 | 0.0060852 | 0.021085 |  |  | 0.0034037 | 0.0060852 | -0.010915 | -0.023915
7079.4578 | -25.9 | -25.907 | -25.9 | -25.879 |  |  | -25.8965 | -25.884 | -25.917 | -25.93 |  |  | -25.9017 | 0.06 | 0.07 | 0.03 | 0.093 |  |  | 0.11 | 0.057 | 0.09 | 0.05 |  |  | 0.0016889 | -0.0053111 | 0.0016889 | 0.022689 |  |  | 0.0051778 | 0.017689 | -0.015311 | -0.028311
7498.9421 | -26.32 | -26.323 | -26.31 | -26.291 |  |  | -26.314 | -26.287 | -26.34 | -26.34 |  |  | -26.3156 | 0.06 | 0.07 | 0.05 | 0.085 |  |  | 0.11 | 0.057 | 0.09 | 0.05 |  |  | -0.0043792 | -0.0073792 | 0.0056208 | 0.024621 |  |  | 0.0016546 | 0.028621 | -0.024379 | -0.024379
7943.2823 | -26.86 | -26.878 | -26.87 | -26.842 | -26.882 | -26.863 | -26.8644 | -26.835 | -26.899 | -26.89 | -26.899 | -26.857 | -26.8699 | 0.06 | 0.07 | 0.07 | 0.081 | 0.05 | 0.12 | 0.14 | 0.057 | 0.1 | 0.05 | 0.07 | 0.1003 | 0.0099467 | -0.0080533 | -5.3266e-05 | 0.027947 | -0.012053 | 0.0069467 | 0.0055859 | 0.034947 | -0.029053 | -0.020053 | -0.029053 | 0.012947
8413.9514 | -27.55 | -27.571 | -27.55 | -27.532 |  |  | -27.5559 | -27.515 | -27.595 | -27.58 |  |  | -27.5561 | 0.11 | 0.08 | 0.1 | 0.078 |  |  | 0.14 | 0.056 | 0.12 | 0.08 |  |  | 0.0061063 | -0.014894 | 0.0061063 | 0.024106 |  |  | 0.00025595 | 0.041106 | -0.038894 | -0.023894
8912.5094 | -28.38 | -28.41 | -28.4 | -28.361 |  |  | -28.3918 | -28.352 | -28.427 | -28.4 |  |  | -28.3902 | 0.11 | 0.09 | 0.1 | 0.081 |  |  | 0.16 | 0.055 | 0.15 | 0.08 |  |  | 0.010224 | -0.019776 | -0.0097756 | 0.029224 |  |  | -0.001571 | 0.038224 | -0.036776 | -0.0097756
9440.6088 | -29.32 | -29.362 | -29.35 | -29.314 |  |  | -29.3501 |  | -29.413 | -29.33 |  |  | -29.3484 | Expanded Uncertainties (dB) UK: 0.11; Expanded Uncertainties (dB) FR: 0.09; Expanded Uncertainties (dB) IT: 0.1; Expanded Uncertainties (dB) CH: 0.09; Expanded Uncertainties (dB) RS: 0.18; Expanded Uncertainties (dB) SE: 0.17; Expanded Uncertainties (dB) DE: 0.08; Level difference relative to estimated sensitivity level (dB) UK: 0.028441; Level difference relative to estimated sensitivity level (dB) FR: -0.013559; Level difference relative to estimated sensitivity level (dB) IT: -0.0015593; Level difference relative to estimated sensitivity level (dB) CH: 0.034441; Level difference relative to estimated sensitivity level (dB) RS: -0.0016442; Level difference relative to estimated sensitivity level (dB) SE: -0.064559; Level difference relative to estimated sensitivity level (dB) DE: 0.018441
10000 | -30.36 | -30.416 | -30.42 | -30.367 | -30.457 | -30.449 | -30.4078 | -30.364 | -30.509 | -30.36 | -30.468 | -30.416 | -30.4161 | 0.11 | 0.1 | 0.16 | 0.107 | 0.06 | 0.14 | 0.2 | 0.108 | 0.2 | 0.08 | 0.08 | 0.1539 | 0.056148 | 0.00014815 | -0.0038519 | 0.049148 | -0.040852 | -0.032852 | 0.0083704 | 0.052148 | -0.092852 | 0.056148 | -0.051852 | 0.00014815

## 811014 Phase
| Sensitivity Phase (4160 811014) degrees | [12 empty] | Uncertainties (degrees) | [11 empty] | Phase difference relative to estimated sensitivity phase (degrees)
Frequency (Hz) | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG | unweighted mean | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG | UK | FR | IT | CH | TR | BG | RS | AT | SE | DE | ES | EG
1.9953 | 175.5 | 175.29 |  | 175.072 | [5 empty] | 175.1 |  |  | 175.2405 | Uncertainties (degrees) UK: 3.3; Uncertainties (degrees) FR: 3.1; Uncertainties (degrees) CH: 0.447; Uncertainties (degrees) DE: 4.5; Phase difference relative to estimated sensitivity phase (degrees) UK: 0.2595; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0495; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.1685; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.1405
2.5119 | 175.79 | 175.64 |  | 175.267 | [5 empty] | 175.3 |  |  | 175.4992 | Uncertainties (degrees) UK: 3.3; Uncertainties (degrees) FR: 2.7; Uncertainties (degrees) CH: 0.371; Uncertainties (degrees) DE: 4.5; Phase difference relative to estimated sensitivity phase (degrees) UK: 0.29075; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.14075; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.23225; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.19925
3.1623 | 175.91 | 175.8 |  | 175.536 | [5 empty] | 175.8 |  |  | 175.7615 | Uncertainties (degrees) UK: 3.3; Uncertainties (degrees) FR: 2.4; Uncertainties (degrees) CH: 0.29; Uncertainties (degrees) DE: 4.5; Phase difference relative to estimated sensitivity phase (degrees) UK: 0.1485; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0385; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.2255; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0385
3.9811 | 176.1 | 176.09 |  | 175.846 | [5 empty] | 175.9 |  |  | 175.984 | Uncertainties (degrees) UK: 1.5; Uncertainties (degrees) FR: 2.1; Uncertainties (degrees) CH: 0.232; Uncertainties (degrees) DE: 2.5; Phase difference relative to estimated sensitivity phase (degrees) UK: 0.116; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.106; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.138; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.084
5.0119 | 176.3 | 176.4 |  | 176.192 | [5 empty] | 176.4 |  |  | 176.323 | Uncertainties (degrees) UK: 1.5; Uncertainties (degrees) FR: 1.9; Uncertainties (degrees) CH: 0.189; Uncertainties (degrees) DE: 2.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.023; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.077; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.131; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.077
6.3096 | 176.62 | 176.73 |  | 176.544 | [5 empty] | 176.6 |  |  | 176.6235 | Uncertainties (degrees) UK: 1.5; Uncertainties (degrees) FR: 1.7; Uncertainties (degrees) CH: 0.149; Uncertainties (degrees) DE: 2.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0035; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.1065; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.0795; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0235
7.9433 | 176.91 | 177.05 |  | 176.895 | [5 empty] | 176.9 |  |  | 176.9387 | Uncertainties (degrees) UK: 0.8; Uncertainties (degrees) FR: 1.6; Uncertainties (degrees) CH: 0.119; Uncertainties (degrees) DE: 1.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.02875; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.11125; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.04375; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.03875
10 | 177.18 | 177.37 |  | 177.224 | [5 empty] | 177.2 |  |  | 177.2435 | Uncertainties (degrees) UK: 0.8; Uncertainties (degrees) FR: 1.4; Uncertainties (degrees) CH: 0.097; Uncertainties (degrees) DE: 1.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0635; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.1265; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.0195; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0435
12.5893 | 177.48 | 177.64 |  | 177.528 | [5 empty] | 177.6 |  |  | 177.562 | Uncertainties (degrees) UK: 0.8; Uncertainties (degrees) FR: 1.2; Uncertainties (degrees) CH: 0.077; Uncertainties (degrees) DE: 1.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.082; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.078; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.034; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.038
15.8489 | 177.72 | 177.89 |  | 177.793 | [5 empty] | 177.8 |  |  | 177.8007 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 1.1; Uncertainties (degrees) CH: 0.062; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.08075; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.08925; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.00775; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.00075
19.9526 | 178 | 178.09 | 178.2 | 178.021 |  |  |  |  | 177.73 | 178 |  |  | 178.0068 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 1; Uncertainties (degrees) IT: 1.6; Uncertainties (degrees) CH: 0.05; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0068333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.083167; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.19317; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.014167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.27683; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0068333
21.1349 | 178.04 | 178.14 | 178.2 | 178.071 |  |  |  |  | 177.78 | 178.1 |  |  | 178.0552 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 1; Uncertainties (degrees) IT: 1.6; Uncertainties (degrees) CH: 0.051; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.015167; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.084833; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.14483; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.015833; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.27517; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.044833
22.3872 | 178.08 | 178.19 | 178.3 | 178.122 |  |  |  |  | 177.82 | 178.1 |  |  | 178.102 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 1; Uncertainties (degrees) IT: 1.5; Uncertainties (degrees) CH: 0.045; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.022; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.088; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.198; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.02; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.282; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.002
23.7137 | 178.13 | 178.23 | 178.3 | 178.162 |  |  |  |  | 177.88 | 178.2 |  |  | 178.1503 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 1; Uncertainties (degrees) IT: 1.4; Uncertainties (degrees) CH: 0.044; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.020333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.079667; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.14967; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.011667; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.27033; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.049667
25.1189 | 178.16 | 178.27 | 178.3 | 178.225 |  |  |  |  | 177.96 | 178.2 |  |  | 178.1858 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.9; Uncertainties (degrees) IT: 1.3; Uncertainties (degrees) CH: 0.06; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.025833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.084167; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.11417; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.039167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.22583; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.014167
26.6073 | 178.2 | 178.3 | 178.4 | 178.238 |  |  |  |  | 178.01 | 178.3 |  |  | 178.2413 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.9; Uncertainties (degrees) IT: 1.3; Uncertainties (degrees) CH: 0.032; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.041333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.058667; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.15867; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.0033333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.23133; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.058667
28.1838 | 178.24 | 178.34 | 178.4 | 178.315 |  |  |  |  | 178.03 | 178.3 |  |  | 178.2708 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.9; Uncertainties (degrees) IT: 1.2; Uncertainties (degrees) CH: 0.052; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.030833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.069167; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.12917; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.044167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.24083; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.029167
29.8538 | 178.27 | 178.37 | 178.4 | 178.321 |  |  |  |  | 178.09 | 178.3 |  |  | 178.2918 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.9; Uncertainties (degrees) IT: 1.1; Uncertainties (degrees) CH: 0.033; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.021833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.078167; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.10817; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.029167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.20183; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0081667
31.6228 | 178.3 | 178.4 | 178.4 | 178.358 |  |  |  |  | 178.13 | 178.4 |  |  | 178.3313 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.8; Uncertainties (degrees) IT: 1.1; Uncertainties (degrees) CH: 0.041; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.031333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.068667; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.068667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.026667; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.20133; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.068667
33.4965 | 178.33 | 178.43 | 178.5 | 178.388 |  |  |  |  | 178.17 | 178.4 |  |  | 178.3697 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.8; Uncertainties (degrees) IT: 1; Uncertainties (degrees) CH: 0.03; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.039667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.060333; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.13033; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.018333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.19967; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.030333
35.4813 | 178.37 | 178.46 | 178.5 | 178.413 |  |  |  |  | 178.2 | 178.4 |  |  | 178.3905 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.8; Uncertainties (degrees) IT: 1; Uncertainties (degrees) CH: 0.029; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0205; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0695; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.1095; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0225; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.1905; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0095
37.5837 | 178.39 | 178.48 | 178.5 | 178.44 |  |  |  |  | 178.23 | 178.4 |  |  | 178.4067 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.8; Uncertainties (degrees) IT: 0.9; Uncertainties (degrees) CH: 0.03; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.016667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.073333; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.093333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.033333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.17667; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0066667
39.8107 | 178.41 | 178.5 | 178.5 | 178.463 |  |  |  |  | 178.26 | 178.5 |  |  | 178.4388 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.8; Uncertainties (degrees) IT: 0.9; Uncertainties (degrees) CH: 0.026; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.028833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.061167; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.061167; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.024167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.17883; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.061167
42.1697 | 178.44 | 178.52 | 178.5 | 178.499 |  |  |  |  | 178.28 | 178.5 |  |  | 178.4565 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.8; Uncertainties (degrees) IT: 0.8; Uncertainties (degrees) CH: 0.031; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0165; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0635; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0435; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0425; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.1765; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0435
44.6684 | 178.45 | 178.54 | 178.6 | 178.506 |  |  |  |  | 178.32 | 178.5 |  |  | 178.486 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.8; Uncertainties (degrees) IT: 0.8; Uncertainties (degrees) CH: 0.023; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.036; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.054; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.114; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.02; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.166; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.014
47.3151 | 178.47 | 178.55 | 178.6 | 178.527 |  |  |  |  | 178.33 | 178.5 |  |  | 178.4962 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.8; Uncertainties (degrees) IT: 0.7; Uncertainties (degrees) CH: 0.03; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.026167; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.053833; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.10383; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.030833; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.16617; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0038333
50.1187 | 178.49 | 178.57 | 178.6 | 178.564 |  |  |  |  | 178.36 | 178.5 |  |  | 178.514 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.7; Uncertainties (degrees) IT: 0.7; Uncertainties (degrees) CH: 0.071; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.024; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.056; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.086; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.05; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.154; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.014
53.0884 | 178.49 | 178.58 | 178.6 | 178.595 |  |  |  |  | 178.38 | 178.6 |  |  | 178.5408 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.7; Uncertainties (degrees) IT: 0.7; Uncertainties (degrees) CH: 0.038; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.050833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.039167; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.059167; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.054167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.16083; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.059167
56.2341 | 178.5 | 178.59 | 178.6 | 178.563 |  |  |  |  | 178.39 | 178.6 |  |  | 178.5405 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.7; Uncertainties (degrees) IT: 0.7; Uncertainties (degrees) CH: 0.02; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0405; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0495; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0595; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0225; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.1505; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0595
59.5662 | 178.51 | 178.59 | 178.6 | 178.568 |  |  |  |  | 178.41 | 178.6 |  |  | 178.5463 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.7; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.019; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.036333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.043667; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.053667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.021667; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.13633; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.053667
63.0957 | 178.51 | 178.6 | 178.6 | 178.575 |  |  |  |  | 178.41 | 178.6 |  |  | 178.5492 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.7; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.017; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.039167; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.050833; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.050833; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.025833; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.13917; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.050833
66.8344 | 178.52 | 178.6 | 178.6 | 178.578 |  |  |  |  | 178.42 | 178.6 |  |  | 178.553 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.7; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.018; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.033; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.047; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.047; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.025; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.133; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.047
70.7946 | 178.52 | 178.6 | 178.6 | 178.577 |  |  |  |  | 178.43 | 178.6 |  |  | 178.5545 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.7; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.016; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0345; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0455; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0455; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0225; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.1245; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0455
74.9894 | 178.52 | 178.59 | 178.6 | 178.576 |  |  |  |  | 178.43 | 178.6 |  |  | 178.5527 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.7; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.016; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.032667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.037333; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.047333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.023333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.12267; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.047333
79.4328 | 178.5 | 178.59 | 178.6 | 178.57 |  |  |  |  | 178.43 | 178.6 |  |  | 178.5483 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.6; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.015; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.048333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.041667; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.051667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.021667; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.11833; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.051667
84.1395 | 178.52 | 178.58 | 178.6 | 178.563 |  |  |  |  | 178.42 | 178.6 |  |  | 178.5472 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.6; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.015; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.027167; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.032833; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.052833; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.015833; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.12717; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.052833
89.1251 | 178.5 | 178.56 | 178.5 | 178.549 |  |  |  |  | 178.41 | 178.5 |  |  | 178.5032 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.6; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.014; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0031667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.056833; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0031667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.045833; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.093167; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0031667
94.4061 | 178.5 | 178.55 | 178.5 | 178.535 |  |  |  |  | 178.4 | 178.5 |  |  | 178.4975 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.6; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.013; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: 0.0025; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0525; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0025; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0375; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.0975; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0025
100 | 178.49 | 178.53 | 178.5 | 178.523 |  |  |  |  | 178.41 | 178.5 |  |  | 178.4922 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.015; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0021667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.037833; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0078333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.030833; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.082167; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0078333
105.9254 | 178.46 | 178.51 | 178.5 | 178.501 |  |  |  |  | 178.37 | 178.5 |  |  | 178.4735 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.015; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0135; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0365; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0265; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0275; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.1035; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0265
112.2018 | 178.42 | 178.49 | 178.5 | 178.479 |  |  |  |  | 178.35 | 178.5 |  |  | 178.4565 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.017; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0365; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0335; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0435; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0225; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.1065; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0435
118.8502 | 178.4 | 178.46 | 178.4 | 178.45 |  |  |  |  | 178.33 | 178.5 |  |  | 178.4233 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.022; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.023333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.036667; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.023333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.026667; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.093333; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.076667
125.8925 | 178.37 | 178.43 | 178.4 | 178.425 |  |  |  |  | 178.3 | 178.4 |  |  | 178.3875 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.015; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0175; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0425; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0125; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0375; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.0875; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0125
133.3521 | 178.33 | 178.39 | 178.4 | 178.393 |  |  |  |  | 178.27 | 178.4 |  |  | 178.3638 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.026; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.033833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.026167; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.036167; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.029167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.093833; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.036167
141.2538 | 178.29 | 178.36 | 178.3 | 178.349 |  |  |  |  | 178.23 | 178.4 |  |  | 178.3215 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.032; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0315; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0385; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0215; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0275; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.0915; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0785
149.6236 | 178.25 | 178.31 | 178.3 | 178.336 |  |  |  |  | 178.2 | 178.3 |  |  | 178.2827 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.095; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.032667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.027333; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.017333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.053333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.082667; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.017333
158.4893 | 178.2 | 178.27 | 178.2 | 178.266 |  |  |  |  | 178.15 | 178.3 |  |  | 178.231 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.016; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.031; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.039; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.031; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.035; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.081; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.069
167.8804 | 178.16 | 178.22 | 178.2 | 178.216 |  |  |  |  | 178.1 | 178.2 |  |  | 178.1827 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.017; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.022667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.037333; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.017333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.033333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.082667; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.017333
177.8279 | 178.09 | 178.16 | 178.1 | 178.161 |  |  |  |  | 178.06 | 178.2 |  |  | 178.1285 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.018; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0385; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0315; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0285; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0325; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.0685; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0715
188.3649 | 178.05 | 178.1 | 178.1 | 178.103 |  |  |  |  | 177.99 | 178.1 |  |  | 178.0738 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.019; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.023833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.026167; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.026167; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.029167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.083833; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.026167
199.5262 | 177.98 | 178.04 | 178 | 178.037 |  |  |  |  | 177.95 | 178 |  |  | 178.0012 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.02; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.021167; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.038833; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0011667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.035833; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.051167; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0011667
211.3489 | 177.92 | 177.97 | 177.9 | 177.972 |  |  |  |  | 177.87 | 178 |  |  | 177.9387 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.021; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.018667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.031333; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.038667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.033333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.068667; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.061333
223.8721 | 177.84 | 177.89 | 177.9 | 177.892 |  |  |  |  | 177.8 | 177.9 |  |  | 177.8703 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.022; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.030333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.019667; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.029667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.021667; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.070333; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.029667
237.1374 | 177.76 | 177.81 | 177.8 | 177.815 |  |  |  |  | 177.73 | 177.8 |  |  | 177.7858 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.023; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.025833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.024167; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.014167; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.029167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.055833; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.014167
251.1886 | 177.67 | 177.72 | 177.7 | 177.734 |  |  |  |  | 177.64 | 177.7 |  |  | 177.694 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.037; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.024; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.026; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.006; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.04; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.054; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.006
266.0725 | 177.57 | 177.63 | 177.6 | 177.637 |  |  |  |  | 177.55 | 177.6 |  |  | 177.5978 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.028; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.027833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.032167; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0021667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.039167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.047833; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0021667
281.8383 | 177.47 | 177.53 | 177.5 | 177.536 |  |  |  |  | 177.45 | 177.5 |  |  | 177.4977 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.027; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.027667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.032333; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0023333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.038333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.047667; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0023333
298.5383 | 177.35 | 177.42 | 177.4 | 177.428 |  |  |  |  | 177.34 | 177.4 |  |  | 177.3897 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.029; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.039667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.030333; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.010333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.038333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.049667; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.010333
316.2278 | 177.24 | 177.31 | 177.3 | 177.316 |  |  |  |  | 177.23 | 177.3 |  |  | 177.2827 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.031; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.042667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.027333; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.017333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.033333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.052667; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.017333
334.9654 | 177.12 | 177.19 | 177.2 | 177.188 |  |  |  |  | 177.11 | 177.2 |  |  | 177.168 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.033; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.048; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.022; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.032; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.02; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.058; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.032
354.8134 | 176.99 | 177.05 | 177 | 177.061 |  |  |  |  | 176.98 | 177 |  |  | 177.0135 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.035; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0235; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0365; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0135; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0475; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.0335; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0135
375.8374 | 176.87 | 176.91 | 176.9 | 176.927 |  |  |  |  | 176.84 | 176.9 |  |  | 176.8912 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.037; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.021167; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.018833; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0088333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.035833; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.051167; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0088333
398.1072 | 176.72 | 176.76 | 176.7 | 176.772 |  |  |  |  | 176.69 | 176.8 |  |  | 176.7403 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.039; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.020333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.019667; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.040333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.031667; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.050333; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.059667
421.6965 | 176.57 | 176.6 | 176.6 | 176.614 |  |  |  |  | 176.53 | 176.6 |  |  | 176.5857 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.041; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.015667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.014333; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.014333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.028333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.055667; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.014333
446.6836 | 176.39 | 176.43 | 176.4 | 176.442 |  |  |  |  | 176.37 | 176.4 |  |  | 176.4053 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.045; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.015333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.024667; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0053333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.036667; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.035333; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0053333
473.1513 | 176.2 | 176.25 | 176.2 | 176.261 |  |  |  |  | 176.19 | 176.2 |  |  | 176.2168 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.047; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.016833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.033167; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.016833; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.044167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.026833; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.016833
501.1872 | 176.03 | 176.06 | 176 | 176.07 |  |  |  |  | 176 | 176 |  |  | 176.0267 | Uncertainties (degrees) UK: 0.4; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.049; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: 0.0033333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.033333; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.026667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.043333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.026667; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.026667
530.8844 | 175.81 | 175.85 | 175.8 | 175.865 |  |  |  |  | 175.79 | 175.8 |  |  | 175.8192 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.053; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0091667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.030833; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.019167; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.045833; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.029167; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.019167
562.3413 | 175.59 | 175.63 | 175.6 | 175.647 |  |  |  |  | 175.57 | 175.6 |  |  | 175.6062 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.056; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.016167; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.023833; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0061667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.040833; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.036167; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0061667
595.6621 | 175.35 | 175.4 | 175.4 | 175.414 |  |  |  |  | 175.35 | 175.4 |  |  | 175.3857 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.059; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.035667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.014333; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.014333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.028333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.035667; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.014333
630.9573 | 175.12 | 175.16 | 175.1 | 175.17 |  |  |  |  | 175.1 | 175.1 |  |  | 175.125 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.063; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.005; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.035; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.025; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.045; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.025; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.025
668.3439 | 174.84 | 174.89 | 174.9 | 174.904 |  |  |  |  | 174.84 | 174.9 |  |  | 174.879 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.067; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.039; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.011; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.021; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.025; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.039; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.021
707.9458 | 174.55 | 174.61 | 174.6 | 174.625 |  |  |  |  | 174.56 | 174.6 |  |  | 174.5908 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.071; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.040833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.019167; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0091667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.034167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.030833; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0091667
749.8942 | 174.26 | 174.31 | 174.3 | 174.321 |  |  |  |  | 174.27 | 174.3 |  |  | 174.2935 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.077; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0335; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0165; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0065; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0275; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.0235; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0065
794.3282 | 173.96 | 173.99 | 174 | 174.008 |  |  |  |  | 173.95 | 174 |  |  | 173.9847 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.08; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.024667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0053333; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.015333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.023333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.034667; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.015333
841.3951 | 173.62 | 173.65 | 173.6 | 173.672 |  |  |  |  | 173.62 | 173.6 |  |  | 173.627 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.085; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.007; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.023; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.027; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.045; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.007; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.027
891.2509 | 173.23 | 173.3 | 173.3 | 173.314 |  |  |  |  | 173.26 | 173.3 |  |  | 173.284 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.091; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.054; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.016; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.016; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.03; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.024; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.016
944.0609 | 172.86 | 172.92 | 172.9 | 172.936 |  |  |  |  | 172.88 | 172.9 |  |  | 172.8993 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.096; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.039333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.020667; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.00066667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.036667; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.019333; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.00066667
1000 | 172.48 | 172.51 | 172.5 | 172.528 |  |  |  |  | 172.47 | 172.5 |  |  | 172.498 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 0.4; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.102; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.018; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.012; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.002; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.03; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.028; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.002
1059.2537 | 172.05 | 172.07 | 172 | 172.094 |  |  |  |  | 172.04 | 172.1 |  |  | 172.059 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.108; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.009; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.011; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.059; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.035; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.019; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.041
1122.0185 | 171.59 | 171.62 | 171.6 | 171.635 |  |  |  |  | 171.59 | 171.6 |  |  | 171.6058 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.115; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.015833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.014167; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0058333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.029167; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.015833; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0058333
1188.5022 | 171.03 | 171.12 | 171.1 | 171.142 |  |  |  |  | 171.09 | 171.1 |  |  | 171.097 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.122; Uncertainties (degrees) SE: 0.4; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.067; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.023; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.003; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.045; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.007; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.003
1258.9254 | 170.57 | 170.6 | 170.6 | 170.621 |  |  |  |  | 170.57 | 170.6 |  |  | 170.5935 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.13; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0235; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0065; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.0065; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0275; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.0235; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0065
1333.5214 | 170 | 170.03 | 170 | 170.063 |  |  |  |  | 170.01 | 170 |  |  | 170.0172 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.138; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.017167; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.012833; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.017167; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.045833; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.0071667; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.017167
1412.5375 | 169.41 | 169.43 | 169.4 | 169.466 |  |  |  |  | 169.42 | 169.4 |  |  | 169.421 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.147; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.011; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.009; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.021; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.045; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.001; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.021
1496.2357 | 168.78 | 168.8 | 168.8 | 168.833 |  |  |  |  | 168.79 | 168.8 |  |  | 168.8005 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.155; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0205; Phase difference relative to estimated sensitivity phase (degrees) FR: -0.0005; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0005; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0325; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.0105; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0005
1584.8932 | 168.08 | 168.13 | 168.1 | 168.149 |  |  |  |  | 168.11 | 168.1 |  |  | 168.1115 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.5; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.165; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0315; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0185; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0115; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0375; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.0015; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0115
1678.804 | 167.32 | 167.4 | 167.4 | 167.426 |  |  |  |  | 167.38 | 167.4 |  |  | 167.3877 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.6; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.175; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.067667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.012333; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.012333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.038333; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.0076667; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.012333
1778.2794 | 166.56 | 166.62 | 166.6 | 166.641 |  |  |  |  | 166.6 | 166.6 |  |  | 166.6035 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.6; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.187; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0435; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0165; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0035; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0375; Phase difference relative to estimated sensitivity phase (degrees) SE: -0.0035; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0035
1883.6491 | 165.7 | 165.78 | 165.7 | 165.809 |  |  |  |  | 165.77 | 165.7 |  |  | 165.7432 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.6; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.198; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.043167; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.036833; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.043167; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.065833; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.026833; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.043167
1995.2623 | 164.8 | 164.88 | 164.8 | 164.911 |  |  |  |  | 164.87 | 164.8 |  |  | 164.8435 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.6; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.211; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0435; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0365; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0435; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0675; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.0265; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0435
2113.489 | 163.88 | 163.93 | 163.9 | 163.946 |  |  |  |  | 163.92 | 163.9 |  |  | 163.9127 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.7; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.225; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.032667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.017333; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.012667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.033333; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.0073333; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.012667
2238.7211 | 162.81 | 162.88 | 162.8 | 162.908 |  |  |  |  | 162.88 | 162.8 |  |  | 162.8463 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.7; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.239; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.036333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.033667; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.046333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.061667; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.033667; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.046333
2371.3737 | 161.7 | 161.77 | 161.7 | 161.786 |  |  |  |  | 161.76 | 161.7 |  |  | 161.736 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.7; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.255; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.036; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.034; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.036; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.05; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.024; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.036
2511.8864 | 160.49 | 160.53 | 160.5 | 160.576 |  |  |  |  | 160.56 | 160.5 |  |  | 160.526 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 0.7; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.27; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.036; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.004; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.026; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.05; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.034; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.026
2660.7251 | 159.16 | 159.21 | 159.2 | 159.269 |  |  |  |  | 159.25 | 159.2 |  |  | 159.2148 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 0.8; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.288; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.054833; Phase difference relative to estimated sensitivity phase (degrees) FR: -0.0048333; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.014833; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.054167; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.035167; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.014833
2818.3829 | 157.74 | 157.82 | 157.7 | 157.839 |  |  |  |  | 157.82 | 157.7 |  |  | 157.7698 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 0.8; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.304; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.029833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.050167; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.069833; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.069167; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.050167; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.069833
2985.3826 | 156.13 | 156.25 | 156.2 | 156.295 |  |  |  |  | 156.29 | 156.2 |  |  | 156.2275 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 0.8; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.324; Uncertainties (degrees) SE: 0.5; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0975; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0225; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.0275; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0675; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.0625; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0275
3162.2777 | 154.51 | 154.59 | 154.5 | 154.606 |  |  |  |  | 154.61 | 154.5 |  |  | 154.5527 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 0.8; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.344; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.042667; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.037333; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.052667; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.053333; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.057333; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.052667
3349.6544 | 152.62 | 152.73 | 152.6 | 152.744 |  |  |  |  | 152.76 | 152.7 |  |  | 152.6923 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 0.9; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.363; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.072333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.037667; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.092333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.051667; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.067667; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0076667
3548.1339 | 150.59 | 150.66 | 150.6 | 150.741 |  |  |  |  | 150.75 | 150.6 |  |  | 150.6568 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 0.9; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.384; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.066833; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0031667; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.056833; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.084167; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.093167; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.056833
3758.374 | 148.35 | 148.47 | 148.4 | 148.514 |  |  |  |  | 148.52 | 148.4 |  |  | 148.4423 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 0.9; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.405; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.092333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.027667; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.042333; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.071667; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.077667; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.042333
3981.0717 | 145.93 | 146.02 | 145.9 | 146.056 |  |  |  |  | 146.08 | 145.9 |  |  | 145.981 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 0.9; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.426; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.051; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.039; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.081; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.075; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.099; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.081
4216.965 | 143.18 | 143.31 | 143.2 | 143.333 |  |  |  |  | 143.38 | 143.2 |  |  | 143.2672 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 1; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.446; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.087167; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.042833; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.067167; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.065833; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.11283; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.067167
4466.8359 | 140.14 | 140.28 | 140.2 | 140.319 |  |  |  |  | 140.38 | 140.2 |  |  | 140.2532 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 1; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.465; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.11317; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.026833; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.053167; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.065833; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.12683; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.053167
4731.5126 | 136.85 | 136.94 | 136.8 | 136.983 |  |  |  |  | 137.05 | 136.9 |  |  | 136.9205 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 1; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.473; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 0.5; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0705; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0195; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.1205; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0625; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.1295; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0205
5011.8723 | 133.15 | 133.23 | 133.1 | 133.28 |  |  |  |  | 133.35 | 133.2 |  |  | 133.2183 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 1; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.49; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.068333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.011667; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.11833; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.061667; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.13167; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.018333
5308.8444 | 128.97 | 129.1 | 128.9 | 129.146 |  |  |  |  | 129.24 | 129.1 |  |  | 129.076 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 1.1; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.481; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.106; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.024; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.176; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.07; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.164; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.024
5623.4133 | 124.45 | 124.53 | 124.3 | 124.56 |  |  |  |  | 124.71 | 124.5 |  |  | 124.5083 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 1.1; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.484; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.058333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.021667; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.20833; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.051667; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.20167; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.0083333
5956.6214 | 119.34 | 119.53 | 119.3 | 119.517 |  |  |  |  | 119.69 | 119.5 |  |  | 119.4795 | Uncertainties (degrees) UK: 0.7; Uncertainties (degrees) FR: 1.1; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.478; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.1395; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0505; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.1795; Phase difference relative to estimated sensitivity phase (degrees) CH: 0.0375; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.2105; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.0205
6309.5734 | 113.92 | 114.04 | 113.8 | 114.012 |  |  |  |  | 114.22 | 114.1 |  |  | 114.0153 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 1.1; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.481; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.095333; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.024667; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.21533; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.0033333; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.20467; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.084667
6683.4392 | 108.02 | 108.14 | 107.9 | 108.069 |  |  |  |  | 108.29 | 108.2 |  |  | 108.1032 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 1.2; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.483; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.083167; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.036833; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.20317; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.034167; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.18683; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.096833
7079.4578 | 101.64 | 101.87 | 101.6 | 101.713 |  |  |  |  | 101.98 | 101.9 |  |  | 101.7838 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 1.2; Uncertainties (degrees) IT: 0.6; Uncertainties (degrees) CH: 0.506; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.14383; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.086167; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.18383; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.070833; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.19617; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.11617
7498.9421 | 95.07 | 95.29 | 95 | 95.118 |  |  |  |  | 95.42 | 95.3 |  |  | 95.1997 | Uncertainties (degrees) UK: 0.6; Uncertainties (degrees) FR: 1.2; Uncertainties (degrees) IT: 0.7; Uncertainties (degrees) CH: 0.531; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.12967; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.090333; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.19967; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.081667; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.22033; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.10033
7943.2823 | 88.49 | 88.62 | 88.4 | 88.423 |  |  |  |  | 88.76 | 88.7 |  |  | 88.5655 | Uncertainties (degrees) UK: 0.5; Uncertainties (degrees) FR: 1.3; Uncertainties (degrees) IT: 0.8; Uncertainties (degrees) CH: 0.564; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.0755; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.0545; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.1655; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.1425; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.1945; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.1345
8413.9514 | 81.9 | 82.05 | 81.9 | 81.802 |  |  |  |  | 82.21 | 82.2 |  |  | 82.0103 | Uncertainties (degrees) UK: 0.8; Uncertainties (degrees) FR: 1.3; Uncertainties (degrees) IT: 0.8; Uncertainties (degrees) CH: 0.615; Uncertainties (degrees) SE: 0.6; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.11033; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.039667; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.11033; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.20833; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.19967; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.18967
8912.5094 | 75.6 | 75.79 | 75.6 | 75.504 |  |  |  |  | 75.89 | 76 |  |  | 75.7307 | Uncertainties (degrees) UK: 0.8; Uncertainties (degrees) FR: 1.3; Uncertainties (degrees) IT: 0.8; Uncertainties (degrees) CH: 0.702; Uncertainties (degrees) SE: 0.7; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.13067; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.059333; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.13067; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.22667; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.15933; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.26933
9440.6088 | 69.89 | 70.09 | 69.8 | 69.701 |  |  |  |  | 69.96 | 70.2 |  |  | 69.9402 | Uncertainties (degrees) UK: 0.8; Uncertainties (degrees) FR: 1.4; Uncertainties (degrees) IT: 0.9; Uncertainties (degrees) CH: 0.83; Uncertainties (degrees) SE: 0.7; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.050167; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.14983; Phase difference relative to estimated sensitivity phase (degrees) IT: -0.14017; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.23917; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.019833; Phase difference relative to estimated sensitivity phase (degrees) DE: 0.25983
10000 | 64.47 | 64.9 | 64.7 | 64.376 |  |  |  |  | 65.07 | 64.6 |  |  | 64.686 | Uncertainties (degrees) UK: 0.8; Uncertainties (degrees) FR: 1.6; Uncertainties (degrees) IT: 1.7; Uncertainties (degrees) CH: 1.011; Uncertainties (degrees) SE: 0.7; Uncertainties (degrees) DE: 1; Phase difference relative to estimated sensitivity phase (degrees) UK: -0.216; Phase difference relative to estimated sensitivity phase (degrees) FR: 0.214; Phase difference relative to estimated sensitivity phase (degrees) IT: 0.014; Phase difference relative to estimated sensitivity phase (degrees) CH: -0.31; Phase difference relative to estimated sensitivity phase (degrees) SE: 0.384; Phase difference relative to estimated sensitivity phase (degrees) DE: -0.086